35457 | 1220648-2023-02278 | 17081006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17081006&pc=OZD | 02/17/2023 | 10/31/2023 |  | 06/07/2023 |  |  | 02/17/2023 | 06/07/2023 |  |  | 12/31/2024 |  | 06/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024285099 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Material Split, Cut or Torn | [5 empty] | Unspecified Infection | Ischemia Stroke | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 02/17/2023 |  |  |  |  | The user facility reported a 24-year-old female presented with acute myocardial infarction/cardiogenic shock and an impella 5.5 was utilized for support awaiting durable ventricular assist device. the patient was on support for 85 days which is beyond the indicated use. there was a concern for infection when the sterile sleeve was compromised. the team gave antibiotics for the infection. additionally, the patient had a cardiovascular accident (cva) which was treated in interventional radiology by thrombectomy. the patient had signs of impairment from the cva in the form of slurred speech and left-sided weakness. | In addition to the original report, the device's anticontamination sleeve "ripped off" during use. the physician cleaned the device and taped it with tegaderm. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: infection: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ischemic cva: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the infection/sepsis and the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the infection was gall bladder patient condition related; the cause of the stroke/cva was not determined since without imaging evidence there is no way to confirm the source of the stroke was due to a clot that passed through the impella device, the relation to impella is unknown. the failure modes will be monitored and trended. | The investigation into the reported "infection/sepsis and stroke/cva" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the cause of the infection was gall bladder patient condition related. the cause of the reported cva was not determined since without imaging evidence there is no way to confirm the source of the stroke was due to a clot that passed through the impella device, the relation to impella is unknown. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ b5 - included report of anticontamination sleeve damage. | The investigation into the infection/sepsis, the stroke/cva and the product damage (sterile sleeve damage) has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the clinical information provided, the root cause of the infection was gall bladder patient condition related. the root cause of the stroke/ cva issue was not determined since without imaging evidence there is no way to confirm the source of the stroke was due to a clot that passed through the impella device, the relation to impella is unknown. the root cause of the sterile sleeve damage issue was not determined since product was not returned for investigation.
35458 | 1220648-2023-02279 | 17083247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17083247&pc=OZD | 04/24/2023 | 06/07/2023 |  | 06/07/2023 |  |  | 04/24/2023 | 06/07/2023 |  |  | 12/31/2024 |  | 06/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321317 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/24/2023 |  |  |  |  | The user facility reported a 68-year-old female presented in acute myocardial infarction and cardiogenic shock. an impella device was implanted for support. the team observed that the 14fr sheath was occlusive to the artery and so a contralateral leg accessed perfusion sheath was placed. it was noted the patient had intermittent non-pulsitility. additionally, the access site was seen to be a place of blood loss that prompted the team to deliver blood products. the following day the family made the patient a dnr. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: limb ischemia: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ access site bleeding: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35459 | 1220648-2023-02280 | 17084142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17084142&pc=OZD | 05/18/2023 | 01/18/2024 |  | 06/07/2023 |  |  | 05/18/2023 | 06/07/2023 |  |  | 01/31/2025 |  | 06/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024330939 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2023 |  |  |  |  | The user facility reported a 76-year-old male with atrial fibrillation presented for electrophysiology ablation. an impella device was inserted for support. the patient had an access site bleed which required two units of blood to make the patient hemodynamically stable. the impella device was removed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the access site bleeding cannot be determined since the complaint device not returned for investigation and lack of clinical details.
35460 | 1220648-2023-02281 | 17084221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17084221&pc=OZD | 05/09/2023 | 01/12/2024 |  | 06/07/2023 |  |  | 05/09/2023 | 06/07/2023 |  |  | 02/29/2024 |  | 06/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200346 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2023 |  |  |  |  | The user facility reported a 70-year-old male noted as high risk percutaneous cardiac insertion and with mitral valve disease presented for impella support. the patient began oozing around the femoral venous swan catheter site. the arterial site had no oozing or bleeding. manual pressure was applied to the venous site. the patient's hematoma worsened after the patient coughed. the venous sheath and swan were removed and manual pressure continued to the site. a ct scan showed a pseudoaneurysm and arterial bleed. it is unclear if pseudoaneurysm was related to the venous or arterial injury and unclear if any surgical intervention was required. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into access site bleeding issue has been completed. the device was not returned for investigation. as per clinical, hematoma at access site had expanded beyond marked borders with cta report showing pseudoaneurysm and arterial bleed. patient with hemolysis and troubleshooting done by decreasing p-level, giving volume. unknown if the hemolysis resolved with volume or reduction of p-level. the cause of the access site bleeding issue was most likely patient condition related due to diseased vessel condition based on the clinical information. the cause of the hemolysis issue was not determined since product and data logs were not returned and with insufficient clinical information.
35461 | 1220648-2023-02271 | 17074912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17074912&pc=OZD | 04/24/2023 | 09/15/2024 |  | 06/06/2023 |  |  | 04/25/2023 | 06/06/2023 |  |  | 01/31/2025 |  | 06/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024313503 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/18/2023 |  |  |  |  | A 60-year-old male with 100% occlusion in left main coronary artery presented for impella support. the user facility reported the patient endured a stroke while on impella support. the cause of the stroke was not known, however, the impella was not ruled out as a factor. | It was reported the patient had noticeable clotting within the impella peel-away sheath as well as other peripheral line locations. the patient then began experiencing stroke-like symptoms and a neurological assessment and computed topography were performed and the stroke was verified. the decision was made to explant the impella cp as a precaution to mitigate and further clotting and possible escalation of clot-related issues. upon explant of the impella deice, large amounts of clots were observed on the impella catheter and within the sheath. it is noted that the survival outcome of procedure that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva has been completed since the original report was filed. since the medical center did not return the impella device, no benchtop analysis was possible, without imaging evidence to confirm the source of the stroke was due to a clot that passed through the impella device or the introducer, the relation to impella is unknown. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02271 was provided.
35462 | 1220648-2023-02272 | 17074952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17074952&pc=OZD | 05/08/2023 | 07/13/2023 |  | 06/06/2023 |  |  | 05/08/2023 | 06/06/2023 |  |  | 12/31/2024 |  | 06/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024308757 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2023 |  |  |  |  | A 46-year-old male post cardiotomy cardiogenic shock presented for impella support. the user facility reported the patient's white blood cell counts increased while on support. the impella device, along with all other devices in patient, were removed due to an unknown source of potential infection. after removal of the impella and surgical graft, an embolectomy was required. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the reported thrombus and infection/sepsis has been completed since the original report was filed. the root cause of thrombus was unable to be determined since the product was not returned, no imaging was available ant there is insufficient clinical information. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35463 | 1220648-2023-02266 | 17058051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17058051&pc=OZD | 02/08/2020 | 06/05/2023 |  | 06/05/2023 |  |  | 02/08/2020 | 06/05/2023 |  |  | 07/31/2021 |  | 06/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1448642 |  | 005042-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/08/2020 |  |  |  |  | The complainant reported a 65 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during an attempted repositioning of the device, it fell out of the ventricle. the device was unable to be advanced past the aortic valve, therefore, removal of pump was deemed necessary. resistance was encountered upon removal of the device back through the peel away sheath and the catheter separated from the cannula and became lodged within the peel away sheath. the component was removed with the sheath and did not require additional intervention. | [5 empty] | The device was returned and an investigation was completed. a confirmed cannula separation was confirmed at the base, near the outflow cage. there were no tears of the polyurethane and no plastic deformation. as such, we believe the root cause of the cannula detachment was most likely a failure of the delo-bead joint due to difficult removal through the kinked introducer.
35464 | 1220648-2023-02265 | 17058064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17058064&pc=PYX | 10/25/2021 | 06/04/2023 |  | 06/05/2023 |  |  | 10/25/2021 | 06/05/2023 |  |  | 03/31/2023 |  | 06/05/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022065827 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/25/2021 |  |  |  |  | The user facility reported after performing serial dilation with the impella rp introducer kit and before loading 23 fr dilator and sheath onto 035 super-stiff wire, the physician and field representative noticed a crack/split in two spots that made a sharp edge that could potentially tear the vein. the 23 fr sheath was replaced and impella rp support proceeded without harm or injury to the 51-year-old male. | [5 empty] | The device was not returned for evaluation. a field service representative was present when the split in the introducer tip was discovered, therefore, this is confirmed. the user facility reported noticing the split prior to insertion into the patient. there is not enough information to fully determine the cause of the introducer tip split.
35465 | 1220648-2023-02267 | 17058400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17058400&pc=OZD | 05/07/2023 | 06/09/2023 |  | 06/05/2023 |  |  | 05/07/2023 | 06/05/2023 |  |  | 01/31/2025 |  | 06/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024309060 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/07/2023 |  |  |  |  | A 49-year-old male with history of cardiogenic shock and cardiomegaly presented for impella support. the user facility reported the patient initially had an impella cp implanted and was escalated to an impella 5.5. the pump stopped during the support. the device was explanted. the patient later expired when the family made the decision to withdraw care. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: try to restart the catheter at previous p-level. if the impella does not restart, try to restart at p-2. if the impella does not restart or stops again, wait 1 minute and try to restart again. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation has been ongoing since the original report was filed. during the investigation, it was determined the bipella patient did have a pump stop, however, it was not the impella 5.5 as originally reported. the pump that stopped was the right-sided impella rp which is a non-reportable event. no report should have been filed for this event.
35466 | 1220648-2023-02268 | 17061939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17061939&pc=OZD | 05/11/2023 | 10/29/2024 |  | 06/05/2023 |  |  | 05/11/2023 | 06/05/2023 |  |  | 01/31/2025 |  | 06/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328977 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/11/2023 |  |  |  |  | The user facility reported a 67-year-old male in ventricular tachycardia arrest presented for impella support. the impella-accessed limb required perfusion and, to reduce limb ischemia, the team placed distal perfusion. the patient later expired with care withdrawn by the family as the patient had underlying ischemic bowel perforation. | Additional information was received from the complainant that the patient was bleeding around all catheter sites and received blood and platelets. it was reported the bleeding improved with added sutures. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was most likely patient condition related. the failure mode will be monitored and trended. | Additional information for the initial mfr 1220648-2023-02268 was provided in sections b5, d6a, d6b, h6 and h10. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
35467 | 1220648-2023-02269 | 17063144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17063144&pc=OZD | 05/10/2023 | 07/12/2023 |  | 06/05/2023 |  |  | 05/10/2023 | 06/05/2023 |  |  | 12/31/2024 |  | 06/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024308767 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2023 |  |  |  |  | A 41-year-old male in myocardial infarction presented for impella support. the user facility reported the patient was to undergo a coronary artery bypass grafting as preparation for left ventricular assist device (lvad) or heart transplant. during the impella support, a diagnosis was made of an embolic cardiovascular accident. limb and facial impairment were noted. the patient remained on impella support for seven days awaiting bridge to lvad. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the stroke/cva was most likely patient condition related. the failure mode will be monitored and trended.
35468 | 1220648-2023-02270 | 17063824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17063824&pc=OZD | 05/13/2023 | 07/25/2023 |  | 06/05/2023 |  |  | 05/16/2023 | 06/05/2023 |  |  | 08/31/2024 |  | 06/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023267291 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2023 |  |  |  |  | The user facility reported a 72-year-old male noted as high risk surgery presented for impella 5.5 support. the device was placed via the axillary artery and graft. the 5.5 was placed to support the patient with heart failure and planned mitraclip procedure. on day seven of the support, the team had to replace the purge cassette due to a leak. there was no patient harm related to purge leak and no potential for patient harm. additionally, there was an access site bleed. the bleed was treated by multiple interventions, of reducing heparin, infusing red blood cells, plasma protein fraction and packed cells. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The medical center in japan did not return for analysis any impella pump product. only the clinical report was investigated. act and aptt levels were shared for analysis. the root cause of access site bleeding was most likely due to anticoagulation management issue.
35469 | 1220648-2023-02235 | 17052960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17052960&pc=OZD | 05/05/2023 | 06/23/2023 |  | 06/02/2023 |  |  | 05/05/2023 | 06/02/2023 |  |  | 12/31/2024 |  | 06/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318008 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/05/2023 |  |  |  |  | A 47 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported a hematoma formed after the patient went to surgery for coronary artery bypass grafting (cabg). transfusion and surgical intervention were required. upon intervention, it was noticed that the vessel was not damaged but the bleeding was coming from around the impella access site. the device was removed and patient remained stable. | [5 empty] | No product was returned. the cause of the access site bleeding was high patient act values based on the provided clinical information. | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35470 | 1220648-2023-02262 | 17054299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17054299&pc=OZD | 10/22/2020 | 08/17/2023 |  | 06/02/2023 |  |  | 10/22/2020 | 06/02/2023 |  |  | 07/31/2022 |  | 06/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 1496179 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2020 |  |  |  |  | A 59-year-old female in cardiogenic shock presented for impella 5.5 support. the user facility reported the device stopped on day 16 of support. this is past the indicated use. the pump stopped after previous purge issues were attempted to be resolved with tissue plasminogen activator (tpa) infusion. the pump was explanted, replaced with a new pump and support proceeded. | [5 empty] | The investigation into the pump stop has been completed since the original report was filed. the user facility has returned the impella device, and the benchtop analysis revealed that the root cause of the pump stop was blood ingress into the motor due to a kinked purge system. | The user facility returned the motor for evaluation. the pump handle was not returned, therefore, the inspector was unable to check for kinks. the pump was confirmed to have blood ingress. the most likely cause of the pump stop was blood ingress into the motor due to a kinked purge system. of note the ifu states: "impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency. maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup."
35471 | 1220648-2023-02263 | 17054605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17054605&pc=OZD | 05/14/2023 | 07/25/2023 |  | 06/02/2023 |  |  | 05/14/2023 | 06/02/2023 |  |  | 11/30/2024 |  | 06/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023280069 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2023 |  |  |  |  | A 40-year-old male in cardiogenic shock presented for impella support. the user facility reported the patient had ongoing support for over 67 days when there was a purge leak in the pump. this was remedied by the use of purge bypass, derma bond, and bone wax. the team then had to replace the pump after it stopped within five hours of the purge leak. there was no harm or injury to the patient. | [5 empty] | Since the original report was submitted the investigation into the purge leak and subsequent pump stop has concluded. the data logs were returned for analysis. part of the pump was returned, but not the purge sidearm. the cause of the pump stop was most likely a leak from the sidearm but the exact location of the leak was unknown. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: purge: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ pump stop: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35472 | 1220648-2023-02264 | 17054606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17054606&pc=OZD | 05/05/2023 | 07/27/2023 |  | 06/02/2023 |  |  | 05/05/2023 | 06/02/2023 |  |  | 12/31/2024 |  | 06/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312837 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/05/2023 |  |  |  |  | An 87-year-old male noted as high risk percutaneous cardiac insertion presented for impella support. the patient was on support for three days when the user facility reported a separation of the cannula during removal of the device. the physician was able to manually grasp and remove the component without surgical intervention. a hematoma developed and was treated with manual pressure, however, reduced distal pulses were reported. vascular surgery was consulted for possible intervention. | [5 empty] | The investigation into the cannula detachment and the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the cannula detachment was most likely patient condition related since resistance was felt in the iliac artery during removal, and the cause of the limb ischemia was most likely use related since the ischemia resulted from the removal complications requiring continual manual pressure. the failure modes will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: cannula detachment: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ impella 5.5 & cp: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ impella rp: ¿handle with care. the impella rp with smartassist system catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ limb ischemia: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35473 | 1220648-2023-02258 | 17037598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17037598&pc=OZD | 05/04/2023 | 08/09/2023 |  | 06/01/2023 |  |  | 05/04/2023 | 06/01/2023 |  |  | 06/30/2023 |  | 06/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022097957 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2023 |  |  |  |  | A 52-year-old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the 14 fr sheath was used for insertion. the sheath had sluggish flow down the impella-accessed limb so the team removed peel-away and placed the reposheath. the team repositioned the impella and discovered a hematoma. the hematoma had manual pressure held and blood products were given to the patient. the pump was explanted on day 3 of support after a successful wean. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the access site bleeding most likely was patient condition related. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35474 | 1220648-2023-02260 | 17043505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17043505&pc=OZD | 05/02/2023 | 07/27/2023 |  | 06/01/2023 |  |  | 05/03/2023 | 06/01/2023 |  |  | 12/31/2024 |  | 06/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024307339 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2023 |  |  |  |  | A 61-year-old male in cardiogenic shock presented form impella support. the user facility reported difficulty inserting the device resulting in a wireless placement via the carotid artery. after impella support, the patient exhibited aortic insufficiency (ai) requiring replacement of the patient's existing bioprosthetic valve. | [5 empty] | The investigation into the heart valve damage has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the root cause of the issue was most likely patient condition related to the prosthetic av. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
35475 | 1220648-2023-02261 | 17043506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17043506&pc=OZD | 05/03/2023 | 09/17/2025 |  | 06/01/2023 |  |  | 05/03/2023 | 06/01/2023 |  |  | 01/31/2024 |  | 06/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190953 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 13 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/03/2023 |  |  |  |  | A 13-year-old male with end-stage heart failure presented for impella support. the user facility reported weak dopplerable pulse and cool extremity in the impella inserted arm. the decision was made to remove out of caution for limb ischemia. | Additional information indicated that, during patient support, the pump experienced placement signal unreliable alarm. |  |  |  |  | The investigation into the limb ischemia and the optical signal issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the limb ischemia was most likely patient condition related, and the cause of the optical signal issue was not determined since product was not returned. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. b5 updated with additional information from the initial manufacturer device report number 1220648-2023-02261. d6a/d6b updated with additional information. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures from the initial manufacturer device report number 1220648-2023-02261. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2023-02261.
35476 | 17030143 | 17030143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17030143&pc=OZD | 05/02/2023 | 05/18/2023 | 05/18/2023 | 05/31/2023 | 05/18/2023 | 05/31/2023 |  | 05/31/2023 |  |  |  |  | 05/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED INC. |  |  | US |  |  |  | 0042-0040-US |  | Yes |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 22995 DA |  | Male |  |  |  | Malfunction | No | Yes | Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | After transfer from bedside commode to bed, noted exposer of distal portion of impella driveline where previously placed tegaderm had separated from the distal hub. overnight provider was immediately notified and came to bedside to visualize. exposed site was re-sterilized with betadine and new tegaderm placed. manufacturer response for temporary non-roller type left heart support blood pump, abiomed impella (per site reporter) they are actively investigating.
35477 | 1220648-2023-02249 | 17032424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17032424&pc=OZD | 05/09/2023 | 06/23/2023 |  | 05/31/2023 |  |  | 05/09/2023 | 05/31/2023 |  |  | 12/31/2024 |  | 05/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024319287 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2023 |  |  |  |  | A 79 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the physician attempted to place the impella cp but lost access to the right femoral artery during percutaneous coronary intervention (pci). the physician re-accessed the right femoral artery and placed the impella cp. the patient had a hematoma present while in the catheterization lab and a growing hematoma in the critical care unit requiring manual pressure, femstop and two units of blood. | [5 empty] | No product was returned for evaluation. the cause of the access site bleeding issue was most likely use related with losing access to rfa. | The impella device was discarded by the customer and therefore, evaluation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, bleeding is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35478 | 1220648-2023-02250 | 17032588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17032588&pc=OZD | 05/15/2023 | 08/11/2023 |  | 05/31/2023 |  |  | 05/15/2023 | 05/31/2023 |  |  | 12/31/2024 |  | 05/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318899 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2023 |  |  |  |  | A 66 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient developed a large hematoma. surgical intervention and a blood transfusion of two units of blood products were required. the device was explanted due to vessel needing repair for bleeding. the patient was reported as stable. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the access site bleeding issue was most likely use related. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35479 | 1220648-2023-02251 | 17032680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17032680&pc=OZD | 04/28/2023 | 08/30/2023 |  | 05/31/2023 |  |  | 04/28/2023 | 05/31/2023 |  |  | 03/31/2024 |  | 05/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204593 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/28/2023 |  |  |  |  | An 80 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the physician opted to wean the patient from the device and explant the device due to concern for disseminated intravascular coagulation (dic) and hemolysis. it is unknown if hemolysis could be confirmed. | [5 empty] | The investigation into the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the hemolysis issue was improper positioning due to improper technique as tte at bedside showed impella ~2cm from av, but did not reposition it. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35480 | 1220648-2023-02252 | 17032728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17032728&pc=OZD | 04/27/2023 | 08/11/2023 |  | 05/31/2023 |  |  | 04/27/2023 | 05/31/2023 |  |  | 12/31/2024 |  | 05/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024319334 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2023 |  |  |  |  | A 57 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs)presented for impella support. the user facility reported the patient developed limb ischemia coinciding with impella support. . the pump was subsequently explanted as a result of the noted condition. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was patient condition related. pulse obtained but unable to maintain a map above 50. due to the patient being hard to ventilate, on max vent settings with persistently oxygen saturations the decision was made to place on ecmo. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35481 | 1220648-2023-02253 | 17032744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17032744&pc=OZD | 05/12/2023 | 08/11/2023 |  | 05/31/2023 |  |  | 05/12/2023 | 05/31/2023 |  |  | 09/30/2024 |  | 05/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278154 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2023 |  |  |  |  | A 71 year old male undergoing left heart catherization had impella cp implanted. the user facility reported that post implant, approximately 24-36 hours, the patient had no distal pulse in the same leg. the physician attempted to reposition the cp and the patient's condition did not improve. the family opted to end support. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia is due to the patient's condition. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35482 | 1220648-2023-02254 | 17032763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17032763&pc=OZD | 05/03/2023 | 10/29/2024 |  | 05/31/2023 |  |  | 05/03/2023 | 05/31/2023 |  |  | 03/31/2025 |  | 05/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024327104 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/03/2023 |  |  |  |  | Additional information was received from the complainant indicating the patient experienced hemolysis which was resolved through repositioning of the pump. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | A 70 year old male with out-of-hospital st elevation myocardial infarction (stemi) presented for impella support. the user facility reported the patient had significant bleeding from the access site. three units of packed red blood cells (pbrc) were administered to the patient. |  |  |  |  | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the access site bleeding was most likely related to use issue. | Additional information for the initial mfr 1220648-2023-02254 was provided in sections b5, d6a, d6b, h6 and h10. the impella device was not returned, however, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm wheth er a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35483 | 1220648-2023-02255 | 17032768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17032768&pc=OZD | 04/27/2023 | 08/31/2023 |  | 05/31/2023 |  |  | 04/27/2023 | 05/31/2023 |  |  | 12/31/2024 |  | 05/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024308773 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2023 |  |  |  |  | A 68 year old male in cardiogenic shock presented for impella cp support. the user facility reported the access site began oozing during support. seven units of packed red blood cells (prbc) were administered. following this intervention, the patient support continued with the impella cp. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the access site bleeding was unable to be determined as the product was not returned and insufficient clinical details were provided. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
35484 | 1220648-2023-02256 | 17032795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17032795&pc=OZD | 05/07/2023 | 08/30/2023 |  | 05/31/2023 |  |  | 05/07/2023 | 05/31/2023 |  |  | 03/31/2025 |  | 05/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024326613 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/07/2023 |  |  |  |  | A 54-year-old male presented for impella support during bridging to heart transplant. the user facility report the patient's axillary insertion site had a hematoma. on day three, the team washed out the hematoma and placed a jackson-pratt (jp) drain. the patient remained on impella support for twelve days until transplant. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical details, the root cause of the hematoma cannot be determined since the device was not returned for investigation and lack of clinical details. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35485 | 1220648-2023-02257 | 17034007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17034007&pc=OZD | 05/09/2023 | 07/13/2023 |  | 05/31/2023 |  |  | 05/09/2023 | 05/31/2023 |  |  | 09/30/2024 |  | 05/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273421 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2023 |  |  |  |  | A 58-year-old male in cardiogenic shock presented for impella cp support. the user facility reported the patient had an access site bleed. the bleeding was not controlled after manual pressure for 45 minutes and forward-tension sutures. the patient underwent cutdown and repair to the right femoral artery prior to escalation to impella 5.5 pump. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ vessel damage: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
35486 | 1220648-2023-02259 | 17035118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17035118&pc=OZD | 05/06/2023 | 08/09/2023 |  | 05/31/2023 |  |  | 05/06/2023 | 05/31/2023 |  |  | 09/30/2024 |  | 05/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023280284 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2023 |  |  |  |  | An 81 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient experienced bleeding from the access site and the hemoglobin (hgb) dropped from 11 to 8.5. the patient received three units of packed red blood cells (pbrc)s. the patient remained on bedrest. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. no benchtop analysis was possible to determine the root cause. since the medical center did not return the impella device, the root cause of the access site bleeding was not determined. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follow ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35487 | 1220648-2023-02243 | 17020908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17020908&pc=PYX | 05/11/2023 | 07/12/2023 |  | 05/30/2023 |  |  | 05/11/2023 | 05/30/2023 |  |  | 01/31/2025 |  | 05/30/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024305484 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2023 |  |  |  |  | A 72 year old male in cardiogenic shock presented for impella rp support. the user facility reported the patient had signs of hemolysis and noted the hemolysis was attributed to the use of the rp pump. after, just 23 hours, the pump was explanted and blood level/labs cleared. there was no further intervention performed. | [5 empty] | The investigation into the reported hemolysis has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35488 | 1220648-2023-02245 | 17024098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17024098&pc=OZD | 05/10/2023 | 07/12/2023 |  | 05/30/2023 |  |  | 05/10/2023 | 05/30/2023 |  |  | 10/31/2024 |  | 05/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023274770 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2023 |  |  |  |  | A 60 year old male of unknown race presented for impella support. the user facility reported an acute motor current spike and then the pump stopped. the pump could not be restarted and was withdrawn into the patient's innominate artery and surgically closed. a replacement pump was not utilized. the patient experienced no harm and is doing well. | [5 empty] | The investigation into the pump stop has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on the clinical details, the cause of the pump stop was biomaterial. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35489 | 1220648-2023-02244 | 17024602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17024602&pc=OZD | 04/28/2023 | 08/30/2023 |  | 05/30/2023 |  |  | 04/28/2023 | 05/30/2023 |  |  | 09/30/2024 |  | 05/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283286 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 04/28/2023 |  |  |  |  | A 61 year old female presented for impella support. the user facility reported the patient's legs were both flow compromised but placed the impella cp. the team placed fem-fem bypass. the patient remained on support for 38 hours when the decision was made to explant the pump due to concerns of distal limb perfusion and hemolysis. the patient was stable after premature explant and physician decided not to escalate the patient to impella 5.5. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the hemolysis and limb ischemia was not determined since product was not returned and lack of clinical data. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: limb ischemia: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ hemolysis: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35490 | 1220648-2023-02246 | 17025830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17025830&pc=OZD | 05/01/2023 | 08/09/2023 |  | 05/30/2023 |  |  | 05/01/2023 | 05/30/2023 |  |  | 01/31/2025 |  | 05/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024294787 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2023 |  |  |  |  | A 69 year old male in cardiogenic shock presented for impella support. the user facility report the patient developed a hematoma at the impella axillary site. the patient was taken to the operating room for a surgical washout, venous suture, wound vac placement, and delivery of 1 unit packed red blood cells (prbc) blood products. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the access site bleeding was patient condition related. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
35491 | 1220648-2023-02247 | 17025913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17025913&pc=OZD | 05/01/2023 | 07/13/2023 |  | 05/30/2023 |  |  | 05/02/2023 | 05/30/2023 |  |  | 12/31/2024 |  | 05/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024318643 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2023 |  |  |  |  | A 68 year old female with post cardiotomy cardiogenic shock (pccs) and low cardiac output syndrome (lcos) presented for impella support. the user facility reported potential air bubbles viewed under echocardiogram imaging on the patient's first day of support. there was no reported injury sustained to the patient and support continued for seven additional days. | [5 empty] | The investigation into the reported air embolism has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the air embolism. no product was returned. there are no air in purge system alarms seen in the lt log. the cause of the embolism was most likely patient condition related. the relation to impella is unknown. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
35492 | 1220648-2023-02248 | 17025985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17025985&pc=OZD | 05/12/2023 | 07/24/2023 |  | 05/30/2023 |  |  | 05/12/2023 | 05/30/2023 |  |  | 12/31/2024 |  | 05/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318830 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Leak/Splash | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2023 |  |  |  |  | A 61 year old male with rapid a-fib presented for impella support. the catheterization lab team applied chloraprep to the site and the sheath sidearm disconnected from the sheath. the team held finger over the side arm to maintain hemostasis. the estimated blood loss was 60ml. there was no replacement of blood products to the patient. the pump was explanted and good hemostasis was achieved. | [5 empty] | The investigation into the introducer issue has been completed since the original report was filed. the user facility has returned the device, and the benchtop analysis of the root cause revealed that the cause of the introducer issue was silicone oil contamination. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿
35493 | 1220648-2023-02229 | 17011644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17011644&pc=OZD | 05/07/2023 | 08/02/2023 |  | 05/26/2023 |  |  | 05/07/2023 | 05/26/2023 |  |  | 12/31/2024 |  | 05/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024306422 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | Company representation |  |  | 88 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/07/2023 |  |  |  |  | An 88 year old male in cardiogenic shock presented for impella support. the user facility reported the patient developed an access site bleed. the device's repositioning sheath and act was adjusted and blood products were delivered. | [5 empty] | The investigation into the access site bleeding and the delivery issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the access site bleeding issue was improper anticoagulation management, and the cause of the delivery issues was use related. the failure modes will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
35494 | 1220648-2023-02230 | 17011668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17011668&pc=OZD | 04/13/2023 | 10/14/2024 |  | 05/26/2023 |  |  | 04/13/2023 | 05/26/2023 |  |  | 08/31/2024 |  | 05/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023261480 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/13/2023 |  |  |  |  | A 69 year old female with post cardiotomy cardiogenic shock (pccs) and low cardiac output syndrome (lcos) presented for impella support. the user facility reported patient was showing signs of hemolysis and 2 units of unknown blood products were given to the patient. the patient's mean arterial pressure (map) was not consistent so the decision was made to explant the device. the patient's condition improved after explant. | It was also reported the patient experienced bleeding and ventricular tachycardia while on support. the patient was sent to the or for the bleeding and was given 2 units of packed red blood cells (prbcs), 1 unit of platelets, and 2 units of fresh frozen plasma; for treatment. ventricular tachycardia occurred and the device level was turned down to eliminate the issue; however, short self-limiting runs of ventricular tachycardia continued. therefore, the pump was explanted. |  |  |  |  | The investigation into the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the hemolysis was most likely patient condition related. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Additional information for the initial mfr was provided in sections b5, d6a, d6b, h1, and h6. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the failure modes will be monitored and trended. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
35495 | 1220648-2023-02231 | 17012318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17012318&pc=OZD | 04/29/2023 | 07/13/2023 |  | 05/26/2023 |  |  | 04/29/2023 | 05/26/2023 |  |  | 12/31/2024 |  | 05/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318192 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2023 |  |  |  |  | A 74 year old female with right ventricle dysfunction presented for impella support. the user facility reported the patient developed an access site bleed after patient repositioning during impella support. manual pressure was applied to stop the bleed and two units of packed red blood cells (prbcs) were transfused to counteract the blood loss. patient support continued following the intervention. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the cause of the access site bleeding was most likely patient movement as the patient was being moved when the bleeding occurred. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35496 | 1220648-2023-02232 | 17012428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17012428&pc=OZD | 05/01/2023 | 07/13/2023 |  | 05/26/2023 |  |  | 05/01/2023 | 05/26/2023 |  |  | 09/30/2024 |  | 05/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283139 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2023 |  |  |  |  | A 64 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient developed hemolysis during support. the patient was escalated from an impella cp to an impella 5.5 pump and the patient's condition improved. | [5 empty] | The investigation into the reported hemolysis has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the data logs and clinical report was evaluated. based on clinical description & data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35497 | 1220648-2023-02233 | 17012449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17012449&pc=OZD | 05/02/2023 | 08/11/2023 |  | 05/26/2023 |  |  | 05/02/2023 | 05/26/2023 |  |  | 08/31/2023 |  | 05/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022124518 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Oxygen Saturation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2023 |  |  |  |  | A male of unknown age with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the patient was being transferred to a different hospital. the impella was locked prior to transfer, but the user facility reported that the impella had been placed too shallow and migrated back into the aorta during the transfer. the ems personnel misinterpreted the function of the pump and the loss in measured pulsatility resulted in a delay in support. the patient was in asystole and cpr was delayed due to confusion in the ambulance. once identified that the pump wasn't supporting the patient, cpr was initiated and the pump was properly repositioning. although the patient stabilized, the lack in support resulted in a decrease in o2 levels. ecmo was subsequently placed to treat the condition. | [5 empty] | The investigation into the positioning issues and patient injury has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the positioning issues was use related, and the cause of the patient injury was patient condition related. the failure modes will be monitored and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
35498 | 1220648-2023-02234 | 17012464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17012464&pc=OZD | 05/09/2023 | 10/15/2024 |  | 05/26/2023 |  |  | 05/09/2023 | 05/26/2023 |  |  | 02/28/2025 |  | 05/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325773 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/09/2023 |  |  |  |  | A 69 year old female in post cardiotomy cardiogenic shock (pccs) and low cardiac output syndrome (lcos) presented for impella support. the user facility reported the patient experienced limb ischemia and they were considering escalating the patient from impella cp to impella 5.5. after the cp was explanted, the patient's condition improved and the impella 5.5 was not used. | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella cp device for mechanical circulatory support. additionally, the pump stopped during treatment, and escalation of care to the impella 5.5 was considered. the impella device was explanted, and the patient survived the procedure. |  |  |  |  | The cause of the limb ischemia was most likely use related since both impella and a dialysis line with in the same access site. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported pump stop has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the pump stop was high purge pressure. the cause of the high purge pressure was unknown. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
35499 | 1220648-2023-02236 | 17015138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17015138&pc=PYX | 05/09/2023 | 12/19/2024 |  | 05/26/2023 |  |  | 05/09/2023 | 05/26/2023 |  |  | 01/31/2025 |  | 05/26/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024310620 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Sepsis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 05/09/2023 |  |  |  |  | A 72 year old male with severe right heart dysfunction presented for impella rp support. the user facility reported the physician utilized a pre-existing swan-ganz catheter for insertion of the impella rp. the swan-ganz catheter was 12 days old. the physician declined to change the swan, noting limited access options. the patient's white blood cell count elevated and sepsis was suspected. on (b)(6) 2023, the swan-ganz catheter was removed slowly while the p-level on the pump reduced to p-2 at bedside. shortly after removal of the catheter, pump flows dropped and pulmonary artery pressures increased on the impella rp. the patient was reintubated and was maxed out on pressors/inotropes. the patient expired the following day while on rp support. | [5 empty] | The investigation into the low pump flow and the infection/sepsis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the low pump flow was biomaterial, and the cause of the infection was determined to be patient condition as the pump was confirmed to be sterilized under validated conditions. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02236 was provided in sections d6 and h1. | This report is being filed as part of a retrospective review of historical records. corrected information/additional code added to h6 medical device problem code for low flow.
35500 | 1220648-2023-02237 | 17015139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17015139&pc=PYX | 05/01/2023 | 07/27/2023 |  | 05/26/2023 |  |  | 05/02/2023 | 05/26/2023 |  |  | 03/31/2024 |  | 05/26/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023198065 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2023 |  |  |  |  | A 60 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the pump stopped suddenly and it was not possible to restart it. at removal of the pump, biomaterial was found in the cannula and near the inflow area. | [5 empty] | The investigation into the pump stop and the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the pump stop and the hemolysis was most likely biomaterial since biomaterial was reported to be observed in the cannula after explant. the failure modes will be monitored and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with psmartassist system catheter whenever a rise in motor current is seen.¿
35501 | 1220648-2023-02238 | 17015140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17015140&pc=OZD | 05/03/2023 | 10/25/2023 |  | 05/26/2023 |  |  | 05/09/2023 | 05/26/2023 |  |  | 03/31/2025 |  | 05/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 55 |  |  | ABIOMED, INC. |  |  | US |  | 2024326609 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Leak/Splash | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia Stroke | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/09/2023 |  |  |  |  | In addition to the originally reported access site bleeding and stroke, damage was sustained to the anticontamination sleeve as evident by blood noted leaking into this sleeve during support. | A 58 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support on (b)(6) 2023. the user facility reported the patient experienced access site bleeding on (b)(6) 2023. the patient received 10 units of packed red blood cells (prbc) over several days and other blood products. on (b)(6) 2023, the user facility reported the patient had left parietal stroke that was not recognized until 5/3 due to being sedated and intubated. there was no treatment. it was noted there was a right-sided deficit. the patient continued on impella support and was successfully weaned on (b)(6) 2023. |  |  |  |  | (h3) the investigation into the reported "access site bleeding - major, stroke/cva and product damage (sterile sleeve damage)" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the cause of the blood leaking into the sterile sleeve was not determined since product was not returned. the cause of the access site bleeding was use related since it was resolved with an additional stitch. the cause of the cva/stroke was not determined due to insufficient clinical details. therefore, the reported events have an unknown relation to impella. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records section d4: unique identifier (udi) #. | The cause of the cva/stroke was not determined due to insufficient clinical details. reported unknown relation to impella. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding: ¿assess access site for bleeding and hematoma.¿ stroke: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35502 | 1220648-2023-02240 | 17015142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17015142&pc=OZD | 05/01/2023 | 07/25/2023 |  | 05/26/2023 |  |  | 05/01/2023 | 05/26/2023 |  |  | 12/31/2024 |  | 05/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024319672 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2023 |  |  |  |  | A 46 year old male with cardiogenic shock presented for impella support. the user facility reported the patient had support for two days and then the impella was explanted. the explant was successful but when imaging was performed, a thrombus was revealed in the right iliac/femoral artery. the team utilized thrombectomy with the penumbra. after thrombectomy, the limb resolved and the patient survived. | [5 empty] | Since the original report was filed, the investigation into the thrombosis has concluded. the medical center did not return for benchtop analysis any impella pump product. the root cause of the thrombosis was most likely patient condition related. | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35503 | 1220648-2023-02239 | 17015143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17015143&pc=OZD | 05/03/2023 | 05/26/2023 |  | 05/26/2023 |  |  | 05/03/2023 | 05/26/2023 |  |  | 01/31/2025 |  | 05/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316079 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Retroperitoneal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/03/2023 |  |  |  |  | A 75 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported a suspected retroperitoneal bleed (rp). the bleed prompted the team to infuse blood products. the pump support was for the hemodynamic support of the ami/cgs patient undergoing left main intervention. the patient expired. | [5 empty] | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35504 | 1220648-2023-02241 | 17015144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17015144&pc=OZD | 04/27/2023 | 08/09/2023 |  | 05/26/2023 |  |  | 04/27/2023 | 05/26/2023 |  |  | 11/30/2024 |  | 05/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295015 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2023 |  |  |  |  | A 64 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the impella cp was explanted due to patient's ischemic limb. the physician noted the limb improved after explant. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was most likely patient condition related due to the listed health condition and medications in the clinical details. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35505 | 1220648-2023-02242 | 17015148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17015148&pc=OZD | 04/26/2023 | 07/13/2023 |  | 05/26/2023 |  |  | 04/27/2023 | 05/26/2023 |  |  | 11/30/2024 |  | 05/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024314740 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2023 |  |  |  |  | A 47 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient developed an access site bleed during support. the site was redressed and hemostasis was achieved by tightening the perclose device, however, 2 units of blood products were required. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the access site bleeding. root cause of the issue was most likely patient condition related as venous access next to cp access also oozing. pulled perclose suture a little and redressed, noted less bleeding. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿
35506 | 1220648-2023-02221 | 17001247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17001247&pc=OZD | 05/10/2023 | 08/09/2023 |  | 05/25/2023 |  |  | 05/10/2023 | 05/25/2023 |  |  | 12/31/2024 |  | 05/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318791 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2023 |  |  |  |  | A 66 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported that the patient developed hemolysis coinciding with impella support. the patient was administered two units of packed red blood cells (prbc) to treat the hemolysis and was considered stable. patient support continued with the implanted pump following the intervention. | [5 empty] | The investigation into the hemolysis has been completed since the original report was filed. no benchtop analysis was possible to determine the root cause. since the medical center did not return the impella device, the root cause of the hemolysis was not determined. the failure mode will be monitored and trended. | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35507 | 1220648-2023-02222 | 17001310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17001310&pc=OZD | 04/26/2023 | 10/23/2023 |  | 05/25/2023 |  |  | 04/26/2023 | 05/25/2023 |  |  | 12/31/2024 |  | 05/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024315831 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thromboembolism | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2023 |  |  |  |  | A 28 year old male with cardiomyopathy presented for impella support. during support, the user facility reported an emergent thrombectomy required, secondary to clot extending from the descending aorta to the iliac arteries. the patient had known heparin-induced thrombocytopenia and coagulopathy issues. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the high purge pressure was not determined since product was not returned; however, the root cause of the thrombus was most likely patient condition related. the failure mode will be monitored and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: purge malfunction: "do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause crack and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." "clean the connector cable with 70% isopropyl alcohol."
35508 | 1220648-2023-02223 | 17001330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17001330&pc=OZD | 04/30/2023 | 10/31/2023 |  | 05/25/2023 |  |  | 04/30/2023 | 05/25/2023 |  |  | 12/31/2024 |  | 05/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024304214 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2023 |  |  |  |  | A 50 year old male with cardiogenic shock presented for impella 5.5 support. the user facility reported the patient experienced hemolysis. troubleshooting was performed such as repositioning, confirming volume status, confirming therapeutic range on act, but the device continued to hemolyze. the decision was made to replace the impella 5.5 with another impella 5.5. the new impella resolved the issued. the patient's hemolysis cleared up and his urine was yellow. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the hemolysis was unable to be determined as no product was returned and limited clinical details were provided. | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35509 | 1220648-2023-02225 | 17001337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17001337&pc=OZD | 04/27/2023 | 06/26/2023 |  | 05/25/2023 |  |  | 04/27/2023 | 05/25/2023 |  |  | 12/31/2024 |  | 05/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024304988 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2023 |  |  |  |  | The user facility reported an impella 5.5 pump exhibited elevated motor current and a stoppage while a patient was ambulating. the pump had been running for approximately 29 days and was removed as a result of the stop. there was no patient impact related to the pump stoppage. | [5 empty] | The investigation into the pump stop has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis of the pump stop; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of pump stop was an open electrical line as a result of patient movement. the failure mode will be monitored and trended. | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35510 | 1220648-2023-02224 | 17001339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17001339&pc=OZD | 05/03/2023 | 06/30/2023 |  | 05/25/2023 |  |  | 05/03/2023 | 05/25/2023 |  |  | 12/31/2024 |  | 05/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024293591 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2023 |  |  |  |  | A 71 year old male with severe aortic insufficiency presented for impella support. the user facility reported the pump stopped shortly after insertion. the pump was explanted and replaced by an intra-aortic balloon pump (iabp). there was no harm or injury to the patient. | [5 empty] | The investigation into the reported pump stop has been completed since the original report was filed. the medical center returned the impella products for benchtop analysis. an electrical resistance test revealed three open lines. a minor kink was observed on the catheter by the motor housing. the red handle was opened and the location of the three open lines was confirmed on the catheter side within the red handle. necking of the motor cables was observed. no other abnormalities were observed. the cause of the pump stop was three open lines. the evaluation revealed that three open lines (technique/use) was the most likely cause of the pump stop. the failure modes will be monitored and trended. | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35511 | 1220648-2023-02226 | 17001351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17001351&pc=OZD | 04/19/2023 | 06/23/2023 |  | 05/25/2023 |  |  | 04/19/2023 | 05/25/2023 |  |  | 11/30/2024 |  | 05/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303575 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2023 |  |  |  |  | A 59 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient lost distal flow. vascular surgery cut down to remove and possibly perform thrombectomy. no clot was found and patient regained distal flow once impella was explanted and right femoral artery was repaired. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center discarded the impella product at the time of explant. no benchtop analysis of the root cause of the limb ischemia was possible; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of limb ischemia was most likely related to patient condition. the failure mode will be monitored and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35512 | 1220648-2023-02227 | 17001686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17001686&pc=OZD | 07/26/2022 | 05/25/2023 |  | 05/25/2023 |  |  | 04/28/2023 | 05/25/2023 |  |  | 02/29/2024 |  | 05/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023189473 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/28/2023 |  |  |  |  | A 65 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support in germany. the impella was implanted on (b)(6) 2022 and was explanted on (b)(6) 2022. the clinical study group reported the patient experienced a pseudoaneurysm of the iliac external and femoral nerve dysfunction on (b)(6) 2022 and noted the event may possibly be related to the impella device. a hematoma evacuation was performed on the right inguinal side and the surgeon performed reconstruction of the puncture site with biviner patch plastic. the patient was stable after surgical reconstruction. | [5 empty] | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35513 | 1220648-2023-02228 | 17007364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17007364&pc=OZD | 05/02/2023 | 10/26/2023 |  | 05/25/2023 |  |  | 05/02/2023 | 05/25/2023 |  |  | 12/31/2023 |  | 05/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176230 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2023 |  |  |  |  | A 65 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient showed signs of hemolysis which was treated with the infusion of one unit of packed red blood cells (pbrc)s and the initiation of continuous renal replacement therapy (crrt) dialysis. the impella pump remained on despite the hemolysis until the patient was weaned. the device was explanted on day six. the patient survived without lasting harm or injury. | [5 empty] | The investigation into the hemolysis, the access site bleeding and the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based off the clinical info available, the cause of the hemolysis was not determined with an unknown relation to impella since insufficient clinical details were provided and it was unknown how the hemolysis was diagnosed or resolved. data logs review showed suction alarms occurred, no motor current devise and no other abnormality showed to indicate the root cause. the cause of the hematoma was use related issue. the cause of the limb ischemia most likely was patient condition related. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35514 | 1220648-2023-02216 | 16991348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16991348&pc=OZD | 04/25/2023 | 09/20/2024 |  | 05/24/2023 |  |  | 04/25/2023 | 05/24/2023 |  |  | 01/31/2025 |  | 05/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024314856 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Death | Yes | Yes | Decreased Pump Speed | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | A 61-year-old male presented with sepsis, pericardial effusion, myocardial infarction and in cardiogenic shock. the plan was to perform a pericardial drainage, place an impella cp with smartassist cardiac pump for hemodynamic support and perform a percutaneous coronary intervention. the impella cp has been successfully placed during cardiopulmonary resuscitation and has been running for 20 minutes when the flow rate of the pump dropped to 0 l/min with the pump motor still running. the pump has been removed and the decision has been made by the attending physician not to insert a new impella device due to the patient¿s condition. according to the cardiologist¿s assessment, the impella cp malfunction was not the cause of the patient¿s death but a coincidence. currently, cause of death is unclear. an autopsy will be performed. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The user facility has returned the impella device, and the benchtop analysis revealed that the root cause of the low pump flow was biomaterial. | Investigation of returned pump and logfiles is ongoing. if possible, additional clinical information will be obtained for investigation. a supplemental report will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6.
35515 | 1220648-2023-02217 | 16992643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16992643&pc=OZD | 04/18/2023 | 06/26/2023 |  | 05/24/2023 |  |  | 04/18/2023 | 05/24/2023 |  |  | 11/30/2024 |  | 05/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024314595 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2023 |  |  |  |  | A 46 year old male with post cardiotomy cardiogenic shock presented for impella support. the user facility reported lack of distal pulses in impella inserted extremity, along with muscle spasms and plantar extension posturing. to avert potentially irreversible damage, the pump was removed and a fasciotomy was performed below the knee to treat compartment syndrome. | [5 empty] | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center discarded the impella product at the time of explant. no benchtop analysis of the root cause of the limb ischemia was possible; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of limb ischemia was most likely related to patient condition. the failure mode will be monitored and trended.
35516 | 1220648-2023-02218 | 16992846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16992846&pc=OZD | 04/26/2023 | 07/27/2023 |  | 05/24/2023 |  |  | 04/26/2023 | 05/24/2023 |  |  | 11/30/2024 |  | 05/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024299520 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2023 |  |  |  |  | A 54 year old male noted as high risk percutaneous cardiac insertion (hrpci) presented for impella support. the user facility reported a pump prolapsing during a failed delivery attempt. a snare was required to remove the fragmented device from the patient. there was no injury or harm to the patient. | [5 empty] | The investigation into the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the hemolysis issue was most likely due to improper positioning of the device based on the data log review. the failure mode will be monitored and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
35517 | 1220648-2023-02219 | 16995777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16995777&pc=OZD | 03/03/2023 | 10/25/2024 |  | 05/24/2023 |  |  | 03/03/2023 | 05/24/2023 |  |  | 09/30/2024 |  | 05/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273675 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2023 |  |  |  |  | A 47 year old male in cardiogenic shock presented for impella support. the user facility reported the development of ischemia in the left leg during ecpella support (combined used of va-ecmo and impella) the pump was explanted from the left femoral artery as a result of the ischemia and an impella 5.5 was implanted via axillary insertion. however, following implant, flow to the limb did not improve and the decision was made to amputate the limb. | It was reported that the procedural outcome is patient died. there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | B2: date of patient death is unknown. this report is being filed as part of a retrospective review of historical records.
35518 | 1220648-2023-02220 | 16996530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16996530&pc=OZD | 04/29/2023 | 09/18/2024 |  | 05/24/2023 |  |  | 04/29/2023 | 05/24/2023 |  |  | 02/28/2025 |  | 05/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321310 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/29/2023 |  |  |  |  | A 71 year old female in cardiogenic shock presented for impella support. the user facility reported the patient developed limb ischemia with no pulses in the left leg, distal to the impella femoral access site. the physician discussed the possibility of fem-fem bypass. the family decided to withdraw care. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was most likely patient condition related as the patient had related medication. the failure mode will be monitored and trended. | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
35519 | 1220648-2023-02213 | 16989886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16989886&pc=OZD | 03/14/2023 | 11/01/2023 |  | 05/23/2023 |  |  | 03/14/2023 | 05/23/2023 |  |  | 12/31/2024 |  | 05/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024301595 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2023 |  |  |  |  | A 79 year old male presented for coronary artery bypass grafting and an impella device was used. the user facility reported that the patient suffered a vascular dissection during impella insertion. the impella 5.5 pump was removed and an aortic repair was completed with a 26mm graft and a coronary artery bypass graft. following intervention, an impella cp pump was implanted for further support. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the vascular damage was unable to be determined as no product was returned and limited clinical details were provided. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
35520 | 1220648-2023-02214 | 16989887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16989887&pc=OZD | 03/19/2023 | 01/18/2024 |  | 05/23/2023 |  |  | 03/19/2023 | 05/23/2023 |  |  | 06/30/2024 |  | 05/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235961 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2023 |  |  |  |  | A 54 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the development of limb ischemia coinciding with impella support. an external bypass was performed to return flow to the lower extremity. however, residual bleeding was noted from the femoral site and bypass site and the decision was made to explant the pump. the patient was considered stable following explant. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding: "assess access site for bleeding and hematoma.¿ limb ischemia: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the access site bleeding was not determined since product was not returned; however, the root cause of the limb ischemia was most likely patient condition related.
35521 | 1220648-2023-02215 | 16989888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16989888&pc=OZD | 04/16/2023 | 10/14/2024 |  | 05/23/2023 |  |  | 04/16/2023 | 05/23/2023 |  |  | 10/31/2024 |  | 05/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024289572 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/16/2023 |  |  |  |  | A 59 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the patient had an impella cp support ongoing when the impella-accessed limb was seen to be "compromised" with distal flow decreased/pulselessness down the leg. the cp was explanted to due limb ischemia concerns and the patient was escalated to the 5.5 pump. the 5.5 pump was via the surgical axillary access. after the cp was explanted the site was surgically repaired and doppler pulses returned and the limb felt less cool. | The patient received total of 1 unit packed red blood cells, 2 units of platelets, 500mls of albumin and 20mg of protamine during the surgical repair of the impella cp site after explant. |  |  |  |  | The investigation into the limb ischemia and the access site bleeding has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis; only the clinical report was evaluated. based on the clinical details, the cause of the limb ischemia was most likely patient condition related since the patient required increases in inotropic support and the cause of the access site bleeding was most likely use related since the bleeding resolved with manual pressure. the failure modes will be monitored and trended. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information for the initial mfr 1220648-2023-02215 was provided in sections b5, d6a, d6b, h1, and h6. the investigations for the reported limb ischemia and access site bleeding have been completed. product was not returned for review. based on clinical description, the cause of the limb ischemia was most likely patient condition related since the patient required increases in inotropic support & the cause of the access site bleeding was most likely use related since the bleeding resolved with manual pressure. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
35522 | 1220648-2023-02205 | 16975347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16975347&pc=OZD | 03/06/2023 | 06/29/2023 |  | 05/22/2023 |  |  | 03/06/2023 | 05/22/2023 |  |  | 11/30/2024 |  | 05/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 55 |  |  | ABIOMED, INC. |  |  | US |  | 2023278007 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2023 |  |  |  |  | A 26 year old male with cardiomyopathy presented for impella support. the user facility reported that the patient developed an access site bleed and hematoma coinciding with impella support. a washout was performed at the right axillary site to treat the noted condition. impella support continued following the intervention. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis. the clinical report was evaluated and it was determined there were insufficient details. the root cause for access site bleeding is undetermined since neither the product nor sufficient clinical information were provided. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35523 | 1220648-2023-02206 | 16975402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16975402&pc=OZD | 02/20/2023 | 08/17/2023 |  | 05/22/2023 |  |  | 02/20/2023 | 05/22/2023 |  |  | 03/31/2024 |  | 05/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212148 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2023 |  |  |  |  | A 37 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported that an iliac dissection occurred during pump removal following impella support. it was further specified that the pigtail pulled part of the iliac during explant. a peripheral stent was subsequently placed to treat the condition. the patient was considered to be stable following the intervention. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information provided, the cause of the vascular damage was not determined since product was not returned.
35524 | 1220648-2023-02207 | 16975497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16975497&pc=OZD | 03/27/2023 | 07/04/2023 |  | 05/22/2023 |  |  | 03/27/2023 | 05/22/2023 |  |  | 11/30/2024 |  | 05/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309037 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2023 |  |  |  |  | A 59 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported that the patient developed ischemia in the impella leg following ecmo cannula placement. an external bypass perfusion catheter was placed to return flow to the limb. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35525 | 1220648-2023-02208 | 16979194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16979194&pc=OZD | 12/28/2022 | 08/15/2023 |  | 05/22/2023 |  |  | 12/28/2022 | 05/22/2023 |  |  | 10/31/2024 |  | 05/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023269621 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/28/2022 |  |  |  |  | A 67 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported bleeding was noted from the anastomoses, but not attributed to the impella device. the patient was provided a number of blood products and was stable. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the access site bleeding issue was not determined since the product was not returned. the failure mode will be monitored and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
35526 | 1220648-2023-02209 | 16979404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16979404&pc=OZD | 02/18/2023 | 06/30/2023 |  | 05/22/2023 |  |  | 02/18/2023 | 05/22/2023 |  |  | 10/31/2024 |  | 05/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285478 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/18/2023 |  |  |  |  | A 75 year old male noted as high risk percutaneous cardiac insertion (hrpci) presented for impella support. the user facility reported the patient developed hemolysis during impella support. it was noted that the patient had a difficult anatomy that made optimal positioning difficult. as a result of the condition, the decision was made to explant the impella pump. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the reported hemolysis has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis. the data logs were evaluated. the evaluation revealed the most likely cause of the hemolysis was improper positioning due to the patient's small anatomy, i.e. technique/use. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35527 | 1220648-2023-02210 | 16979435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16979435&pc=OZD | 02/19/2023 | 05/31/2023 |  | 05/22/2023 |  |  | 02/19/2023 | 05/22/2023 |  |  | 06/30/2024 |  | 05/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236797 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/19/2023 |  |  |  |  | A 75 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient developed leg ischemia. it was documented that a contra-lateral distal perfusion catheter was placed in an attempt to treat the condition, however, the decision was ultimately made to explant the pump. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure modes will be monitored and trended. instructions for use for the related event are as follows: ¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35528 | 1220648-2023-02211 | 16979443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16979443&pc=OZD | 03/25/2023 | 07/14/2023 |  | 05/22/2023 |  |  | 03/25/2023 | 05/22/2023 |  |  | 11/30/2024 |  | 05/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024297966 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2023 |  |  |  |  | A 58 year old male noted as high risk percutaneous cardiac insertion presented for impella support. the user facility reported that the patient developed an access site bleed during impella support. the patient was provided two units of packed red blood cells (prbc) and pressure was held on the site to treat the bleed. impella support continued following the intervention. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the access site bleeding was most likely patient condition related, specifically the patient's history of bleeding issues. the failure mode will be monitored and trended.
35529 | 1220648-2023-02212 | 16979533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16979533&pc=OZD | 04/24/2023 | 01/23/2024 |  | 05/22/2023 |  |  | 04/24/2023 | 05/22/2023 |  |  | 02/28/2025 |  | 05/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024320996 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Injury | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 74-year-old male with known cad has been treated for ami / cgs. the patient received support from an impella 5.5 cardiac pump and a va-ecmo. anticoagulation has been low due to bleeding problems in the past. a sudden stop of the impella 5.5 occurred, with the pump unable to restart. the patient has been stable on va-ecmo and has been transferred to cardiovascular icu, receiving respiratory support from a ventilator. | [5 empty] | The impella device has been returned to the manufacturer. when the investigation is complete, a supplemental report will be filed. as noted in the impella ifu: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop has been completed since the original report was filed. the user facility has returned the impella device, and the benchtop analysis revealed that the root cause of the pump stop was an open electrical motor cable lines due to patient movement excessive force on the catheter.
35530 | 1220648-2023-02191 | 16960567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16960567&pc=OZD | 02/17/2023 | 05/31/2023 |  | 05/19/2023 |  |  | 02/17/2023 | 05/19/2023 |  |  | 10/31/2024 |  | 05/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024287285 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2023 |  |  |  |  | A 53 year old male with cardiomyopathy presented for impella support. the user facility reported the patient experienced acute limb ischemia following the placement of the impella. the distal pulses were not found. the physician performed external fem-fem bypass for distal perfusion. | [5 empty] | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure modes will be monitored and trended. instructions for use for the related event are as follows: ¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35531 | 1220648-2023-02192 | 16963073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16963073&pc=OZD | 05/04/2023 | 07/06/2023 |  | 05/19/2023 |  |  | 05/04/2023 | 05/19/2023 |  |  | 03/31/2024 |  | 05/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206035 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2023 |  |  |  |  | A 92 year old male noted for high risk percutaneous cardiac insertion (hrpci) presented for impella support. the user facility reported the patient developed hemolysis coinciding with the impella support. a nitroglycerin drip was administered until the mean arterial pressure lowered and the observed hematuria began to clear. patient support continued following the intervention. | [5 empty] | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis; only the clinical report was evaluated. based on the clinical details, the cause of the hemolysis was most likely related to the patient condition. the failure mode will be monitored and trended. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35532 | 1220648-2023-02185 | 16963215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16963215&pc=OZD | 04/21/2023 | 07/26/2023 |  | 05/19/2023 |  |  | 04/21/2023 | 05/19/2023 |  |  | 07/31/2024 |  | 05/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED INC. |  |  | US |  | 2023253744 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/21/2023 |  |  |  |  | The us complainant had placed an impella cp for support of a high risk coronary angioplasty (hrpci) patient who had been denied surgery due to underlying condition. the team placed the cp and experienced limb ischemia. the patient had the hpci completed, and immediately afterwards due to the ischemia they prematurely explanted the cp and placed a balloon pump (iabp). even with the smaller diameter iabp the limb was ischemic. | [5 empty] | The medical center discarded the impella cp and introducer sheath at the time of explant. there can be no benchtop analysis to root cause. only the clinical report can be reviewed. the fact that even with the smaller 8fr iabp there was observed ischemia, does bring patient condition into review as a possible cause of the ischemia. should more information be shared, and root cause be determined, a final report will be forthcoming. the failure mode will be monitored and trended. of not the ifu states: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the hemolysis issue was patient condition related (the small vessel size was due to vasopressors (high dose of epi and levo) given to treat the perforated in the lad).
35533 | 1220648-2023-02194 | 16963941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16963941&pc=OZD | 04/20/2023 | 10/25/2023 |  | 05/19/2023 |  |  | 04/20/2023 | 05/19/2023 |  |  | 12/31/2024 |  | 05/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024294355 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Pressure Problem | [5 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/20/2023 |  |  |  |  | A 22 year old male with cardiomyopathy presented for impella support. the user facility reported high purge pressure during use of the device. the pump was removed and an embolectomy was performed to remove residual clot. a new pump (serial number 435334) was opened and delivered. the patient was showing no signs of embolic complications. | [5 empty] | The device has been received by the manufacturer for investigation. at the date of this report, the investigation has not been completed. a supplemental mdr will be filed upon completion of the investigation. | The investigation for the reported purge pressure has been completed. the data log was reviewed and confirmed the failure mode. about 631 hours into the case, the purge pressure increased gradually from 500 mmhg to 700 mmhg, then 1100 mmhg & remained to the rest of the case. the root cause of high purge pressure was due to biomaterial built up & blocked the purge gap. the instructions for use provides troubleshooting for high pressure issues. it states unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported thrombus has been completed since the original report was filed. the medical center returned the impella products for benchtop analysis. visual inspection found blood clot blocked whole outflow cages& purge gap. blood clot was also found in inflow cage, along the cannula & around the impeller. no other issues were found on the rest of the pump. the data logs were reviewed: data confirmed the failure mode of high purge pressure. about 631 hours into the case, pp increased gradually from 500 mmhg to 700 mmhg, then 1100 mmhg & remained to the rest of the case. a clinical review was performed: the root cause of thrombus was most likely due to patient condition the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35534 | 1220648-2023-02195 | 16965126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16965126&pc=OZD | 04/21/2023 | 01/18/2024 |  | 05/19/2023 |  |  | 04/21/2023 | 05/19/2023 |  |  | 08/31/2024 |  | 05/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, CANADA |  |  | ABIOMED, INC. |  |  | US |  | 2023268391 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/21/2023 |  |  |  |  | The complainant reported a 65 year old female patient presenting for high risk percutaneous coronary intervention for impella support. after insertion and initiation of the device, bleeding was noted at the insertion site. the patient was taken to the operating room in order for the physician to clean the site and tighten sutures, however, the bleed persisted. the pump was removed and blood products were required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿ | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical information provided, the root cause of the access site bleeding was patient condition related since best practices were followed and bleeding continued.
35535 | 1220648-2023-02197 | 16965827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16965827&pc=OZD | 02/10/2023 | 07/11/2023 |  | 05/19/2023 |  |  | 02/10/2023 | 05/19/2023 |  |  | 07/31/2024 |  | 05/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023281642 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/10/2023 |  |  |  |  | A 69 year old male with right heart failure presented for impella rp support. the user facility reported the patient experienced hemolysis due to the stiffness of the guidewire used for implant. the patient underwent a blood replacement of unknown quantity to treat the condition. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the patient injury has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on the clinical details, the root cause of the patient injury, specifically hemoptysis was determined to be use related as the wire was not properly shaped per ifu and insufficient traction was placed on the wire during delivery which allowed the advancement into the lung. the failure mode will be monitored and trended.
35536 | 1220648-2023-02199 | 16965852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16965852&pc=OZD | 04/22/2023 | 05/19/2023 |  | 05/19/2023 |  |  | 05/01/2023 | 05/19/2023 |  |  | 11/30/2024 |  | 05/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024291150 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/01/2023 |  |  |  |  | A 63 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient developed an access site bleed during support. as treatment, the pump was removed and a blood transfusion of unknown quantity was performed. | [5 empty] | The customer reported the impella device was to be returned, however, it has not been returned at the date of this report. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
35537 | 1220648-2023-02198 | 16965871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16965871&pc=OZD | 02/16/2023 | 06/07/2023 |  | 05/19/2023 |  |  | 02/16/2023 | 05/19/2023 |  |  | 03/31/2024 |  | 05/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212193 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2023 |  |  |  |  | A 63 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility report the patient was experiencing runs of ventricular fibrillation (vt) so the physician made the decision to remove the impella. two percloses were deployed. there was no bleeding or hematoma. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | The investigation into the ventricular tachycardia (vt) has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the vt. the failure mode will be monitored and trended.
35538 | 1220648-2023-02202 | 16965905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16965905&pc=OZD | 02/03/2023 | 06/02/2023 |  | 05/19/2023 |  |  | 02/03/2023 | 05/19/2023 |  |  | 07/31/2024 |  | 05/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023248398 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/03/2023 |  |  |  |  | A 53 year old male presented for impella support. the user facility reported that the cutdown technique for impella insertion caused a hematoma. although no allegation was made against the impella, a contribution cannot be completely ruled out. drainage was placed and blood products were administered. | [5 empty] | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the technique used to perform the cutdown was the most likely cause of the access site bleeding. the failure modes will be monitored and trended. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿
35539 | 1220648-2023-02201 | 16965936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16965936&pc=OZD | 04/20/2023 | 08/09/2023 |  | 05/19/2023 |  |  | 04/20/2023 | 05/19/2023 |  |  | 08/31/2024 |  | 05/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2023263383 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2023 |  |  |  |  | The complainant reported an 85-year-old female patient requiring high risk percutaneous coronary intervention was presented for impella placement. during support, it was documented that the patient developed a decrease in flow to their impella leg. the patient noted having pain in the left groin and leg which coincided with a lack of flow to the limb. as a result of the noted condition, the decision was made to explant the pump. following removal, flow returned to the leg and the patient regained movement in their toes. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the limb ischemia was most likely patient condition related.
35540 | 1220648-2023-02200 | 16965945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16965945&pc=OZD | 05/03/2023 | 05/22/2023 |  | 05/19/2023 |  |  | 05/03/2023 | 05/19/2023 |  |  | 06/30/2024 |  | 05/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247256 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2023 |  |  |  |  | A 61 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient suffered vascular damage during insertion and subsequently experienced a retroperitoneal bleed. the patient went into cardiac arrest and support was withdrawn. the patient passed away as a result of the bleed. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | Correction to h6 evaluation codes: correct investigation conclusions: from "19 cause traced to user". to: "4315 cause not established". correct h10 additional manufacturer narrative: from: "the impella device was discarded by the customer". to: "the impella device was not returned by the customer".
35541 | 1220648-2023-02203 | 16965975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16965975&pc=OZD | 05/08/2023 | 08/09/2023 |  | 05/19/2023 |  |  | 05/08/2023 | 05/19/2023 |  |  | 11/30/2024 |  | 05/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2024286223 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2023 |  |  |  |  | The complainant reported a 61-year-old male patient with acute myocardial infarction and cardiogenic shock presenting for impella placement. during support, it was documented that the patient developed hemolysis. there was no noted suction associated with the pump and troubleshooting steps failed to resolve the issue. the pump was not explanted in response, however, the patient was administered haptoglobin to treat the hemolysis. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, hemolysis is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the hemolysis was most likely patient condition related.
35542 | 1220648-2023-02204 | 16966027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16966027&pc=OZD | 02/01/2023 | 05/26/2023 |  | 05/19/2023 |  |  | 02/01/2023 | 05/19/2023 |  |  | 10/31/2024 |  | 05/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290458 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2023 |  |  |  |  | A 76 year old female of unknown race presented for impella support after repeated events of coding. the user facility reported that the patient developed an access site bleed during impella support. the patient was provided multiple units of blood and an additional stitch was placed at the groin site to treat the bleed. patient support continued following the event. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that use issue was the most likely cause of the access site bleeding. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support:¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35543 | 1220648-2023-02193 | 16968627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16968627&pc=OZD | 04/20/2023 | 10/25/2023 |  | 05/19/2023 |  |  | 04/20/2023 | 05/19/2023 |  |  | 02/28/2025 |  | 05/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024308558 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Pressure Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/20/2023 |  |  |  |  | A 22 year old male with cardiomyopathy presented for impella support. the user facility reported high purge pressure during use of the device. the pump was removed and an embolectomy was performed to remove residual clot. a new pump was opened and delivered. the patient showed no sign of embolic complications. | [5 empty] | The device has been returned for evaluation but the evaluation is incomplete on date of this report. a supplemental mdr will be filed upon completion of investigation. | The investigation into the reported high purge pressure and thrombus has been completed since the original report was filed. the medical center returned the impella products for benchtop analysis. biomaterial was found in the inlet, outlet and purge gap. the pump was purged and the high purge pressure reproduced. the pump would not run. the catheter was cut by the motor housing and purge fluid exited quickly at the location of the cut. the location of the blockage was isolated to the remaining portion of the pump. the cause of the high purge pressure was biomaterial. the lt log shows multiple suction alarms throughout the case and a rise in purge pressure once the pump is started. the cause of the thrombus was most likely patient condition related based on the provided clinical information. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | H6-medical device problem code, component code, type of investigation, investigation findings, investigation conclusion.
35544 | 1220648-2023-02196 | 16968965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16968965&pc=OZD | 02/03/2023 | 06/08/2023 |  | 05/19/2023 |  |  | 02/03/2023 | 05/19/2023 |  |  | 10/24/2024 |  | 05/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024287289 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/03/2023 |  |  |  |  | A 67 year old female noted as high risk percutaneous cardiac insertion (hrpci) presented for an impella implant. the user facility reported the patient experience an access site bleed coinciding with the explant. it was noted that the patient had bleeding following attempts to secure the perclose around the site. an up-and-over balloon tamponade was utilized, however, the resulting hemostasis was only temporary. as a result of the persistent bleed, a covered stent was placed to treat the condition. a definitive cause for the bleed was unknown. distal flow was verified following the intervention. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the root cause for access site bleeding remains undetermined since neither the product nor sufficient clinical information were provided. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35545 | 1220648-2023-02184 | 16952065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16952065&pc=OZD | 05/04/2023 | 09/22/2024 |  | 05/18/2023 |  |  | 05/04/2023 | 05/18/2023 |  |  | 11/30/2024 |  | 05/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023276351 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2023 |  |  |  |  | A 63 year old male with acute myocardial infarction and cardiogenic shock presented for impella support. the user facility reported that the pump stopped unexpectedly during patient support. the pump was restarted multiple times, at different support levels, however the device stopped working each time. as a result of the ongoing failure, the pump was subsequently explanted. given patient terrible prognosis, both devices were turned off and patient expired. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02184 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6.
35546 | 1220648-2023-02186 | 16955243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16955243&pc=OZD | 02/09/2023 | 06/06/2023 |  | 05/18/2023 |  |  | 02/09/2023 | 05/18/2023 |  |  | 08/31/2024 |  | 05/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270992 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2023 |  |  |  |  | A 66 year old male of unknown race with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient experienced a left ventricle perforation coinciding with impella support. while examining the patient for pericardial effusion, it was discovered that the pigtail of the impella was through the left ventricle. as a result of the noted issue, the patient was scheduled to go back to the operating room for pump explant and ventricle repair. the physician noted unsure if there was a free wall rupture during cpr of during mi. the patient was considered stable at the time of the report. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported ventricle perforation has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the root cause for ventricle perforation remains undetermined since neither the product nor sufficient clinical information were provided. it is unknown if the perforation occurred during cardiopulmonary resuscitation or myocardial infarction. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35547 | 1220648-2023-02187 | 16955726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16955726&pc=OZD | 02/09/2023 | 09/20/2024 |  | 05/18/2023 |  |  | 02/09/2023 | 05/18/2023 |  |  | 09/30/2024 |  | 05/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273565 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/09/2023 |  |  |  |  | A 77 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient developed ischemia in the impella leg during support. the patient was scheduled for placement of an external fem-fem bypass, however the patient's family decided to withdraw care as a result of the patient's declining condition. | Additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was patient condition related. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02187 was provided in sections b2, b5, d6a, d6b, h1, and h6.
35548 | 1220648-2023-02188 | 16956521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16956521&pc=OZD | 02/07/2023 | 05/26/2023 |  | 05/18/2023 |  |  | 02/07/2023 | 05/18/2023 |  |  | 10/31/2024 |  | 05/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024288206 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2023 |  |  |  |  | A 65 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the development of hemolysis coinciding with the support. as the physician was unable to adequately reposition the pump, the decision was made to explant the device | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the reported hemolysis has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that improper positioning was the most likely cause of the hemolysis. the failure modes will be monitored and trended.
35549 | 1220648-2023-02189 | 16956523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16956523&pc=OZD | 04/18/2023 | 08/09/2023 |  | 05/18/2023 |  |  | 04/18/2023 | 05/18/2023 |  |  | 12/31/2024 |  | 05/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2024320262 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2023 |  |  |  |  | The complainant reported a 71-year-old male patient with acute myocardial infarction and cardiogenic shock presenting for impella placement. during support, it was documented that the patient experienced a recurring hematoma around the access site. the hematoma was determined to be related to an access site bleed. the pump was subsequently explanted and a covered stent was placed to resolve the bleed. the patient was considered stable following the event. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, bleeding is a known potential adverse event associated with impella support: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding has been completed since the original report was filed. no benchtop analysis was possible to determine the root cause. since the medical center did not return the impella device, the root cause of the access site bleeding was not determined. the failure mode will be monitored and trended.
35550 | 1220648-2023-02190 | 16956531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16956531&pc=OZD | 02/10/2023 | 06/07/2023 |  | 05/18/2023 |  |  | 02/10/2023 | 05/18/2023 |  |  | 12/31/2024 |  | 05/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024287857 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/10/2023 |  |  |  |  | A 61 year old male with low cardiac output syndrome requiring coronary artery bypass grafting presented for impella support. during the insertion of the impella 5.5 direct via aortic graft, the tip of the impella catheter perforated the left ventricle (lv) during insertion and heart manipulation. immediate surgical repair was initiated and successfully completed without further complications. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation into the ventricle perforation has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that use issue was the most likely cause of the ventricle perforation. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
35551 | 1220648-2023-02176 | 16945054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16945054&pc=OZD | 04/21/2023 | 07/11/2023 |  | 05/17/2023 |  |  | 04/21/2023 | 05/17/2023 |  |  | 01/31/2025 |  | 05/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024309116 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 62-year-old male in the us with non ischemic cardiomyopathy required increased pressor support. decision was made for mechanical circulatory support. as a bridge to mcs, first an impella cp has been inserted and later escalated to an impella 5.5. the patient developed pain in the right hand, with doppler pulses undetectable. impella 5.5 removed without incident, pulses reestablished in hand. no deficits noted. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on the clinical details, the cause of the limb ischemia was most likely patient condition related. the failure mode will be monitored and trended. | The impella device has not been received by the manufacturer for investigation. should any new information be returned, a supplemental report will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35552 | 1220648-2023-02177 | 16946634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16946634&pc=OZD | 05/01/2023 | 05/17/2023 |  | 05/17/2023 |  |  | 05/01/2023 | 05/17/2023 |  |  | 11/30/2024 |  | 05/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024308533 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/01/2023 |  |  |  |  | A 70 year old male in cardiogenic shock with a pre-existing left ventricle apical thrombus presented for impella support. the patient became mobile during support and then the thrombus was found at the cage of the pump. the pump was removed and a thrombectomy of the right femoral artery was subsequently performed to remove the thrombus. the patient was considered stable following the intervention. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35553 | 1220648-2023-02178 | 16949954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16949954&pc=OZD | 04/16/2023 | 06/23/2023 |  | 05/17/2023 |  |  | 04/17/2023 | 05/17/2023 |  |  | 11/30/2024 |  | 05/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024297303 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2023 |  |  |  |  | The complainant reported a female patient of unknown age presenting with acute myocardial infarction and cardiogenic shock. during support, the patient developed ischemia in the impella inserted leg. the decision was made to remove the device and replace with a new impella via axillary access. after removal, blood flow was restored. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center discarded the impella product at the time of explant. no benchtop analysis of the root cause of the limb ischemia was possible; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of limb ischemia was most likely related to patient condition. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35554 | 1220648-2023-02179 | 16950267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16950267&pc=OZD | 05/02/2023 | 07/10/2023 |  | 05/17/2023 |  |  | 05/02/2023 | 05/17/2023 |  |  | 12/31/2024 |  | 05/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320846 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2023 |  |  |  |  | A 35 year old female in cardiogenic shock presented for impella support. the user facility reported that the impella pump stopped functioning during impella support. the pump failed to restart and the decision was made to explant the pump. the patient was noted to be prothrombotic at the time of the failure and vascular surgery was completed to fix the occluded femoral artery. | [5 empty] | The investigation into the pump stop has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis; only the clinical report was evaluated. based on the clinical details, the cause of the pump stop was bearing wear. pump used beyond design life. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.
35555 | 1220648-2023-02180 | 16950283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16950283&pc=OZD | 04/20/2023 | 07/18/2023 |  | 05/17/2023 |  |  | 04/20/2023 | 05/17/2023 |  |  | 12/31/2024 |  | 05/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320363 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2023 |  |  |  |  | A 59 year old male had an impella cp placed via the femoral artery for pre-operation for a coronary artery bypass grafting (cabg). the patient's limb became ischemic and the impella cp was explanted. the impella was replaced with an intra-aortic balloon pump (iabp). the team additionally performed thrombectomy to the limb. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the limb ischemia issue was most likely due to patient condition related based on the clinical information provided. the failure mode will be monitored and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35556 | 1220648-2023-02182 | 16950285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16950285&pc=OZD | 04/20/2023 | 07/11/2023 |  | 05/17/2023 |  |  | 04/20/2023 | 05/17/2023 |  |  | 01/31/2025 |  | 05/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024297864 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2023 |  |  |  |  | A 59 year old female with cardiomyopathy presented for impella support. the user facility reported the inability to deliver the pump into the patient's left ventricle despite various techniques and force applied. the axillary artery ultimately tore and bled. the pump was removed and the graft site was sutured and repaired successfully. | [5 empty] | The investigation into the vascular damage has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on the clinical details, the cause of the vascular damage was most likely due to use related as excessive force was applied. the failure mode will be monitored and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
35557 | 1220648-2023-02181 | 16950286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16950286&pc=OZD | 04/20/2023 | 10/15/2024 |  | 05/17/2023 |  |  | 04/20/2023 | 05/17/2023 |  |  | 11/11/2023 |  | 05/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166397 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | Life Threatening | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2023 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. additionally, patient went into ventricular tachycardia. cardiopulmonary resuscitation (cpr) was administered for about 2 minutes. patient survived the procedure. | A 42 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient developed hemolysis during the support and was transferred between hospitals with ecmo cannulation. |  |  |  |  | Additional data: b5, d6a, d6b and h6. the investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the hemolysis and ventricular tachycardia was not determined since product was not returned, extracorporeal membrane oxygenation (ecmo) was also in place, and removing impella did not resolve the hemolysis indicating unknown relation with impella. this report is being filed as part of a retrospective review of historical records. corrected data: d6a, d6b, f6, and f8. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35558 | 1220648-2023-02183 | 16950291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16950291&pc=OZD | 04/21/2023 | 07/10/2023 |  | 05/17/2023 |  |  | 04/21/2023 | 05/17/2023 |  |  | 09/30/2024 |  | 05/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023257920 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/21/2023 |  |  |  |  | A 60 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) was to undergo a right heart catheterization (rhc). the decision was made to place an impella 5.5 via an axillary insertion. it was noted the vessel size was inappropriate. the team made the decision to explant the impella 5.5 and place an impella cp instead. upon imaging with contrast, it was noted that axillary artery had been dissected during the explant of the impella 5.5 and the impella placement was aborted. the team had to perform a cutdown and surgical repair of the axillary artery. the patient received 2u packed red blood cells (prbc)s as well as fresh frozen plasma (ffp) and platelets. | [5 empty] | The investigation into the delivery issues has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis; only the clinical report was evaluated. based on the clinical details, the root cause of the delivery issues is due to the patient's anatomy. the failure mode will be monitored and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
35559 | 1220648-2023-02170 | 16938290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16938290&pc=OZD | 04/14/2023 | 08/09/2023 |  | 05/16/2023 |  |  | 04/17/2023 | 05/16/2023 |  |  | 10/31/2023 |  | 05/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 2022148659 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 53-year-old male presented with ami / cgs and has been treated with pci. first, an iabp and va-ecmo have been inserted, later the iabp has been replaced for an impella cp for hemodynamic support. after va-ecmo removal and progressive low output syndrome, hypoperfusion of the leg, ischemia and acidosis have been noted. decision has been made to remove the impella device and insert a new impella cp via contralateral femoral artery. pump removal resolved the ischemia. the patient has been critical and has been sent to icu. within three hours of admission to unit, the patient expired. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the ischemia was progressive general hypoperfusion patient condition related. | The impella device has not been received by the manufacturer for investigation. should any new information be returned, a supplemental report will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35560 | 1220648-2023-02171 | 16938604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16938604&pc=OZD | 04/17/2023 | 09/20/2024 |  | 05/16/2023 |  |  | 04/17/2023 | 05/16/2023 |  |  | 01/31/2025 |  | 05/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024317680 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | A 61-year-old patient with known coronary artery disease was treated for acute myocardial infarction and cardiogenic shock. an impella cp with smart assist cardiac pump was inserted for hemodynamic support. the patient required cardiopulmonary resuscitation while the impella cp was inserted. the patient's anticoagulation status was inadequate due to compliance problems with heparin. the pump was removed after approximately 20 hours. the patient was in poor neurologic condition. ct scans of the brain showed thrombotic arteria cerebri media, which was several hours old at that time. the patient was transferred to the palliative care unit. the cause of the ischemic stroke is unclear. |  |  |  |  | Since the medical center did not return the impella device, no benchtop analysis was possible. without further clinical details or imaging, the cause of the stroke is undetermined and the relationship to impella is unknown. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02171 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The impella device has not been received by the manufacturer for investigation. should any new information be returned, a supplemental report will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35561 | 1220648-2023-02172 | 16943449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16943449&pc=OZD | 04/18/2023 | 06/23/2023 |  | 05/16/2023 |  |  | 04/19/2023 | 05/16/2023 |  |  | 12/31/2024 |  | 05/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024304162 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | No | Other | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2023 |  |  |  |  | A 52 year old male presented for impella support. mitral regurgitation (mr) was confirmed by transesophageal echocardiography on admission to the operating room. when the heart was opened, a chordae rupture was found. the physician judged the pump may have caused or contributed to the mr/chordae rupture. the impella was removed and closure surgery was performed. | [5 empty] | The investigation into the heart valve damage has been completed since the original report was filed. the medical center discarded the impella product at the time of explant. no benchtop analysis of the root cause of the heart valve damage was possible; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of heart valve damage was most likely related to improper positioning of the device deep in the ventricle since the pigtail was reported in contact with the mitral apparatus. the failure mode will be monitored and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
35562 | 1220648-2023-02173 | 16943450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16943450&pc=OZD | 04/19/2023 | 07/25/2023 |  | 05/16/2023 |  |  | 04/19/2023 | 05/16/2023 |  |  | 01/31/2025 |  | 05/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024296598 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2023 |  |  |  |  | A 59 year old male underwent elective placement of impella for coronary artery bypass grafting (cabg), mitral valve replacement (mvr), and maze procedure to treat atrial fibrillation. the pump was repositioned twice with suction alarms and hemolysis after second repositioning. the pump was weaned and removed. the user facility expressed concerns of hemolysis with low pump flow. | [5 empty] | Since the original report was filed, the investigation into the hemolysis has concluded. data logs were analyzed however, the pump itself was not returned for benchtop investigation. the logs revealed the pump to have come out of proper position but, the positioning issue was resolved. the cause of the hemolysis was not determined since product was not returned. | The user facility has notified abiomed that the impella 5.5 will be returned for evaluation. as of the date of this report, it has not been received. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35563 | 1220648-2023-02174 | 16943455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16943455&pc=OZD | 02/19/2023 | 06/26/2023 |  | 05/16/2023 |  |  | 02/19/2023 | 05/16/2023 |  |  | 10/31/2024 |  | 05/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024287238 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/19/2023 |  |  |  |  | An 89 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient developed an access-site hematoma during the support. the repositioning sheath was re-sutured in place and five (5) units of blood was delivered. the relationship to the impella device and the need for blood products could not be denied. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center discarded the impella product at the time of explant. no benchtop analysis of the root cause of the limb ischemia was possible; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of access site bleeding was most likely related to use issue. the failure mode will be monitored and trended. | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
35564 | 1220648-2023-02175 | 16943464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16943464&pc=OZD | 04/19/2023 | 07/11/2023 |  | 05/16/2023 |  |  | 04/19/2023 | 05/16/2023 |  |  | 12/31/2024 |  | 05/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024294358 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | Yes | Pumping Stopped | Material Split, Cut or Torn | [5 empty] | Cardiac Arrest | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2023 |  |  |  |  | A 61 year old male of unknown race with cardiomyopathy presented for impella support. the user facility reported damage to the pump and cable. the pump stopped with retrograde flow alarms. they proceeded to code the patient and the patient required ecmo. eventually, the pump was removed after ecmo cannulated. the pump was sequestered by the hospital. | [5 empty] | The investigation into the pump stop has been completed since the original report was filed. the user facility has returned the impella device, and the benchtop analysis of the root cause revealed that the cause of the pump stop was most likely an open electrical line due to excessive force applied to the catheter while moving the patient. the failure mode will be monitored and trended. | The user facility has reported the impella device will be returned. abiomed has not received it as of the date of this report. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35565 | 1220648-2023-02164 | 16932597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16932597&pc=OZD | 02/02/2023 | 08/15/2023 |  | 05/15/2023 |  |  | 02/02/2023 | 05/15/2023 |  |  | 11/30/2024 |  | 05/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023281927 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  |  |  | Injury | Yes | Yes | Device Dislodged or Dislocated | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/02/2023 |  |  |  |  | A 55 year old female of unknown race with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the impella was successfully introduced. the impella dislodged during the switch from introducer to repositioning sheath. the physician had handling issues with the position of the repo sheath on the impella catheter. impella was then kinked in aorta or lv (uncertain). impella was later removed and the patient was placed under va ecmo in cpr condition. due to the dislodgement, this medical intervention was required. another impella cp placed later, finally patient stable and successfully weaned. although patient had been unstable before the incident, it cannot be ruled out that the malfunction did not additionally contribute to the patient's critical condition. | [5 empty] | The investigation into the positioning issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the positioning issues was use related. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly." ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
35566 | 1220648-2023-02165 | 16934775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16934775&pc=OZD | 01/31/2023 | 06/09/2023 |  | 05/15/2023 |  |  | 03/06/2023 | 05/15/2023 |  |  | 08/31/2024 |  | 05/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269565 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2023 |  |  |  |  | A 79 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient developed lower limb ischemia coinciding with the impella support. a fem-fem bypass was performed to treat the noted ischemia. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the root cause for limb ischemia remains undetermined since neither the product nor sufficient clinical information were provided. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35567 | 1220648-2023-02166 | 16934785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16934785&pc=PYX | 05/10/2023 | 05/15/2023 |  | 05/15/2023 |  |  | 05/11/2023 | 05/15/2023 |  |  | 01/31/2025 |  | 05/15/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024322358 |  | 0046-0035 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/11/2023 |  |  |  |  | A 79 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support post cath lab procedures. the patient had 100% occlusion to the right coronary artery and three stents were placed. the patient received impella rp support for four (4) days when there were signs of hemolysis. the patient was supported by both rp and intra-aortic balloon pump (iabp). the patient had minimal urine output and was red in color. the medical team made the choice to explant the impella rp. the patient had care withdrawn and subsequently expired. | [5 empty] | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿
35568 | 1220648-2023-02167 | 16935336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16935336&pc=OZD | 04/15/2023 | 07/26/2023 |  | 05/15/2023 |  |  | 04/15/2023 | 05/15/2023 |  |  | 11/30/2024 |  | 05/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2024310502 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2023 |  |  |  |  | The complainant reported a 65-year-old male patient with acute myocardial infarction and cardiogenic shock presenting for impella placement. during support, it was documented that the patient lost distal flow in the impella leg. the patient was escalated from an impella cp to an impella 5.5 as a result of the condition, however an access site bleed developed. the patient was provided an unknown quantity of blood products to counteract the bleed. support with the new pump continued following the intervention and original pump explant. | [5 empty] | The investigation into the access site bleeding and the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the access site bleeding was most likely use related and the cause of the limb ischemia was most likely patient condition related. the failure modes will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, ischemia and bleeding are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿
35569 | 1220648-2023-02169 | 16935343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16935343&pc=OZD | 04/19/2023 | 06/15/2023 |  | 05/15/2023 |  |  | 04/19/2023 | 05/15/2023 |  |  | 09/30/2024 |  | 05/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2023264183 |  | 0048-0024 JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2023 |  |  |  |  | A 63 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient was experiencing lower limb ischemia. the limb lost color and the team explanted the device to perform angioplasty with thrombectomy. the patient had know peripheral artery disease (pad) with a narrow blood vessel in the proximal part of the superficial femoral artery (sfa). | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the patient condition was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35570 | 1220648-2023-02168 | 16935735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16935735&pc=OZD | 02/10/2023 | 06/07/2023 |  | 05/15/2023 |  |  | 02/10/2023 | 05/15/2023 |  |  | 10/31/2024 |  | 05/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283337 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/10/2023 |  |  |  |  | A 76 year old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported they successfully placed the impella in the right femoral artery post arrest. the patient developed a hematoma after removal of 14 fr peel-away and advancement of the impella repositioning sheath. manual pressure was held and two (2) units of blood products were delivered to treat the hematoma-related blood volume loss | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that use issue was the most likely cause of the access site bleeding. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
35571 | 1220648-2023-02155 | 16918315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16918315&pc=OZD | 04/14/2023 | 10/26/2023 |  | 05/12/2023 |  |  | 04/14/2023 | 05/12/2023 |  |  | 10/31/2024 |  | 05/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2023272116 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age and gender presented with acute myocardial infarction and cardiogenic shock. for hemodynamic support, an impella cp with smartassist cardiac pump has been inserted. after one day, hemolysis has been observed. correct positioning of the pump and sufficient volume status of the patient has been confirmed. thrombolysis using tpa has only been partially effective, so after two days the decision was made to remove the impella cp. a new impella cp has successfully been inserted. the patient has been stable afterwards. | [5 empty] | The investigation into the hemolysis issue has been completed since the original report was filed. the product was returned for investigation. no abnormalities found on returned pump. pump was run and purged, and the purge pressure was within normal range. the flow was slightly saturated. the pump was run through a hemolysis test and the hemolysis did not reproduce. the cause of the hemolysis was most likely patient condition related as the hemolysis did not reproduce on the returned pump and no abnormalities were observed on the pump. | The impella device has been returned to the manufacturer. when the investigation is complete, a supplemental report will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35572 | 1220648-2023-02156 | 16918336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16918336&pc=OZD | 04/13/2023 | 05/12/2023 |  | 05/12/2023 |  |  | 04/13/2023 | 05/12/2023 |  |  | 02/28/2025 |  | 05/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024313449 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 65-year-old male presented with a week of chest pain. stemi alert called and brought to cardiac catheter lab for pci procedure. severe mitral regurgitation noted. post pci an impella cp has been placed. during the case the decision has been made to escalate the patient to an impella 5.5. the patient suffered from bleeding that required transfusion of an unknown amount of blood products. location of the bleeding remained unknown and involvement of the axillary impella 5.5 access site cannot be excluded. the patient recovered the following days. finally, the impella 5.5 has been successfully weaned and explanted. | [5 empty] | The impella device has not been received by the manufacturer for investigation. should any new information be returned, a supplemental report will be filed. of note the ifu states: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35573 | 1220648-2023-02158 | 16919947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16919947&pc=OZD | 04/19/2023 | 05/12/2023 |  | 05/12/2023 |  |  | 04/25/2023 | 05/12/2023 |  |  | 07/31/2024 |  | 05/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2023269018 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/25/2023 |  |  |  |  | A 50 year old male with cardiomyopathy presented for impella support. the user facility reported the patient endured hemolysis during the support. haptoglobin was administered and support was escalated to an impella 5.5. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35574 | 1220648-2023-02157 | 16920451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16920451&pc=OZD | 04/13/2023 | 07/18/2023 |  | 05/12/2023 |  |  | 04/13/2023 | 05/12/2023 |  |  | 11/30/2024 |  | 05/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298207 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2023 |  |  |  |  | A 39 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient developed an access site bleed around the 14 fr impella introducer. manual pressure was held to achieve hemostasis and multiple units of blood were delivered. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the issue was high patient act values. the failure mode will be monitored and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
35575 | 1220648-2023-02159 | 16924280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16924280&pc=PYX | 04/12/2023 | 07/26/2023 |  | 05/12/2023 |  |  | 04/12/2023 | 05/12/2023 |  |  | 01/31/2025 |  | 05/12/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024320058 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 68-year-old male has been treated for right heart failure. the patient has been transferred from another hospital with an iabp already in place. va-ecmo has been placed. unsuccessful attempt made to place a protek duo rvad. difficult but successful impella rp placement. during first attempts of femoral insertion, it was not possible to advance the pump past the tricuspid valve. to facilitate pump insertion, a sternotomy has been performed and the heart has been manipulated. the impella pump worked as intended. after 30 minutes of support blood pressure in pulmonary artery increased and impella flows lowered. pericardial effusion has been suspected. pulmonary artery has been opened and a embolectomy has been performed. impella position has been confirmed. impella low flows have not been resolved. hematuria and hemolysis has been noted. since the attending physician considered it possible that the impella rp has been causing or contributing to the hemolysis and the patient has also been doing better hemodynamically, the impella rp has been removed. va- ecmo has also been decannulated and patient has been sent back to the icu. the patient has been stable afterwards. | [5 empty] | The investigation into the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause. based on clinical information provided, the cause of the hemolysis was not determined since neither product nor data logs were returned. the failure mode will be monitored and trended. | The impella device will be returned to the manufacturer. when the investigation is complete, a supplemental report will be filed.
35576 | 1220648-2023-02160 | 16924289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16924289&pc=OZD | 04/16/2023 | 06/30/2023 |  | 05/12/2023 |  |  | 04/16/2023 | 05/12/2023 |  |  | 11/30/2024 |  | 05/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311235 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/16/2023 |  |  |  |  | A 54 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient developed ischemia at the impella-inserted limb. the team noted the patient had vascular dissection of superficial femoral artery (sfa), and profunda with 4-compartment fasciotomy, thrombectomy of common femoral artery (cfa) due to ischemic right leg with compartment syndrome. the patient support was for 2 days when pump was explanted. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that biomaterial was the most likely cause of the limb ischemia. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35577 | 1220648-2023-02161 | 16924296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16924296&pc=OZD | 04/17/2023 | 07/20/2023 |  | 05/12/2023 |  |  | 04/19/2023 | 05/12/2023 |  |  | 12/31/2024 |  | 05/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316127 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2023 |  |  |  |  | A 56 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported the patient had an impella cp inserted via the left femoral after replacing a prior impella. the new impella was inserted via the same left femoral. the patient's right femoral had a bleed from the percutaneous coronary intervention (pci), but this pump site was intact without any bleeding. the limb, however, was ischemic. there was no pedal pulse but above the popliteal was fine with flow. the physician consulted with vascular for the removal of the pump. the impella cp was removed on (b)(6) 2023. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the limb ischemia was patient condition related due to pad/pvd vessel conditions. the failure mode will be monitored and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35578 | 1220648-2023-02162 | 16924322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16924322&pc=OZD | 05/02/2023 | 07/18/2023 |  | 05/12/2023 |  |  | 05/10/2023 | 05/12/2023 |  |  | 10/31/2024 |  | 05/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023266157 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2023 |  |  |  |  | A 58 year old male with cardiomyopathy presented for impella support. the patient was on impella support for eighty-one (81) days awaiting a heart transplant. on day 80, the team had to troubleshoot to remedy a purge leak at the sidearm. the team intervened by applying a wrap with tegaderm and sutures as there was no bone wax available. the team was successful and the support ended the following day as planned. there was no early explant. the impella had been used well beyond indication. there was no serious injury to the patient. | [5 empty] | The investigation into the purge leak has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the issue was a leak from the sidearm, but the exact location of the leak was unknown since product was not returned. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35579 | 16910200 | 16910200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16910200&pc=OZD | 04/07/2023 | 05/08/2023 | 05/08/2023 | 05/11/2023 | 05/08/2023 | 05/11/2023 |  | 05/11/2023 |  |  |  |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED INC. |  |  | US | [5 empty] | No |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 22265 DA |  | Male | 101 KG |  | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Patient's impella device was repositioned by team. attending noticed during repositioning that the impella sheath has been torn. impella was retracted and then advanced by attending. manufacturer response for temporary non-roller type left heart, impella (per site reporter) manufacturer is working on re-education and investigating further.
35580 | 16910219 | 16910219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16910219&pc=OZD | 04/09/2023 | 05/08/2023 | 05/08/2023 | 05/11/2023 | 05/08/2023 | 05/11/2023 |  | 05/11/2023 |  |  |  |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED INC. |  |  | US | [5 empty] | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 16060 DA |  | Male | 110 KG |  |  | Malfunction | No | Yes | Improper Flow or Infusion | Protective Measures Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella purge alarming for "purge pressure high" and "purge system blocked". purge flow noted to be trending down for ~1 hr prior to the event. the rn change purge cassette without resolution of alarm. abiomed called and reported that tpa needed to be given through purge. pharmacy consulted and reviewed abiomed policy for tpa administration. tpa prepared and administered as purge solution. patient hemodynamically stable with no drops in impella blood flow/pressure. manufacturer response for temporary non-roller type left heart support blood pump, impella (per site reporter) under active investigation by the vendor. in constant communication.
35581 | 16910246 | 16910246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16910246&pc=OZD | 03/08/2023 | 05/08/2023 | 05/08/2023 | 05/11/2023 | 05/08/2023 | 05/11/2023 |  | 05/11/2023 |  |  |  |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED INC. |  |  | US | [5 empty] | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 23725 DA |  | Male | 94 KG |  |  | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | After fellow did adjustments to the catheter it was noted the sterile sheath covering the impella was ripped at the hub entering the patient. it was unclear if this was present before the patient arrived to [redacted name] hospital or before. this is the second noticed torn sheath. manufacturer response for temporary non-roller type left heart support pump, impella (per site reporter): manufacturer is currently investigating.
35582 | 16910269 | 16910269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16910269&pc=OZD | 02/13/2023 | 05/08/2023 | 05/08/2023 | 05/11/2023 | 05/08/2023 | 05/11/2023 |  | 05/11/2023 |  |  |  |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA PURGE CASSETTE |  |  | ABIOMED INC. |  |  | US |  |  |  | 0042-0040-US |  | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 22630 DA |  | Male | 72 KG |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella purge cassette was found to be leaking after a low-pressure alarm sounded. the cassette was replaced per the manufacturer¿s guidelines with no adverse outcome. manufacturer response for temporary non-roller type left heart support blood pump, impella (per site reporter) manufacturer is actively investigating.
35583 | 16910297 | 16910297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16910297&pc=OZD | 04/10/2023 | 05/08/2023 | 05/08/2023 | 05/11/2023 | 05/08/2023 | 05/11/2023 |  | 05/11/2023 |  |  |  |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED INC. |  |  | US |  |  |  | 0042-0040-US |  | No answer provided |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 16060 DA |  | Male | 110 KG |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella purge cassette leaked. bag and cassette changed. the cassette was gen 1 and replaced with gen 2 when available. consistent with previous purge cassette failures. manufacturer response for temporary non-roller type left heart support blood pump., impella (per site reporter) the vendor is actively investigating.
35584 | 16910348 | 16910348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16910348&pc=OZD | 05/02/2023 | 05/08/2023 | 05/08/2023 | 05/11/2023 | 05/08/2023 | 05/11/2023 |  | 05/11/2023 |  |  |  |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED INC. |  |  | US |  |  |  | 0042-0040-US |  | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 22995 DA |  | Male | 98 KG |  |  | Malfunction | No | Yes | Detachment of Device or Device Component | Protective Measures Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Proximal hub/connector of protective sheath (farthest away from patient) on impella drive line was not connected to plastic part of sheath exposing drive line to room/patient. the nurse came into the room and the two were detached. device is still implanted, pending transplant. manufacturer response for temporary non-roller type left heart support blood pump, abiomed impella (per site reporter) they are actively investigating.
35585 | 1220648-2023-02143 | 16910660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16910660&pc=OZD | 04/12/2023 | 07/14/2023 |  | 05/11/2023 |  |  | 04/12/2023 | 05/11/2023 |  |  | 01/31/2025 |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024285062 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 48-year-old male has been treated for ami/cgs. initially, an impella cp cardiac pump as well as an va-ecmo has been placed. subsequently, the patient required additional hemodynamic support, so the decision was made to escalate the patient from impella cp to an impella 5.5 pump. axillary insertion of impella 5.5 proved difficult. the attending physician noted a lot of resistance while inserting the pump. during the intervention, an aortic dissection type b has been found. the physician reviewed the case and thinks that the dissection happened during the femoral insertion of the initial impella cp. conservative treatment of the aortic dissection. patient is on impella and ecmo support. | [5 empty] | The investigation into the reported access site bleeding and vascular dissection has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the cause of the vascular damage was not determined since product was not returned and there is insufficient clinical information regarding the cp insertion. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device has not been received by the manufacturer for investigation. should any new information be returned, a supplemental report will be filed. as noted in the impella ifu: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿
35586 | 1220648-2023-02144 | 16910693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16910693&pc=PYX | 01/19/2023 | 05/11/2023 |  | 05/11/2023 |  |  | 01/19/2023 | 05/11/2023 |  |  | 12/31/2023 |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023164857 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/19/2023 |  |  |  |  | A 46 year old male with acute myocardial infarction and cardiogenic shock presented for impella support. the user facility reported the patient experienced an access site bleed during impella support. the patient was given three units of packed red blood cells (prbc) to treat the noted bleed/ooze. patient support continued following the event. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
35587 | 1220648-2023-02145 | 16911307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16911307&pc=OZD | 01/20/2023 | 10/14/2024 |  | 05/11/2023 |  |  | 01/20/2023 | 05/11/2023 |  |  | 10/31/2024 |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283247 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Tachycardia | [10 empty] | Other | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/20/2023 |  |  |  |  | A 63 year old male with acute myocardial infarction and cariogenic shock presented for impella support. the patient had some notable mottling to right lower extremity (rle) and, once in the unit, the primary nurse noted patient to have hemolysis. the catheter was repositioned and the issue was resolved. | It was reported that the patient had also experienced ventricular tachycardia. action taken to resolve the issue are unknown. |  |  |  |  | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the hemolysis issue was not determined since product was not returned; however, the limb ischemia was determined to be patient condition related since the ischemia improved upon removal of the introducer. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: hemolysis: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ limb ischemia: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information for the initial mfr 1220648-2023-02145 was provided in sections b5, d6a, d6b, h1, and h6. the investigation for the reported ventricular tachycardia has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of ventricular tachycardia could not be determined. this report is being filed as part of a retrospective review of historical records.
35588 | 1220648-2023-02146 | 16912091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16912091&pc=OZD | 01/30/2023 | 08/17/2023 |  | 05/11/2023 |  |  | 01/30/2023 | 05/11/2023 |  |  | 07/31/2024 |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250080 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2023 |  |  |  |  | A 57 year old female with acute myocardial infarction and cardiogenic shock presented for impella support. during support, a hematoma was noted in the right groin area. two units of packed red blood cells (prbc) were administered to the patient. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information provided, the cause of the access site bleeding was not determined since the product was not returned. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿
35589 | 1220648-2023-02147 | 16912720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16912720&pc=OZD | 04/25/2023 | 07/14/2023 |  | 05/11/2023 |  |  | 04/25/2023 | 05/11/2023 |  |  | 12/31/2024 |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321808 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2023 |  |  |  |  | A 62 year old female with acute myocardial infarction and cardiogenic shock presented for impella support. the user facility reported that the superficial femoral artery of a patient with peripheral vascular disease became occluded following placement of the impella pump. the pump was subsequently explanted and the patient's condition improved. the patient was considered to be stable following the event. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the cause of the issue was patient condition related due to peripheral artery/peripheral vascular disease conditions based on the clinical information provided. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35590 | 1220648-2023-02148 | 16916570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16916570&pc=OZD | 09/26/2022 | 06/09/2023 |  | 05/11/2023 |  |  | 09/26/2022 | 05/11/2023 |  |  | 11/30/2023 |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161896 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2022 |  |  |  |  | A 52 year old male patient was transferred to uab hospital for escalation of care. it was reported that the impella cp device was explanted on (b)(6) 2022 for an apparent hemolysis problem. | [5 empty] | The investigation into the reported hemolysis has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report and data logs were evaluated. data logs showed there were no abnormal pump performance metrics or trends on the day of the reported hemolysis. no relevant alarms occurred. in conclusion, the cause of the hemolysis and its relation to the impella were unable to be determined since neither the product nor sufficient clinical information were provided. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35591 | 1220648-2023-02149 | 16916572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16916572&pc=OZD | 04/14/2023 | 05/11/2023 |  | 05/11/2023 |  |  | 04/14/2023 | 05/11/2023 |  |  | 05/31/2024 |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023217470 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Rupture | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/14/2023 |  |  |  |  | The complainant reported a male patient with post cardiotomy cardiogenic shock was presenting for impella placement. while weaning from the cardiopulmonary bypass following tvr, mitral valve replacement, and cox-maze procedures, the patient began to experience left and right ventricular failure. the impella pump was subsequently implanted for support and fluids were given to achieve acceptable flow levels. however, at that point in time, the left ventricle ruptured. attempts were made to repair the rupture, but the tissue was fragile and more ruptures occurred. the decision was made to withdraw care and the patient passed away. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
35592 | 1220648-2023-02150 | 16916575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16916575&pc=OZD | 12/23/2022 | 08/21/2023 |  | 05/11/2023 |  |  | 12/23/2022 | 05/11/2023 |  |  | 08/31/2024 |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023266125 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/23/2022 |  |  |  |  | A 78 year old male patient with impella cp was having hemolysis. the staff was unable to resolve by repositioning. the impella cp was removed and physician successfully inserted impella 5.5. the patient was stable afterwards. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information provided, the root cause of the hemolysis was not determined since the product was not returned and insufficient clinical details were provided. the failure mode will be monitored and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35593 | 1220648-2023-02152 | 16916579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16916579&pc=OZD | 04/14/2023 | 05/11/2023 |  | 05/11/2023 |  |  | 04/14/2023 | 05/11/2023 |  |  | 11/30/2024 |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED INC. |  |  | US |  | 2024302464 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/14/2023 |  |  |  |  | The medical center had an impella cp support ongoing along with ecmo when there were signs of hemolysis. pump position was assessed for standard troubleshooting measures. the team explanted the pump due to the hemolysis concerns and were considering dialysis. the patient remained on ecmo support. | [5 empty] | The medical center did not yet return any impella pump product for investigation or benchtop analysis and hemolysis testing. the root cause of the hemolysis has not been determined. the concurrent use of ecmo may or may not have contributed to the hemolysis. should new information be shared by the medical center, and root cause determined, a final report will be filed. of note the ifu states the following: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ "patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35594 | 1220648-2023-02151 | 16916580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16916580&pc=OZD | 01/18/2023 | 10/29/2024 |  | 05/11/2023 |  |  | 01/18/2023 | 05/11/2023 |  |  | 06/30/2024 |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236409 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | Renal Impairment | [10 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/18/2023 |  |  |  |  | A 38 year old male with complete left main coronary artery blockage presented for impella support. the user facility reported the patient developed hemolysis during the use of the device. the physician opted to continue impella support, however continuous renal replacement therapy (crrt) was required. the patient expired two days later. | Additional information was received from the complainant that reported the patient had episodes of atrial tachycardia. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information provided, the cause of the hemolysis was not determined since no product or data was returned and the site noted that there was no known relationship between hemolysis and impella. the failure mode will be monitored and trended. | The impella device was not returned by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Additional information for the initial mfr 1220648-2023-02151 was provided in sections b5, d6a, d6b, h3, h6 and h10. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records.
35595 | 1220648-2023-02153 | 16916592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16916592&pc=OZD | 04/27/2023 | 09/20/2024 |  | 05/11/2023 |  |  | 04/27/2023 | 05/11/2023 |  |  | 11/30/2024 |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024295825 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 34 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/27/2023 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | A 34 year old male presented for impella support along with ecmo in japan. the patient was found down after ventricular fibrillation and arrest. the pump support was ongoing for multiple days when the pump was seen to be protruding from the left ventricle (lv) via a perforation in the tissue. the inlet part was in the lv therefore impella continued to support it. the patient condition was poor but, life-saving surgery considered and discussed. the team instead made the patient an dnr and the patient expired with underlying hypoxic encephalopathy. |  |  |  |  | The investigation into the ventricle perforation and the patient death has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the ventricle perforation was most likely patient condition related and related to the patient's vertical heart and fragile myocardium; the patient death was not related to the abiomed device. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
35596 | 1220648-2023-02154 | 16916597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16916597&pc=OZD | 04/12/2023 | 05/11/2023 |  | 05/11/2023 |  |  | 04/12/2023 | 05/11/2023 |  |  | 02/28/2025 |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024311633 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/12/2023 |  |  |  |  | A 77 year old male with cardiomyopathy presented for impella support. the user facility reported multiple pump stops with successful restarts during the course of support with the device. the decision was made to remove the pump after approximately 15 days of support. the device was explanted as the patient was deemed not a candidate for left ventricle assist device (lvad) and the goal is to eventually be discharged home on milrinone. after the pump was removed, the patient's hemodynamics were stable. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35597 | MW5117484 | 16919520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16919520&pc=OZD | 04/26/2023 | 05/09/2023 |  | 05/11/2023 |  |  |  | 05/11/2023 |  |  | 12/31/2024 |  | 05/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA 5.5 |  |  | OZD |  |  | 0500-0008 |  |  | ABIOMED INC. |  |  | US | [5 empty] | Yes | [6 empty] | Health Professional |  |  | Voluntary report | Yes | Other Health Care Professional | [8 empty] | Male |  | Non Hispanic |  | Injury | No | Yes | Break | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Required Intervention | Life Threatening | Hospitalization | [26 empty] | Initial submission | [9 empty] | Abiomed impella 5.5 tubing came apart between gray and red and wire broken upon patient was transferred back to bed. no apparent or obvious pull or snag. the impella stopped working and then the patient became hypotensive and coded.
35598 | 1220648-2023-02105 | 16901738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16901738&pc=OZD | 02/20/2023 | 07/18/2023 |  | 05/10/2023 |  |  | 02/20/2023 | 05/10/2023 |  |  | 10/31/2024 |  | 05/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2024288293 |  | 008-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2023 |  |  |  |  | A 74-year-old present with chest pain and in cardiogenic shock. impella removed due to loss of right lower extremity (rle) distal pulses and concern for limb ischemia. the patient remains on levophed at 30 mcg/min and vasopressin at .03 units/min. explanted without notice to field team. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the limb ischemia was most likely patient condition related. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella information for use related to this event: impella 5.5: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ impella cp: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35599 | 1220648-2023-02112 | 16901740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16901740&pc=OZD | 04/05/2023 | 06/12/2023 |  | 05/10/2023 |  |  | 04/06/2023 | 05/10/2023 |  |  | 12/31/2024 |  | 05/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024302275 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 69-year-old male patient was admitted in cardiogenic shock due to acute myocardial infarction. there was known coronary artery disease, which was treated by percutaneous coronary intervention without impella support. the patient was then transferred to the intensive care unit, where his hemodynamic condition worsened. ventricular tachycardia required cardiopulmonary resuscitation. an impella cp cardiac pump was successfully inserted to unload the heart. the patient was moving in bed, coughing, and applying pressure. bleeding occurred at the access site. the anticoagulation status at the time of the incident is unknown. the impella cp was removed, and the femoral artery was surgically closed. an unknown number of red blood cell concentrates, and additional catecholamine were administered to hemodynamically stabilize the patient. the attending physician is uncertain whether introducer sheath was inserted far enough into the femoral artery. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that use error was the mostly like cause of the access site bleeding because the repo sheath was not advanced deep enough into the arterial femoral artery. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device has not been received by the manufacturer for investigation. should any new information be returned, a supplemental report will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿
35600 | 1220648-2023-02104 | 16901741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16901741&pc=OZD | 04/05/2023 | 06/06/2023 |  | 05/10/2023 |  |  | 04/05/2023 | 05/10/2023 |  |  | 02/28/2025 |  | 05/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC |  |  | US |  | 2024310470 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2023 |  |  |  |  | The complainant reported a 29-year-old female patient presenting for post cardiotomy cardiogenic shock. during initial impella placement, the pump was placed deeper in the small ventricle as a result of the patient's small aorta and subsequently punctured through the ventricular wall. the pump was removed and an lv repair was completed to treat the perforation. following the intervention, the patient was stable and preparations for va ecmo support were initiated. | [5 empty] | The investigation into the ventricle perforation has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that poor ventricle condition and impella position deep in the ventricle was the cause of the ventricle perforation. the failure mode will be monitored and trended. ventricle perforation ifu quotes: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. of note the ifu states: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
35601 | 1220648-2023-02106 | 16901746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16901746&pc=OZD | 02/10/2023 | 08/15/2023 |  | 05/10/2023 |  |  | 02/10/2023 | 05/10/2023 |  |  | 10/31/2024 |  | 05/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290493 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 02/10/2023 |  |  |  |  | A 63-year-old female presented for impella support. the physician performed percutaneous axillary impella insertion in the right upper extremity (rue) for high risk percutaneous coronary intervention (hrpci). it was discussed that ideally device would be left in place but also noted that limb ischemia to rue was present. because of this, device was weaned to assess patient and if patient would tolerate explant. device was weaned and hemodynamics remained stable for approximately 5-10 minutes and the physicians decided to go ahead and explant impella. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was most likely patient condition related. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. instructions for use related to this event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35602 | 1220648-2023-02113 | 16901747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16901747&pc=OZD | 02/10/2023 | 06/06/2023 |  | 05/10/2023 |  |  | 02/10/2023 | 05/10/2023 |  |  | 01/31/2024 |  | 05/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED |  |  | US |  | 2023181551 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/10/2023 |  |  |  |  | A 58-year-old male with decompensating non-ischemic cardiomyopathy (nicmp) of unknown origin presented with signs of infection at the impella's axillary insertion site. culture revealed proteus bacteria strain. the impella was removed and a new impella was placed in the opposing axillary artery. treatment consisted of site washout, debridement, wound vac and iv antibiotics. the site was showing good signs of tissue granulation and patient is currently awaiting a heart transplant. | [5 empty] | The investigation into the reported infection/sepsis has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the infection/sepsis. abiomed device is sterilized under validated conditions. the failure modes will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. instructions for use for this event are as follows: impella 5.5 & cp: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ "arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿ ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿
35603 | 1220648-2023-02138 | 16902500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16902500&pc=OZD | 04/12/2023 | 06/15/2023 |  | 05/10/2023 |  |  | 04/12/2023 | 05/10/2023 |  |  | 09/30/2024 |  | 05/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023281443 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2023 |  |  |  |  | The medical center had an impella cp inserted via the femoral artery to support a patient during a surgical work-up . the patient developed limb ischemia and an access site bleed. the limb lost pulses and the hematoma prompted vascular repair. the placement was noted to be done emergently during cpr, which may or may not have contributed to the vascular issue. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the patient condition; patient's calcified vessels was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The medical center discarded all impella pump and introducer sheath product at the time of explant. no benchtop analysis to root cause is possible. the contribution of the active cpr during impella insertion can not be denied. the causal relationship is not known. the failure mode will be monitored and trended. should more information be shared that leads to root cause, a final report will be filed. of note in the ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
35604 | 1220648-2023-02139 | 16902605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16902605&pc=OZD | 04/11/2023 | 05/10/2023 |  | 05/10/2023 |  |  | 04/11/2023 | 05/10/2023 |  |  | 06/30/2024 |  | 05/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023247611 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/11/2023 |  |  |  |  | The complainant reported an 84-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was documented that the patient developed an access site bleed and a hematoma. manual pressure was applied and an unknown quantity of blood products were administered to treat the condition. the pump was subsequently explanted following the intervention. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, bleeding is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35605 | 1220648-2023-02140 | 16902617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16902617&pc=OZD | 04/13/2023 | 08/09/2023 |  | 05/10/2023 |  |  | 04/13/2023 | 05/10/2023 |  |  | 01/31/2025 |  | 05/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC |  |  | US |  | 2024316615 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2023 |  |  |  |  | The complainant reported a 79-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was documented that the patient developed limb ischemia. the pump was explanted as a result of the condition. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the limb ischemia was most likely patient condition related. | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35606 | 1220648-2023-02141 | 16903039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16903039&pc=OZD | 12/16/2022 | 06/07/2023 |  | 05/10/2023 |  |  | 12/16/2022 | 05/10/2023 |  |  | 09/30/2024 |  | 05/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274887 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Twisted/Bent | Ejection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2022 |  |  |  |  | A 61 year old male presented for impella support. the user facility reported the pump became entangled/prolapsed within the patient during a repositioning attempt. a snare was required as intervention to straighten the device and re-cross the valve. the patient was stable with no complications. | [5 empty] | The investigation into the positioning issues has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that use issue was the most likely cause of the positioning issues. the failure mode will be monitored and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
35607 | 1220648-2023-02142 | 16906537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16906537&pc=OZD | 01/10/2023 | 06/25/2023 |  | 05/10/2023 |  |  | 01/10/2023 | 05/10/2023 |  |  | 08/31/2024 |  | 05/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270289 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2023 |  |  |  |  | A 76 year old female presented for impella support. the user facility reported leg ischemia during support that prompted intervention via an antegrade sheath. there was no failure with the pump automated impella controller (aic). | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center discarded the impella product at the time of explant. no benchtop analysis of the root cause of the limb ischemia was possible; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of limb ischemia was most likely related to patient condition. the failure mode will be monitored and trended. | The impella device was not returned from the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35608 | 1220648-2023-02126 | 16897140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16897140&pc=OZD | 04/09/2023 | 07/12/2023 |  | 05/09/2023 |  |  | 04/10/2023 | 05/09/2023 |  |  | 12/31/2024 |  | 05/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC |  |  | US |  | 2024309139 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2023 |  |  |  |  | The complainant reported a 25-year-old male patient presenting for cardiomyopathy. during impella support, it was noted that the patient experienced an ischemic stroke. the patient's vitals returned to baseline levels after 48 hours and support continued with the implanted pump. the patient was considered stable following the event. | [5 empty] | The investigation into the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the stroke/cva was most likely patient condition related. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, stroke is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35609 | 1220648-2023-02127 | 16897154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16897154&pc=OZD | 04/23/2023 | 08/09/2023 |  | 05/09/2023 |  |  | 04/23/2023 | 05/09/2023 |  |  | 12/31/2024 |  | 05/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321391 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2023 |  |  |  |  | A 51 year old male presented with hemolysis. the catheter was underneath the papillary muscle during support and developed hemolysis. the pump was repositioned in an attempt to troubleshoot the issue, however, the condition remained. the pump was subsequently explanted and replaced to resolve the hemolysis. | [5 empty] | The investigation into the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the hemolysis was most likely improper positioning of the device. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35610 | 1220648-2023-02128 | 16898470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16898470&pc=OZD | 04/10/2023 | 08/04/2023 |  | 05/09/2023 |  |  | 04/10/2023 | 05/09/2023 |  |  | 11/30/2024 |  | 05/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC |  |  | US |  | 2024297842 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2023 |  |  |  |  | The complainant reported a 72-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was documented that the patient developed hemolysis. the pump was explanted as a result of the hemolysis and the patient's condition stabilized. | [5 empty] | The investigation into the hemolysis has been completed since the original report was filed. the user facility has returned the impella device, and the benchtop analysis of the root cause revealed that biomaterial was the cause of the hemolysis. | The device was returned and an evaluation is pending. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results. as noted in the impella ifu, hemolysis is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
35611 | 1220648-2023-02129 | 16898935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16898935&pc=OZD | 03/30/2023 | 07/18/2023 |  | 05/09/2023 |  |  | 03/30/2023 | 05/09/2023 |  |  | 10/31/2024 |  | 05/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285865 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2023 |  |  |  |  | A 47-year-old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) developed distal limb ischemia after leaving the 14 fr peel-away introducer in the site. an antegrade sheath had been placed, however, the physician implanted the sheath in the reverse direction, which exacerbated the ischemia. the patient was taken to the or for surgical explant, closure, and escalation to an impella 5.5. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the limb ischemia was use related, specifically the improper implantation of the reperfusion sheath by the physician. the failure mode will be monitored and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35612 | 1220648-2023-02120 | 16898963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16898963&pc=OZD | 02/06/2023 | 06/07/2023 |  | 05/09/2023 |  |  | 02/06/2023 | 05/09/2023 |  |  | 12/31/2024 |  | 05/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024289688 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2023 |  |  |  |  | A 70-year-old male presented with post cardiotomy cardiogenic shock for impella support. during support, discharge developed from the impella access site. anti-biotics were administered to the patient. | [5 empty] | The investigation into the infection/sepsis has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the infection/sepsis. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿
35613 | 1220648-2023-02130 | 16898993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16898993&pc=OZD | 03/30/2023 | 08/04/2023 |  | 05/09/2023 |  |  | 03/30/2023 | 05/09/2023 |  |  | 06/30/2024 |  | 05/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023222104 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2023 |  |  |  |  | A 76 year old male with a coronary artery bypass graft (cabg) presented for impella support. while advancing the pump wirelessly into position, the patient's left ventricle was perforated. the injury was repaired and impella support proceeded. | [5 empty] | The investigation into the ventricle perforation has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the ventricle perforation was use related and related to manual manipulation of the heart. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
35614 | 1220648-2023-02131 | 16900038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16900038&pc=OZD | 04/10/2023 | 05/26/2023 |  | 05/09/2023 |  |  | 04/10/2023 | 05/09/2023 |  |  | 11/30/2024 |  | 05/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2024286189 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | HEART LUNG BYPASS | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2023 |  |  |  |  | The medical center in japan had an impella cp support in which there was positional issues and hemolysis. the hemolysis was confirmed by the team by the urine color change. the pump was prematurely explanted and replaced/escalated to the impella 5.5 pump. the support was for 5 days. | [5 empty] | Since the original report was filed the investigation into the reported hemolysis has concluded. the medical center in japan did not return any impella product for benchtop analysis and hemolysis testing. only the clinical report could be evaluated. the evaluation determined the hemolysis to be due to improper pump positioning, and as such is user related. the failure mode will be monitored and trended. of note the ifu states the following: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The complainant in japan discarded the impella cp at the time of explant and care escalation to the 5.5 pump. there can be no benchtop analysis to root cause via hemolysis testing. as such, no root cause is determined, but should the team in japan furnish more clinical details that lead to root cause, a final report will be forthcoming. the failure mode will be monitored and trended. of note in the ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿
35615 | 1220648-2023-02132 | 16900040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16900040&pc=OZD | 04/10/2023 | 06/15/2023 |  | 05/09/2023 |  |  | 04/10/2023 | 05/09/2023 |  |  | 12/31/2024 |  | 05/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2024300780 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2023 |  |  |  |  | The medical team in japan had an impella cp support for 5 days. the pump was then surgically explanted. the team observed aortic thrombosis cascading to/down the limbs. the team observed the thrombosis to cause lost pulses to the limbs and thrombectomy and the fogarty embolectomy was performed. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that biomaterial was the cause of the limb ischemia. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The medical team in japan discarded the impella pump product at the time of explant. no benchtop analysis to root cause is possible. the team did not furnish further clinical details or imaging for review. no root cause has been determined. should the team in japan relay details that lead to root cause a final report will be filed. the failure mode will be monitored and trended. of note the ifu states the following: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35616 | 1220648-2023-02133 | 16900041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16900041&pc=OZD | 03/28/2023 | 05/16/2023 |  | 05/09/2023 |  |  | 03/28/2023 | 05/09/2023 |  |  | 10/31/2024 |  | 05/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285929 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2023 |  |  |  |  | The user facility reported a 70 year old male undergoing impella cp support experienced a bleed from the access site. the patient was with activated clotting time (act) supra-therapeutic and was also on ecmo. the bleeding was treated with blood products. the blood products were given with the blood loss as well as volume replacement. the team sutured the site as well. | [5 empty] | Since the original report was filed, the investigation into the access site bleeding has concluded. no product was returned for benchtop analysis. only the clinical report was evaluated. the bleed was determined to be patient condition related, as the act values were too high and excessive blood loss occurred. the failure mode will be monitored and trended. as noted in the ifu: ¿assess access site for bleeding and hematoma.¿ | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿
35617 | 1220648-2023-02135 | 16900044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16900044&pc=OZD | 04/10/2023 | 06/15/2023 |  | 05/09/2023 |  |  | 04/10/2023 | 05/09/2023 |  |  | 07/31/2024 |  | 05/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023255125 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2023 |  |  |  |  | The us complainant had a failed delivery of the 5.5 despite all troubleshooting measures and angioplasty of the vascular/arterial path. the graft then pulled away from the artery and vascular repair was needed. the axillary artery was repaired with surgical sutures. the 5.5 delivery was aborted. | [5 empty] | The investigation into the delivery issues has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the patient condition was the delivery issues. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The medical center discarded the impella product at the time of explant. there can be no benchtop analysis to root cause. should more information be shared that leads to root cause, a final report will be forthcoming. the failure mode will be monitored and trended. of note the ifu states: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
35618 | 1220648-2023-02134 | 16900047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16900047&pc=OZD | 04/19/2023 | 08/09/2023 |  | 05/09/2023 |  |  | 04/19/2023 | 05/09/2023 |  |  | 01/31/2025 |  | 05/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312456 |  | 10000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Material Twisted/Bent | Ejection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2023 |  |  |  |  | A 78 year old male with accute myocardial infarction (ami) and cardiogenic shock (cgs) presented for impella support. the user facility reported a delivery issue involving the impella cp pump and guidewire. it was stated that the guidewire had wrapped around the cannula and caused a kink at the junction of the cannula and outlet cage. as a result, the pump would not advance across the aortic valve. the guidewire could not be manipulated and a snare was utilized to straighten out the cannula for explant. both the guidewire and pump were removed and the implant was aborted. there was no patient harm. | [5 empty] | The investigation into the delivery issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the delivery issues was patient condition related. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
35619 | 1220648-2023-02137 | 16900048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16900048&pc=OZD | 04/01/2023 | 06/15/2023 |  | 05/09/2023 |  |  | 04/01/2023 | 05/09/2023 |  |  | 08/31/2024 |  | 05/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270616 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2023 |  |  |  |  | A 47 year old female with acute myocardial infarction (ami) and cardiogenic shock (cgs) had diminished distal pules after arrival to the icu. the impella inserted leg became mottled. as a result, the device required premature removal. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the patient condition was the cause of the limb ischemia. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35620 | 1220648-2023-02136 | 16900049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16900049&pc=OZD | 04/10/2023 | 08/04/2023 |  | 05/09/2023 |  |  | 04/10/2023 | 05/09/2023 |  |  | 08/31/2024 |  | 05/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023262261 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2023 |  |  |  |  | The complainant reported a 70-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that an impella position in aorta alarm triggered after the patient began shifting around. during attempted repositioning of the pump, the device folded on itself in two different locations. a snare was used to capture the pigtail and straighten out the device, however the decision was made to explant the pump as the patient's pressure had stabilized. there was no patient harm. | [5 empty] | The investigation into the positioning issues has been completed since the original report was filed. the user facility has returned the impella device, and the benchtop analysis of the root cause of the positioning issue was patient movement as he was very fidgety and unable to lay that impacted impella and led to sudden drop in motor current and flow. | The device was returned for further evaluation. upon visual inspection, it was verified that the cannula and pigtail on the pump were damaged and creased. the investigation is still ongoing regarding a potential cause for the noted issue. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results.
35621 | 1220648-2023-02117 | 16884936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16884936&pc=OZD | 04/08/2023 | 08/04/2023 |  | 05/08/2023 |  |  | 04/08/2023 | 05/08/2023 |  |  | 01/31/2025 |  | 05/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024301187 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  |  | Injury | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 56-year-old male has been treated for cardiomyopathy / cardiogenic shock. for hemodynamic support, the patient received an impella 5.5 cardiac pump. on day 10, the device stopped and could not be restarted. after removal, the impeller of the device has been found coagulated. the stopped impella 5.5 has successfully replaced with a new impella 5.5 device. the patient has been stable afterwards. | [5 empty] | The investigation into the pump stop has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the pump stop was biomaterial. the failure mode will be monitored and trended. | The impella device has been returned to the manufacturer. when the investigation is complete, a supplemental report will be filed. as noted in the impella ifu: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: try to restart the catheter at previous p-level. if the impella does not restart, try to restart at p-2. if the impella does not restart or stops again, wait 1 minute and try to restart again. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35622 | 16885498 | 16885498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16885498&pc=OZD | 04/10/2023 | 05/01/2023 | 05/01/2023 | 05/08/2023 | 05/01/2023 | 05/08/2023 |  | 05/08/2023 |  |  |  |  | 05/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED INC. |  |  | US |  | 1698396 |  | 0043-0001 |  | No |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 16060 DA |  | Male | 108 KG |  |  | Malfunction | No | Yes | Break | Crack | Leak/Splash |  |  |  |  | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The impella purge cassette was noted to be leaking again and upon investigation the side arm on the purge cassette also broke off. the company was called to help find a solution with the use of a needle in line with the side arm. the solution was decided to be a lower risk than the alternative. the impella cassette is sequestered for evaluation but the implanted line with cracked side arm is still with the patient. manufacturer response for temporary non-roller type left heart support blood pump., impella (per site reporter) they were on the phone instructing staff how to mitigate risks and move forward.
35623 | 1220648-2023-02118 | 16885661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16885661&pc=OZD | 02/06/2023 | 06/15/2023 |  | 05/08/2023 |  |  | 02/06/2023 | 05/08/2023 |  |  | 09/30/2024 |  | 05/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281103 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2023 |  |  |  |  | A 62-year-old male presented for peripheral venoarterial extracorporeal membrane oxygenation (va ecmo) and distal reperfusion catheter was placed. the patient was experiencing a lack of dopplerable pulses in the impella limb which lead to the use of antegrade sheath. after the distal perfusion catheter was place, the patient's blood flow returned. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of the limb ischemia was most likely patient condition. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35624 | 1220648-2023-02119 | 16885690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16885690&pc=OZD | 02/06/2023 | 10/15/2024 |  | 05/08/2023 |  |  | 02/06/2023 | 05/08/2023 |  |  | 09/30/2024 |  | 05/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274605 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | An 81-year-old female presented for impella support. during removal of the device, arterial damage was found and required a patch to repair. the physician believes the artery was damaged during initial access for placing the device. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. additionally, groin site started bleeding with large hematoma noted and the team was consulted for surgical groin management. additionally, purse string suture was placed to control the bleeding. patient survived the procedure. |  |  |  |  | The investigation into the vascular damage has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that use issue was the most likely cause of the vascular damage. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation...¿ | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use related to this event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | Device was not returned. based on the clinical account, the cause of the vascular damage was due to damage of the artery during initial access due to the access technique. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
35625 | 1220648-2023-02121 | 16887125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16887125&pc=OZD | 04/08/2023 | 07/11/2023 |  | 05/08/2023 |  |  | 04/08/2023 | 05/08/2023 |  |  | 07/31/2024 |  | 05/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2023253080 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2023 |  |  |  |  | The complainant reported a 69-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during initial impella support, it was noted that the patient experienced a decrease in bilateral distal pulses. the pump was explanted as a result of the noted issue and a fogarty catheter was utilized to perform a clot aspiration for an observed thrombus. the patient was subsequently escalated to an impella 5.5 pump for ongoing support. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on the clinical details, the cause of the limb ischemia was most likely due to biomaterial. the failure mode will be monitored and trended.
35626 | 1220648-2023-02122 | 16887641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16887641&pc=OZD | 04/08/2023 | 01/18/2024 |  | 05/08/2023 |  |  | 04/08/2023 | 05/08/2023 |  |  | 12/31/2024 |  | 05/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC |  |  | US |  | 2024305722 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Obstruction of Flow | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2023 |  |  |  |  | The complainant reported a 46-year-old female patient presenting for cardiomyopathy. during impella support, it was noted that a purge system blocked alarm was triggered and the pump stopped unexpectedly. the pump was able to restart, however, saturated flows were noted and both purge system blocked and purge system not recognized alarms were identified. the pump was scheduled to be replaced as a result of the failure. there was no patient harm resulting from the failures. | [5 empty] | The investigation into the pump stop has been completed since the original report was filed. data logs showed the purge pressure gradually increased by about 100 mmhg over a 6-day period before the pump was disconnected from the first console. one hour later, the pump was connected to the second console. immediately upon starting the pump with the second console, high purge pressure and flow saturation occurred. about 18 hours later, the motor current rapidly increased following a motor current spike until the pump stopped with the "impella stopped - motor current high" alarm. visual inspection of the returned product revealed the presence of biomaterial in the purge gap and blood in the purge lumen from the motor to the patient cable. in conclusion, it was determined that the cause of the pump stop was blood ingress most likely due to sustained lack of purge during a console change. the root cause of the pump stop was blood ingress most likely due to sustained lack of purge during a console change. | The device was returned and an evaluation is pending. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results. as noted in the impella ifu: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35627 | 1220648-2023-02123 | 16892355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16892355&pc=OZD | 01/28/2023 | 05/25/2023 |  | 05/08/2023 |  |  | 01/28/2023 | 05/08/2023 |  |  | 09/30/2024 |  | 05/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023279629 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2023 |  |  |  |  | A 31-year-old male with acute myocardial infarction (ami) and cardiogenic shock (cgs) presented with damage sustained to the patient's impella inserted artery after patient sat up in bed and vomited. the pump was removed, artery was repaired, and a new impella was placed in the opposing artery. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient movement was the most likely cause of the access site bleeding. the mode of failure will be tracked and trended. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿
35628 | 1220648-2023-02124 | 16892362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16892362&pc=OZD | 01/28/2023 | 06/07/2023 |  | 05/08/2023 |  |  | 01/28/2023 | 05/08/2023 |  |  | 10/31/2024 |  | 05/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024286760 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2023 |  |  |  |  | A 42-year-old male with acute myocardial infarction (ami)/cardiogenic shock (cgs) presented with lack of distal perfusion that prompted an antegrade sheath. vascular surgery also performed a cutdown to remove the impella device, extract a clot, and perform a fasciotomy. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. limb ischemia ifu quote: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instruction for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35629 | 1220648-2023-02125 | 16892372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16892372&pc=OZD | 01/31/2023 | 08/15/2023 |  | 05/08/2023 |  |  | 01/31/2023 | 05/08/2023 |  |  | 09/30/2024 |  | 05/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281617 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2023 |  |  |  |  | A 60-year-old male with high risk percutaneous cardiac insertion (hrcpi) proceeded to use the impella device despite the patient having poor circulation issues and a previous toe amputation. pulses were not present pre/post impella insertion. the decision was made to remove the device due to ischemia, rather than using side port of repositioning sheath to check distal flow. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was most likely patient condition related. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35630 | 1220648-2023-02114 | 16880812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16880812&pc=OZD | 04/07/2023 | 06/12/2023 |  | 05/05/2023 |  |  | 04/07/2023 | 05/05/2023 |  |  | 02/29/2024 |  | 05/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023193698 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2023 |  |  |  |  | The complainant reported a 62 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. the pump stopped after approximately 10 days. attempts were made to restart the pump, to no avail. the device was removed and the patient was pharmacologically supported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure, which triggers the ¿purge pressure high¿ alarm message, could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported malfunction of pump stop has been completed since the original report was filed. the medical center returned the product for benchtop analysis. the returned pump was visually inspected and a foreign blue catheter was observed tangled in the cannula and pinched against the impeller. the root cause of the pump stop was an interaction with another device resulting in ingested blue catheter entraining the impeller. the failure modes will be monitored and trended.
35631 | 1220648-2023-02115 | 16881624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16881624&pc=OZD | 04/08/2023 | 06/27/2023 |  | 05/05/2023 |  |  | 04/08/2023 | 05/05/2023 |  |  | 11/30/2024 |  | 05/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED INC. |  |  | US |  | 2024315180 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2023 |  |  |  |  | The medical center had an impella cp placed at the femoral artery for support. on the second day of support the team observed signs of limb ischemia. the team placed a reperfusion sheath to restore flow to the impella limb that had a cold foot. after 2 more days of support the team made choice to withdraw care and the patient expired. | [5 empty] | The medical center discarded all product at the time of care withdrawal. there can be no benchtop analysis of the pump to root cause of ischemia. there will be a final report filed should more information be shared that leads to root cause. failure mode will be monitored and trended. of note in the ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis of the limb ischemia; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of the limb ischemia issue was most likely related to patient condition. the failure mode will be monitored and trended.
35632 | 1220648-2023-02116 | 16881631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16881631&pc=OZD | 04/08/2023 | 10/14/2024 |  | 05/05/2023 |  |  | 04/08/2023 | 05/05/2023 |  |  | 10/31/2024 |  | 05/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2024288547 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2023 |  |  |  |  | The medical center had an impella cp support for the stemi dtu protocol and support. the patient sat up and moved her leg with vomiting. the patient developed a hematoma at the site. decision made to treat the patient with blood products, inotropic support and fluid bolus. the team made decision to explant the pump and consult vascular surgery. the surgeon assessed the vessel and placed 3 percloses. | It was also reported the patient experienced vascular damage while on support. due to the patient extensive bleeding, the decision was made to take the patient to surgery and perform an angiogram and repair the artery. once the angiogram was performed, the surgeon decided to turn off the device and place percloses along with manual pressure which hemostasis was achieved. |  |  |  |  | The medical center discarded all the impella product at the time of explant. no benchtop analysis to root cause is possible. the known patient movement could have caused/contributed to the bleed and hematoma. should more information be shared that leads to root cause, a final report will be forthcoming. the failure mode will be monitored and trended. of note in the ifu: "assess access site for bleeding and hematoma.¿ | Additional information for the initial mfr was provided in sections b5,d6a, d6b, h1, and h6. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. the failure modes will be monitored and trended. instructions for use for the related event are as follows: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
35633 | 1220648-2023-02102 | 16867250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16867250&pc=OZD | 02/18/2023 | 07/14/2023 |  | 05/04/2023 |  |  | 04/05/2023 | 05/04/2023 |  |  | 10/31/2024 |  | 05/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2024284969 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 78-year-old white male admitted with chest pain and shortness of breath. pci has been performed without impella support. subsequently, an impella cp has been placed. noted slow flow to lower extremity. external bypass performed. distal pulses present via doppler. patient was sent to unit to rest and recover. | [5 empty] | The impella device has not been received by the manufacturer for investigation. should any new information be returned, a supplemental report will be filed. of note the ifu states the following: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the limb ischemia since a bypass resolved the issue and there was no mention of a thrombus. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35634 | 1220648-2023-02103 | 16867706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16867706&pc=OZD | 04/06/2023 | 06/15/2023 |  | 05/04/2023 |  |  | 04/06/2023 | 05/04/2023 |  |  | 12/31/2024 |  | 05/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED INC. |  |  | US |  | 2024285908 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2023 |  |  |  |  | The medical center had an impella 5.5 support ongoing for a patient escalated from the cp support. the patient was escalated prior to bypass surgery/cabg. on day 15 the pump stopped and was explanted. the pump was able to to be restarted after the stop, that had occurred while patient was vomiting, but still explanted. | [5 empty] | The investigation is on-going at this time. the impella 5.5 pump has been returned for benchtop analysis. upon completion of the investigation, a final report will be filed. of note in the ifu: "impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported pump stop has been completed since the original report was filed. the medical center returned the device for evaluation. in addition, the data logs were reviewed to determine the root cause. data logs showed the pump stopped immediately upon the occurrence of the "impella stopped - motor current high" alarm. in the 5 minutes leading up to the pump stop, a sudden motor current spike occurred (with a simultaneous dip in motor speed), immediately followed by increasingly elevated motor current and significant flow saturation with no change in p-level. after 2 minutes of stoppage, the pump successfully restarted and the alarm resolved. visual inspection of the returned product revealed a large mass of biomaterial on and around the impeller. no other abnormalities were found. when the pump was activated, the biomaterial was ejected out of the outflow and the device ran with no issues. in conclusion, it was determined that the cause of the pump stop was ingested biomaterial. the evaluation revealed the cause of the pump stop was ingested biomaterial. the failure modes will be monitored and trended.
35635 | 1220648-2023-02107 | 16873332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16873332&pc=OZD | 04/03/2023 | 07/26/2023 |  | 05/04/2023 |  |  | 04/03/2023 | 05/04/2023 |  |  | 09/30/2024 |  | 05/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED INC. |  |  | US |  | 2023273524 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2023 |  |  |  |  | The medical center had an impella cp inserted for hemodynamic support of a patient with cardiac history and electrophysiology intervention. the pump was explanted after just over 6 hours of support due to an access site bleed and associated hematoma. the team had attempted to achieve hemostasis with a femostop and manual hold. the pump was explanted, 3 units of rbcs, 1 unit of ffp, and kcentra, and the team had vascular surgery close the groin site. | [5 empty] | The medical center discarded all impella pump product at the time of explant. no benchtop analysis to root cause is possible. should more information be shared that leads to root cause, a final report will be forthcoming. the failure mode will be monitored and trended. of note the ifu states: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the access site bleeding was most likely high patient act values. the failure mode will be monitored and trended.
35636 | 1220648-2023-02109 | 16873334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16873334&pc=OZD | 04/05/2023 | 05/16/2023 |  | 05/04/2023 |  |  | 04/05/2023 | 05/04/2023 |  |  | 11/30/2024 |  | 05/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2024300816 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2023 |  |  |  |  | The medical center had placed an impella cp via the femoral access for a patient admitted with multiple cardiac and systemic comorbidities. the patient had small vessels and after placement of the impella the team placed a infusion catheter to have flow to the impella limb. additionally there was an access site ooze and associated hemoglobin drop that prompted prbc infusion to the patient. the family made choice to withdraw care. | [5 empty] | The medical center discarded all impella pump product at the time of explant when the family made the choice to withdraw care. root cause can not be determined, though should more information be shared which leads to root cause, a final report will be forthcoming. of note in the ifu: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Since the original report was filed, the investigation into the ischemia and access site bleeding has concluded. the product was not returned for analysis, and only the clinical details were reviewed for root cause determination. the root cause of the ischemia was patient condition. the vessels were of small diameter. the root cause of the bleeding was most the high patient's act values. the failure mode will be monitored and trended.
35637 | 1220648-2023-02108 | 16873335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16873335&pc=OZD | 04/05/2023 | 05/04/2023 |  | 05/04/2023 |  |  | 04/05/2023 | 05/04/2023 |  |  | 12/31/2024 |  | 05/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED INC. |  |  | US |  | 2024304629 |  | 10000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Transient Ischemic Attack | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/05/2023 |  |  |  |  | The medical center had an impella 5.5 support ongoing for over a month bridging to heart transplant, when there was an observation made of a tia from a embolic clot. the event did prompt the team to explant the embolus in the ir lab. there has been no noted lasting harm or injury/sided deficits. the pump remained on for support for another day, was weaned, and explanted. at the time of explant the team noted no biomaterial attached to the impella pump. the contribution of the impella to the tia could not be determined or denied. | [5 empty] | The medical center discarded all impella pump product at the time of explant. no benchtop analysis to root cause is possible. should more clinical information be shared that leads to root cause, a final report will be forthcoming. failure mode will be monitored and trended. of note the ifu states: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35638 | 1220648-2023-02110 | 16873337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16873337&pc=OZD | 04/26/2023 | 07/18/2023 |  | 05/04/2023 |  |  | 04/26/2023 | 05/04/2023 |  |  | 10/31/2024 |  | 05/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED INC. |  |  | US |  | 2024290431 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2023 |  |  |  |  | The medical center had an impella cp support ongoing for 2 days when the team observed signs of limb ischemia in the cp limb. there was no pulse and so the pump was explanted and surgery consulted. patient will go to the or for cut down artery repair. | [5 empty] | The medical center did not return the impella pump or introducer sheath for any benchtop analysis and testing. the root cause of the reported ischemia is not known at this time from the clinical report alone. should more information be shared, and a root cause determined, a final report will be forthcoming. the failure mode will be monitored and trended. of note the ifu states: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the limb ischemia was most likely patient condition related. the failure mode will be monitored and trended.
35639 | 1220648-2023-02111 | 16873339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16873339&pc=OZD | 04/05/2023 | 07/27/2023 |  | 05/04/2023 |  |  | 04/05/2023 | 05/04/2023 |  |  | 11/30/2024 |  | 05/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2024317947 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2023 |  |  |  |  | The complainant reported a 57-year-old female patient presenting for high risk percutaneous coronary intervention. during impella support, it was noted that the patient developed ischemia. the left femoral artery distal to the impella was observed to be mottled and a doppler pulse could not be obtained. the pump was subsequently explanted as a result of the reported condition. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the root cause of the limb ischemia was most likely patient condition related. the failure mode will be monitored and trended.
35640 | 1220648-2023-02095 | 16859688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16859688&pc=OZD | 04/04/2023 | 06/16/2023 |  | 05/03/2023 |  |  | 04/04/2023 | 05/03/2023 |  |  | 12/31/2024 |  | 05/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2024316431 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Collapse | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 66-year-old male patient admitted with nstemi, chf, htn, known cad. he had refused medical intervention in the past but agreed for cabg. three vessel high risk surgery performed with support from an iabp. due to lcos, decision was made to change the iabp for an impella cp cardiac pump. the impella has been placed successfully. during removal of peelaway sheath and advancing repositioning sheath, the pump has been pushed completely into the left ventricle, causing the pump to kink. attempts to straighten the pump have not been successful. the impella device has been pulled back to the access site and removed using a surgical cutdown of the right femoral artery. a new impella cp could not be placed via the left femoral artery due to calcified left aorta. therefore, another iabp has been placed. the patient survived the incident. | [5 empty] | The impella device has been returned to the manufacturer. when the investigation is complete, a supplemental report will be filed. | The investigation into the reported positioning issues has been completed since the original report was filed. the medical center returned the impella products for benchtop analysis. no abnormalities were found with the returned product. based on the clinical account, the root cause of the positioning issues was due to the migration of the impella into the ventricle when the 14fr was peeled away and the gwru was advanced. this is a result of improper technique. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35641 | 1220648-2023-02096 | 16859715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16859715&pc=OZD | 04/03/2023 | 08/11/2023 |  | 05/03/2023 |  |  | 04/03/2023 | 05/03/2023 |  |  | 01/31/2025 |  | 05/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024307418 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 60-year-old white male presented with heart failure and cardiogenic shock. iabp has been placed and cabg procedure planned. post cabg patient deteriorated on iabp and was escalated to central va ecmo. an impella 5.5 has been inserted with a difficult delivery. the attending physician noted that the insertion took an unusually high force to get the pump across the aortic valve and that the catheter had to be torqued to move the impella inlet away from the mitral apparatus. there were low purge flow alarms and the pump stopped. the impella 5.5 has been removed. subsequently, the patient was stable on central va ecmo. | [5 empty] | The impella device has been returned to the manufacturer. when the investigation is complete, a supplemental report will be filed. as noted in the impella ifu: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the pump stop was blood ingress due to purge lumen kinking along the length of the catheter.
35642 | 1220648-2023-02097 | 16861155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16861155&pc=OZD | 04/04/2023 | 09/19/2024 |  | 05/03/2023 |  |  | 04/04/2023 | 05/03/2023 |  |  | 01/31/2025 |  | 05/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024296619 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | 51-year-old male white patient with ventricular septal defect has been treated for cardiogenic shock. peripheral va ecmo has been place. an attempt was made to place an impella 5.5 cardiac pump via right axillary cutdown. due to the angulation of the ascending aorta physician had difficulty trying to manipulate wire/catheter into the left ventricle. decision made to do left axillary cutdown, resulting in successful impella 5.5 insertion. patient has been improving. 14 days after implantation, patient status slowly declined after the patient coughed. tee revealed a type ii aortic dissection. it is undetermined if the dissection is related to the impella device. two days later care was withdrawn. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was has not been received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the vascular damage has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the vascular damage was most likely use related since resistance was felt during the insertion process due to the difficult patient anatomy. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
35643 | 1220648-2023-02098 | 16862499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16862499&pc=OZD | 04/03/2023 | 07/19/2023 |  | 05/03/2023 |  |  | 04/03/2023 | 05/03/2023 |  |  | 10/31/2024 |  | 05/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2024287188 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2023 |  |  |  |  | The complainant reported a 79 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. an access site bleed developed after removal of peel away sheath and advancement of repositioning sheath. the bleed could not be controlled with manual pressure, therefore, the pump was removed and vascular surgery was required to repair/close the site. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿ | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the bleeding was most likely a lack of vessel recoil onto the sheath. the failure mode will be monitored and trended.
35644 | 1220648-2023-02099 | 16862598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16862598&pc=OZD | 02/20/2023 | 09/26/2024 |  | 05/03/2023 |  |  | 02/20/2023 | 05/03/2023 |  |  | 09/30/2024 |  | 05/03/2023 |  | PUMP 371 14F LT CMR SET |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023281403 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Perforation | [11 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/20/2023 |  |  |  |  | A 72-year-old male presented for impella support in the catheterization lab during an st-elevation myocardial infarction (stemi). pericardiocentesis was performed and a ventricle perforation was discovered. emergent repair was required. the repair was completed, however, the patient expired the next day following continued bleeding for chest tubes and profound cardiogenic shock. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. instructions for use: related to this event: impella 5.5 & cp: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ impella 5.5, cp, rp, & rp flex: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The following fields have been corrected. d1. corrected brand name. e1. doctor added. remove abiomed from first name. remove physician from middle name. added phone number. g1. corrected manufacturer name, street address, city, state, country and fax number. h3. was the device evaluated by the manufacturer? added no. if the device was not evaluated... added other (code unspecified, describe in h10). | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-01954 was provided in sections b2, b5, d6a, d6b, h1, and h6.
35645 | 1220648-2023-02100 | 16862999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16862999&pc=OZD | 02/27/2023 | 08/09/2023 |  | 05/03/2023 |  |  | 02/27/2023 | 05/03/2023 |  |  | 01/31/2025 |  | 05/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024309850 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2023 |  |  |  |  | The user facility reported a 53-year-old male presented for impella support. while the physician was removing the impella, some resistance was felt and the impella came out through the graft. the impella was removed. however, massive bleeding ensued. graft was clamped, but bleeding continued. the patient was sedated and put under and chest re-opened. bleeding brought under control with clamp. the physician found that there was a tear in the innominate artery about 2 cm from graft anastomosis. the physician was not sure what caused the tear during removal. patient was stabilized. chest closed and sent to icu under 2 pressors. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. the following statements are found in the instructions for use in the related event: impella 5.5 & cp ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events ((b)(6)): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the vascular damage has been completed since the original report was filed. no benchtop analysis was possible to determine the root cause. since the medical center did not return the impella device, the root cause of the vascular damage was not determined. the failure mode will be monitored and trended.
35646 | 1220648-2023-02101 | 16865009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16865009&pc=OZD | 04/04/2023 | 06/28/2023 |  | 05/03/2023 |  |  | 04/04/2023 | 05/03/2023 |  |  | 11/30/2024 |  | 05/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2024313511 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2023 |  |  |  |  | The complainant reported a 68 year old male patient presenting with high risk percutaneous coronary intervention. during support, the patient developed ischemia in the impella inserted limb due to physician opting to leave the 14fr introducer in the access site. an antegrade sheath was configured to restore blood flow. | [5 empty] | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis of the access limb ischemia; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of the limb ischemia issue was related to patient condition; the small size of the patient's vessels. the failure mode will be monitored and trended.
35647 | 1220648-2023-02091 | 16852639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16852639&pc=OZD | 04/02/2023 | 07/11/2023 |  | 05/02/2023 |  |  | 04/02/2023 | 05/02/2023 |  |  | 01/31/2025 |  | 05/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024313574 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An 81-year-old female who suffered from chest pain for two days presented with stemi. coronary artery disease has been found. during pci, the patient became hypotensive and required neo-synephrine medication. decision was made to insert an impella cp to provide hemodynamic support during the high-risk pci. the pci procedure has been successfully completed with the patient hemodynamically stable throughout the pci procedure. decision made to leave impella in overnight. after removal of peel-away-sheath and placement of repositioning sheath, the patient became unstable. a large hematoma was found in her right groin below the sheath. the impella device has been removed and replaced with a fogarty balloon to maintain hemostasis while the patient has been prepared for femoral cutdown. the patient required cpr, vascular repair of the femoral injury and transfusion of multiple units of blood and platelets. the patient has been in critical condition afterwards and has been transferred to icu. the attending physician stated he is unsure of the cause and noted patient movement as a contributing factor. he assessed the adverse event as unlikely due to incorrect peel-away process but possibly related to perclose device. | [5 empty] | The impella device has not been received by the manufacturer for investigation. should any new information be returned, a supplemental report will be filed. as noted in the impella ifu: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on the clinical details, the root cause of the access site bleeding was most likely patient condition. the failure mode will be monitored and trended.
35648 | 16852687 | 16852687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16852687&pc=OZD | 04/19/2023 | 04/21/2023 | 04/21/2023 | 05/02/2023 | 04/21/2023 | 05/02/2023 |  | 05/02/2023 |  |  |  |  | 05/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SET WITH SMARTASSIST |  |  | OZD |  |  | IMPELLA CP SET WITH SMARTASSIST |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003 |  | No |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | No | Risk Manager | [6 empty] | 28470 DA |  | Male | 86 KG |  | White | Injury | Yes | No | Physical Resistance/Sticking | [6 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Hypoxia | [9 empty] | Hospitalization | Required Intervention | [23 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | This elderly male with pmh (past medical history) of hypertension presented to the ed (emergency department) with complaints of worsening shoulder pain and back pain. ekg revealed st elevation mi (myocardial infarction) in rca (right coronary artery) distribution. pt (patient) was taken to the ccl (cardiovascular catheterization lab). a coronary arteriography was performed and revealed totally occluded lad (left anterior descending artery) in its proximal segment. the distal left main revealed 60% stenosis, ostial circumflex has 60% stenosis, and the rca was totally occluded. following a cardiovascular consultation, the decision was made to proceed with rca intervention, stent placement and hemodynamic support with left ventricular device and reassess need for urgent surgery versus elective surgery for lad distribution based upon the patient¿s clinical status in the hospital. the right coronary artery was successfully opened with a balloon, but unable to advance the stents because of calcification. two drug-eluting stents position, one in distal rca, other to the rv (right ventricular) branch and the second one is proximal rca before the rv branch and its proximal segment. excellent flow noted in entire lad distribution, however the patient continued to have persistent st elevations in all the leads and borderline low blood pressure. the decision was made to proceed with a left ventricular assist device placement. utilizing the existing arterial sheath, which was exchanged to a 9-french impella sheath; and the left ventricle was accessed with a pigtail, and this was exchanged with impella wire and over the wire. the impella device was advanced; however, it was unable to advance to position in left ventricle. under fluoroscopy, the impella device was noted to have a kink in the distal segment of port and unable to advance the balloon and wire as a single unit. attempts were made to unkink the impella device, which was unsuccessful. the patient was becoming hypotensive and shortness of breath and cough. the anesthesia team was called in to intubate the patient. the patient was breathing over the ventilator and was very stable on ventilator support. the arteriotomy site was bleeding and the hypotension was corrected with vasopressors, which include vasopressin, levophed and epinephrine. the patient developed a vasovagal bradycardia, which was managed with a temporary pacemaker and subsequently maintained in sinus rhythm. albumin and blood was given. a vascular surgery consultation was obtained and reviewed the impella status when in the groin and it was felt for him to take it out using a snare. a snare was used in an attempt to remove the impella device. the kink was partially released and subsequently the vascular surgeon was able to remove from the groin as a single unit and arteriotomy site was compressed with a femostop and complete hemostasis achieved. he received a unit of blood and a liter of albumin and fluids and his foley catheter was placed. the patient was then taken to the or for repair of the arteriotomy site and to introduce the impella device through the subclavian artery approach. the patient was subsequently transferred to the cvicu (cardiovascular intensive care unit). manufacturer response for impella cp, impella cp set with smart assist (per site reporter) clinical representative was present in the ccl during the event in question. the regional manager was contacted by rep during the event in question.
35649 | 1220648-2023-02005 | 16854349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16854349&pc=OZD | 03/12/2023 | 06/27/2023 |  | 05/02/2023 |  |  | 03/12/2023 | 05/02/2023 |  |  | 11/30/2024 |  | 05/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2024303001 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 42-year-old male in the us presented with nstemi in cardiogenic shock. an impella cp has been inserted. patient developed leg ischemia below impella cp femoral access site. ct scan showed thrombus around impella repositioning sheath. impella cp has been replaced for axially placed impella 5.5. the cp was explanted and thrombectomy was performed in iliac and femoral arteries. a moderate sized thrombus was removed, as well as tissue that appeared to be atheroma. surgeon suspects an atheromatous plaque was dislodged with impella access and thrombus formed at the lesion. repair of vessel was successful with distal pulses restored. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. the impella ifu notes the following statements, in relationship to the reported symptom: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis of the access limb ischemia; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of the limb ischemia issue was verified to be the presence of biomaterial in the patient vessel, since the thrombectomy resolved the limb ischemia. the failure mode will be monitored and trended.
35650 | 1220648-2023-02092 | 16855590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16855590&pc=OZD | 04/11/2023 | 05/02/2023 |  | 05/02/2023 |  |  | 04/18/2023 | 05/02/2023 |  |  | 08/31/2024 |  | 05/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC |  |  | US |  | 2023267266 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Ejection Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/18/2023 |  |  |  |  | The complainant reported a 58-year-old female patient presenting for cardiomyopathy. following initial difficulty with placing the pump, the device was withdrawn and a part of the intima of the blood vessel peeled off. an artificial v-plasty was performed to treat the injury. direct access was subsequently utilized to deliver the pump and provide support. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, vascular injury is a known potential adverse event associated with impella support: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
35651 | 1220648-2023-02094 | 16856970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16856970&pc=OZD | 04/11/2023 | 07/06/2023 |  | 05/02/2023 |  |  | 04/14/2023 | 05/02/2023 |  |  | 10/31/2024 |  | 05/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023276455 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 26 year old male was presented for impella cp support. an antegrade sheath was used because the patient was prone to lower limb ischemia on the impella side. despite this, the pump required premature removal and escalation to an impella 5.0 in the axillary artery. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. instructions for use related to this event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis; only the clinical report was evaluated. based on the clinical details, the cause of the limb ischemia was most likely related to the patient condition. the failure mode will be monitored and trended. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35652 | 1220648-2023-02093 | 16857187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16857187&pc=OZD | 04/11/2023 | 06/15/2023 |  | 05/02/2023 |  |  | 04/11/2023 | 05/02/2023 |  |  | 02/28/2025 |  | 05/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2024313701 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2023 |  |  |  |  | A 68-year-old female presented for cardiomyopathy. after removal of the impella device, pericardial effusion and cardiac tamponade were discovered. a ventricle perforation was identified and repaired. blood products were also delivered. | [5 empty] | The impella device was not received and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. instructions for use related to this event are as follows: impella 5.5 & cp: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ impella 5.5, cp, rp, & rp flex: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation into the ventricle perforation has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that improper positioning was the most likely cause of the ventricle perforation. the failure mode will be monitored and trended.
35653 | 1220648-2023-02086 | 16847816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16847816&pc=OZD | 04/01/2023 | 06/06/2023 |  | 05/01/2023 |  |  | 04/01/2023 | 05/01/2023 |  |  | 01/31/2025 |  | 05/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 380 (US) |  |  | ABIOMED, INC |  |  | US |  | 2024313604 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2023 |  |  |  |  | The complainant reported a 63-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. upon attempted delivery of the impella pump, the cannula folded over and kinked. the catheter was pulled back, however, the cannula remained kinked and the folded device came back across the valve. a vascular surgeon was subsequently required to remove the pump and repair the vessel with a surgical graft. another pump was then implanted for ongoing support. | [5 empty] | The investigation into the reported delivery issues and vascular damage has been completed since the original report was filed. the medical center returned the impella product for benchtop analysis. the root cause for delivery issue and resulting vascular damage was determined to be user error. the impella device went deep in the left ventricle upon advancing of the repo unit. catheter positioning is intended to be held and maintained upon advancing of the repo unit. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The pump was returned for further evaluation. a visual inspection was performed and it was noted that the catheter was cut and kinked, the cannula was kinked, the repo sheath was damaged, and a biomaterial was found in the outlfow/inflow. an investigation into the evaluation is ongoing. upon completion, a supplemental report will be submitted with any relevant information pertaining to the findings. as noted in the impella ifu, vascular injury is a known potential adverse event associated with impella support: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
35654 | 1220648-2023-02087 | 16850428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16850428&pc=OZD | 04/12/2023 | 07/11/2023 |  | 05/01/2023 |  |  | 04/17/2023 | 05/01/2023 |  |  | 01/31/2025 |  | 05/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2024312448 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Necrosis | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2023 |  |  |  |  | The complainant reported an 87-year-old female patient presenting for pre-op placement of the impella pump. during support, the patient developed limb ischemia and a thrombus was discovered. the pump was explanted as a result of the ischemia and the thrombus was removed with a fogarty catheter. however, a follow up ct was performed and multiple thrombi were discovered. the patient developed unilateral kidney necrosis and dialysis was initiated. thrombolysis treatment was maintained to treat the condition. | [5 empty] | The investigation into the limb ischemia and the renal/kidney failure has been completed since the original report was filed. the medical center in japan did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on the clinical details, the root cause of the limb ischemia was biomaterial as the thrombus was confirmed and blood flow decreased in both lower extremities, and the root cause of the kidney failure was the patient condition related because of thrombolysis and pre-existing stenosis in both lower extremities, which could have resulted in blood stasis and thrombus formation and let to renal necrosis and dialysis was initiated the failure modes will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35655 | 1220648-2023-02089 | 16850470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16850470&pc=OZD | 04/02/2023 | 05/01/2023 |  | 05/01/2023 |  |  | 04/02/2023 | 05/01/2023 |  |  | 01/31/2025 |  | 05/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 380 (US) |  |  | ABIOMED INC. |  |  | US |  | 2024314164 |  | 100080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/02/2023 |  |  |  |  | The medical center had an impella cp pump placed as a replacement for the first pump that had low flows. the 2nd pump could not advance and fully deliver. when the pump was removed the team observed an iliac artery dissection. the dissection was treated by iliac artery angioplasty. | [5 empty] | The medical center discarded all impella pump product at the time of the explant. no benchtop analysis to root cause is possible. should more information be shared that leads to root cause of the injury, a final report will be be forthcoming. of note in the ifu: ¿potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation...¿ ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿
35656 | 1220648-2023-02088 | 16850472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16850472&pc=OZD | 03/27/2023 | 05/01/2023 |  | 05/01/2023 |  |  | 03/27/2023 | 05/01/2023 |  |  | 11/30/2024 |  | 05/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED INC. |  |  | US |  | 2024313255 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/27/2023 |  |  |  |  | The us complainant had a patient admitted with an acute myocardial infarction. the plan was for cardiac bypass (cabg) and aortic valve repair (avr). the impella cp and ecmo were placed in the or to give hemodynamic support. the impella limb had signs of ischemia and so the surgeon had an antegrade sheath placed. the need for the intervention to perfuse the limb prompts reporting. the pump remained on for 7 days beyond the intervention for the limb. | [5 empty] | The medical center has not returned any impella pump product for analysis or benchtop testing. no root cause can be determined with the clinical information alone shared to date. should more be shared, and root cause determined, a follow-up report will be forthcoming. of note the ifu states: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35657 | 1220648-2023-02090 | 16850473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16850473&pc=OZD | 04/02/2023 | 07/26/2023 |  | 05/01/2023 |  |  | 04/02/2023 | 05/01/2023 |  |  | 12/31/2024 |  | 05/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 380 (US) |  |  | ABIOMED INC. |  |  | US |  | 2024298228 |  | 100080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2023 |  |  |  |  | The medical center had an impella cp support of a patient with a diagnosis of acute myocardial infarction. the cp placement via the sheath (to the left iliac artery) caused a vascular dissection. the dissection was treated by the vascular surgeon and stenting of the iliac. | [5 empty] | The investigation into the vascular damage has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the dissection was use related since a non-indicated guide wire was used and resistance was felt during delivery. the failure mode will be monitored and trended. | The medical center did not return for investigation any of the impella product for any benchtop analysis or testing. the clinical report did not furnish details regarding the vascular dissection. should any further details be forthcoming which leads to root cause, a final report will be forthcoming. of note the ifu states: ¿potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation...¿ ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿
35658 | 1220648-2023-02080 | 16832176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16832176&pc=OZD | 03/30/2023 | 04/28/2023 |  | 04/28/2023 |  |  | 03/30/2023 | 04/28/2023 |  |  | 10/31/2024 |  | 04/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023283357 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hypervolemia | Vascular Dissection | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/30/2023 |  |  |  |  | The complainant reported a 76-year-old male patient presenting for impella support. following impella placement, bleeding at the access site was noted. vascular intervention was required but resulted in an additional tear. surgical repairs were then performed on the site, however, the patient became hypovolemic. blood and fluids were administered, but hemostasis could not be maintained. the patient passed away after sixty minutes of cpr. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, bleeding, vascular injury, and death are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35659 | 1220648-2023-02081 | 16833091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16833091&pc=OZD | 03/20/2023 | 10/07/2023 |  | 04/28/2023 |  |  | 03/30/2023 | 04/28/2023 |  |  | 11/30/2024 |  | 04/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023278631 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 33-year-old female presented with peripartum cardiomyopathy and received hemodynamic support by an impella cp cardiac pump. the patient was transferred from (b)(6) university hospital to (b)(6) heart diseases center. the patient suffered from vascular damage at the impella access site which has been repaired by surgery. the patient has been stable afterwards. | [5 empty] | The investigation into the low pump flow and the vascular damage has been completed since the original report was filed. since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the low pump flow and the vascular damage was unable to be determined because no product was returned for evaluation and no data logs, insufficient clinical information was provided. the failure mode will be monitored and trended. | Benchtop investigation of the returned pump will be performed after return of the device. when the investigation is complete, a supplemental report will be filed.
35660 | 1220648-2023-02082 | 16833099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16833099&pc=OZD | 03/31/2023 | 10/25/2023 |  | 04/28/2023 |  |  | 03/31/2023 | 04/28/2023 |  |  | 10/31/2024 |  | 04/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED INC. |  |  | US |  | 2023274556 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Pressure Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/31/2023 |  |  |  |  | It was reported that there were high purge pressures. tissue plasminogen activator (tpa) was attempted but the pump ultimately had to be removed to avoid a pump stop while in use. | The medical center had an impella 5.5 supporting a patient awaiting transplant for over 44 days when due to rising purge pressures the pump was removed. this was to explant prior to any chance of an unexpected pump stop. they had additionally troubleshot with tpa in purge. at the time of explant the team observed thrombosis at the graft insertion site of the pump. the team consulted with vascular surgery and conducted a thrombectomy. |  |  |  |  | Since the original report was filed the investigation has concluded. the cause of the thrombus observed at the explant of the pump was patient condition related. there was biomaterial buildup within the pump. the failure mode will be monitored and trended. d4: udi information corrected. | The investigation for the reported high purge pressures has been completed. data logs showed high purge pressure alarms triggered after a 3-hour purge pressure rise, and the condition persisted for the remainder of support. during the purge rise, pump flow became saturated and the motor current became less pulsatile, indicative of biomaterial buildup on the impeller and within the purge gap. the device was returned for evaluation. during the device evaluation, the high purge pressure was reproduced when running the purge when the biomaterial was removed from the purge gap, the pump purged successfully. therefore, it was determined that the cause of the high purge pressure was biomaterial buildup occluding the purge gap. this report is being filed as part of a retrospective review of historical records. h6-medical device problem code, component code, investigation conclusions. | The medical center has returned the impella pump product for the benchtop analysis and testing. the investigation is on-going at this time. upon completion of the investigation, a final report will be forthcoming. of note in the ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35661 | 1220648-2023-02083 | 16838364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16838364&pc=OZD | 03/30/2023 | 07/05/2023 |  | 04/28/2023 |  |  | 03/30/2023 | 04/28/2023 |  |  | 11/30/2024 |  | 04/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2024303039 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2023 |  |  |  |  | The complainant reported a 81-year-old male patient presenting for impella implant. during impella support, it was noted that the patient had a persistent access site bleed. the patient was provided nine units of blood to counteract the bleed. the patient was stable following the event. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis; only the clinical report was evaluated. based on the clinical details, the cause of the access site bleeding was use related. the failure mode will be monitored and trended. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, bleeding is a known potential adverse event associated with impella support: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35662 | 1220648-2023-02084 | 16838453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16838453&pc=OZD | 04/03/2023 | 10/14/2024 |  | 04/28/2023 |  |  | 03/31/2023 | 04/28/2023 |  |  | 11/30/2024 |  | 04/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2024312740 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/31/2023 |  |  |  |  | It was also reported the patient experienced thrombocytopenia and vascular damage while on support. systemic heparin was turned off to resolve thrombocytopenia, as well as discussion of removing heparin from the purge. the patient's hemoglobin levels dropped and blood was provided due to a gastrointestinal bleeding which resolved. | The complainant reported a 60-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. following impella placement, it was noted that the patient had a small vasculature and developed ischemia. a fem-fem bypass was subsequently performed to treat the noted ischemia. the patient's condition improved following the intervention and impella support continued. |  |  |  |  | The investigation into the limb ischemia has been completed since the original report was filed. the medical center discarded the impella product at the time of explant. no benchtop analysis of the root cause of the limb ischemia was possible; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of limb ischemia was related to patient condition, the size of the patient's vessels. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. the failure modes will be monitored and trended. instructions for use for the related event are as follows: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35663 | 1220648-2023-02085 | 16838518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16838518&pc=OZD | 03/27/2023 | 06/01/2023 |  | 04/28/2023 |  |  | 03/31/2023 | 04/28/2023 |  |  | 11/30/2024 |  | 04/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2024309981 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/31/2023 |  |  |  |  | The complainant reported a 78-year-old male patient presenting for high risk percutaneous coronary intervention. following impella placement, it was noted that the patient experienced ventricular tachycardia. amiodarone was administered to the patient via iv to treat the condition. patient support continued with the implanted impella following the intervention. | [5 empty] | The investigation into the reported ventricular fibrillation, ventricular tachycardia, atrial fibrillation, and atrial tachycardia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that underlying electrophysiologic patient condition was the most likely cause of the ventricular fibrillation, ventricular tachycardia, atrial fibrillation, atrial tachycardia occurred during device support. the failure modes will be monitored and trended. | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed.
35664 | 1220648-2023-02068 | 16822940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16822940&pc=OZD | 03/28/2023 | 08/09/2023 |  | 04/27/2023 |  |  | 03/28/2023 | 04/27/2023 |  |  | 01/31/2024 |  | 04/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023202526 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 72-year-old patient in switzerland received hemodynamic support from an impella 5.0 heart pump to give the heart some rest prior to a high-risk surgery for treatment of a ventricular septal defect (vsd). a perforation of the left ventricle at the rear wall has been noted, requiring surgical closure of the ventricular perforation. the impella 5.0 was removed with the patient being critical. the patient was subsequently supported by va-ecmo and stable. | [5 empty] | The investigation into the ventricle perforation has been completed since the original report was filed. the user facility has returned the impella device, and the benchtop analysis revealed that the root cause of the lv perforation was most likely improper positioning of the device since a kink was observed on the returned device indicative of an impella position deep in the ventricle. | The device has been returned for investigation. when the investigation is complete, a supplemental report will be filed.
35665 | 1220648-2023-02070 | 16827503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16827503&pc=OZD | 03/28/2023 | 04/27/2023 |  | 04/27/2023 |  |  | 03/29/2023 | 04/27/2023 |  |  | 12/31/2024 |  | 04/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024299895 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/29/2023 |  |  |  |  | The complainant reported a 61 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, an access site bleed/hematoma developed that prompted vascular surgery to remove/repair. the hematoma impacted lower limb perfusion and required use of an antegrade sheath. | [5 empty] | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35666 | 1220648-2023-02069 | 16828135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16828135&pc=OZD | 03/29/2023 | 08/09/2023 |  | 04/27/2023 |  |  | 03/29/2023 | 04/27/2023 |  |  | 10/31/2024 |  | 04/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2024287156 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2023 |  |  |  |  | The complainant reported a 51 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support, during which, the patient developed an access site hematoma/bleed. attempts to achieve hemostasis were not successful, therefore, the pump was removed. two units of blood products were required. | [5 empty] | The investigation into the access site bleeding and the limb ischemia has been completed since the original report was filed. no benchtop analysis was possible to determine the root cause. since the medical center did not return the impella device, the root cause of the access site bleeding and the limb ischemia was not determined. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿
35667 | 1220648-2023-02072 | 16828316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16828316&pc=OZD | 06/10/2022 | 04/27/2023 |  | 04/27/2023 |  |  | 06/10/2022 | 04/27/2023 |  |  | 03/31/2024 |  | 04/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023201454 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/10/2022 |  |  |  |  | The medical center had an impella 5.5 support for 2 weeks that was successful for the patient admitted with cardiomyopathy, despite a purge sidearm leak and bypass being performed during the support. there was no harm or injury. the purge was inclusive of sodium bicarbonate. the need for the intervention in the form of a purge bypass prompts fda reporting. | [5 empty] | The medical center did return the pump and data logs for investigation. the logs did reveal purge alarms during support. visual inspection of the pump revealed a significant lengthwise crack on the sidearm yellow luer. no other damage or abnormalities were found. therefore, it was determined that the cause of the purge leak was sidearm yellow luer damage. the failure mode will be monitored and trended. there is a capa open for the failure.
35668 | 1220648-2023-02071 | 16828317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16828317&pc=OZD | 06/08/2022 | 04/27/2023 |  | 04/27/2023 |  |  | 06/08/2022 | 04/27/2023 |  |  | 12/31/2022 |  | 04/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022026898 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/08/2022 |  |  |  |  | The us complainant reported a side arm leak during support with the impella 5.5 pump. the issue occurred after more than 25 days of support. the leaking/malfunctioning pump was explanted. no further impella support was required and there was no injury sustained as a result of this issue. the malfunction is reportable to the fda. | [5 empty] | The medical center returned the impella pump for benchtop analysis. the clinical information was analyzed as well. sodium bicarbonate was used in the purge for 5 days of the support. heparin used for the remainder of the pump support. when trouble shooting measures failed the pump was explanted earlier than anticipated. this explant was due to sidearm damage. the failure mode will be monitored and trended.
35669 | 1220648-2023-02073 | 16828319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16828319&pc=OZD | 08/05/2022 | 04/27/2023 |  | 04/27/2023 |  |  | 08/05/2022 | 04/27/2023 |  |  | 03/31/2024 |  | 04/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023206389 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/05/2022 |  |  |  |  | The complainant reported a purge side arm leak with a crack found on the yellow luer of the device. this occurred after a month of support, 36+ days of support for the patient admitted with cardiomyopathy. the pump was removed and replaced with no patient harm. the explant/intervention prompts fda reporting. | [5 empty] | The medical center did return the pump product for benchtop analysis. product was visually inspected and a clear substance was noticed around the yellow luer, placed by the medical team and most likely was wax to plug up the leak. the medical team did disclose sodium bicarbonate was used during this case in the purge. the cause of the purge leak was sidearm yellow luer damage. the failure mode will be monitored and trended. there is a capa open for the failure.
35670 | 1220648-2023-02075 | 16828324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16828324&pc=OZD | 10/17/2022 | 04/27/2023 |  | 04/27/2023 |  |  | 10/17/2022 | 04/27/2023 |  |  | 05/31/2024 |  | 04/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023226504 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/17/2022 |  |  |  |  | The us complainant had an impella 5.5 support of 19 days when there was a purge sidearm leak. the leak was remedied by bypass placement. the purge had bicarb running at the time of leakage. the pump remained on for support and beyond indication, after the purge was bypassed. reportable for the intervention needed. the family made the choice to withdraw care the following day. | [5 empty] | The medical center returned some, but not all impella pump product, and the data logs. the data logs showed purge pressure alarms. visual inspection of the returned sidearm revealed the presence of significant cracks on the yellow luer. the pump had sodium bicarbonate running in the purge. it was determined that the cause of the purge leak was sidearm luer damage. the failure mode will be monitored and trended. there is a capa open for the failure.
35671 | 1220648-2023-02076 | 16828325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16828325&pc=OZD | 10/27/2022 | 04/27/2023 |  | 04/27/2023 |  |  | 10/27/2022 | 04/27/2023 |  |  | 06/30/2024 |  | 04/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023237419 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/27/2022 |  |  |  |  | The complainant reported a leak from the clear sidearm during impella 5.5 support. the pump had supported the cardiomyopathy patient for 27 days. a sidearm bypass was performed to counter the leak, however, shortly after the bypass, the waveforms began trending to a saturated state with no resolution. as a result, the pump was removed and patient support continued with an iabp. there was no harm to the patient. the intervention in the form of the purge bypass and explant prompts fda reporting. | [5 empty] | The medical center did return for analysis the impella pump and data logs. the logs confirmed the purge alarms. the inspection of the pump revealed damage to the sidearm yellow luer. no other abnormalities were observed. the cause of the purge leak was sidearm yellow luer damage. the purge had sodium bicarbonate running. the cause of the purge leak was the luer damage, and this failure will be monitored and trended. there is a capa open for the failure.
35672 | 1220648-2023-02078 | 16828340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16828340&pc=OZD | 12/02/2022 | 04/27/2023 |  | 04/27/2023 |  |  | 12/02/2022 | 04/27/2023 |  |  | 09/30/2024 |  | 04/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023257936 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/02/2022 |  |  |  |  | The medical center reported a fluid leak, on day 7, from the yellow luer during impella 5.5 support, for a patient admitted with acute myocardial infarction. a purge sidearm bypass was performed to resolve the noted issue. there was no harm to the patient and the pump remained on for support for another 6 days til explanted. the malfunction and necessary intervention in the form of the purge bypass prompts fda reporting. | [5 empty] | No product was returned for investigation. only the clinical report was evaluated for investigation. the medical team had reported upon a leak from the yellow luer. after reviewing the clinical information, it was determined that the cause of the purge leak was most likely sidearm luer damage due to the presence of sodium bicarbonate in the purge solution. the failure mode will be monitored and trended. a capa is open for the event.
35673 | 1220648-2023-02079 | 16828342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16828342&pc=OZD | 03/29/2023 | 07/05/2023 |  | 04/27/2023 |  |  | 03/29/2023 | 04/27/2023 |  |  | 09/30/2024 |  | 04/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023274020 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2023 |  |  |  |  | The complainant reported an 81-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. following impella implant, it was noted that the physician was unable to get a doppler pulse. attempts were made to place a distal perfusion catheter, however access could not be obtained. after being transferred to another unit, a 4 fr sheath was placed, but flow to the leg remained poor. as a result of the ongoing condition, the decision was eventually made to explant the pump. the patient was stable following the events. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis; only the clinical report was evaluated. based on the clinical details, the root cause of the limb ischemia was patient condition related due to her disease; patient had peripheral arterial disease (pad). the failure mode will be monitored and trended. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35674 | 1220648-2023-02074 | 16828394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16828394&pc=OZD | 08/23/2022 | 04/27/2023 |  | 04/27/2023 |  |  | 08/23/2022 | 04/27/2023 |  |  | 05/31/2024 |  | 04/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023224005 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2022 |  |  |  |  | The complainant reported a cracked purge filter during impella 5.5 support. this occurred on day 6 of the support of the patient admitted for mitraclip. the team remedied the situation by wrapping 2 iv tourniquets around the filter and securing it with tape and support continued. there was no harm to the patient. the need for the intervention prompts fda reporting for the product malfunction. | [5 empty] | The medical center did not return for analysis the impella pump product. only the clinical report was analyzed. the team remarked there had been cleanser applied to the impella due to the surgical nature of the support. based on the clinical account, the root cause of the purge leak was sidearm filter damage. this is likely a result of cleaning the sidearm. the failure mode will be monitored and trended. a capa is open for the failure.
35675 | 1220648-2023-02077 | 16828395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16828395&pc=OZD | 11/01/2022 | 10/13/2023 |  | 04/27/2023 |  |  | 11/01/2022 | 04/27/2023 |  |  | 05/31/2024 |  | 04/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023224019 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2022 |  |  |  |  | After more than 20 days of impella 5.5 support, the medical center observed a purge leak. the leak was on the purge sidearm's yellow luer lock. to troubleshoot a bypass system was put in place. alarm resolved. patient stable and successfully weaned the following day from the 5.5 pump. reportable for the intervention needed at the purge line. | [5 empty] | The investigation into the reported "purge leak - pump and high purge pressure" has been completed since the original report was filed. product was returned for investigation. the device was inspected and found the lt log confirms purge pressure low alarms. a visual inspection of the pump revealed damage to the yellow luer. the cause of the purge leak was sidearm yellow luer damage. the pump was attached to the automated impella controller (aic) and a purge cassette and the high purge pressure alarm was reproduced. the pump was purged with tergazyme and no fluid was observed exiting the purge gap. the pump would not run. the catheter was cut to look for a blockage in the purge line. dark yellow fluid was observed in the purge line proximal to the motor indicating that blood had ingressed into the motor as a result of the yellow luer leak and was causing a blockage in the line. fluid in the purge line distal to the motor was clear. | The medical center returned the pump and data logs for analysis. the logs and visual inspection of the pump revealed purge pressure alarms and damage to the yellow luer. when the pump was taken down there was an observation made of blood ingress into the pump motor. the failure mode will be monitored and trended. there is a capa open for the failure.
35676 | 1220648-2023-02060 | 16817812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16817812&pc=OZD | 03/27/2023 | 07/05/2023 |  | 04/26/2023 |  |  | 03/27/2023 | 04/26/2023 |  |  | 11/30/2024 |  | 04/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2024302424 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Entrapment of Device | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2023 |  |  |  |  | The complainant reported an 81-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, the complainant reported a pump positioning issue and a subsequent damaged leaflet. following the initial implant, it was noted that the pigtail and cannula were caught in the mitral valve apparatus. repositioning attempts were initially unsuccessful and mitral regurgitation (mr) was observed via echocardiogram. after pulling the pump far enough back in the aorta, the pigtail was finally freed from the mitral valve apparatus. the mitral regurgitation appeared to have resolved and the decision was made to explant the pump. however, the following day, the mr recurred, and it was discovered that there was damage to the anterior leaflet. mitra-clips were placed and the mr was resolved. | [5 empty] | The investigation into the heart valve damage has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the cause of the heart valve damage was improper positioning of the device. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. of note the ifu states the following: ¿impella catheter in papillary muscle if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿
35677 | 1220648-2023-02061 | 16820781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16820781&pc=OZD | 03/27/2023 | 05/26/2023 |  | 04/26/2023 |  |  | 03/27/2023 | 04/26/2023 |  |  | 10/31/2024 |  | 04/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023269555 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2023 |  |  |  |  | The complainant reported a 47 year old male patient presenting for impella support. less than 24 hours into support, a "code stroke" was called due to patient enduring an embolic stroke. support proceeded, however, an allegation was made towards the impella as a causal/contributory factor. | [5 empty] | The investigation into the reported stroke/cva has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that biomaterial due to low patient acts was the most likely cause of the stroke/cva. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35678 | 1220648-2023-02062 | 16820802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16820802&pc=OZD | 02/23/2023 | 04/26/2023 |  | 04/26/2023 |  |  | 03/27/2023 | 04/26/2023 |  |  | 12/31/2024 |  | 04/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  |  |  | 2024300722 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 03/27/2023 |  |  |  |  | The complainant reported a 51 year old male patient presenting with cardiomyopathy. damage to the patient's iliac artery was endured upon insertion of the dilator, used to insert the 14 fr introducer. a balloon tamponade was deployed, a stent was placed, and blood products were delivered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35679 | 1220648-2023-02065 | 16820808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16820808&pc=OZD | 03/29/2023 | 07/25/2023 |  | 04/26/2023 |  |  | 03/29/2023 | 04/26/2023 |  |  | 12/31/2024 |  | 04/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED INC. |  |  | US |  | 2024291950 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2023 |  |  |  |  | The medical center had an impella 5.5 support ongoing for a patient with cardiomyopathy who was awaiting heart transplant/vad. on day 9 of the support the team observed air within the echo imaging. abiomed medical affairs was consulted and troubleshooting measures taken. there has been no harm or embolization to date. the malfunction is reportable as we can not deny the causality of air within the vasculature/left ventricle. the pump remains on for support to date, on day 35 of support. | [5 empty] | The investigation into the presumed air emboli (air bubbles) observed by the medical team has concluded. the cause of the "air bubbles" and its relation to the impella were not determined since the pump performed per specification on the day of the alleged issue, as observed in data logs, and the nature of the "bubbles" could not be verified clinically. | The pump remains on for support without any noted harm or embolization of air. the causal link between the air seen on echo imaging and the use of impella can not yet been proven or disproven. upon return of the impella pump and benchtop analysis, a final report will be filed with root cause. the failure mode will be monitored and trended. of note the ifu states the following: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿infusion filter: the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿
35680 | 1220648-2023-02063 | 16820809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16820809&pc=OZD | 03/28/2023 | 04/26/2023 |  | 04/26/2023 |  |  | 03/28/2023 | 04/26/2023 |  |  | 10/31/2024 |  | 04/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 380 (US) |  |  | ABIOMED INC. |  |  | US |  | 2024291763 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/28/2023 |  |  |  |  | The medical center had an impella cp support ongoing for 2 days when the team explanted and escalated to the 5.5 pump. the premature explant was due to the team observing signs of hemolysis, which they presumed was due to the use of the impella. the ldh levels were rising but no definitive plasma free hemoglobin (pfhg) was drawn to confirm hemolysis. the 5.5 was placed and supported the patient for over 8 days. | [5 empty] | The medical center discarded the impella pump product at the time of explant. data logs were reviewed and they revealed the pump was operating to specification. the cause of the issue was not determined since product was not returned for any benchtop testing. should the medical team forward new clinical details, that lead to root cause, a final report will be forthcoming. the failure mode will be monitored and trended. of note the ifu states: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿
35681 | 1220648-2023-02064 | 16820810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16820810&pc=OZD | 03/29/2023 | 06/15/2023 |  | 04/26/2023 |  |  | 03/29/2023 | 04/26/2023 |  |  | 11/30/2024 |  | 04/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2024311735 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2023 |  |  |  |  | The medical center had a patient with peripheral arterial disease (pad) and a complicated sheath and impella cp pump delivery. the team swapped sheaths, and performed shockwave to the anatomy to be able to eventually deliver a gore 16fr x33 cm sheath and the cp pump. the long 14fr abiomed provided sheath was noted to be cracked in the mid-portion. there was no harm or injury from the crack, though the possibility of harm from the cracked sheath is present and the md questions the malfunction. the gore sheath additionally kinked in the pad diseased anatomy. when the cp was explanted the groin access needed closure by the manta device to achieve successful hemostasis. | [5 empty] | The investigation into the introducer issue has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the patient¿s condition, the patient¿s calcified vessels, was the cause of the limb ischemia. the failure mode will be monitored and trended. | The medical center had returned product for benchtop analysis and testing. upon completion of the investigation, a final report will be forthcoming. of note the ifu states: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35682 | 1220648-2023-02066 | 16820811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16820811&pc=OZD | 03/29/2023 | 10/25/2023 |  | 04/26/2023 |  |  | 03/29/2023 | 04/26/2023 |  |  | 12/31/2024 |  | 04/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023284002 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Air Embolism | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 03/29/2023 |  |  |  |  | Roughly seven days after the report of air bubbles, it was reported that multiple "impella in aorta" alarms were triggered but the pump was confirmed to be in the proper position. due to the possibility of a clot, the decision was made to replace the pump. | The medical center had an impella 5.5 support ongoing for a patient with cardiomyopathy. the patient was being evaluated for lvad. on day 8 of the support the team observed air within the echo imaging. abiomed medical affairs was consulted and troubleshooting measures taken with purge lines/cassette changed. there has been no harm or embolization to date. the malfunction is reportable as we can not deny the causality. later on day 15 the team explanted the pump due to alarms and concerns for thrombus. |  |  |  |  | B1, h1, and h6 were updated/corrected to report the air embolism as a serious injury. | The investigation for the impella in aorta alarm has been completed. regarding the optical signal issue, data logs showed persistent "impella position in aorta" alarms without any abnormal placement signal metrics. about 16 hours prior to alarms, a sudden motor current spike (and simultaneous motor speed drop) occurred, immediately followed by suction and flow saturation. only the motor housing and cannula were returned. the rest of the pump was cut off. visual inspection of the returned product revealed a significant mass of biomaterial on and around the impeller. biomaterial was also found on the motor housing. in conclusion, it was determined that the cause of the optical signal issue was ingested biomaterial. this report is being filed as part of a retrospective review of historical records. | Since the original report was filed, the investigation into the presumed air emboli (air bubbles) has concluded. the cause of the "air bubbles" and its relation to the impella were not determined since the pump performed per specification on the day of the alleged issue and the nature of the "bubbles" could not be verified clinically. | The pump remained on for support without any noted harm or embolization of air. the pump was returned for benchtop analysis and testing, which is currently ongoing. the causal link between the air seen on echo imaging and the use of impella can not yet been proven or disproven. upon completion of the investigation and benchtop analysis, a final report will be filed with root cause. the failure mode will be monitored and trended. of note the ifu states the following: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿infusion filter: the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿
35683 | 1220648-2023-02067 | 16820813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16820813&pc=OZD | 03/30/2023 | 07/11/2023 |  | 04/26/2023 |  |  | 03/30/2023 | 04/26/2023 |  |  | 10/31/2024 |  | 04/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023276231 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 7 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2023 |  |  |  |  | The medical center in (b)(6) had an impella cp support that was escalated to the 5.5 pump. the cp was for just under 10 hours when explanted due to a left ventricle (lv) perforation attribute to the cp. the question the pump was improperly pushed/advanced into the lv when making adjustments at the femoral insertion site. the medical team in (b)(6) stated the relationship to the use of impella and lv perforation is unknown, at is may be the "original state" of the lv, but the causality can not be denied. | [5 empty] | The investigation into the ventricle perforation has been completed since the original report was filed. the medical center in japan discarded all product at the time of explant. no benchtop analysis was possible to root cause; only the clinical report was evaluated. the root cause of the ventricle perforation was most likely use related as the clinical details specify that the pump was likely pushed deeper when placing the peel away or repositioning unit. the failure modes will be monitored and trended. | The (b)(6) center in (b)(6) discarded all product at the time of explant. no benchtop analysis is possible to root cause. no imaging has been shared by the team in (b)(6). no root cause or causal relationship between the impella cp and the lv perforation can be made. the failure mode will be monitored and trended. the ifu notes the following: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral. vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
35684 | 1220648-2023-02056 | 16806095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16806095&pc=OZD | 03/26/2023 | 05/15/2023 |  | 04/25/2023 |  |  | 03/26/2023 | 04/25/2023 |  |  | 01/31/2025 |  | 04/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024307348 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 47-year-old male has been treated for acute myocardial infarction / cardiogenic shock. for hemodynamic support, an impella 5.5 cardiac pump has been inserted. after 46 hours runtime, the pump has been unplugged from the automated impella controller to remove some torque from the patient cable. subsequently, a device defective alarm occurred. the pump stopped and could not be restarted. another impella 5.5 pump has successfully been inserted. no harm has been done to the patient. | [5 empty] | The logfiles recorded from the automated impella controller showed a high motor current pump stop and after disconnecting and reconnecting the impella 5.5 to the aic, an impella defective alarm was confirmed to trigger. the returned pump was visually inspected, and the catheter was observed to be cut. the returned purge cassette had a detached purge disc and a clamp over the tubing next to the pressure line one-way valve. the red handle was opened, and blood ingress was observed in the red handle. a hole was observed in the catheter at the 29 cm mark. the cause of the impella defective alarm was blood ingress due to a hole in the catheter resulting from improper use. | The pump has been returned for investigation. when the investigation is complete, a supplemental report will be filed. as noted in the impella ifu: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿impella stopped if the impella catheter has stopped suddenly: try to restart the catheter at previous p-level. if the impella does not restart, try to restart at p-2. if the impella does not restart or stops again, wait 1 minute and try to restart again. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35685 | 16806300 | 16806300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16806300&pc=OZD | 03/12/2023 | 04/12/2023 | 04/12/2023 | 04/25/2023 | 04/12/2023 | 04/25/2023 |  | 04/25/2023 |  |  |  |  | 04/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0042-0040-US |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 50 YR |  | Male |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | Followup | Followup |  |  |  | Y | HOSPITAL |  |  | Impella critical alarm heard, went to inspect alarm, alarm shown "battery failure". noted battery was at 50%. outlet changed with no resolution of battery failure alarm. abiomed support number called immediately for assistance. back up controller obtained. np notified. with the assistance of the support from abiomed over the phone, patient transferred to new impella controller. patient remained hemodynamically stable throughout transfer. original impella controller taken out of commission. the abiomed support person over the phone said she would notify the impella rep via e-mail. impella device information: sn: (b)(6). rf: (b)(4). | Impella critical alarm heard, went to inspect alarm, alarm shown "battery failure". noted battery was at 50%. outlet changed with no resolution of battery failure alarm. abiomed support number called immediately for assistance. back up controller obtained. np notified. with the assistance of the support from abiomed over the phone, patient transferred to new impella controller. patient remained hemodynamically stable throughout transfer. original impella controller taken out of commission. the abiomed support person over the phone said she would notify the impella rep via e-mail. impella device information: sn: [redacted]. rf: (b)(4). | Impella critical alarm heard, went to inspect alarm, alarm shown "battery failure". noted battery was at 50%. outlet changed with no resolution of battery failure alarm. abiomed support number called immediately for assistance. back up controller obtained. np notified. with the assistance of the support from abiomed over the phone, patient transferred to new impella controller. patient remained hemodynamically stable throughout transfer. original impella controller taken out of commission. the abiomed support person over the phone said she would notify the impella rep via e-mail. impella device information: sn: [redacted]. rf: 0042-0040-us.
35686 | 1220648-2023-02057 | 16806381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16806381&pc=OZD | 03/22/2023 | 04/25/2023 |  | 04/25/2023 |  |  | 03/29/2023 | 04/25/2023 |  |  | 10/31/2023 |  | 04/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  |  |  | 2023276464 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | [5 empty] | ECMO | [18 empty] | A |  |  |  | Initial submission |  |  |  |  | 03/29/2023 |  |  |  |  | The complainant reports a 58 year old male receiving support from impella for lv unloading that experienced insertion site bleeding with the formation of a large hematoma. angle adjustment and compression did not maintain hemostasis so the pump was removed and replaced. the patient was sent to icu to recover. this is a reportable complaint. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding ¿assess access site for bleeding and hematoma.¿
35687 | 1220648-2023-02058 | 16812215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16812215&pc=OZD | 03/27/2023 | 06/06/2023 |  | 04/25/2023 |  |  | 03/27/2023 | 04/25/2023 |  |  | 10/31/2024 |  | 04/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2024288775 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 74 year-old male presented with chest pain / shortness of breath. coronary artery disease has been diagnosed. during complicated pci procedure, patient required cpr. impella cp successfully placed for hemodynamic support. patient regained pulse once pump started. after peel-away sheath has been removed and repositioning sheath placed, serious access site bleeding noted. bleeding could not be stopped by applying pressure. 5 units of prbcs given. femoral angiogram performed which showed side wall stick with extravasation. impella removed and covered stent placed. patient remained stable. | [5 empty] | The investigation into the reported vascular damage has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the root cause for vascular damage remains undetermined since neither the product nor sufficient clinical information were provided. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device has not been received by the manufacturer for investigation. should any new information be returned, a supplemental report will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿
35688 | 1220648-2023-02059 | 16812247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16812247&pc=OZD | 03/28/2023 | 08/03/2023 |  | 04/25/2023 |  |  | 03/28/2023 | 04/25/2023 |  |  | 11/30/2024 |  | 04/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023283390 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 62-year-old male presented with ami/cgs. the patient received hemodynamic support from an impella cp during a cabg procedure. limb ischemia has been noted below the impella access site, requiring device removal, cut down, repair, thrombectomy, and fasciotomy. an impella 5.5 for increased hemodynamic support successfully been inserted. the patient has been stable afterwards. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the cause of the limb ischemia was calcified vessel patient condition related. the failure mode will be monitored and trended. | The impella device has not been received by the manufacturer for investigation. should any new information be returned, a supplemental report will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35689 | 1220648-2023-02052 | 16797941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16797941&pc=OZD | 11/27/2022 | 04/24/2023 |  | 04/24/2023 |  |  | 11/27/2022 | 04/24/2023 |  |  | 07/31/2024 |  | 04/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023250348 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/31/2022 |  |  |  |  | The medical center reported a fluid leak from the purge cassette (pc) during an impella 5.5 support. the pc and tubing were replaced to resolve the leak, and the purge system was bypassed by standard practice. roughly 30 days after implant, the pump was reported to have stopped. the stop occurred beyond the intended use life of the pump and was attributed to there being no purge solution in the system as the purge bypass had been removed. the pump was removed as a result of the stoppage and a new impella 5.5 pump was implanted for continued support. there was no harm to the patient. | [5 empty] | The medical center discarded the impella pump product at the time of explant and replacement. the cause of the pump stop was most likely blood ingress due to sidearm yellow luer damage stemming from sodium bicarbonate in the purge. the failure mode will be monitored and trended. a capa is open at this time. as noted in the ifu: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35690 | 1220648-2023-02053 | 16803874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16803874&pc=OZD | 12/11/2021 | 04/24/2023 |  | 04/24/2023 |  |  | 12/11/2021 | 04/24/2023 |  |  | 05/31/2023 |  | 04/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC |  |  | US |  | 2022088358 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/11/2021 |  |  |  |  | The medical center had a 5.5 impella placed to replace/escalate care from the impella cp. after 19 days the pump stopped. the pump stop occurred after a purge sidearm leak was attempted to be troubleshot by bypassing it. the pump stop occurred beyond indicated use, however the event is reportable for the full support pump. | [5 empty] | The medical center discarded the impella pump product at the time of explant. the data logs and clinical report alone were analyzed to achieve root cause determination. the cause of the pump stop was most likely blood ingress as a result of a broken purge sidearm due to sodium bicarbonate in the purge solution. the failure mode will be monitored and trended. a capa is open for the failure. of note in the ifu: ¿impella stopped if the impella catheter has stopped suddenly: try to restart the catheter at previous p-level. if the impella does not restart, try to restart at p-2. if the impella does not restart or stops again, wait 1 minute and try to restart again. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
35691 | 1220648-2023-02054 | 16803963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16803963&pc=OZD | 09/06/2021 | 10/25/2023 |  | 04/24/2023 |  |  | 09/06/2021 | 04/24/2023 |  |  | 11/30/2021 |  | 04/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC |  |  | US |  | 1465760 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 65 YR |  | Female |  |  |  | Malfunction | No | Yes | Crack | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/06/2021 |  |  |  |  | The medical center had an impella 5.5 support ongoing for care escalated from an impella cp. the patient was suffering from post-cardiotomy cardiogenic shock. after 18 days of support the 5.5 pump stopped. the surgical team made decision to explant the pump and monitor the patient's hemodynamics without impella. the pump was not replaced. | [5 empty] | (h3) the investigation into the reported 'pump stop" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. review of data logs showed that the pump stopped immediately upon the sudden occurrence of the "impella stopped - motor current high" alarm. the pump stop was immediately preceded by multiple "purge system blocked" alarms accompanied by a rapid increase in purge pressure. about two days prior to the pump stop, the purge pressure had suddenly decreased. there were no other abnormalities noted. in conclusion, after reviewing the clinical information, it was determined that the cause of the pump stop was blood ingress as a result of a broken purge sidearm due to the use of sodium bicarbonate in the purge solution. | The medical center discarded all product at the time of explant. no benchtop analysis to root cause is possible. only the clinical details were reviewed for investigation. the cause of the pump stop was blood ingress as a result of a broken purge sidearm due to the use of sodium bicarbonate in the purge solution.
35692 | 1220648-2023-02055 | 16803965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16803965&pc=OZD | 12/21/2021 | 09/15/2024 |  | 04/24/2023 |  |  | 12/21/2021 | 04/24/2023 |  |  | 06/30/2023 |  | 04/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022100797 |  | 0550-00008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [5 empty] | VA-ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/20/2021 |  |  |  |  | The medical center placed an impella 5.5 for care escalation from the impella cp and va-ecmo support. the 5.5 was placed and supported the patient for just under 19 days when the pump stopped. the pump was explanted but not replaced due to the patient's poor prognosis. the patient expired with underlying cardiogenic and septic shock. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported "pump stop" has been completed since the original report was filed. the returned purge sidearm was microscopically inspected and a crack on the male portion of the yellow luer was observed along the weld line. although the pump was not returned, the reported pump stop was most likely due to the confirmed sidearm damage leading to blood ingress as a result of sodium bicarbonate usage in the purge. a risk assessment has been initiated to escalate this issue. section d4: unique identifier (udi) #. | The medical center did return the impella sidearm, but not the pump, for case investigation upon the benchtop. although the pump was not returned, the reported pump stop was most likely due to the confirmed sidearm damage leading to blood ingress as a result of sodium bicarbonate usage in the purge. the failure mode will be monitored and trended. a capa is open for the failure. of note in the ifu; ¿impella stopped if the impella catheter has stopped suddenly: try to restart the catheter at previous p-level. if the impella does not restart, try to restart at p-2. if the impella does not restart or stops again, wait minute and try to restart again. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02055 was provided.
35693 | MW5117020 | 16809890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16809890&pc=OZD |  | 04/20/2023 |  | 04/24/2023 |  |  |  | 04/24/2023 |  |  |  |  | 04/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA PURGE CASSETTE |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  |  |  | 2023274900 |  |  |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Other Health Care Professional | [8 empty] | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | Leak noted in the impella purge tubing. tubing changed and therapy continued. no patient harm/injury.
35694 | 1220648-2023-02045 | 16786149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16786149&pc=OZD | 03/22/2023 | 04/21/2023 |  | 04/21/2023 |  |  | 03/22/2023 | 04/21/2023 |  |  | 12/31/2024 |  | 04/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC |  |  | US |  | 2024304230 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/22/2023 |  |  |  |  | The complainant reported a 67-year-old female patient presenting for cardiomyopathy. during impella support, it was documented that the patient experienced an embolic stroke. patient support continued following the event and the patient was considered to be stable until the planned time of explant. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu, stroke is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35695 | 1220648-2023-02046 | 16792410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16792410&pc=OZD | 03/23/2023 | 06/15/2023 |  | 04/21/2023 |  |  | 03/23/2023 | 04/21/2023 |  |  | 12/31/2024 |  | 04/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024291926 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Difficult to Remove | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/23/2023 |  |  |  |  | The complainant reported a 49 year old male patient presenting with post cardiotomy cardiogenic shock for impella support. difficulty removing the impella device was reported, after which the chest was opened to find bleeding from the aortic graft suture line. new sutures were applied, graft was removed and chest washed out, then 2 units of blood products were delivered. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the most likely cause of the access site bleeding was use related as increased sutures were needed. the failure mode will be monitored and trended. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
35696 | 1220648-2023-02047 | 16792429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16792429&pc=OZD | 03/23/2023 | 04/21/2023 |  | 04/21/2023 |  |  | 03/23/2023 | 04/21/2023 |  |  | 12/31/2024 |  | 04/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 380 (US) |  |  | ABIOMED, INC |  |  | US |  | 2024298241 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/16/2023 |  |  |  |  | The complainant reported a 64-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was documented that the patient lost doppler pulse to the impella leg. a distal perfusion catheter was placed in the left leg to treat the noted condition. patient support continued following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35697 | 1220648-2023-02048 | 16792431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16792431&pc=OZD | 03/13/2023 | 04/21/2023 |  | 04/21/2023 |  |  | 03/23/2023 | 04/21/2023 |  |  | 11/30/2024 |  | 04/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2024299659 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/23/2023 |  |  |  |  | The complainant reported a 67-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was documented that the pump was unable to maintain adequate flows while implanted without experiencing suction issues. it was noted that the patient's ventricle appeared small during observation and that the guidewire was unable to advance far across the aortic valve before deflecting due to the patient's anatomy. the pump was removed, flushed, and re-inserted and fluid was given in an attempt to troubleshoot the issue, however, the flow rate remained low. the pump was subsequently removed and an intra-aortic balloon pump was inserted to continue with patient support. following the initial report, the pump was returned for evaluation and it was discovered that the flow issues were caused by broken impeller blades. there was no patient harm. | [5 empty] | The device was returned and an evaluation was completed. a visual inspection of the pump indicated that the blades on the impeller were broken. because the fluoroscopic image of the pump was not returned, a definitive root cause of the fracture could not be determined. should additional relevant information become available, a supplemental report will be submitted. as noted in the impella ifu: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿.
35698 | 1220648-2023-02049 | 16792433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16792433&pc=OZD | 03/22/2023 | 08/04/2023 |  | 04/21/2023 |  |  | 03/22/2023 | 04/21/2023 |  |  | 10/31/2024 |  | 04/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024310813 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Female |  |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An 88-year-old female patient was treated for coronary artery stenosis. during percutaneous coronary intervention, hemodynamic support was provided by an impella cp with smartassist cardiac pump. after insertion of the impella cp, reduced pump flow of the impella cardiac pump was observed and the patient suffered from increasing shortness of breath. pericardial effusion was noted, which was treated by pericardial drainage. the patient was hemodynamically stable. contributing factors to the perforation of the ventricle were a very thin apical wall and a small cavum. | [5 empty] | Benchtop investigation of the returned pump will be performed after return of the device. when the investigation is complete, a supplemental report will be filed. | The investigation into the ventricle perforation has been completed since the original report was filed. the user facility has returned the impella device, and the benchtop analysis of the cause of the ventricle perforation was most likely improper positioning of the device.
35699 | 1220648-2023-02051 | 16792435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16792435&pc=OZD | 03/25/2023 | 07/05/2023 |  | 04/21/2023 |  |  | 03/25/2023 | 04/21/2023 |  |  | 10/31/2024 |  | 04/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2024285535 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2023 |  |  |  |  | The complainant reported a 62-year-old male patient presenting for high risk percutaneous coronary intervention. during impella support, it was documented that that the 14 fr introducer was noted to be occlusive and was swapped. sluggish flow was still noted in the leg and the decision was eventually made to explant the pump. the patient was stable following explant. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the impella device was returned for evaluation and the benchtop analysis of the root cause revealed that the cause of the issue was patient condition related. the failure mode will be monitored and trended. | The device was returned and an evaluation is pending. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35700 | 1220648-2023-02050 | 16792436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16792436&pc=OZD | 03/26/2023 | 05/05/2023 |  | 04/21/2023 |  |  | 03/26/2023 | 04/21/2023 |  |  | 11/30/2024 |  | 04/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2024303640 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2023 |  |  |  |  | The medical center placed an impella cp for care escalation from an iabp. on the day after placement there were signs of limb ischemia, as pulses were lost, which prompted the team to have surgery intervene bedside. the patient had a bedside fasciotomy performed and an antegrade sheath was placed for leg perfusion. | [5 empty] | The medical center has not yet returned for investigation the pump product. upon completion of the investigation, a final report with root cause will be submitted. of note the ifu states: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Since the original report was filed, the investigation into the reported limb ischemia has concluded. no product was returned for analysis. only the clinical report was reviewed for investigation. the cause of the limb ischemia was most likely patient condition. this was determined as the medical team provided clinical information that the patient had small vessels. the failure mode will be monitored and trended.
35701 | 1220648-2023-02040 | 16777975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16777975&pc=OZD | 04/04/2023 | 04/20/2023 |  | 04/20/2023 |  |  | 04/04/2023 | 04/20/2023 |  |  | 10/31/2024 |  | 04/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC |  |  | US |  | 2023277214 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Other | Required Intervention | Hospitalization | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/04/2023 |  |  |  |  | The complainant reports that a 92 year old male patient presented with acute myocardial infarction and cardiogenic shock requiring impella support. during support, a left ventricular perforation was discovered and the device was removed. the perforation was repaired and a intra-aortic pump was used to continue hemodynamic support. the patient was sent to icu to recover. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35702 | 1220648-2023-02041 | 16778060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16778060&pc=OZD | 03/22/2023 | 04/27/2023 |  | 04/20/2023 |  |  | 03/22/2023 | 04/20/2023 |  |  | 12/31/2024 |  | 04/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED INC. |  |  | US |  | 2024300125 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2023 |  |  |  |  | The medical center had an impella cp explanted for care escalation to the impella 5.5, when there was an observation made of no distal flow or pulses in the cp inserted limb. the cp was explanted and the thrombus was observed and so the limb had thrombus removal which restored flow. the pulses returned after the thrombus was removed and a patch was made to the right common femoral artery. the patient had known peripheral vascular disease. the patient survived. | [5 empty] | The medical center discarded the pump at the time of explant. no benchtop analysis to root cause is possible. the patient had known peripheral arterial disease which may or may not have caused or contributed to the limb ischemia. the manufacturer will file a final report should more information be shared that leads to root cause. the failure mode will be monitored and trended. of note, in the ifu the following is listed: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | Since the original report was filed, the investigation into the limb ischemia has concluded. the product was not returned, but the clinical report was analyzed to root cause. the root cause was determined to be biomaterial. large thrombus/biomaterial was removed from from the femoral artery and the pulses returned post intervention. no issue with the patient's limb post procedure. the failure mode will be monitored and trended.
35703 | 1220648-2023-02043 | 16781957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16781957&pc=OZD | 03/25/2023 | 04/20/2023 |  | 04/20/2023 |  |  | 03/25/2023 | 04/20/2023 |  |  | 11/30/2024 |  | 04/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2024313613 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | PROTEK DUO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/25/2023 |  |  |  |  | The medical center had an impella cp support ongoing in which there was limb ischemia. the ischemia prompted the team to place fem-fem bypass to reperfuse the limb. the patient was known to have peripheral arterial disease. the pump remained on for 5 days of support, along with a protek duo for right sided heart support. | [5 empty] | The medical center discarded the impella pump at the time of explant. no benchtop analysis to root cause is possible. the data logs were returned and show no pump malfunction. the known peripheral arterial disease could have contributed to the ischemia. the manufacturer will close the investigation but, should more information be shared, a final report can follow. the failure mode will be monitored and trended. the ifu states the following; ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35704 | 1220648-2023-02042 | 16781958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16781958&pc=OZD | 03/25/2023 | 06/15/2023 |  | 04/20/2023 |  |  | 03/25/2023 | 04/20/2023 |  |  | 07/31/2024 |  | 04/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023252326 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death |  |  |  |  | ECMO AND PROTEK DUO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2023 |  |  |  |  | The medical center had an impella 5.5 delivered via direct access for a surgical bypass (cabg) patient who was failing to come off the bypass. as the patient's condition was poor, the team also delivered ecmo to the patient and a protek duo rvad. the patient was found to have a pulmonary artery injury and was taken to the or. while in the or the team was manipulating the heart and found to have a left ventricle (lv) perforation. the allegation of the 5.5 pump causing or contributing to the lv perforation could not be denied. the patient expired. | [5 empty] | The medical center has not yet returned for analysis the impella pump product. upon completion of the investigation, a final report will be forthcoming. of note in the ifu: "potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the reported death has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that unintended user error was the most likely cause of the death. based on the discussion with the clinical staff and medical affairs, the cause of the ventricle perforation was most likely use related and related to manipulation of the heart. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35705 | 1220648-2023-02044 | 16781959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16781959&pc=OZD | 03/15/2023 | 09/19/2024 |  | 04/20/2023 |  |  | 03/15/2023 | 04/20/2023 |  |  | 09/30/2024 |  | 04/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED INC. |  |  | US |  | 2023277390 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/15/2023 |  |  |  |  | The medical center had placed an impella cp for support. the cp site developed a hematoma. the team infused back 4 units of blood to the patient after the team estimated the patient to have lost 2 units of blood. they reviewed the situation and believed the team to have damaged the bifurcation of the profunda and sfa, or possibly caused a retroperitoneal bleed. the team placed a femostop and life-flighted the patient to the next level of tertiary care. | The procedural outcome noted patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The medical center did not retain the impella pump product for investigation and benchtop analysis. further clinical details have been asked, but not yet furnished. should more information be shared that leads to root cause, a final report will be forthcoming. the failure mode will be monitored and trended. of note the ifu states the following: ¿assess access site for bleeding and hematoma.¿ | Since the original report was filed, the investigation has concluded. no product was returned for investigation, and so only the clinical report was analyzed. the cause of the access site bleeding was most likely a lack of vessel recoil onto the sheath since bleeding occurred after sheath exchange and best practices were followed. the failure mode will be monitored and trended. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02044 was provided in sections b2, b5, h1, and h6.
35706 | 1220648-2023-02039 | 16771267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16771267&pc=OZD | 03/20/2023 | 04/27/2023 |  | 04/19/2023 |  |  | 03/20/2023 | 04/19/2023 |  |  | 11/30/2024 |  | 04/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED INC. |  |  | US |  | 2024304003 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/20/2023 |  |  |  |  | The medical center placed an impella cp via the femoral artery to support a patient with multivessel coronary artery disease. the team sent the patient to the or for bypass surgery (cabg) and upon admission to the or the patient was found to have bent up both legs at the knees. with the impella in leg, there was profuse bleeding and the pump had come out of appropriate position. vascular surgery performed a cutdown and arterial repair. the pump was explanted after only just under 4 hours. | [5 empty] | The medical center discarded the impella pump product at the time of explant in the or. the product can not be analyzed on the benchtop. the patient bending at the knees could have caused/contributed to the bleed and need for vascular repair. further clinical details are being asked, but not yet returned. should a root cause be determined, a final report will follow. the failure mode will be monitored and trended. of note the ifus states: "be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿. "assess access site for bleeding and hematoma.¿. ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | Since the original report was followed, the investigation has concluded. the cause of the vascular damage was use related since the pump's repositioning sheath was observed to have been pulled back the failure mode will be monitored and trended.
35707 | 1220648-2023-02038 | 16761774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16761774&pc=OZD | 03/18/2023 | 05/25/2023 |  | 04/18/2023 |  |  | 03/18/2023 | 04/18/2023 |  |  | 10/31/2024 |  | 04/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024290669 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2023 |  |  |  |  | The complainant reported a 69 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed an access site hematoma. two units of red blood cells were delivered and surgical debridement/repair to the access site was completed. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient movement was the most likely cause of the access site bleeding. the mode of failure will be monitored and trended.
35708 | 1220648-2023-02035 | 16753315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16753315&pc=OZD | 03/24/2023 | 04/17/2023 |  | 04/17/2023 |  |  | 03/24/2023 | 04/17/2023 |  |  | 04/30/2024 |  | 04/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023217721 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/24/2023 |  |  |  |  | The medical center had an impella cp supporting a patient who was sat up-right in bed by the radiology team for a chest x-ray on the day of cp implant. after the patient was sat up-right the patient had signs of limb ischemia. the patient lost distal pulses and the foot became dusky. the team explanted the pump to remedy the signs of limb ischemia. the team made the choice not to replace the pump. the support was for just over 11 hours. | [5 empty] | The medical center discarded the impella pump product at the time of explant. as such, no benchtop analysis is possible. the clinical report alone was available for the investigation. the report of the patient being sat up to >75 degrees in bed was a use related issue that caused the ischemia. the ifu states the following: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35709 | 1220648-2023-02036 | 16756913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16756913&pc=PYX | 03/20/2023 | 07/11/2023 |  | 04/17/2023 |  |  | 03/20/2023 | 04/17/2023 |  |  | 11/30/2024 |  | 04/17/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023281800 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 55-year-old female presented with right heart failure. to rest the right heart, an impella rp has successfully been placed. high purge pressure/purge system blocked alarm occurred, subsequently hemolysis has been noted. decision was made to remove the impella rp, as the physician team assessed that the patient had recovered enough. | [5 empty] | The investigation into the hemolysis has been completed since the original report was filed. the user facility has returned the impella device, and the benchtop analysis of the root cause revealed that the cause of the hemolysis was biomaterial. the failure mode will be monitored and trended. | The impella device has been returned to the manufacturer. when the investigation is complete, a supplemental report will be filed.
35710 | 1220648-2023-02037 | 16758683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16758683&pc=OZD | 03/17/2023 | 04/26/2023 |  | 04/17/2023 |  |  | 03/17/2023 | 04/17/2023 |  |  | 11/30/2024 |  | 04/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED INC. |  |  | US |  | 2024303605 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/17/2023 |  |  |  |  | The medical center had an impella cp placed for support of a patient after the surgical team had turned her down for bypass (cabg) surgery. the cp was to support the patient for coronary angioplasty via femoral artery access. the pump was placed and was supporting the patient when she became disorientated in the icu and was grabbing at her impella pump. the team made choice to explant the pump. at the time of explant they discovered thrombosis at the impella access site. the thrombosis was treated by penumbra thrombectomy and balloon angioplasty. | [5 empty] | Since the original report was filed the investigation has concluded. the product was not returned, only the data logs and clinical report were available for investigation. the logs showed the pump operated to specification. there was no spike or changes that would be indicative of biomaterial ingestion. the cause of the thrombosis was determined to be patient condition related. the failure mode will be monitored and trended. | The medical center discarded the impella pump product at the time of explant. there can be no benchtop analysis to root cause. the vessel thrombosis cause is not known at this time. further clinical details have not yet been shared. should root cause be determined, a final report will be forthcoming. of note the ifu states the following: "potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
35711 | 1220648-2023-02033 | 16742126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16742126&pc=OZD | 03/18/2023 | 10/06/2025 |  | 04/14/2023 |  |  | 03/18/2023 | 04/14/2023 |  |  | 11/30/2024 |  | 04/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024285047 |  | 0550-0002 |  | Yes |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | Immediately leading up to the reported pump stop, it was further noted that the placement signal from the pump was not accurate. | A male patient of unknown age in the eu presented with de-novo cardiomyopathy. a va-ecmo has been in place before impella support. an impella 5.5 cardiac pump has been inserted to provide hemodynamic support. the patient has been on bridge to lvad. on day 18 of support, after the purge pressure has been rising, a pump stop occurred. a new impella 5.5 pump has successfully been inserted. patient has been alive and on icu afterwards. | Additional information indicated the aic showed a ps-lv about 330/160mmhg with no adjustment possible. decision was made to withdraw care and the patient expired. |  |  |  | The investigation into the optical signal issue and pump stop has been completed since the original report was filed. the user facility has returned the impella device, and the benchtop analysis of the root cause revealed that the cause of the optical signal issue was most likely biomaterial, and the cause of the pump stop was most likely blood ingress. the cause of the blood ingress was unknown. the failure modes will be monitored and trended. device available for eval in section d8 updated to yes. | Section h6 updated to correspond with the report of a placement signal issue. this report is being filed as part of a retrospective review of historical records. | The impella device has been returned to the manufacturer. when the investigation is complete, a supplemental report will be filed. as noted in the impella ifu: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: try to restart the catheter at previous p-level. if the impella does not restart, try to restart at p-2. if the impella does not restart or stops again, wait 1 minute and try to restart again. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no data logs were returned/available for review. the product was returned and examined. biomaterial was found in the outflow cage and in the purge gap. a high motor current pump stop and high purge pressure were reproduced. the biomaterial was removed however the high motor current pump stop was still reproducing. the catheter was cut by the motor to expose the purge lumen and blood was found to have ingressed into the purge lumen. the root cause of the pump stop was determined to be blood ingress. the cause of the blood ingress is unable to be determined with no data logs and limited clinical details. the 5s parameters of the optical sensor on the returned product were taken and found to be within specification. the root cause of the offset placement signal was most likely biomaterial as a large amount of biomaterial was found around the optical sensor. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
35712 | 1220648-2023-02034 | 16746607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16746607&pc=OZD | 03/17/2023 | 05/26/2023 |  | 04/14/2023 |  |  | 03/17/2023 | 04/14/2023 |  |  | 11/30/2024 |  | 04/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC |  |  | US |  | 2023278001 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/17/2023 |  |  |  |  | The complainant reported a 71-year-old female patient presenting for cardiomyopathy. during impella support, it was documented that the aic displayed an impella stopped alarm after the patient was transferred from the or to their bed. after multiple unsuccessful attempts to restart the pump, the decision was made to explant the device and implant another impella 5.5 pump. patient support continued following the event. | [5 empty] | Since the original report was filed, the investigation into the pump stop has concluded. the impella pump and data logs were returned and analyzed to root cause. the returned pump was visually inspected and biomaterial was observed in the outflow. the data logs were reviewed and the sudden pump stop during patient movement indicates that most likely an open motor current line occurred due to excessive force on the catheter. the excessive force is a user related. the failure mode will be monitored and trended. | The device was returned and an evaluation is pending. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results. as noted in the impella ifu, precautions and instruction pertaining to pump stops include: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35713 | 1220648-2023-02026 | 16731214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16731214&pc=OZD | 03/14/2023 | 05/25/2023 |  | 04/13/2023 |  |  | 03/14/2023 | 04/13/2023 |  |  | 10/31/2024 |  | 04/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2024288217 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2023 |  |  |  |  | The complainant reported a 69 year old male patient presenting for high risk percutaneous coronary intervention using the impella for hemodynamic support. the 14fr introducer was left in the access site due to oozing, however, ischemia developed. the pump was removed. | [5 empty] | The investigation into the reported limb ischemia and access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the limb ischemia and access site bleeding. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support:¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35714 | 1220648-2023-02027 | 16731945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16731945&pc=OZD | 03/14/2023 | 07/11/2023 |  | 04/13/2023 |  |  | 03/14/2023 | 04/13/2023 |  |  | 11/30/2024 |  | 04/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2024300045 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2023 |  |  |  |  | The complainant reported a 54-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient developed lower limb ischemia. it was documented that the patient lost pulse distal to the impella insertion site. vascular surgery was performed to execute a cutdown of the femoral artery and to remove the pump. following explant, a double fasciotomy of the lower left limb was initiated to remove a clot from the left iliac artery. a stent was subsequently placed after the pump was removed. the patient was deemed stable following the intervention. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on the clinical details, the root cause of the limb ischemia was most likely thrombus in the patient's vessel. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35715 | 1220648-2023-02028 | 16732102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16732102&pc=OZD | 03/14/2023 | 06/23/2023 |  | 04/13/2023 |  |  | 03/14/2023 | 04/13/2023 |  |  | 12/31/2024 |  | 04/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC |  |  | US |  | 2024305146 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2023 |  |  |  |  | The complainant reported a 68-year-old male patient presenting for elective placement. during impella support, it was noted that the pump showed signs of flow saturation and subsequently stopped. the halt in function occurred following a spike in the motor current. the pump was removed as a result of the failure and a visual inspection identified biomaterial within the pump next to the impeller. there was no patient harm as a result of the failure. another pump was implanted for ongoing support. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu regarding pump stops: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop has been completed since the original report was filed. the medical center discarded the impella product at the time of explant. no benchtop analysis of the root cause of the pump stop was possible; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of the pump stop was most likely related to biomaterial. the failure mode will be monitored and trended.
35716 | 1220648-2023-02029 | 16732116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16732116&pc=OZD | 03/16/2023 | 04/13/2023 |  | 04/13/2023 |  |  | 03/16/2023 | 04/13/2023 |  |  | 10/31/2024 |  | 04/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023272119 |  | 0048-0014 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 73-year-old male presented with acute myocardial infarction. he received support by a va-ecmo and an impella cp cardiac pump. during the procedure, the impella cp pump was replaced for another impella cp. after the pump exchange, the patient suffered from severe access site bleeding. several interventions have been required to stop the bleeding: additional stitches, multiple new pressure bandages, blood products (unknown number) and correction of the anticoagulation. finally, the impella cp pump has been removed. at this point, the patient did not require unloading any longer, so no new impella device has been inserted. | [5 empty] | The impella device has not been received by the manufacturer for investigation. should any new information be returned, a supplemental report will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
35717 | 1220648-2023-02030 | 16734132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16734132&pc=OZD | 03/28/2023 | 09/26/2024 |  | 04/13/2023 |  |  | 03/31/2023 | 04/13/2023 |  |  | 11/30/2024 |  | 04/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023278899 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | Thrombosis/Thrombus | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/31/2023 |  |  |  |  | The complainant reported a 79-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient experienced an ischemic stroke and later passed away. leading up to the event, it was documented that the patient¿s left ventricular lumen was narrow and limited the impella pump to low support levels. a thrombus formed in the left atrium and caused a cerebral infarction when it shifted location. patient support continued following the event, however due to the patient's subsequently poor neurological prognosis and dilated pupils, dnar status was established. the patient passed away three days after the stroke. | [5 empty] | The investigation into the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the strove/cva was due to a thrombus forming in the left atrium during impella support. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-02030 was provided in sections d6 and h1. | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu, a stroke is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35718 | 1220648-2023-02032 | 16734477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16734477&pc=OZD | 03/17/2023 | 08/09/2023 |  | 04/13/2023 |  |  | 03/17/2023 | 04/13/2023 |  |  | 11/30/2024 |  | 04/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2024302433 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/17/2023 |  |  |  |  | The complainant reported a 59 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. the patient developed an impella access site hematoma after removal (during site attempted closure). a covered stent was required to achieve hemostasis. a 30-day mdr is required. | [5 empty] | The investigation into the access site bleeding and the limb ischemia has been completed since the original report was filed. no benchtop analysis was possible to determine the root cause. since the medical center did not return the impella device, the root cause of the access site bleeding and the limb ischemia was not determined. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿.
35719 | 1220648-2023-02031 | 16734837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16734837&pc=OZD | 03/27/2023 | 06/04/2023 |  | 04/13/2023 |  |  | 03/30/2023 | 04/13/2023 |  |  | 09/30/2024 |  | 04/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023269048 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2023 |  |  |  |  | The complainant reported a patient of unknown age and gender presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was documented that the pigtail of the impella cp pump perforated the patient's left ventricle. a surgical closure was performed to repair the perforation and the pump was replaced with an impella 5.5 pump. the patient was considered to be stable following the intervention. | [5 empty] | The investigation into the ventricle perforation has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that a use issue was the cause of the ventricle perforation. the failure mode will be monitored and trended. ventricle perforation ifu quotes: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35720 | 1220648-2023-02024 | 16724021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16724021&pc=OZD | 02/13/2023 | 07/04/2023 |  | 04/12/2023 |  |  | 03/13/2023 | 04/12/2023 |  |  | 11/30/2024 |  | 04/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023282617 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2023 |  |  |  |  | The complainant reported a 68-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient experienced a ventricle perforation. it was documented that the catheter perforated the apex of the left ventricle while attempting to position the pump. the pump was pulled back into position and sutures were applied to treat the perforation. the patient was considered stable following the intervention. | [5 empty] | The investigation into the ventricle perforation has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the cause of the ventricle perforation was use related and occurred during positioning of the device in the ventricle. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35721 | 1220648-2023-02025 | 16728869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16728869&pc=OZD | 01/28/2023 | 05/26/2023 |  | 04/12/2023 |  |  | 01/28/2023 | 04/12/2023 |  |  | 10/31/2024 |  | 04/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 380 (US) |  |  | ABIOMED INC. |  |  | US |  | 2024291749 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2023 |  |  |  |  | The medical center had an impella cp support ongoing for a patient admitted with multiple cardiac and systemic risk factors/comorbidities. after insertion the patient developed signs of limb ischemia. the ischemia resulted in lost pulses that were treated by fem-fem bypass placement. the support was for 4 days til wean and explant. | [5 empty] | Since the original report was filed, the investigation into the reported limb ischemia has concluded. the medical center did not return for benchtop analysis any impella product, only the clinical report was evaluated. the evaluation determined that most likely the cause of the limb ischemia was underlying patient condition. the failure mode will be monitored and trended. | The medical center discarded all impella pump and sheath product at the time of explant. further clinical details regarding the limb ischemia are not known. should more information be shared that leads to a root cause, a final report will be forthcoming. the failure mode will be monitored and trended. of note in the ifu listed is the following: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
35722 | 1220648-2023-02022 | 16716222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16716222&pc=OZD | 03/14/2023 | 04/11/2023 |  | 04/11/2023 |  |  | 03/14/2023 | 04/11/2023 |  |  | 10/31/2024 |  | 04/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2024287046 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/14/2023 |  |  |  |  | The medical center had an impella cp and ecmo support ongoing for young 25 year old patient. the patient had a history of right heart failure and pulmonary hypertension. there was bleeding from the ecmo and cp sites. the team infused blood products and questions if the patient may be suffering for dic. vascular consulted but no further intervention needed. the pump support was successful for 3 days and then explanted, the ecmo remained on for continued hemodynamic monitoring. | [5 empty] | The medical center discarded the impella pump at the time of explant. no benchtop analysis to root cause is possible. there was however the assumption of dic contributing to the bleed noted by the medical team. patient condition may have been the contributing factor for the bleed and blood product infusion. should more information be shared that leads to the root cause, a final report will be forthcoming. of noted the ifu notes: "assess access site for bleeding and hematoma.¿
35723 | 1220648-2023-02023 | 16716229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16716229&pc=OZD | 04/10/2023 | 07/11/2023 |  | 04/11/2023 |  |  | 04/10/2023 | 04/11/2023 |  |  | 09/30/2024 |  | 04/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023282554 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2023 |  |  |  |  | The medical center had an impella cp support ongoing for a patient admitted in cardiogenic shock (cgs). the team had placed a femoral bypass due to the lost distal pulses in the impella inserted leg. the pump remained on for support of 8 days, was weaned and explanted. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on the clinical details, the root cause of the limb ischemia was most likely patient condition related. the failure mode will be monitored and trended. | The medical center discarded the impella pump and sheath product at the time of explant. no benchtop analysis to root cause is possible. the failure mode will be monitored and trended. should new information be shared that leads to a root cause, a final report will be forthcoming.
35724 | 1220648-2023-02012 | 16706666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16706666&pc=OZD | 09/15/2022 | 08/17/2023 |  | 04/10/2023 |  |  | 09/15/2022 | 04/10/2023 |  |  | 03/31/2024 |  | 04/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023212487 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/15/2022 |  |  |  |  | The complainant reported a 61 year old male patient presenting with cardiomyopathy for impella support. the patient exhibited signs of hemolysis, despite optimal positioning. liver profile labs and creatinine became elevated. ultimately, the decision was made to remove the device and escalate to an impella 5.0 via trans-caval access, due to hemolysis and acute kidney injury. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information provided, the cause of the hemolysis was not determined since product was not returned. | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35725 | 1220648-2023-02013 | 16706814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16706814&pc=OZD | 01/09/2023 | 06/26/2023 |  | 04/10/2023 |  |  | 01/09/2023 | 04/10/2023 |  |  | 09/30/2024 |  | 04/10/2023 |  | 3)TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023259457 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/09/2023 |  |  |  |  | The complainant reported a 71 year old male patient presenting with cardiomyopathy for impella support. oozing was noted from the impella access site, post device insertion. as a result, 3 units of blood products were delivered. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis of the access site bleeding; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of access site bleeding was most likely related to patient condition of esophagus bleeding. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35726 | 1220648-2023-02014 | 16708723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16708723&pc=OZD | 03/12/2023 | 04/10/2023 |  | 04/10/2023 |  |  | 03/12/2023 | 04/10/2023 |  |  | 10/31/2024 |  | 04/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2024286606 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/12/2023 |  |  |  |  | The medical center had an impella cp support of 2 days for a patient admitted with cardiomyopathy and in need of hemodynamic support. during the entire support there was signs of hemolysis despite all troubleshooting measures taken. the team explanted prematurely due to the hemolysis. | [5 empty] | The medical center has not returned any impella pump product for benchtop analysis. the pump can not be run through hemolysis testing. the root cause can not be determined with known clinical information alone. should more information be shared, and root cause determined, a final report will be forthcoming. of note, within the ifu there is mentioned: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿
35727 | 1220648-2023-02015 | 16708733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16708733&pc=OZD | 03/12/2023 | 10/13/2025 |  | 04/10/2023 |  |  | 03/12/2023 | 04/10/2023 |  |  | 11/30/2024 |  | 04/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 2024303633 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [5 empty] | VV ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2023 |  |  |  |  | The medical center had an impella cp placed via the right femoral after the vv ecmo was also placed to the right leg, at the femoral vein. the team on the next day saw signs of limb ischemia. to treat the cold foot external bypass was placed. the limb improved. the patient remained on the impella support for two days, was then weaned and explanted. | [5 empty] | The medical center discarded the impella pump at the time of explant. no benchtop analysis is possible to root cause. should the product be returned and root cause determined, a final report will be forthcoming. of noted the ifu states the following: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35728 | 1220648-2023-02016 | 16708750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16708750&pc=OZD | 02/16/2023 | 07/11/2023 |  | 04/10/2023 |  |  | 02/16/2023 | 04/10/2023 |  |  | 09/30/2024 |  | 04/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023284591 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2023 |  |  |  |  | The medical team had an impella cp supporting a percutaneous coronary angioplasty (pci) patient. the patient had known peripheral arterial disease (pad) but, had the femoral placement of the patient's impella despite the known disease. the team performed a groin shot after the pci and noted sluggish flow down the limb. the team prematurely removed the pump due to the limb ischemia. the team placed a contralateral leg accessed balloon pump (iabp) to continue support for the patient. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on the clinical details, the root cause of the limb ischemia was most likely patient condition related. the failure mode will be monitored and trended. | The medical center discarded all product at the time of explant. no benchtop analysis is possible to root cause. the patient had known peripheral arterial disease which may have contributed to the sluggish flow and ischemia concerns. the failure mode will be monitored and trended. of note the ifu states the following: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35729 | 1220648-2023-02017 | 16708785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16708785&pc=OZD | 02/05/2023 | 04/14/2023 |  | 04/10/2023 |  |  | 02/05/2023 | 04/10/2023 |  |  | 07/31/2024 |  | 04/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023257974 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2023 |  |  |  |  | The medical center had an impella cp support ongoing in which there was an access site bleed that prompted the team to intervene by infusing blood products and stitching the site. the pump supported the patient for 6 days and was weaned and explanted. | [5 empty] | Since the original report was filed the investigation into the access site bleeding has concluded. the impella pump and sheath product were not returned but, clinical details were reviewed. the physician noted that the patient was moving which resulted in bleeding form the access site. the failure mode will be monitored and trended. | The us complainant has not returned any product for benchtop analysis to root cause. should more information be shared that leads to root cause, a final report will be forthcoming. of note in the ifu there is the following noted: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
35730 | 1220648-2023-02018 | 16709270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16709270&pc=OZD | 03/12/2023 | 09/21/2024 |  | 04/10/2023 |  |  | 03/12/2023 | 04/10/2023 |  |  | 09/30/2024 |  | 04/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023284598 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Hispanic | American Indian Or Alaskan Native | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2023 |  |  |  |  | The complainant reported an 86-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. during insertion of a 12 fr dilator for impella implant, it was noted that a vascular dissection had occurred. it was specified that once the dilator was inserted, it appeared as if the dilator would not advance. further evaluation discovered that the dilator was in a different plain for the artery. the dilator was removed and a balloon tamponade was implemented via an inserted sheath. a covered stent was then deployed to treat the dissection. the impella sheath was subsequently placed without further issue. | Survival outcome at explanted noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu, vascular dissection is a known potential adverse event associated with impella support: "cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
35731 | 1220648-2023-02019 | 16709489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16709489&pc=OZD | 03/10/2023 | 04/10/2023 |  | 04/10/2023 |  |  | 03/12/2023 | 04/10/2023 |  |  | 11/30/2024 |  | 04/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023282441 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/12/2023 |  |  |  |  | The complainant reported a 62-year-old male patient presenting for cardiomyopathy. during impella support, it was documented that the patient developed a bleed around the graft at the insertion site. a washout was performed and the patient received four units of prbc to counteract the blood loss. patient support continued following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu, bleeding is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿
35732 | 1220648-2023-02020 | 16709501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16709501&pc=OZD | 01/29/2023 | 07/06/2023 |  | 04/10/2023 |  |  | 01/29/2023 | 04/10/2023 |  |  | 10/31/2024 |  | 04/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023268086 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2023 |  |  |  |  | The complainant reported an 62-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was documented that the patient experienced a ventricle perforation. following placement of the impella pump, it was documented that the patient experienced pericardial effusion. the patient's chest was opened and it was discovered that the left ventricle had been perforated. a surgical repair was performed to treat the perforation. roughly five days later, the pump was explanted as a result of ongoing ventricular tachycardia. the patient was defibrillated multiple times as a result of the vt and the pump was repositioned, however the condition continued. following explant, the vt stopped occurring. the patient was stable following pump removal. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the ventricle perforation has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis; only the clinical report was evaluated. based on the clinical details, the cause of the ventricle perforation was most likely improper positioning of the device. the failure mode will be monitored and trended.
35733 | 1220648-2023-02021 | 16709527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16709527&pc=OZD | 03/13/2023 | 04/24/2023 |  | 04/10/2023 |  |  | 03/13/2023 | 04/10/2023 |  |  | 09/30/2024 |  | 04/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023263108 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Hospitalization | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2023 |  |  |  |  | The medical center had an impella 5.5 support ongoing for 34 days, for a 59 year old cardiomyopathy patient, when there was suction and an observation made of a gastrointestinal (gi) bleed. there was concern from the medical team that the 5.5 support and continuous flow caused the avm (ateriovenous malformation) and the injury was due to the impella. the team adjusted the purge fluid and heparin. there is an allegation of impella causing harm to the patient that at this time can not be proven or disproven. the patient is awaiting transplant and has survived. | [5 empty] | Since the original report was filed, the investigation has concluded. the cause was patient condition related. the patient's arteriovenous malformation was investigated by the review of the data logs, which simply revealed suction, which indicated poor patient condition. the failure mode will be monitored and trended. | The medical center has not returned any pump product yet for analysis. the clinical report is being evaluated. the medical center has revealed that the bleed diagnosis was difficult in the scoping of the patient. the team has infused multiple units of blood and adjusted the purge. the investigation is ongoing at this time. upon completion of the investigation, a final report will be filed.
35734 | 1220648-2023-02002 | 16700315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16700315&pc=OZD | 03/10/2023 | 05/16/2023 |  | 04/07/2023 |  |  | 03/10/2023 | 04/07/2023 |  |  | 01/31/2025 |  | 04/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED INC. |  |  | US |  | 2024296611 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/10/2023 |  |  |  |  | The medical center had an impella 5.5 and ecmo support ongoing for 14 days. the team had been able to remove ecmo after 8 days. the 5.5 pump suddenly stopped on day 14. the team troubleshot but, decision made to explant the pump. there was no lasting harm or injury. the pump was not replaced. at the time of explant biomaterial was noticed adhered to the pump despite assurances that the patient was effectively anticoagulated. | [5 empty] | Since the original report was filed, the investigation into the pump stop has concluded. the pump was returned for benchtop analysis. the pump was seen to have biomaterial on the inlet an outlet of the pump. additionally, there were open lines within the red handle of the pump. the open lines caused an electrical failure of the pump. this would have been due to excessive force. the failure will be monitored and trended. as noted in the ifu: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The medical center has returned for benchtop analysis the impella pump product. the investigation into the pump stopping is ongoing. upon completion of the investigation into the sudden pump stop, a final report will be forthcoming. of note per the ifu "if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35735 | 1220648-2023-02008 | 16700915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16700915&pc=OZD | 03/08/2023 | 05/10/2023 |  | 04/07/2023 |  |  | 03/08/2023 | 04/07/2023 |  |  | 10/31/2024 |  | 04/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2024286752 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2023 |  |  |  |  | The complainant reported a 68-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient experienced an embolic stroke. the source for the stroke could not be definitively determined. patient support continued following the event. | [5 empty] | Following the initial report, the pump was returned for further evaluation. a visual inspection identified biomaterial on the impeller, however, the investigation was unable to establish a definitive relation between the impella pump and the noted embolic stroke. should additional relevant information become available, a supplemental report will be submitted. | The device was returned and an evaluation is pending. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results. as noted in the impella ifu, this is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35736 | 1220648-2023-02009 | 16702715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16702715&pc=OZD | 03/10/2023 | 05/26/2023 |  | 04/07/2023 |  |  | 03/10/2023 | 04/07/2023 |  |  | 08/31/2024 |  | 04/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023272626 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/10/2023 |  |  |  |  | The complainant reported a 58-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient developed limb ischemia. it was documented that the peel away sheath was left in the groin following impella implant per the physician's preference. roughly four days later, the patient's right foot looked mottled and a pulse was difficult to find. as a result of the condition, the pump was explanted and the peel away sheath was removed. a fasciotomy was subsequently performed to return blood flow to the patient's foot. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that a use issue was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support:¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35737 | 1220648-2023-02010 | 16704052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16704052&pc=OZD | 03/09/2023 | 06/02/2023 |  | 04/07/2023 |  |  | 03/09/2023 | 04/07/2023 |  |  | 11/30/2024 |  | 04/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2024299508 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/09/2023 |  |  |  |  | The complainant reported an 80 year old female patient presenting for high risk percutaneous coronary intervention using the impella cp for hemodynamic support. a difficult insertion was noted that required multiple attempts, traversing both the inanimate artery and ascending aorta using the supplied 14 fr introducer and a subsequent cook sheath. the pump was ultimately delivered, however, much resistance was noted. upon explant, the patient was slow to respond, indicating a likely stroke due to complicated delivery attempts. | [5 empty] | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported stroke/cva has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the stroke/cva. the failure modes will be monitored and trended. instructions for use for the related event are as follows: ¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35738 | 1220648-2023-02011 | 16704279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16704279&pc=OZD | 03/10/2023 | 05/26/2023 |  | 04/07/2023 |  |  | 03/10/2023 | 04/07/2023 |  |  | 09/30/2024 |  | 04/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023257930 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/10/2023 |  |  |  |  | The complainant reported a 63-year-old male patient presenting for cardiomyopathy. during impella support, it was noted that the patient began to feel numbness in their right arm. a vascular analysis was performed and it was discovered that there was a dissection in the axillary artery through which the pump had been implanted. as a result of the ischemia, the pump was removed and the patient¿s condition stabilized. | [5 empty] | The investigation into the reported limb ischemia and vascular damage has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the limb ischemia. the root cause for vascular damage remains undetermined since neither the product nor sufficient clinical information were provided. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35739 | 1220648-2023-02003 | 16688629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16688629&pc=OZD | 03/09/2023 | 07/26/2023 |  | 04/06/2023 |  |  | 03/09/2023 | 04/06/2023 |  |  | 10/31/2024 |  | 04/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2024296940 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/09/2023 |  |  |  |  | The complainant reported a 63-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. following impella explant, it was noted that the patient developed bleeding at the access site. manual pressure was not enough to manage the bleed, and four perclose coinciding with placement of an 8fr angioseal were necessary to achieve hemostasis. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. the impella ifu notes the following statements, in relationship to the reported symptom: "assess access site for bleeding and hematoma.¿ | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the access site bleeding was most likely patient condition related. the failure mode will be monitored and trended. as noted in the impella ifu, ischemia is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35740 | 1220648-2023-02004 | 16688630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16688630&pc=OZD | 03/08/2023 | 07/12/2023 |  | 04/06/2023 |  |  | 03/08/2023 | 04/06/2023 |  |  | 12/31/2024 |  | 04/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 380 (US) |  |  | ABIOMED, INC |  |  | US |  | 2024303734 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2023 |  |  |  |  | The complainant reported a 66-year-old male patient presenting for high risk percutaneous cardiac insertion. upon completion of support and explant of the pump, excessive bleeding was noted at the access site. the utilized preclose sutures were inadequate to treat the bleed and a stent was subsequently placed to achieve hemostasis. the patient was stable following the event. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. the impella ifu notes the following statements, in relationship to the reported symptom: "assess access site for bleeding and hematoma.¿ | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on the clinical details, the root cause of the access site bleeding was determined to be use related and the pre-close sutures failed; there was no allegation against impella. the failure mode will be monitored and trended.
35741 | 16688853 | 16688853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16688853&pc=OZD | 03/06/2023 | 03/14/2023 | 03/27/2023 | 04/06/2023 | 03/27/2023 | 04/06/2023 |  | 04/06/2023 |  |  |  |  | 04/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Yes |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 17520 DA |  | Male | 96 KG |  | Black Or African American | Malfunction | No | Yes | Infusion or Flow Problem | Output Problem | [5 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Device placed on [date redacted]. on [date redacted], device alarmed following physical therapy session. flows had dropped from 2.5-3l to 1l. staff turned down performance level to 0 and back up to 5 without any change in flow. waveforms also not normal. patient asymptomatic. impella rep and cardiologist informed. chest xray and echo ordered. cardiologist and rep came to bedside and worked to reposition impella under echo for approximately 30 minutes. placement was obtained (46.5cm per measurement) visually but flows never got above 1.2l and waveforms remained abnormal. patient continued to be asymptomatic. both cardiologist and rep said would watch over weekend and make decision [date redacted] about removal/exchange/etc. rep did mention could be a clot in/on device. patient on iv heparin. aptts variable, being adjusted per protocol. no heparin in impella purge at this time. on [date redacted], the impella alarmed. staff turned impella performance level down to 0 then up to 4. flows went to 4l. decision made by cardiologist and rep to remove the device.
35742 | 1220648-2023-02006 | 16695996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16695996&pc=OZD | 03/07/2023 | 04/06/2023 |  | 04/06/2023 |  |  | 03/07/2023 | 04/06/2023 |  |  | 08/31/2024 |  | 04/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023263792 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/07/2023 |  |  |  |  | The complainant reported a patient of unknown age and gender presenting for impella support. during removal of the guidewire, it was noted that the component uncoiled and was missing its tip upon explant. the location of the missing tip could not be identified. the pump was removed as a result of the noted issue and a new impella cp was implanted for patient support. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35743 | 1220648-2023-02007 | 16696659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16696659&pc=OZD | 06/23/2022 | 10/24/2023 |  | 04/06/2023 |  |  | 03/07/2023 | 04/06/2023 |  |  | 03/31/2024 |  | 04/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023205512 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2023 |  |  |  |  | The complainant reported a 62 year old male patient presenting with cardiomyopathy for impella support. approximately 6 days into support, the patient developed ischemia in the impella inserted extremity. a thrombus was also noted both in the jugular vein and left ventricle. intervention was required, however, it's unclear exactly what that intervention was and what factor the thrombus played in the event. | [5 empty] | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation has completed. since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the thrombus issue was most likely due to patient condition related, and the root cause of the limb ischemia issue was unable to be determined since product was not returned and with insufficient clinical information, no data logs being provided. the failure mode will be monitored and trended.
35744 | 1220648-2023-01998 | 16681539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16681539&pc=OZD | 02/16/2023 | 04/27/2023 |  | 04/05/2023 |  |  | 02/16/2023 | 04/05/2023 |  |  | 09/30/2024 |  | 04/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023283276 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2023 |  |  |  |  | The medical center had placed an impella cp via the axillary insertion site. during the support the team noted bleeding from the axillary site of the cp. the team attempts of troubleshooting by sponges and reinforcement failed. the team had to infuse blood products back to the patient. | [5 empty] | The medical center discarded the impella product at the time of explant. no benchtop analysis to root cause is possible. further clinical questions have been asked of the medical team. should a root cause be determined, a final report will be forthcoming. the failure mode will be monitored and trended. | Since the original report was filed, the investigation has concluded. without product return, the clinical details alone was analyzed. root cause was most likely patient condition related as the high blood pressure increased the bleeding and made it tough for hemostasis to be achieved. failure mode will be monitored and trended.
35745 | 1220648-2023-02001 | 16682148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16682148&pc=OZD | 03/07/2023 | 07/14/2023 |  | 04/05/2023 |  |  | 03/07/2023 | 04/05/2023 |  |  | 11/30/2024 |  | 04/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2024300642 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2023 |  |  |  |  | The complainant reported that a patient developed hemolysis during impella support. due to the noted condition, the pump was explanted and the patient was escalated to an impella 5.5 pump for ongoing support. the patient was considered stable following the event. | [5 empty] | The device has been returned and an evaluation is pending. upon conclusion of the investigation, a supplemental report will be filed with a summary of the findings. | The investigation into the reported hemolysis has been completed since the original report was filed. multiple suction alarms are seen in the lt log during the beginning of the case before the rise in purge pressure. the pump was investigated and biomaterial was found in the purge gap. no other abnormalities were observed. the cause of the hemolysis was biomaterial. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35746 | 1220648-2023-02000 | 16677994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16677994&pc=OZD | 03/08/2023 | 06/23/2023 |  | 04/04/2023 |  |  | 03/08/2023 | 04/04/2023 |  |  | 08/31/2024 |  | 04/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023270106 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2023 |  |  |  |  | The medical center had an impella cp placed via the femoral for support of a critical patient. when in the icu the team observed the access site to be saturated and treated by the infusion of 2 units of blood. of note the patient was on multiple anticoagulants and had an elevated ptt lab value reported. the pump remained on for successful hemodynamic support for 11 days until care was withdrawn by the family. | [5 empty] | The medical center has not yet returned for analysis any impella pump or introducer sheath product. the return is expected. upon completion of the investigation, a final report will be forthcoming. | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the access site bleeding with bleeding occurring from multiple sites. the failure mode will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35747 | 1220648-2023-01999 | 16678004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16678004&pc=OZD | 03/08/2023 | 09/24/2025 |  | 04/04/2023 |  |  | 03/08/2023 | 04/04/2023 |  |  | 01/31/2025 |  | 04/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED INC. |  |  | US |  | 2024297784 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/08/2023 |  |  |  |  | The medical center had an impella 5.5 support ongoing, via an axillary graft insertion site, for a 39 year old patient with congenital disease history awaiting a heart transplant. the axillary site developed a hematoma that was washed out on two separate occasions in the operating room. the washouts were all that was needed, as no blood products were infused back to the patient. the patient had his pump explanted after 19 days when the patient bridged to transplant. | The complainant also reported that the patient's pump had an irregular placement signal and false readings during the support. |  |  |  |  | The medical center has not yet returned for analysis and benchtop testing the impella 5.5 pump. the return is expected. upon completion of the investigation, a final report will be forthcoming. | Since the original report was filed the investigation has concluded. the pump was returned for analysis. the cause of the access site bleeding was most likely use related since it was believed to be related to the glue the site used at the anastomosis. the failure mode will be monitored and trended. of note the ifu stated the following: ¿assess access site for bleeding and hematoma." | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the root cause of the optical signal issue was most likely biomaterial related. no change to 5s parameters during the shift in placement signal indicates the issue was an interaction with the patient. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2023-01999. b7 other relevant history, including preexisting medical conditions updated. d4 model number updated. d4 unique identifier (udi) # updated. d6a / d6b updated. d10 concomitant medical products and therapy dates updated. f6 and f8 have erroneous entries in the initial report - should have been left blank. g1 reporting contact email updated.
35748 | 16663166 | 16663166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16663166&pc=OZD | 03/23/2023 | 03/24/2023 | 03/24/2023 | 04/03/2023 | 03/24/2023 | 04/03/2023 |  | 04/03/2023 |  |  |  |  | 04/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED INC. |  |  | US |  | 1707810 |  | 0042-0040-US |  | Yes |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 16060 DA |  | Male | 113 KG |  | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | Product Quality Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Staff noticed the sodium bicarb was leaking from the cassette in the abiomed impella circuit. this has happened half a dozen times so far, some of the incidents lead to purge pressure alarm. this one did not alarm but the staff replaced the cassette after noticing fluid leaking. manufacturer response for impella controller & cassette, impella (per site reporter) they are going to investigate upon receiving the cassettes. they have not seen this issue in any other sites.
35749 | 1220648-2023-01997 | 16671094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16671094&pc=OZD | 03/06/2023 | 07/12/2023 |  | 04/03/2023 |  |  | 03/06/2023 | 04/03/2023 |  |  | 02/29/2024 |  | 04/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023199833 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Valvular Insufficiency/ Regurgitation | Aortic Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2023 |  |  |  |  | In (b)(6) 2022 the medical center had an impella cp support for a patient admitted in cardiogenic shock for 5 days when explanted successfully. later in (b)(6) 2023 physicians involved in the care of the patient reported in an acc abstract poster that there had been valve damage. the valve damage was observed a week after pump explant on echo. the team presumed the aortic valve regurgitation and possible perforation of the non-coronary cusp to have occurred at the time of explant. abiomed personnel reached out to team in (b)(6) 2023 and discovered the valve was surgically replaced. | [5 empty] | The root cause of the valve damage is not known at this time. the authors are being contacted and further clinical details requested. upon completion of the investigation a final report will be forthcoming. the failure mode will be monitored and trended. | The investigation into the reported heart valve damage has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the root cause for damage remains undetermined since neither the product nor sufficient clinical information were provided. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35750 | 1220648-2023-01996 | 16671095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16671095&pc=OZD | 03/06/2023 | 04/13/2023 |  | 04/03/2023 |  |  | 03/06/2023 | 04/03/2023 |  |  | 11/30/2024 |  | 04/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED INC. |  |  | US |  | 2023275812 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2023 |  |  |  |  | The medical center had placed an impella 5.5 for support while awaiting heart transplant. during the support after approximately 10 days there were signs of a stroke (cva). the patient had aphasia, blurred vision, and limb weakness. the team removed the patient from transplant list due to the cva and performed thrombectomy in interventional radiology. the patient has survived the event and the patient is awaiting re-eval for the transplant list. | [5 empty] | The medical center has not returned the impella pump product or any additional clinical details or imaging. the root cause is not yet determined for the cva. upon completion of the investigation, a final report will be forthcoming. the failure mode will be monitored and trended. | Since the original report was filed the investigation has concluded. only data logs were returned for investigation. the root cause of the cva could not be determined without more clinical details and imaging. there is no evidence to confirm that the stroke was due to impella as imaging was not provided and no biomaterial interactions can be confirmed in the data logs. of note in the ifu is the following statement regarding cva: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35751 | 1220648-2023-01992 | 16660495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16660495&pc=OZD | 11/09/2021 | 10/17/2025 |  | 03/31/2023 |  |  | 03/03/2023 | 03/31/2023 |  |  | 11/30/2022 |  | 03/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1452229 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/03/2023 |  |  |  |  | Additional information noted that patient experienced bleeding and heparin was withheld. | The complainant reported a 60 year old male patient presenting for impella support. the patient developed thrombocytopenia in relationship to the impella device. as treatment, a platelet transfusion was delivered. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2023-01992. d6a and d6b revised from the initial manufacturer device report 1220648-2023-01992 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-01992 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-01992 in accordance with updated procedures. h4 and h5 was inadvertently omitted on the initial manufacturer device report number 1220648-2023-01992. h6 component code revised from the initial submission in accordance with updated procedures. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35752 | 1220648-2023-01993 | 16660547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16660547&pc=OZD | 03/03/2023 | 07/13/2023 |  | 03/31/2023 |  |  | 03/03/2023 | 03/31/2023 |  |  | 08/31/2024 |  | 03/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED INC. |  |  | US |  | 2023267244 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2023 |  |  |  |  | The medical center in japan had an impella 5.5 placed for a post operative patient who had undergone aortic arch replacement surgery and was in shock. the team had added 5.5 support and va ecmo. during the support the team did not fix the pump appropriately and the pump was "loosely fixed" and the team additionally did not peel away the sheath correctly per abiomed ifu. there was an left ventricle (lv) perforation that may or may not have been caused by the impella pump moving out of ideal position. | [5 empty] | The investigation into the reported ventricle perforation has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed the cause of the ventricle perforation was use related and was related to the insertion technique and handling of the peel-away sheath. in addition, there was not proper fixation. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The medical center in japan has not returned for analysis or benchtop testing. the product was discarded at the time of explant with bridge to transplant. the team has asked for further clinical details which should help determine causality of the lv perforation. should more information be shared and root cause be determined, a final report will be forthcoming. the failure mode will be monitored and trended.
35753 | 1220648-2023-01994 | 16660551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16660551&pc=OZD | 03/03/2023 | 05/05/2023 |  | 03/31/2023 |  |  | 03/03/2023 | 03/31/2023 |  |  | 11/30/2024 |  | 03/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2024295722 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Material Deformation | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2023 |  |  |  |  | The medical center had an impella cp support via femoral artery placement to a support a patient admitted suffering from a myocardial infarction. the patient developed an access site injury/vessel perforation at the site. the team had to achieve hemostasis by contralateral balloon access and removal of the impella. the vascular perforation timing is unknown, however the team did note the dilator had kinked upon insertion, and still used by the operator. they estimated 1 unit of blood was lost by the patient. | [5 empty] | The investigation has concluded since the original report was filed. the root cause of the access site bleeding was most likely patient condition. the medical team reported that gaining access to the vessel was difficult and a dilator kinked in the process likely leading to the vessel damage. no product was returned for this investigation. the failure mode will be monitored and trended. | The medical center has not returned for analysis any impella product, neither the pump nor the introducer sheath which was noted to kink at time of insertion. the investigation into the case is ongoing at this time and upon completion, should a root cause be determined without product return, a final report will be forthcoming. the failure mode will be monitored and trended.
35754 | 1220648-2023-01995 | 16660578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16660578&pc=OZD | 03/03/2023 | 06/23/2023 |  | 03/31/2023 |  |  | 03/03/2023 | 03/31/2023 |  |  | 01/31/2025 |  | 03/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED INC. |  |  | US |  | 2024297795 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2023 |  |  |  |  | The medical center had an impella 5.5 support initiated for a patient with coronary artery disease and emergent surgery. the pump was placed but within hours the pump stopped and was restarted. the pump restarted with low flows and was explanted. there was a question from the medical team of thrombus ingestion causing the pump to stop. | [5 empty] | The investigation into the reported pump stop has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the data logs were evaluated. data logs showed that the pump stopped immediately upon the sudden occurrence of the "impella stopped - motor current high" alarm, accompanied by multiple motor current spikes. the motor current was non-pulsatile and the pump flow was low with erratic values leading up to the stoppage, which occurred within 30 minutes of pump activation. the placement signal waveform also fluctuated. the pump was able to be successfully restarted about one minute later. however, low flow and suction alarms persisted following the pump stop. in conclusion, it was determined that the cause of the pump stop was most likely ingested biomaterial. the evaluation revealed that ingested biomaterial was the most likely cause of the limb ischemia. the failure modes will be monitored and trended. | The medical center has not yet returned the impella pump for investigation and benchtop analysis. return is expected. upon completion of the investigation, a final report will be forthcoming. the failure mode will be monitored and trended.
35755 | 1220648-2023-01983 | 16642722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16642722&pc=OZD | 02/28/2023 | 03/30/2023 |  | 03/30/2023 |  |  | 02/28/2023 | 03/30/2023 |  |  | 07/31/2024 |  | 03/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023250513 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/28/2023 |  |  |  |  | The complainant reported a 64-year-old male patient presenting with cardiomyopathy. during impella support, it was documented that the patient developed a hematoma at the right groin. up until that point in time, the patient had instances of oozing around the access site that resolved via repositioning or manual pressure. the patient was provided 1 unit prbc and the hematoma was subsequently evacuated. the patient was considered stable following the intervention and impella support continued. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35756 | 1220648-2023-01984 | 16642736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16642736&pc=OZD | 01/20/2023 | 07/13/2023 |  | 03/30/2023 |  |  | 01/20/2023 | 03/30/2023 |  |  | 10/31/2024 |  | 03/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2024285938 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2023 |  |  |  |  | The complainant reported a 52-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was decided that the patient required escalation to an impella 5.5 pump for continued support. however, upon explant of the already in place impella cp, the patient developed an access site bleed for which the physician had difficulty achieving hemostasis. a surgical cutdown was subsequently performed to treat the bleed and patient support continued with the impella 5.5 pump. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. root cause of the issue cannot be determined since the complaint device was not returned for investigation and there is a lack of clinical information the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35757 | 1220648-2023-01985 | 16642745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16642745&pc=OZD | 01/24/2023 | 03/30/2023 |  | 03/30/2023 |  |  | 01/25/2023 | 03/30/2023 |  |  | 10/31/2024 |  | 03/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023268881 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/25/2023 |  |  |  |  | The complainant reported a 77-year-old male patient presenting for multiple chronic conditions. during the initial implant attempt for an impella 5.5 pump, it was documented that the device punctured the left ventricle. the perforation was surgically repaired and the pump was subsequently implanted successfully for patient support. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35758 | 1220648-2023-01986 | 16642759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16642759&pc=OZD | 03/05/2023 | 10/25/2023 |  | 03/30/2023 |  |  | 03/05/2023 | 03/30/2023 |  |  | 12/31/2024 |  | 03/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC |  |  | US |  | 2024292752 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | No | Yes | Crack | Pumping Stopped | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/05/2023 |  |  |  |  | Additionally, information was received noting a cracked peel away introducer that was resolved by removal/replacement prior to initiation of the pump. this issue was resolved by removing and replacing with a new introducer & axillary insertion kit. | The complainant reported a 48-year-old male patient presenting for cardiomyopathy. during impella support, it was noted by the physician that the device was experiencing suction and positional alarms coinciding with a flat motor current. the pump subsequently stopped but re-started immediately. however, the saturated flow issue continued and the pump stopped a second time. the pump could not be restarted following the second stop. as a result of the failure, the pump was explanted and it was noted that a clot was visible around the cannula. the patient was considered to be stable following explant and an intra-aortic balloon pump was implanted to continue with patient support. |  |  |  |  | The investigation into the pump stop has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis; only the clinical report was evaluated. based on the clinical details, the cause of the pump stop was biomaterial. the failure mode will be monitored and trended. | Additional information added to b5 and h6 related to the report of a cracked introducer. the investigation into the pump stop, the thrombus, and the saturated flow has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the lt log shows a suction event on the second day of support. at this time the flow drops, and the motor current loses pulsatility. two days later it can be seen in the lt log that there is a motor current spike and the motor current increases with increased resistance. the purge pressure also rises, and the flow becomes saturated at the time of the initial stop. the pump can be restarted until the pump stops again and multiple restart attempts are made until the pump is explanted. the cause of the pump stop was biomaterial based on the trends seen in the data logs. a cause for the cracked introducer was not established. | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35759 | 1220648-2023-01988 | 16642804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16642804&pc=OZD | 01/26/2023 | 06/21/2023 |  | 03/30/2023 |  |  | 01/26/2023 | 03/30/2023 |  |  | 09/30/2024 |  | 03/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023276819 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2023 |  |  |  |  | The complainant reported a 59-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient experienced an access site bleed coinciding with a drop in hgb levels. to treat the bleed, the patient was provided two units of prbc. the patient was subsequently escalated to an impella 5.5 pump to continue with support. | [5 empty] | The investigation into the positioning issues has been completed since the original report was filed. the medical center discarded the impella product at the time of explant. no benchtop analysis of the root cause of the access site bleeding was possible; only the clinical report was evaluated. based on the clinical data, use issue was the most likely cause of the positioning issues because the pump migrated too deep to the ventricle. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35760 | 1220648-2023-01987 | 16642805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16642805&pc=OZD | 03/03/2023 | 07/12/2023 |  | 03/30/2023 |  |  | 03/03/2023 | 03/30/2023 |  |  | 05/31/2024 |  | 03/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023275967 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 50-year-old male suffered from heart failure and cardiogenic shock. he was implanted with an impella 5.0 cardiac pump for hemodynamic support. after 4.5 hours and problems with the purge system, the pump stopped. the pump could not be restarted and was replaced with a new impella 5.0 device. subsequently, the patient has been stable. | [5 empty] | The investigation into the purge leak has been completed since the original report was filed. the user facility has returned the impella device, and the benchtop analysis of the root cause revealed that the cause of the purge leak was failure of the thf bond at the cartridge cap. the failure mode will be monitored and trended. | The impella device has been returned to the manufacturer. when the investigation is complete, a supplemental report will be filed.
35761 | 16643460 | 16643460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16643460&pc=OZD | 02/23/2023 | 03/08/2023 | 03/08/2023 | 03/30/2023 | 03/08/2023 | 03/30/2023 |  | 03/30/2023 |  |  |  |  | 03/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | PURGE CASSETTES |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0043-0003 |  | Yes |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 9125 DA |  | Female |  |  |  | Malfunction | No | Yes | Leak/Splash | [6 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella purge set leaking. equipment saved for manufacturer to follow-up. no other info has been received. submitter stated the equipment had been sequestered for manufacturer to examine. uncertain if that has occurred.
35762 | 1220648-2023-01989 | 16651651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16651651&pc=OZD | 02/20/2023 | 06/16/2023 |  | 03/30/2023 |  |  | 02/28/2023 | 03/30/2023 |  |  | 09/30/2024 |  | 03/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023284569 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2023 |  |  |  |  | The complainant reported an 89 year old female patient presenting for high risk percutaneous coronary intervention using the impella for hemodynamic support. after the procedure, post impella removal, an access site bleeding complication occurred. the perclose closure device suture snapped and manual pressure/femstop was applied. upon follow up with the physician, it was revealed that the patient had expired as a result of the bleeding complication. the physician alleged the bleeding complication was related to the introducer, however, no further explanation could be provided. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report and real time data logs were evaluated. there were no abnormalities found in the real time data logs. the cause of the access site bleeding was not determined since product was not returned. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35763 | 1220648-2023-01990 | 16651689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16651689&pc=OZD | 02/20/2023 | 08/21/2023 |  | 03/30/2023 |  |  | 02/28/2023 | 03/30/2023 |  |  |  |  | 03/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | UNSPECIFIED IMPELLA PUMP |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2023 |  |  |  |  | The complainant reported an impella patient of unknown age or gender presenting for impella support on an unknown date. a groin bleed had developed post impella removal and the patient expired on an unknown date. the introducer supplied in the impella kit was alleged in relationship to the bleed/death. attempts were made to obtain additional details from the complainant without success. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information provided, the cause of the access site bleeding was not determined since neither product nor data logs were returned and there is insufficient clinical information. | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35764 | 1220648-2023-01991 | 16652005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16652005&pc=OZD | 03/02/2023 | 03/30/2023 |  | 03/30/2023 |  |  | 03/02/2023 | 03/30/2023 |  |  | 08/31/2024 |  | 03/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023256209 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/02/2023 |  |  |  |  | The (b)(6) in japan had an impella cp support ongoing for a patient in cardiogenic shock and having a coronary angioplasty. the support was for a day when the pump was explanted. on the day of explant an embolic stroke (cva) was noted. the causal relationship of the cva to the pump explant could not be proven or denied. | [5 empty] | The (b)(6) in japan discarded the impella pump product after explant. the cva was diagnosed after pump explant. the exact time of the stroke could not be determined as the patient was sedated during the entire time of the impella support, and as such the causality could not be proven or denied. further clinical details have been asked but, not yet furnished. root cause can not yet be determined. should a cause be determined, a final report will be forthcoming. the failure mode will be monitored and trended.
35765 | 1220648-2023-01979 | 16627815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16627815&pc=OZD | 02/27/2023 | 06/16/2023 |  | 03/28/2023 |  |  | 02/27/2023 | 03/28/2023 |  |  | 01/31/2025 |  | 03/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024297888 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An 18-year-old male patient suffered from dizziness, shortness of breath and lack of energy. patient showed heart failure with ejection fraction of 9%. an impella 5.5 device has successfully been inserted for hemodynamic support. during transport in the hospital, the impella was caught in a bedrail and pulled back. the device has not been dislocated inside the patient. subsequently, the pump stopped and could not be restarted. a new impella 5.5 device has successfully been inserted. patient stable afterwards, on bridge to lvad. | [5 empty] | The investigation into the reported pump stop has been completed since the original report was filed. the data logs confirmed a sudden pump stop. the medical center returned impella products for benchtop analysis. electrical resistance testing showed all 3 motor cable lines were open. the red handle was opened and the motor cable lines were found to be broken on the catheter side; evidence of necking of the cables was present. the root cause of the pump stop was determined to be an electrical open in all 3 motor cables caused by excessive force that was applied when the pump got caught on the bedrail during transport. the cause of the issue was an open electrical line. the failure modes will be monitored and trended. | Benchtop analysis of the returned pump is ongoing. when the investigation is complete, a supplemental report will be filed.
35766 | 1220648-2023-01981 | 16632663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16632663&pc=PYX | 02/28/2023 | 04/17/2023 |  | 03/28/2023 |  |  | 02/28/2023 | 03/28/2023 |  |  | 08/31/2024 |  | 03/28/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2024289365 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2023 |  |  |  |  | The medical center was placing an rp flex for support of a patient on bypass post valve replacement surgery for the mitral and tricuspid valves. positioning was complicated and the rp flex coming back across the pulmonic valve into the right ventricle (rv). the pump positioning may have contributed to the perforation of the rv. the pump was seen to be protruding through the ventricular tissue. the rv was repaired, pump positioned, and the patient survived the event. | [5 empty] | The investigation into the right ventricle perforation has concluded since the original report was filed. product was returned for analysis. the root cause was determined to be user related. the pump was over delivered into the ventricle causing the cannula to kink and allow for the more rigid length of the rp pump to puncture the wall of the ventricle. additionally, as the ventricle volume was lacking this allowed for additional resistance on the rp pump. the user additionally lost fluoroscopic sight of the pump when delivering. the failure mode will be monitored and trended. the ifu notes the following: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | The medical center has returned the impella pump product for investigation and benchtop analysis. the analysis is on-going at this time. upon completion of the investigation, a final report will be forthcoming.
35767 | 1220648-2023-01980 | 16632665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16632665&pc=OZD | 02/12/2023 | 06/15/2023 |  | 03/28/2023 |  |  | 02/12/2023 | 03/28/2023 |  |  | 07/31/2024 |  | 03/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023255913 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2023 |  |  |  |  | The medical center had an impella cp placed for support via the right femoral in a patient with known peripheral arterial disease and small diameter vessels. the team noted limb ischemia and monitored the limb, considered thrombectomy, and finally explanted the pump and instead placed an iabp by the contralateral leg. beyond explant no other measures were performed. the thrombectomy was not deemed necessary. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the patient condition was the most likely cause of the access site bleeding. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The medical center discarded the impella pump product and introducer sheath product at the time of explant. no benchtop analysis to root cause is possible. additional clinical details and imaging have been asked for but, to date not furnished. the manufacturer will file a final report if root cause is determined with new details being shared for investigation. the failure mode will be monitored and trended.
35768 | 1220648-2023-01982 | 16632681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16632681&pc=OZD | 02/27/2023 | 03/28/2023 |  | 03/28/2023 |  |  | 02/27/2023 | 03/28/2023 |  |  | 12/31/2024 |  | 03/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024297194 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 68-year-old patient was treated with percutaneous coronary intervention for coronary artery disease. during the high-risk procedure, the patient received hemodynamic support from an impella cp cardiac pump. insertion of the pump was described as complicated because of severe calcification of the femoral artery. after successful completion of the coronary intervention and removal of the impella cp, a large hematoma formed at the access site. vascular reconstruction was performed at the access site, and the access site was surgically closed. the patient was subsequently stable and received ongoing support in the intensive care unit. | [5 empty] | The impella device has not been received by the manufacturer for investigation. should any new information be returned, a supplemental report will be filed.
35769 | 1220648-2023-01978 | 16622296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16622296&pc=OZD | 02/27/2023 | 03/27/2023 |  | 03/27/2023 |  |  | 02/27/2023 | 03/27/2023 |  |  | 07/31/2024 |  | 03/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023246684 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/27/2023 |  |  |  |  | The medical center in japan had an impella cp support ongoing when on day 6 there was a diagnosis of hemolysis. the team saw hematuria and proceeded to infuse haptoglobin for treatment. the pump remained on for support despite the presumed hemolysis til wean and explant the following day. | [5 empty] | The medical center in japan discarded all impella pump product at the time of explant. no benchtop analysis to root cause or hemolysis testing is possible. further clinical details have not been forthcoming, inclusive of how many units of haptoglobin was infused. the failure mode will be monitored and trended. should a root cause be determined, a report will be forthcoming.
35770 | 1220648-2023-01972 | 16614430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16614430&pc=OZD | 02/23/2023 | 10/25/2023 |  | 03/24/2023 |  |  | 02/23/2023 | 03/24/2023 |  |  | 01/31/2025 |  | 03/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC |  |  | US |  | 2024294354 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2023 |  |  |  |  | The complainant reported a 38-year-old male patient presenting for cardiomyopathy. during impella support, it was documented that the physician had unplugged the catheter to remove twists in the component and, upon re-insertion, they were unable to re-start the device. multiple attempts were made, however, the issue continued and an "impella catheter defective" alarm subsequently appeared on the aic. as a result of the ongoing issue, the pump was replaced with another impella 5.5 pump. there was no harm to the patient. | [5 empty] | (h3) the investigation into the reported "pump stop" has been completed since the original report was filed. product was returned for investigation. the device was visually inspected and found that there was no abnormalities found. the imc logs were investigated and showed that upon re- plug into the console the console was unable to read the eeprom of the impella 5.5. multiple crc can be seen. the root cause of the eeprom corruption is unable to be determined. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records | The pump was returned and an evaluation of the device is pending. upon completion of the investigation, a supplemental report will be filed.
35771 | 1220648-2023-01973 | 16614484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16614484&pc=OZD | 02/09/2023 | 03/24/2023 |  | 03/24/2023 |  |  | 02/09/2023 | 03/24/2023 |  |  | 08/31/2024 |  | 03/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023269894 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/09/2023 |  |  |  |  | The medical center had an impella cp support placed via the right femoral artery. the cp was placed at the tertiary center after patient transfer from the community facility. the cp limb was ischemic and a bypass externally was placed to return perfusion to the limb. the team later made choice to explant and escalate care to the axillary placed impella 5.5 pump. the cp had supported the pump for just over 11 hours. | [5 empty] | The medical center did not retain the impella pump or introducer sheath for any investigation or benchtop analysis to root cause. clinical details have been asked of the medical center, as has the request for imaging. we have learned that the patient had no known underlying peripheral arterial disease that may have contributed. should more details be shared that leads to root cause, a final report will be forthcoming.
35772 | 1220648-2023-01975 | 16614599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16614599&pc=OZD | 02/01/2023 | 03/24/2023 |  | 03/24/2023 |  |  | 02/01/2023 | 03/24/2023 |  |  | 09/03/2024 |  | 03/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023280177 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/01/2023 |  |  |  |  | The medical center had a 14fr sheath placed for insertion of the impella cp via the right femoral artery. the product placement was occlusive to the right limb and blood flow and so the team placed a reperfusion catheter. the flow restored and the pump remained on for support for 2 more days until wean and explant. the bypass catheter was explanted and flow was restored. no noted lasting harm or injury. | [5 empty] | The medical center discarded the impella pump and introducer sheath product at the time of explant. there can be no benchtop analysis of the product. further clinical details and anatomic imaging have been requested, but not yet furnished. should new information be shared and root cause determined, a final report will be forthcoming. the manufacturer will monitor and trend the failure mode.
35773 | 1220648-2023-01974 | 16614688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16614688&pc=OZD | 02/13/2023 | 03/24/2023 |  | 03/24/2023 |  |  | 02/13/2023 | 03/24/2023 |  |  | 09/30/2024 |  | 03/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023259151 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/13/2023 |  |  |  |  | The medical center had an impella 5.5 pump placed for support in which the team noted an access site hematoma observed which prompted surgical intervention/washout and placement of a jp drain. the pump remained on for support for the patient for over 15 days despite the harm. the patient then bridged to lvad. | [5 empty] | The medical center has discarded all impella pump and introducer product. no benchtop analysis to root cause is possible at this time. further clinical details and imaging have been requested but, not yet furnished. should more information be shared that leads to root cause of the hematoma, a final report will be forthcoming. the failure mode will be monitored and trended.
35774 | 1220648-2023-01976 | 16616461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16616461&pc=OZD | 02/24/2023 | 09/06/2023 |  | 03/24/2023 |  |  | 02/24/2023 | 03/24/2023 |  |  | 02/29/2024 |  | 03/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023199813 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2023 |  |  |  |  | The complainant reported a 47 year old male patient presenting with cardiomyopathy for impella support. during support, the patient developed hemolysis. as a result, the pump was removed and replaced with an impella 5.5. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information provided, the cause of the hemolysis was not determined since the product was not returned. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35775 | 1220648-2023-01977 | 16616465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16616465&pc=OZD | 01/29/2023 | 06/07/2023 |  | 03/24/2023 |  |  | 01/29/2023 | 03/24/2023 |  |  | 04/30/2024 |  | 03/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023214989 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2023 |  |  |  |  | The complainant reported an 84 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed ischemia, as evident by absence of distal pulses in the impella inserted limb. during use. the pump was removed and patient was escalated to an impella 5.5. after explant of the cp, the femoral artery in which the device traversed required vascular repair due to initial stick. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35776 | 1220648-2023-01971 | 16598348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16598348&pc=OZD | 02/23/2023 | 06/14/2023 |  | 03/23/2023 |  |  | 02/23/2023 | 03/23/2023 |  |  | 10/31/2024 |  | 03/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023284188 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  |  |  | Malfunction | Yes | Yes | Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Hold js 3/23. a 18-year-old male with no known past medical history developed symptoms of severe weakness and has been diagnosed anemic. he presented to the hospital and has been diagnosed with septic shock. echo found global biventricular heart failure. patient was emergency placed on va-ecmo. then, an impella cp has successfully been introduced. during change from peel-away introducer to repositioning sheath the impella device has been pushed into the left ventricle, which caused the cannula to kink. during several unsuccessful attempts to straighten and reposition the impella cp again, the cannula broke apart, staying partially connected by the nitinol wire. impella catheter and pump has been removed and nitinol wire was cut. the remaining inlet portion of the catheter could not be removed using a snare, so an aortotomy has been performed. nitinol wire was visually confirmed to be entangled/looped around the anterior mitral leaflet. attending physician reported that patient has anomalous papillary attached to septum which increased risk of entanglement. left femoral access site surgically closed per cutdown. patient has been escalated to an impella 5.5 and has been in a critical condition. | [5 empty] | Images received from the field revealed the cannula detached in the middle at the area where the cannula was significantly kinked within the lv. the detachment appeared rough with the nitinol unwinding and the polyurethane torn. there was no evidence of bond failure. in conclusion, after reviewing the clinical information, it was determined that the cause of the cannula detachment was most likely use related, specifically the excessive force used to unkink the cannula after the pump was inadvertently advanced too deep in the lv. | Benchtop investigation of the returned pump will be performed after return of the device. when the investigation is complete, a supplemental report will be filed.
35777 | 16582064 | 16582064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16582064&pc=OZD | 02/24/2023 | 03/13/2023 | 03/13/2023 | 03/21/2023 | 03/13/2023 | 03/21/2023 |  | 03/21/2023 |  |  |  |  | 03/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP WITH SMARTASSIST |  |  | OZD |  |  | IMPELLA CP WITH SMARTASSIST |  |  | ABIOMED, INC. |  |  | US |  | 2023284188 |  | 0048-0024 |  | Yes |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | No | Risk Manager | [6 empty] | 6570 DA |  | Male |  | Hispanic |  | Malfunction | No | Yes | Material Fragmentation | [6 empty] | Foreign Body In Patient | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | OTHER |  |  | At some point during the deployment of an impella cp from the left groin, the internalized pump fractured into two pieces. the distal piece appeared in mitral apparatus, and the leaflets were no longer coapting, with no wide-open mr, no ejection, and frothy pulmonary edema in the vent circuit. at this point, vascular surgery and ct surgery came to evaluate. the plan at this time was to try remove the proximal piece through the groin sheath, and a sternotomy and open-heart retrieval of the distal piece. the surgeon emergently removed the remaining impella device fragments from the heart and placed a direct aortic impella device. patient taken to critical care in critical condition.
35778 | 1220648-2023-01968 | 16582570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16582570&pc=OZD | 02/21/2023 | 06/20/2023 |  | 03/21/2023 |  |  | 02/21/2023 | 03/21/2023 |  |  | 01/31/2025 |  | 03/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024294798 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 30-year-old male with previous history of bi-ventricular heart failure of lv thrombosis that was thought to have resolved presented in cardiogenic shock. the attending physician was aware of the contraindications but felt the benefit outweighed the risk. the patient was to have an impella 5.5 inserted. insertion proved difficult due to small vessel size but has been successful. there have been suction alarms after starting the pump. due to high right heart pressures and hypotension, the patient has been escalated to a va ecmo. the patient subsequently stabilized. further investigation by echo revealed a thrombus in the left ventricle. the physician opted against removing the impella due to small vessel size and planned to remove the pump at the time of transplant. later, the impella 5.5 was removed due to hemolysis from suspected clot on the inlet which was confirmed after explant. patient survived. | [5 empty] | The impella device has not been returned to the manufacturer. when the investigation is complete, a supplemental report will be filed. | No product was returned for investigation. the lt log confirms the reported impella position in aorta alarms. the lt log also shows one impella position unknown alarm and several suction alarms in the beginning of the case. a returned photo from the field representative shows biomaterial after explant of the device. the cause of the hemolysis was most likely biomaterial based on the provided clinical information.
35779 | 1220648-2023-01969 | 16585673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16585673&pc=PYX | 02/21/2023 | 07/27/2023 |  | 03/21/2023 |  |  | 02/21/2023 | 03/21/2023 |  |  | 10/31/2024 |  | 03/21/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023279157 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Malfunction | Yes | No | Unable to Obtain Readings | [6 empty] | Liver Damage/Dysfunction | [12 empty] | Hospitalization | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 57-year-old male has been treated for right heart failure and received support by a vv-ecmo. additionally, an impella rp heart pump has been inserted to provide hemodynamic support. after 8 hours of support, the differential pressure sensor of the impella rp failed, therefore there was no flow calculation displayed. correct position of the impella device has been confirmed by tee. doppler echo also confirmed that the pump was still running. flow rate of the impella device unknown. the attending physician suspects that the flows provided by the impella rp were insufficient, causing liver disfunction. no pump exchange or other intervention known. total impella rp runtime 49 hours. patient survived, but condition critical, staying on ecmo support. | [5 empty] | The impella device has not been returned to the manufacturer yet. when the investigation is complete, a supplemental report will be filed. | The investigation into the placement signal issue and the patient injury has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the root cause of the placement signal issue was determined to be dp sensor drift, and the cause of the patient injury was most likely patient condition related. the failure modes will be monitored and trended.
35780 | 1220648-2023-01970 | 16585866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16585866&pc=OZD | 03/07/2023 | 03/21/2023 |  | 03/21/2023 |  |  | 03/07/2023 | 03/21/2023 |  |  | 09/30/2024 |  | 03/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023261037 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 29 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/07/2023 |  |  |  |  | The complainant reported a 29-year-old female patient presenting for cardiomyopathy. during impella support, it was documented that the patient developed. the hemolysis was verified via hematuria. the patient was provided four units of packed red blood cells as a result of the hemolysis and impella support continued. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35781 | 1220648-2023-01967 | 16571839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16571839&pc=OZD | 02/18/2023 | 06/30/2023 |  | 03/20/2023 |  |  | 02/18/2023 | 03/20/2023 |  |  | 03/31/2024 |  | 03/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023209923 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 12 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/18/2023 |  |  |  |  | The complainant reported a 12 year old male patient presenting for impella support. a bleed was noted around the access site, post insertion. further examination noted a vascular injury sustained to the celiac artery. the device was removed and a stent was placed as intervention. the physician could not rule out the 0.018" guidewire as a causal/contributory factor. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the reported vascular damage has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the vascular damage. the vascular damage was determined to be unrelated to the impella because the device does not contact the celiac artery with proper insertion technique. the failure modes will be monitored and trended.
35782 | 1220648-2023-01964 | 16569856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16569856&pc=PYX | 02/15/2023 | 02/20/2024 |  | 03/18/2023 |  |  | 02/15/2023 | 03/18/2023 |  |  | 07/31/2024 |  | 03/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023281618 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/15/2023 |  |  |  |  | The complainant reported a 54 year old female patient presenting for hemodynamic support using the impella rp flex device. during support, there was suspicion of hemolysis, as evident by elevated ldh labs. out of concern, the pump was removed. | Additional information received via loqi (abiomed¿s long-term outcome and quality indicator (loqi) impella registry, pertaining to this case. it was noted that, in addition to hemolysis, thrombocytopenia occurred with a possible relationship to the impella rp device. |  |  |  |  | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Since the original report, an investigation has been completed with regards to the allegation of hemolysis and the impella device. the product was returned with the catheter cut. no abnormalities on the pump were found. the pump was decontaminated and the motor cable lines were soldered back together so a hemolysis test could be conducted. the hemolysis test was completed and the resulting mih was a passing value of 3.08. the data logs were reviewed and found spikes in placement signal with increased p-levels, likely due to the pump migrating toward the vasculature., the positioning of the pump was confirmed through imaging, the root cause of the hemolysis was determined to be patient condition as the device placed after the rp flex, protek, also produced elevated ldh values and the returned rp flex pump passed hemolysis testing ruling out a malfunction of the device. since this report, an allegation of thrombocytopenia with a possible relationship to the impella was received. an additional investigation will be performed and an additional supplemental will be filed. b(5) - additional information received via loqi database pertaining to this patient enduring thrombocytopenia. added to b5. | The investigation into the hemolysis and the thrombocytopenia issues has been completed since the original report was filed. the device was returned and tested after arriving in fs. the root cause of the hemolysis was determined to be patient condition as the device placed after the rp flex, protek, also produced elevated ldh values and the returned rp flex pump passed hemolysis testing ruling out a malfunction of the device. however, the root cause of the thrombocytopenia was unable to be determined as limited clinical details were provided.
35783 | 1220648-2023-01965 | 16569886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16569886&pc=OZD | 02/17/2023 | 05/26/2023 |  | 03/18/2023 |  |  | 02/17/2023 | 03/18/2023 |  |  | 07/31/2024 |  | 03/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET APAC |  |  | ABIOMED, INC. |  |  | US |  | 2023257862 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Pseudoaneurysm | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2023 |  |  |  |  | The complainant reported a 63 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. an impella access site hematoma/pseudoaneurysm developed, along with a lack of pulses in the impella inserted limb. the pump was removed and vascular intervention was performed. | [5 empty] | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the access site bleeding. the failure mode will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support:¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35784 | 1220648-2023-01966 | 16569887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16569887&pc=OZD | 02/17/2023 | 03/17/2023 |  | 03/18/2023 |  |  | 02/17/2023 | 03/18/2023 |  |  | 11/30/2023 |  | 03/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC |  |  | US |  | 2023196582 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Eject | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/17/2023 |  |  |  |  | The complainant reported a 72-year-old male patient presenting for high risk percutaneous coronary intervention. upon attempting to explant the pump following support, the device got stuck in the sheath and could not be advanced or removed. as a result, a balloon needed to be placed above the access site and inflated to allow for removal. the patient was stable following the event. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35785 | 1220648-2023-01963 | 16563951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16563951&pc=OZD | 01/19/2023 | 03/17/2023 |  | 03/17/2023 |  |  | 01/19/2023 | 03/17/2023 |  |  | 09/30/2024 |  | 03/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023283193 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 72-year-old white male with a past medical history of stage 4 cancer with metastases presented with stemi, chest pain in the past 2 weeks and shortness of breath in the past 4 days. the patient decompensated, requiring intubation. an impella cp cardiac pump has been inserted for hemodynamic support. access site bleeding occurred the repositioning sheath. the surgeon successfully performed a cutdown to control the bleeding. the patient has been stable afterwards and has been transferred from the ccl to icu. | [5 empty] | Manufacturer has not received pump for investigation. root cause cannot be determined. should any new information be returned, a supplemental report will be filed.
35786 | 1220648-2023-01961 | 16554792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16554792&pc=OZD | 02/15/2023 | 06/23/2023 |  | 03/16/2023 |  |  | 03/02/2023 | 03/16/2023 |  |  | 10/31/2024 |  | 03/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023276325 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2023 |  |  |  |  | The complainant reported a 52-year-old male patient presenting for cardiomyopathy. during impella support, it was noted that the patient developed upper limb ischemia. it was specified that the patient lost pulse in the upper extremity on the impella side after the pump was implanted through the axillary. to treat the ischemia, a sheath was inserted from the radial artery and blood flow was provided via ecmo. patient support continued following the intervention. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center discarded the impella product at the time of explant. no benchtop analysis of the root cause of the limb ischemia was possible; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of the limb ischemia was related to patient condition. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35787 | 1220648-2023-01962 | 16560711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16560711&pc=OZD | 02/14/2023 | 06/30/2023 |  | 03/16/2023 |  |  | 02/14/2023 | 03/16/2023 |  |  | 08/31/2024 |  | 03/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023269541 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/14/2023 |  |  |  |  | The complainant reported a 72 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. the pump was removed prematurely due to distal limb ischemia. a dissection was noted via the impella access site in which the ischemia was attributed to. a balloon tamponade was required to achieve hemostasis. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35788 | 1220648-2023-01958 | 16547861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16547861&pc=OZD | 03/03/2023 | 09/22/2023 |  | 03/15/2023 |  |  | 02/21/2023 | 03/15/2023 |  |  | 11/30/2024 |  | 03/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023279527 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | A 32-year-old pregnant female presented to hospital with shortness of breath. cardiomyopathy was diagnosed. an impella 5.5 cardiac pump has been inserted to provide hemodynamic support and extend pregnancy. insertion of the impella device proved difficult due to patient anatomy. the impella has been removed and successfully reinserted, after mineral oil has been applied as lubricant. on day six after insertion the patient developed a worsening hematoma with swelling extending to her neck. the hematoma has been successfully cleaned out and a jp drain has been placed. the next day, the patient continued to have significant pain and swelling at site, but less firm. patient and baby have been stable all the time. two days later, baby successfully delivered per c-section and stable. seven days later patient became febrile with decreased level of consciousness. fungal sepsis has been suspected and confirmed. source of the infection remains unknown. contribution of the impella device cannot be ruled out. infection has been successfully treated with antifungal treatment. patient stabilized again and remained on impella support. case is ongoing and patient is being prepared for a durable lvad. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical details the root cause of the access site bleeding and the infection cannot be determined since the device was not returned for investigation and lack of clinical information. the physician was unable to advance cath when pushing blue release button, root cause of the catheter anchor issue cannot be determined since the complaint device not returned for investigation. | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical details, the root cause of the access site bleeding and the infection cannot be determined since the device was not returned for investigation and there is a lack of clinical information. the physician was unable to advance cath when pushing blue release button, root cause of the cath anchor issue cannot be determined since the complaint device not returned for investigation.the failure mode will be monitored and trended. | Patient is still on impella support and investigation ongoing. a supplemental report will be filed when case and investigation is complete.
35789 | 1220648-2023-01959 | 16548241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16548241&pc=OZD | 02/14/2023 | 07/12/2023 |  | 03/15/2023 |  |  | 02/14/2023 | 03/15/2023 |  |  | 08/31/2024 |  | 03/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023263806 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/14/2023 |  |  |  |  | The complainant reported a 55-year-old male patient presenting for elective placement. during impella support, it was noted that the patient developed hemolysis. an echocardiogram was performed and it was discovered that the impella had gotten caught around the papillary muscle. repositioning of the impella was completed successfully, however the patient's ldh level continued to rise. the decision was subsequently made to replace the device with a new impella cp pump. | [5 empty] | The investigation into the hemolysis has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on the clinical details, the root cause of the hemolysis was ingested biomaterial. the failure mode will be monitored and trended. | The pump was returned and an evaluation of the device is pending. upon completion of the investigation, a supplemental report will be filed.
35790 | 1220648-2023-01960 | 16550599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16550599&pc=OZD | 02/13/2023 | 06/15/2023 |  | 03/15/2023 |  |  | 02/13/2023 | 03/15/2023 |  |  | 10/31/2024 |  | 03/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2024290683 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2023 |  |  |  |  | The complainant reported a 48 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient exhibited limb ischemia and a vascular dissection. the pump was removed, a drug eluding stent was placed, and vascular repair was completed. perfusion was successfully restored, thereafter. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the most likely cause of limb ischemia was patient condition. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35791 | 1220648-2023-01954 | 16539850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16539850&pc=OZD | 02/27/2023 | 09/16/2024 |  | 03/14/2023 |  |  | 02/28/2023 | 03/14/2023 |  |  | 12/31/2024 |  | 03/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC |  |  | US |  | 2024298576 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Impairment | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2023 |  |  |  |  | It was reported the patient experienced a gastrointestinal bleed while on impella support. as a result of the bleed, heparin was withheld. prior to impella insertion, the patient had renal disfunction, however the patient's renal function may have decreased due to the hemolysis the patient experienced. the patient's prognosis was poor, and the decision was made to withdraw care of the patient. the pump was explanted, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a 67-year-old male patient presenting for impella implant. during support, it was noted that the patient developed hemolysis. as a result of the condition, the patient was administered an unknown quantity of blood products. patient support continued following the intervention. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-01954 was provided in sections b2, b5, d6a, d6b, h1, and h6. | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35792 | 1220648-2023-01955 | 16539911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16539911&pc=OZD | 02/18/2023 | 03/14/2023 |  | 03/14/2023 |  |  | 02/18/2023 | 03/14/2023 |  |  | 10/31/2023 |  | 03/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022156242 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 42-year-old male presented post cardiac arrest. patient required cpr several times and received support from a va ecmo and later from an impella cp. lhc procedure performed. during repositioning the impella cp, it has been pulled back into the aorta. the impella pump was successfully reincerted into the left ventricle. after removal of the impella, severe aortic insufficiency has been noted. according to the attending physician it is unclear whether the damage of the aortic valve leaflet has been caused by heart failure or by impella repositioning. the patient survived the event and has been sent to icu for recovery. | [5 empty] | Manufacturer has not received pump for investigation. root cause cannot be determined. should any new information be returned, a supplemental report will be filed.
35793 | 1220648-2023-01956 | 16539926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16539926&pc=OZD | 02/17/2023 | 06/12/2023 |  | 03/14/2023 |  |  | 02/17/2023 | 03/14/2023 |  |  | 01/31/2025 |  | 03/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024294357 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 72-year-old female patient has been admitted to hospital in cardiogenic shock. an impella 5.5 cardiac pump has been inserted to stabilize the patient. cad diagnosed and treated with cabg. during the cabg procedure the surgeon noted that the impella device perforated the left ventricle. the ventricle has been repaired with pledgeted sutures. the 4 vessel cabg procedure has been successfully completed and the patient has been stable afterwards. | [5 empty] | No product was returned for investigation. the cause of the ventricle perforation was most likely use related and occurred while the pump was in surgical mode during a cabg procedure. | The impella device has not been returned to the manufacturer. when the investigation is complete, a supplemental report will be filed.
35794 | 1220648-2023-01957 | 16541306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16541306&pc=OZD | 02/15/2023 | 05/26/2023 |  | 03/14/2023 |  |  | 02/15/2023 | 03/14/2023 |  |  | 09/30/2024 |  | 03/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023281098 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 91-year-old female in the us presented for elective hrpci. after successful completion of the pci procedure, impella cp has been weaned and removed. impella sheath removed, and closure of the access site attempted. several attempts failed. patient had an act of 237 s and had hematoma rapidly forming. dryseal used to tamponade vessel perforation via contralateral access until vascular surgery. patient lost 2 liters of blood. three units of rbc given. cutdown performed and left cfa repaired. contributing factor for vessel perforation: calcified vessels. patient survived. contribution of impella cannot be completely ruled out. | [5 empty] | The investigation into the reported vascular damage has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the vascular damage. the failure mode will be monitored and trended. | The impella device has not been received by the manufacturer for investigation. root cause cannot be determined. should any new information be returned, a supplemental report will be filed.
35795 | 1220648-2023-01950 | 16529084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16529084&pc=OZD | 02/11/2023 | 03/13/2023 |  | 03/13/2023 |  |  | 02/11/2023 | 03/13/2023 |  |  | 11/30/2024 |  | 03/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023281951 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/11/2023 |  |  |  |  | The complainant reported a 66 year old male patient presenting with acute myocardial infarction and cardiogenic shock. during support, the patient developed signs of hemolysis. the decision was made to remove the pump prematurely as a result. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35796 | 1220648-2023-01951 | 16534946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16534946&pc=OZD | 02/14/2023 | 03/13/2023 |  | 03/13/2023 |  |  | 02/14/2023 | 03/13/2023 |  |  | 10/31/2024 |  | 03/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2024290762 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/14/2023 |  |  |  |  | The medical center placed an impella cp for support of elderly patient with left main coronary artery disease as it was 99% occluded. the pump was placed after surgery declined to intervene and operate upon the patient. the patient deteriorated on support and a left ventricle (lv) angiogram found the patient to have suffered from a lv perforation with causality to the use of impella. the pump was explanted. | [5 empty] | The medical center did not return the impella product for investigation and benchtop analysis. further clinical details and imaging have been requested, but not yet furnished. should more be shared and root cause of the lv perforation be determined, a final report will be following. the failure mode will be monitored and trended.
35797 | 1220648-2023-01952 | 16534947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16534947&pc=OZD | 02/15/2023 | 03/13/2023 |  | 03/13/2023 |  |  | 02/15/2023 | 03/13/2023 |  |  | 05/31/2023 |  | 03/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022083757 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/15/2023 |  |  |  |  | The medical team had an impella cp inserted via the right femoral artery at a community medical center prior to transfer to tertiary center for continued monitoring and care for the cardiogenic shock. the pump only remained on for support for 1 day due to limb ischemia concerns in the impella accessed limb. the pump was explanted due to lost pulses and the vascular repair was done to the femoral access. after surgical repair the pulses returned. | [5 empty] | The medical center discarded the impella pump product at the time of explant. clinical details and imaging have been requested but not yet furnished. should the medical center relay more information that leads to root cause determination, a final report will follow. the failure mode will be monitored and trended.
35798 | 1220648-2023-01953 | 16534950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16534950&pc=OZD | 02/20/2023 | 03/13/2023 |  | 03/13/2023 |  |  | 02/20/2023 | 03/13/2023 |  |  | 07/31/2024 |  | 03/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023255545 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/20/2023 |  |  |  |  | The medical center had an impella cp support ongoing in which there was pulses lost on the right lower extremity, which was accessed by and for the impella. the team made choice to explant the pump after just over 1 hour of support. no further measures were taken for the limb ischemia. | [5 empty] | The medical center discarded the impella pump and introducer sheath product at the time of explant. no benchtop analysis to root cause is possible. further clinical details and imaging have been requested, but not yet furnished. should root cause be determined via new information being shared, a final report will follow. the failure mode will be monitored and trended.
35799 | 1220648-2023-01948 | 16521807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16521807&pc=OZD | 01/13/2023 | 06/12/2023 |  | 03/10/2023 |  |  | 02/08/2023 | 03/10/2023 |  |  | 08/31/2024 |  | 03/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023257458 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2023 |  |  |  |  | The complainant reported a 67 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed leg ischemia on the access side of the impella pump. as a result, the pump was removed. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the patient condition was the cause of the limb ischemia. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35800 | 1220648-2023-01949 | 16522136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16522136&pc=OZD | 01/25/2023 | 08/17/2023 |  | 03/10/2023 |  |  | 02/08/2023 | 03/10/2023 |  |  | 09/30/2024 |  | 03/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023275247 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2023 |  |  |  |  | The complainant reported an 81 year old female patient presenting for high risk percutaneous coronary intervention using the impella for hemodynamic support. an access site hematoma developed and a drop in hemoglobin was noted. two units of blood products was required. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information provided, the cause of the access site bleeding was not determined since the product was not returned. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35801 | 1220648-2023-01941 | 16515784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16515784&pc=OZD | 02/07/2023 | 03/09/2023 |  | 03/09/2023 |  |  | 02/07/2023 | 03/09/2023 |  |  | 01/31/2024 |  | 03/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP SET JP |  |  | ABIOMED, INC. |  |  | US |  | 2023179465 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 38 YR |  | Male |  |  | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | Low Cardiac Output | [12 empty] | Death | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 02/07/2023 |  |  |  |  | The complainant reported a 38 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. a dnr order (do not resuscitate) was executed after the patient's condition had not improved. it was decided to run the pump until it stopped. after 144 days of use, the pump stopped and the patient expired. | [5 empty] | The impella device was discarded by the customer. the investigation is in process and a supplemental will be filed upon completion.
35802 | 1220648-2023-01942 | 16518623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16518623&pc=OZD | 02/13/2023 | 03/09/2023 |  | 03/09/2023 |  |  | 02/13/2023 | 03/09/2023 |  |  | 11/30/2023 |  | 03/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED INC. |  |  | US |  | 2023196831 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/13/2023 |  |  |  |  | The medical team was trying to deliver an impella 2.5 for support for a patient undergoing high risk coronary angioplasty. when the 2.5 was delivered through the 13fr sheath the pump would not get past the native vessel bifurcation and so the team treated the vessels with shockwave. afterwards the team observed vessel damage. the team evaluated the damage with ultrasound (ivus) and stented and post dilated with balloons the iliac arteries. the vessel damage was remedied and the patient survived. | [5 empty] | The medical center did not return any impella pump or sheath product for investigation. the causal link between the vessel damage and the use of impella could not be determined. the causal link between the vessel damage and the use of shockwave balloons also could not be determined. should the medical center furnish further details that lead to a definitive root cause, a final report will be filed. the failure mode will be monitored and trended.
35803 | 1220648-2023-01944 | 16518724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16518724&pc=OZD | 12/14/2020 | 03/09/2023 |  | 03/09/2023 |  |  | 12/14/2020 | 03/09/2023 |  |  | 08/31/2022 |  | 03/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1497068 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/14/2020 |  |  |  |  | The medical center had an impella 5.5 running for hemodynamic support for 15 days when the purge sidearm failed, was bypassed, and eventually the pump stopped. the pump was explanted and not noted to be replaced. the patient survived the event. | [5 empty] | The medical center returned the impella 5.5 pump for analysis and benchtop testing. the investigation revealed the cause of the pump stop to be blood ingress as a result of the broken purge sidearm. the failure mode will be monitored and trended.
35804 | 1220648-2023-01945 | 16518727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16518727&pc=OZD | 10/17/2021 | 03/09/2023 |  | 03/09/2023 |  |  | 10/17/2021 | 03/09/2023 |  |  | 05/31/2023 |  | 03/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC |  |  | US |  | 2022088775 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/17/2021 |  |  |  |  | The medical reported a damaged purge sidearm/leak with an eventual pump stop after over 20 days of use. the pump was removed with no additional pump support provided. there was no adverse impact to the patient as a result of either the pump stop or the side arm break. | [5 empty] | The medical center did not return for benchtop analysis any impella product. only the clinical report could be analyzed for cause. the presumption is that with the broken purge sidearm, there was blood ingress which caused the pump stop. the failure mode will be monitored and trended.
35805 | 1220648-2023-01946 | 16518753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16518753&pc=OZD | 04/22/2021 | 03/09/2023 |  | 03/09/2023 |  |  | 04/22/2021 | 03/09/2023 |  |  | 02/28/2023 |  | 03/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022047547 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/22/2021 |  |  |  |  | The medical center had an impella 5.5 pump stop on day 17 of use. the pump stopped after prior purge issues and purge bypass troubleshooting measures taken . the pump was explanted and replaced as the patient bridges to a permanent ventricular assist device. | [5 empty] | The medical center did return the impella pump for benchtop analysis and investigation. the investigation revealed the root cause of the pump stop to be blood ingress due to the damage to the purge sidearm. the failure mode will be monitored and trended.
35806 | 1220648-2023-01947 | 16518761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16518761&pc=OZD | 01/09/2021 | 03/09/2023 |  | 03/09/2023 |  |  | 01/09/2021 | 03/09/2023 |  |  | 09/30/2022 |  | 03/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2021003934 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/09/2021 |  |  |  |  | The medical center had an impella 5.5 supporting a patient for 17 days, 11 of which the anticoagulation was not optimal and per ifu recommendations. the patient had bleeding concerns and anticoagulants were withheld due to clots within the patient's lungs and surgical retrieval of them. after 17 days the pump stopped and was explanted. | [5 empty] | The medical center did return the impella 5.5 pump for benchtop analysis and investigation. the investigation revealed the cause of the pump stop to be blood ingress due to the broken purge sidearm. the failure mode will be monitored and trended.
35807 | 1220648-2023-01943 | 16519116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16519116&pc=OZD | 02/10/2023 | 03/09/2023 |  | 03/09/2023 |  |  | 02/10/2023 | 03/09/2023 |  |  | 11/30/2024 |  | 03/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024289168 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 77 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 77-year-old female patient was treated with percutaneous coronary intervention due to acute myocardial infarction/cardiogenic shock as part of a clinical trial. an impella cp cardiac pump was inserted for hemodynamic support. after removal of the impella cp, a hematoma formed at the access site. two units of red blood cell concentrate were administered to treat the blood loss. the patient recovered without residual effects. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental report will be filed.
35808 | 1220648-2023-01939 | 16501234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16501234&pc=OZD | 02/07/2023 | 06/06/2023 |  | 03/08/2023 |  |  | 02/07/2023 | 03/08/2023 |  |  | 11/30/2024 |  | 03/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023278051 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 73-year-old male patient with cad, shortness of breath and mitral regurgitation has been scheduled for a cabg and mitral valve replacement procedure. to provide hemodynamic support, an impella 5.5 heart pump has been inserted. there have been difficulties getting the impella 5.5 across the aortic valve, but insertion has been successful. cabg procedure has been successful, mitral valve could not be replaced due to patients¿ anatomy. after removal of the impella 5.5 tte echo showed moderate to severe aortic insufficiency. attending physician believes it may have been caused by the impella 5.5 insertion. the patient had a tavr implanted and improved to where he did not need a mitraclip anymore. patient stable. | [5 empty] | The cause of the patient injury, specifically aortic insufficiency, was not determined since product was not returned and there is insufficient clinical information. | Investigation is ongoing based on available clinical information. when the investigation is complete, a supplemental report will be filed.
35809 | 16501668 | 16501668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16501668&pc=OZD | 01/16/2023 | 02/13/2023 | 02/13/2023 | 03/08/2023 | 02/13/2023 | 03/08/2023 |  | 03/08/2023 |  |  |  |  | 03/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | No answer provided |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 25185 DA |  | Female |  |  |  | Malfunction | No | Yes | Air/Gas in Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella alarmed "air in purge line". purge fluid bag with >250 mls in bag. attempted to re-prime purge line x2 though no fluid would come through line. purge cassette changed & primed with no further alarms.
35810 | 1220648-2023-01940 | 16505451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16505451&pc=OZD | 02/19/2023 | 09/24/2024 |  | 03/08/2023 |  |  | 02/21/2023 | 03/08/2023 |  |  | 10/31/2023 |  | 03/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2022144127 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Vascular Dissection | Thrombosis/Thrombus | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/21/2023 |  |  |  |  | The complainant reported an event involving a male patient presenting for impella support. during support, the patient was noted to have experienced an embolic stroke. a thrombectomy was performed in an attempt to remove the clot, however, additional complications associated with the patient¿s condition prevented further action from being taken. there is currently no causal relationship alleged between the impella and the other noted conditions encountered during the thrombectomy. | It was reported that the patient was believed to have been experiencing an embolic stroke, and was placed on impella for hemodynamic support. after the patient had the impella cp inserted, and support was started, the physician discovered there was no thrombus causing a lack of blood flow to the brain. angiography was performed that revealed an aortic dissection that extended to the carotid artery. the physician concluded the patient experienced a dissection that caused the lack of the blood flow to the brain. it is unknown if the impella contributed to the aortic dissection while being inserted. the patient continued to decline, further treatment could not be performed, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the vascular damage and stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause vascular damage was most likely patient condition related due to carotid artery dissection and aortic dissection and may have occurred prior to support, and the cause of the stroke was patient condition related to the dissection. the failure modes will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-01940 was provided in sections b2, b5, d6a, d6b, h1, and h6.
35811 | 16494925 | 16494925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16494925&pc=OZD | 01/25/2023 | 01/25/2023 | 01/25/2023 | 03/07/2023 | 01/25/2023 | 03/07/2023 |  | 03/07/2023 |  |  |  |  | 03/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No answer provided |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | No | Risk Manager | [8 empty] | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | In evening, ecmo coordinator received call from dayshift bedside cvicu rn. bedside rn stated that she discovered sticky, leaking fluid from the front of the impella aic. ecmo coordinator went immediately to bedside to investigate. five minutes later, it appeared the leaking fluid was coming from the purge cassette compartment. a new bag of purge solution was ordered from pharmacy, and a new cassette kit was obtained. once supplies were prepared, purge cassette and purge solution were exchanged for a brand new set. once leaking purge cassette was removed, it was inspected for source of leak. a specific source was not detected. however, the entirety of the cassette (box-shaped portion) was wet with purge solution. some purge solution had even dripped onto the floor from the cassette compartment door. cassette was thoroughly inspected for lot#, but could not be located. a call was placed to the abiomed hotline to ascertain if there were any means possible to track down the lot#. ecmo coordinator was informed that the lot# would only be available from the original outer packaging, which was no longer available. prior cassette and tubing were collected and kept if needed. i did document the new cassette kit lot#.
35812 | 16495196 | 16495196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16495196&pc=OZD | 02/10/2023 | 02/13/2023 | 02/14/2023 | 03/07/2023 | 02/14/2023 | 03/07/2023 |  | 03/07/2023 |  |  |  |  | 03/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA PURGE CASSETTE |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1702028 |  |  |  | Yes |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 21170 DA |  | Male | 106 KG |  | White | Malfunction | No | Yes | Leak/Splash | [6 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Bedside rn observed sticky, clear fluid leaking out of impella purge cassette compartment and dripping onto the floor. ecmo coordinator and perfusionist arrived and confirmed that the purge cassette was indeed leaking purge solution. a new purge cassette was obtained and ecmo coordinator and perfusionist assisted bedside rn with changing to the new purge cassette. the leaking purge cassette was removed and inspected however there was no singular, definitive source of leaking. the fluid seemed to be leaking from the box-shaped cassette portion in multiple places. of note, this purge cassette was only inserted the day prior.
35813 | 16495245 | 16495245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16495245&pc=OZD | 01/12/2023 | 02/13/2023 | 02/13/2023 | 03/07/2023 | 02/13/2023 | 03/07/2023 |  | 03/07/2023 |  |  |  |  | 03/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0042-0010-US |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No answer provided |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 18980 DA |  | Male |  |  |  | Malfunction | No | Yes | Communication or Transmission Problem | Activation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella cassette changed per protocol. shortly after change out, impella began alarming disc not detected. this charge rn came to assist bedside rn and assisted in taking out and replacing current cassette with no improvement. this rn suggested trying a new cassette. together we changed out the cassette without improvement of alarm. bedside rn on phone with impella rep throughout process. after second cassette change, this rn suggested changing out the impella console and the rep agreed. this fixed the problem. no harm to patient.
35814 | 1220648-2023-01937 | 16497004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16497004&pc=OZD | 02/07/2023 | 06/26/2023 |  | 03/07/2023 |  |  | 02/07/2023 | 03/07/2023 |  |  | 03/31/2023 |  | 03/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022056615 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2023 |  |  |  |  | The medical team had an impella cp-ecmo support ongoing for a patient admitted in cardiogenic shock, when on day 3 the access site ooze prompted intervention. the team infused 2 units of blood products and made plans to remove and escalate to an axillary placed impella 5.5 pump. | [5 empty] | The medical center has not yet returned any impella pump or sheath product for investigation and benchtop analysis. should product be returned, and root cause determined, a final report will follow. the failure mode will be monitored and trended. | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis of the access site bleeding; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of access site bleeding was most likely related to patient condition of esophagus bleeding. the failure mode will be monitored and trended.
35815 | 1220648-2023-01938 | 16498560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16498560&pc=OZD | 02/06/2023 | 05/25/2023 |  | 03/07/2023 |  |  | 02/06/2023 | 03/07/2023 |  |  | 02/28/2023 |  | 03/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022050238 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2023 |  |  |  |  | The complainant reported a patient of unknown age and gender presenting with post cardiotomy cardiogenic shock for impella support. the patient developed an access site bleed that prompted a blood transfusion. | [5 empty] | The investigation into the reported access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that a lack of vessel recoil was the most likely cause of the access site bleeding. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35816 | 1220648-2023-01925 | 16485083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16485083&pc=OZD | 02/27/2023 | 05/05/2023 |  | 03/06/2023 |  |  | 02/27/2023 | 03/06/2023 |  |  | 09/30/2024 |  | 03/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023276281 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2023 |  |  |  |  | The medical center had an impella cp placement delayed due to the md loosing arterial access at the right groin the team needed to place imeplla for support of the patient suffering from an acute myocardial infarct, hypotension, and ventricular tachycardia. the left groin was then accessed and the pump placed there successfully. the patient did develop a hematoma at both access sites and 2 units of red blood cells (rbcs) were infused. the rbcs were given after the patient had cpr due to deteriorating condition. | [5 empty] | The medical center did not return for analysis any impella product, neither the pump nor sheath. a benchtop analysis to root cause is not possible. the user loss of access at the femoral site could have contributed to the blood loss and hematoma. should more information be shared or product returned, and a definitive root cause determined, a final report will follow. | Since the original report was filed, the investigation has concluded. the access site bleeding was found to be due to patient condition. this is due to the fact that both groin sites contributed to the blood loss, not only the impella limb's access site. the failure mode will be monitored and trended.
35817 | 1220648-2023-01926 | 16486132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16486132&pc=OZD | 06/21/2022 | 06/12/2023 |  | 03/06/2023 |  |  | 06/21/2022 | 03/06/2023 |  |  | 03/31/2024 |  | 03/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023203662 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/21/2022 |  |  |  |  | The complainant reported an 88 year old female patient presenting for high risk percutaneous coronary intervention using the impella for hemodynamic support. during support, the patient developed leg ischemia. as a result, the pump was removed and replaced with an intra-aortic balloon pump. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the patient condition was the cause of the limb ischemia. the failure mode will be monitored and trended.
35818 | 1220648-2023-01927 | 16486163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16486163&pc=OZD | 09/09/2021 | 03/06/2023 |  | 03/06/2023 |  |  | 09/09/2021 | 03/06/2023 |  |  | 07/31/2023 |  | 03/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022112631 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/09/2021 |  |  |  |  | The complainant reported a 72 year old male patient presenting for an ep/ablation procedure using the impella for hemodynamic support. an access site bleed developed when accidentally bumping cardiac catheterization lab's c-arm into the impella's axillary introducer kit. the patient required 2 units of blood and 2 units of fresh frozen plasma in response. surgical intervention was also required to achieve hemostasis. the device remained inserted and the procedure continued. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35819 | 1220648-2023-01928 | 16486650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16486650&pc=OZD | 02/06/2023 | 03/06/2023 |  | 03/06/2023 |  |  | 02/06/2023 | 03/06/2023 |  |  | 08/31/2024 |  | 03/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023260385 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 66-year-old male presented with acute stemi and cad. the patient then decompensated and required cpr. impella cp placement successfully stabilized the patient. due to need of increased hemodynamic support, the patient has been escalated to an impella 5.5 heart pump. patient was slightly bleeding from femoral impella cp access site. inability to hold pressure during cab procedure, therefore a cutdown has been performed to stop the bleeding. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental report will be filed.
35820 | 1220648-2023-01929 | 16489243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16489243&pc=OZD | 01/09/2023 | 06/01/2023 |  | 03/06/2023 |  |  | 01/09/2023 | 03/06/2023 |  |  | 07/31/2024 |  | 03/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023249995 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/09/2023 |  |  |  |  | The complainant reported a 35 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. a dissection was noted in the common femoral artery post removal of the 14fr peel away sheath. as a result, intervention was required (angioplasty). | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the reported vascular damage has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition, specifically the patient's vomiting and pancreatitis creating excessive pressure on the groin was the most likely cause of the vascular damage. the failure modes will be monitored and trended. instructions for use for the related event are as follows: ¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35821 | 1220648-2023-01931 | 16489365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16489365&pc=OZD | 01/06/2023 | 08/15/2023 |  | 03/06/2023 |  |  | 01/06/2023 | 03/06/2023 |  |  | 07/31/2024 |  | 03/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  |  |  | 2023253287 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2023 |  |  |  |  | The complainant reported a 64 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during use, the patient developed an access site bleed. a blood transfusion of 2 units was required as a result. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the access site bleeding was not determined since product was not returned and there is insufficient clinical information.
35822 | 1220648-2023-01932 | 16489434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16489434&pc=OZD | 01/05/2023 | 06/06/2023 |  | 03/06/2023 |  |  | 01/05/2023 | 03/06/2023 |  |  | 09/30/2024 |  | 03/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023276529 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2023 |  |  |  |  | The complainant reported a 68 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during use, the patient developed hemolysis. dialysis was required. | [5 empty] | The investigation into the hemolysis has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the hemolysis. the failure mode will be monitored and trended. hemolysis ifu quotes: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35823 | 1220648-2023-01930 | 16489550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16489550&pc=OZD | 01/06/2023 | 08/17/2023 |  | 03/06/2023 |  |  | 01/06/2023 | 03/06/2023 |  |  | 08/31/2024 |  | 03/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023272377 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2023 |  |  |  |  | The complainant reported a 46 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during use, the patient developed ischemia in the impella inserted leg. the pump was removed as a result. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that there are insufficient clinical details to determine the root cause. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35824 | 1220648-2023-01933 | 16489950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16489950&pc=OZD | 01/04/2023 | 06/06/2023 |  | 03/06/2023 |  |  | 01/04/2023 | 03/06/2023 |  |  | 07/31/2024 |  | 03/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023257882 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2023 |  |  |  |  | The complainant reported a 66 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed leg ischemia in the impella inserted limb. as treatment, the pump was removed and replaced with an intra-aortic balloon pump. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that stenotic/small patient vessel patient condition was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. limb ischemia ifu quotes: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
35825 | 1220648-2023-01934 | 16491492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16491492&pc=OZD | 12/12/2022 | 05/16/2025 |  | 03/06/2023 |  |  | 12/11/2022 | 03/06/2023 |  |  | 09/30/2024 |  | 03/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273580 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/11/2022 |  |  |  |  | The complainant reported a 63 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed ischemia in the impella inserted leg. as treatment, an antegrade sheath was configured. | New information received 2023-09-14 via loqi registry indicating the patient endured an aortic dissection, however, it was unclear if the impella was a factor. | The patient expired while on support. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation of limb ischemia and vascular damage - great vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-01934. b5 patient outcome with mso statement added. d4 model number. d4 unique identifier (udi). d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email revised in accordance with updated procedures. g2 report source; study was missing. h6 component code revised from the initial submission in accordance with updated procedures. h10 instructions for use revised from the initial submission in accordance with updated procedures.
35826 | 1220648-2023-01935 | 16492729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16492729&pc=OZD | 02/06/2023 | 03/06/2023 |  | 03/06/2023 |  |  | 02/06/2023 | 03/06/2023 |  |  | 09/30/2024 |  | 03/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023264284 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/06/2023 |  |  |  |  | The (b)(6) team in japan had an impella cp support in which there were signs of hemolysis. this prompted the team to treat the patient with dialysis. additionally the team in japan states the patient suffered renal failure that prompte dialysis (chdf) which was additionally noted as a treatment of the hemolysis. | [5 empty] | The (b)(6) center in japan discarded the impella pump product at the time of explant. no benchtop analysis to root cause is possible. no hemolysis testing was performed on the pump. the manufacturer will file a final report if the team in japan submits more information that leads to root cause. the failure mode will be monitored and trended.
35827 | 1220648-2023-01936 | 16492735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16492735&pc=OZD | 02/04/2023 | 09/06/2023 |  | 03/06/2023 |  |  | 02/04/2023 | 03/06/2023 |  |  | 03/31/2024 |  | 03/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023205195 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/04/2023 |  |  |  |  | The complainant reported a 75 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during use, the patient developed an access site bleed. as treatment, a blood transfusion was required. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information provided, the root cause of the access site bleeding was not determined due to insufficient clinical details. the failure mode will be monitored and trended.
35828 | 1220648-2023-01923 | 16477799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16477799&pc=OZD | 02/01/2023 | 06/07/2023 |  | 03/03/2023 |  |  | 02/01/2023 | 03/03/2023 |  |  | 10/31/2024 |  | 03/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2024286243 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2023 |  |  |  |  | The complainant reported a 46 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. the patient developed limb ischemia in the impella inserted limb post insertion. a fem-fem bypass was configured as treatment. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. limb ischemia ifu quote: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35829 | 1220648-2023-01924 | 16478593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16478593&pc=OZD | 02/01/2023 | 03/03/2023 |  | 03/03/2023 |  |  | 02/01/2023 | 03/03/2023 |  |  | 03/31/2024 |  | 03/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023207376 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/01/2023 |  |  |  |  | The medical center had an impella cp support ongoing in which there was an access site ooze. the ooze was due to the patient movement. the team took the patient to the operating room and believed the patient to be suffering from hypovolemic/hemorrhagic shock. the team gave 3 units of blood products to remedy the situation. the plan in the operating room was to place an axillary accessed impella 5.5 for new access and care escalation. | [5 empty] | The medical center discarded the impella pump and sheath product. no benchtop analysis to root cause is possible. the contribution of the patient movement to the ooze and access site bleeding can not be denied. should more information be shared that leads to root cause, a final report will follow. the failure mode will be monitored and trended.
35830 | 1220648-2023-01917 | 16468132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16468132&pc=OZD | 02/15/2023 | 07/20/2023 |  | 03/02/2023 |  |  | 02/16/2023 | 03/02/2023 |  |  | 10/31/2024 |  | 03/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023269480 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2023 |  |  |  |  | The complainant reported a 78-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. following removal of an impella pump and the subsequent explant of a utilized unapproved 16 fr sheath, the patient developed lower limb ischemia. further observation noted a thrombus and a dissection within the external iliac artery. thrombolytics were provided and a stent was placed to treat the dissection. a definitive cause for the events could not be determined. following the intervention, flow returned to the limb. the patient was considered stable following the events. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the limb ischemia and the vascular damage has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to root cause; only the clinical report was evaluated. based on clinical information provided, the cause of the limb ischemia was most likely biomaterial since ischemia occurred after the removal; the cause of the vascular damage was not determined since product was not returned. the failure modes will be monitored and trended.
35831 | 1220648-2023-01918 | 16468174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16468174&pc=OZD | 02/01/2023 | 03/02/2023 |  | 03/02/2023 |  |  | 02/01/2023 | 03/02/2023 |  |  | 08/31/2024 |  | 03/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023267619 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/01/2023 |  |  |  |  | The medical team had an impella cp inserted for support of patient undergoing coronary angioplasty, when there was an access site bleed and hematoma. the bleed was treated by pump removal and placement of a new sheath for hemostasis and the team consulted surgery for repair of the femoral artery. | [5 empty] | The medical center discarded the impella sheath and pump product. no benchtop analysis to root cause is possible. should more information be shared and a definative root cause determined, a final report will be filed. the failure mode will be monitored and trended.
35832 | 1220648-2023-01919 | 16469347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16469347&pc=OZD | 09/28/2022 | 06/21/2023 |  | 03/02/2023 |  |  | 02/02/2023 | 03/02/2023 |  |  | 06/30/2024 |  | 03/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023237404 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/02/2023 |  |  |  |  | The complainant reported a patient of unknown gender and age presenting for high risk percutaneous coronary intervention. following impella support, it was noted that the patient developed a hematoma after the pump was explanted. the patient was subsequently provided two units of packed red blood cells to treat the hematoma. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the access site bleeding has been completed since the original report was filed. the medical center discarded the impella product at the time of explant. no benchtop analysis of the root cause of the access site bleeding was possible; only the clinical report was evaluated. the clinical evaluation revealed that the patient's condition was most likely cause of the access site bleeding. the failure mode will be monitored and trended.
35833 | 1220648-2023-01920 | 16470080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16470080&pc=OZD | 02/02/2023 | 03/02/2023 |  | 03/02/2023 |  |  | 02/02/2023 | 03/02/2023 |  |  | 12/31/2024 |  | 03/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC |  |  | US |  | 2024291872 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/02/2023 |  |  |  |  | The complainant reported a 79-year-old male patient presenting for post cardiotomy cardiogenic shock. during impella insertion, it was noted that the left ventricle wall was perforated. it was documented that the ventricle wall was already thin as a result of dilation and that little effort was needed to perforate the wall. three plegeted stitches were used to fix the rupture and patient support continued with the implanted 5.5. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35834 | 1220648-2023-01921 | 16470372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16470372&pc=OZD | 10/20/2020 | 08/17/2023 |  | 03/02/2023 |  |  | 01/31/2023 | 03/02/2023 |  |  | 03/31/2022 |  | 03/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1481387 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2023 |  |  |  |  | The spouse of an impella patient called abiomed collections claiming that his wife died in 2020 as a result of a malfunctioning impella. based on the abiomed on-site representative, the patient's cause of death was related to a failed perclose with uncontrolled hemostasis. protamine was delivered, which lead to clotting, arrhythmia, and eventual death. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. the device data logs were analyzed and no abnormalities were found. should any new information be returned, a supplemental mdr will be filed. | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information provided, the root cause of the access site bleeding was not determined since product was not returned. the relation to impella is unknown.
35835 | 1220648-2023-01922 | 16470414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16470414&pc=OZD | 01/27/2023 | 08/17/2023 |  | 03/02/2023 |  |  | 01/31/2023 | 03/02/2023 |  |  | 11/30/2024 |  | 03/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023276291 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2023 |  |  |  |  | The complainant reported a 52 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. the patient sustaining an embolic stroke upon explant of the impella device. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Since the medical center did not return the impella device, no benchtop analysis was possible. however, data logs were returned and no trends in motor current, purge pressure, or flow were noted that would indicate any biomaterial/clot interaction with the pump. therefore, the root cause of the cva is most likely related to patient condition and the relationship to impella is unknown.
35836 | 16459290 | 16459290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16459290&pc=OZD | 12/25/2022 | 02/13/2023 | 02/13/2023 | 03/01/2023 | 02/13/2023 | 03/01/2023 |  | 03/01/2023 |  |  |  |  | 03/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0042-0010-US |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 26280 DA |  | Male |  |  |  | Malfunction | No | Yes | Leak/Splash | Infusion or Flow Problem | [5 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella alarming for increased purge flow. recognized fluid leaking from door of impella pump at the impella cartridge. caused additional cassette change.
35837 | 16459489 | 16459489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16459489&pc=OZD | 02/04/2023 | 02/06/2023 | 02/06/2023 | 03/01/2023 | 02/06/2023 | 03/01/2023 |  | 03/01/2023 |  |  |  |  | 03/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | 0550-0008 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | No answer provided |  |  | Not Available |  |  |  | Health Professional |  |  | User Facility report | No | Risk Manager | [6 empty] | 20440 DA |  | Female |  |  |  | Malfunction | No | Yes | Device Alarm System | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Patient started having suction alarms on impella console for anything above p-4 despite receiving 250cc albumin and 500cc plasmalyte. cvp adequate (12). position was confirmed on bedside echo. when impella was turned back to p-7 the flow remained stable and 4lpm but suction alarms continued. motor current readings remained the same. ao placement signal did not correlate with bp. discussed with impella rep and it was determined that due to the above information the placement signal sensor was no longer working. the suction alarms were disabled on the console and impella turned back up to p-7. if ao signal returns to normal the suction alarms would be able to be turned back on, but the most likely result will be the impella console will stop reading placement. impella rep advised to not change the impella console as it may not restart on a new console.
35838 | 1220648-2023-01909 | 16459876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16459876&pc=OZD | 01/27/2023 | 09/19/2024 |  | 03/01/2023 |  |  | 01/30/2023 | 03/01/2023 |  |  | 10/31/2023 |  | 03/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2023199519 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2023 |  |  |  |  | The complainant reported a 74 year old male patient presenting with cardiomyopathy. during support, a cerebral infarction occurred with unknown relationship to the impella device. the patient then experienced a cerebral hemorrhage and expired (again, with unknown relationship). | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
35839 | 1220648-2023-01910 | 16460136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16460136&pc=OZD | 12/27/2022 | 03/01/2023 |  | 03/01/2023 |  |  | 12/27/2022 | 03/01/2023 |  |  | 03/31/2024 |  | 03/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023209267 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/27/2022 |  |  |  |  | The complainant reported a 74-year-old female patient presenting for cardiomyopathy. it was noted that during impella support, the patient began to experience numbness in their leg. the pump was subsequently removed as a result of the condition. the patient was considered stable following explant. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35840 | 1220648-2023-01911 | 16460347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16460347&pc=OZD | 01/31/2023 | 09/20/2024 |  | 03/01/2023 |  |  | 01/31/2023 | 03/01/2023 |  |  | 11/30/2024 |  | 03/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023278028 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | A female patient of unknown age in the eu was treated for subacute myocardial infarction and cardiogenic shock. an impella cp heart pump was inserted for hemodynamic support. the impella cp heart pump perforated the pericardium, resulting in a hemorrhagic pericardial effusion. the pericardial effusion was punctured for relief and the perforation was surgically closed. anticoagulation was briefly suspended, and coagulation support medications were administered. the impella cp pump was able to continue to provide hemodynamic support. the patient was subsequently transferred to another hospital and remained in the icu for several days. due to severe cardiogenic shock, the patient eventually died of multiple organ failure. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. after completion of the investigation based on clinical information, a supplemental mdr will be filed. | The root cause of the ventricle perforation was unable to be determined as limited clinical details were provided. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-01911 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
35841 | 1220648-2023-01912 | 16461787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16461787&pc=OZD | 01/31/2023 | 06/12/2023 |  | 03/01/2023 |  |  | 01/31/2023 | 03/01/2023 |  |  | 04/30/2024 |  | 03/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023223677 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 93 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2023 |  |  |  |  | The complainant reported a 93-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient began to develop limb ischemia on the impella side. the cause for the ischemia was subsequently found to be an occlusion within the sheath in the right femoral artery. the pump was explanted as a result of the ischemia and the patient¿s condition improved. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of the leg ischemia was determined to be patient condition as the patient had calcified vessels. the failure mode will be monitored and trended.
35842 | 1220648-2023-01915 | 16463352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16463352&pc=OZD | 01/10/2023 | 03/01/2023 |  | 03/01/2023 |  |  | 01/10/2023 | 03/01/2023 |  |  | 11/30/2024 |  | 03/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC |  |  | US |  | 2023279017 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/10/2023 |  |  |  |  | The complainant reported an 63-year-old male patient presenting for post cardiotomy cardiogenic shock. following impella explant from direct access, bleeding was identified at the graft at the aorta site. as a result of the bleed, the patient was reopened and the site was repaired. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35843 | 1220648-2023-01914 | 16463428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16463428&pc=OZD | 01/30/2023 | 03/01/2023 |  | 03/01/2023 |  |  | 01/30/2023 | 03/01/2023 |  |  | 09/30/2024 |  | 03/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023275339 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/30/2023 |  |  |  |  | The complainant reported an 81-year-old male patient presenting for elective placement. during impella support, it was noted that the patient developed an access site bleed coinciding with a hematoma following removal of the peel-away sheath and placement of the repo-sheath. a femoral exploration identified that the bleeding was coming from around the repositioning sheath at the arteriotomy. a purse string suture was subsequently placed on the vessel to achieve hemostasis. patient support continued following the event. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35844 | 1220648-2023-01916 | 16463461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16463461&pc=OZD | 01/31/2023 | 06/15/2023 |  | 03/01/2023 |  |  | 01/31/2023 | 03/01/2023 |  |  | 08/31/2024 |  | 03/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023262760 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2023 |  |  |  |  | The medical team had placed an impella cp for support via the left femoral artery, after removing and escalating from the iabp, for a patient being considered for valvuloplasty. the patient had signs of limb ischemia in the impella limb and had distal perfusion catheter placed to address the issue. the pump remained on for support for 17 hours and was then explanted as support was deemed insufficient and care was escalated to the impella 5.5. at the time of the cp explant the team observed vascular damage and reconstructed the vascular anatomy with a patch. | [5 empty] | The medical center discarded the impella product at the time of explant and care escalation. no benchtop analysis to the root cause of the ischemia and vessel dissection is possible. no imaging has been shared for investigation. the failure mode will be monitored and trended. a final report can follow should new information be shared by the medical center that leads to definative root cause. | The investigation into the limb ischemia has been completed since the original report was filed. the medical center discarded the impella product at the time of explant and care escalation. no benchtop analysis of the root cause of the ischemia and vessel dissection was possible; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of the limb ischemia was most likely patient condition. the failure mode will be monitored and trended.
35845 | 1220648-2023-01913 | 16464022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16464022&pc=OZD | 02/01/2023 | 06/29/2023 |  | 03/01/2023 |  |  | 02/01/2023 | 03/01/2023 |  |  | 12/31/2024 |  | 03/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC |  |  | US |  | 2024291899 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2023 |  |  |  |  | The complainant reported a 55-year-old male patient presenting for impella implant. during impella support, it was documented that the physician was squeezing volume out of the left ventricle (lv) and subsequently perforated the free wall of the lv when the heart was rotated. the impella was pulled back into place and the free wall was repaired with sutures. patient support continued following the incident. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the reported vascular damage has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the most likely root cause for vascular damage was use error, specifically the manipulation of the patient's heart. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35846 | 1220648-2023-01905 | 16453063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16453063&pc=OZD | 02/13/2023 | 02/28/2023 |  | 02/28/2023 |  |  | 02/14/2023 | 02/28/2023 |  |  | 06/30/2024 |  | 02/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023233660 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/14/2023 |  |  |  |  | The complainant reported a 75-year-old male patient presenting for cardiomyopathy. it was noted that during insertion of a 14fr peel-away long sheath, the vessel was perforated and an artificial blood vessel replacement was performed. the 14fr sheath was being used to dilate the vessel prior to a planned insertion of a 16fr sheath for impella implant. following the intervention, the impella was introduced from the peripheral side of the artificial blood vessel. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35847 | 1220648-2023-01906 | 16453782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16453782&pc=OZD | 01/29/2023 | 06/06/2023 |  | 02/28/2023 |  |  | 01/29/2023 | 02/28/2023 |  |  | 05/31/2023 |  | 02/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022090907 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2023 |  |  |  |  | The complainant reported a 74 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. while the patient had pre-existing pvd, the pump was removed due to absence of pulses in the foot of the impella inserted limb. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that pre-existing pvd patient condition was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. limb ischemia ifu quote: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35848 | 1220648-2023-01908 | 16455848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16455848&pc=OZD | 01/31/2023 | 02/28/2023 |  | 02/28/2023 |  |  | 01/31/2023 | 02/28/2023 |  |  | 01/31/2024 |  | 02/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023186698 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/31/2023 |  |  |  |  | The medical center had placed an impella cp for support of a paraplegic with bilateral aka amputee and right hand amputee in septic shock. the patient was overly anticoagulated for the weight of the patient and began to bleed from the impella access site. the patient had been given heparin bolus in the ccl and systemic heparin. the team infused blood products, but then made choice to withdraw care after just 22 hours of support. the patient expired. | [5 empty] | The medical center discarded the impella product when care was withdrawn and the patient expired. no benchtop analysis to root cause is possible. the patient condition and excessive anticoagulation for a patient of this weight was stated to have been a contributing factor to the access site bleeding. the failure mode will be monitored and trended.
35849 | 1220648-2023-01907 | 16455880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16455880&pc=OZD | 02/05/2023 | 02/28/2023 |  | 02/28/2023 |  |  | 02/05/2023 | 02/28/2023 |  |  | 10/31/2024 |  | 02/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023276307 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/05/2023 |  |  |  |  | The medical center in japan had an impella cp support initiated for a patient admitted suffering from an acute myocardial infarction (ami). the team observed a left ventricle (lv) perforation attributed to perhaps the cp pump's pigtail puncturing the already injured lv heart tissue that was weakened by the ami. the team patched the perforation and the patient has survived. | [5 empty] | The medical center in japan did not retain the impella pump for any benchtop analysis. additionally, the team did not forward any imaging for analysis. no root cause can be determined, though the medical team in japan did allege the underlying patient condition of friable infarcted tissue could have contributed to the perforation. the failure mode will be monitored and trended.
35850 | 1220648-2023-01902 | 16443840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16443840&pc=OZD | 01/29/2023 | 05/25/2023 |  | 02/27/2023 |  |  | 01/29/2023 | 02/27/2023 |  |  | 10/31/2024 |  | 02/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023257452 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Death | Yes | No | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Perforation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The medical center in switzerland placed an impella cp heart pump for support of a patient suffering from cardiogenic shock. the impella cp was placed within the left ventricle, but due to the patient's obesity the visualization of its position was complicated. echo imaging was not optimal. during the support there were suction alarms that prompted the team to attempt pump repositioning. during that repositioning there was a diagnosis made of a tamponade. the team began cpr and during the attempted cpr the pump stopped and the patient expired. | [5 empty] | Data logs showed that the pump stopped immediately with a sudden motor current spike accompanied by the occurrence of the "impella stopped - motor current high" alarm; no other pump stop alarms triggered. the pump was unable to be restarted. the pump stop was preceded by a gradual decrease in motor current and placement signal over 1.5 hours. about 10 minutes after the pump stop, a brief "impella in ventricle" alarm triggered. the pump was in proper position prior to the pump stop. in conclusion, and after reviewing the clinical information, it was determined that the cause of the ventricle perforation was most likely improper positioning of the device. the cause of the pump stop was not determined since product was not returned. | No product has been returned for analysis as it was discarded at the time of explant. further clinical details have been requested but, not yet relayed to manufacturer. upon completion of the investigation, a final report will be filed regarding the tamponade and pump stop.
35851 | 1220648-2023-01903 | 16449212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16449212&pc=OZD | 01/23/2023 | 02/27/2023 |  | 02/27/2023 |  |  | 02/03/2023 | 02/27/2023 |  |  | 07/31/2023 |  | 02/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022110895 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/03/2023 |  |  |  |  | The medical center in japan had an impella cp support in which there was an access site bleed. the bleed was treated by application of the sandbag and additionally 8 units of blood products were infused. the pump remained on for support until wean and explant. | [5 empty] | The complainant in japan discarded the impella product at the time of explant. no benchtop analysis to root cause is possible. the manufacturer will close the case, though should more information be shared by the medical team in japan that leads to root cause, a final report will be filed. the failure mode will be monitored and trended.
35852 | 1220648-2023-01904 | 16449215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16449215&pc=OZD | 02/03/2023 | 09/19/2024 |  | 02/27/2023 |  |  | 02/03/2023 | 02/27/2023 |  |  | 05/31/2024 |  | 02/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED INC. |  |  | US |  | 2023227110 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/03/2023 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant in the eu had an impella cp placed for support when there was "slight resistance" felt in positioning of the pump. the team later observed that there was a drop in blood pressure after the peel away was removed. the team observed there to be a left ventricle perforation which was causing the deterioration in the patient's condition. the patient expired while in the cardiac catheterization lab despite the surgical team being called to repair the perforation. the perforation being caused by either the pump or guidewire is not yet known. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The medical center discarded the pump product at the time of explant. no benchtop analysis can be performed. no imaging was shared for the investigation. root cause can not be determined. should more information be shared, and cause determined, a final report will follow. the failure mode will be monitored and trended.
35853 | 1220648-2023-01898 | 16435417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16435417&pc=OZD | 01/25/2023 | 02/24/2023 |  | 02/24/2023 |  |  | 01/25/2023 | 02/24/2023 |  |  | 09/30/2024 |  | 02/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023281707 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The medical center had an impella cp supporting a patient when after 26 minutes the pump was explanted due to flow/suction alarms. upon inspection on the benchtop the impeller blade was found to have breakage. this was not observed by the medical center team. | [5 empty] | The returned pump was visually inspected and no abnormalities were observed. the pump was run in a basin of water and pump flow was 3.0 l/min with suction. a broken impeller blade was observed. the cause of the issue was fractured impeller blade(s). the fracture was characteristic of an interaction with another device but that was unable to be definitively determined since fluoroscopic imaging and/or additional clinical details were not provided. possible interaction with pci catheter.
35854 | 1220648-2023-01899 | 16436001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16436001&pc=OZD | 11/08/2022 | 08/17/2023 |  | 02/24/2023 |  |  | 01/26/2023 | 02/24/2023 |  |  | 08/31/2024 |  | 02/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023260858 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2023 |  |  |  |  | The complainant reported a male patient of unknown age presenting for with acute myocardial infarction and cardiogenic shock for impella support. after approximately 5 days, the impella was removed due to sepsis. the physician alleged the source of the infection was likely the impella device. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information provided, the root cause of the sepsis was not determined since there was no evidence to confirm that the infection was due to impella. abiomed device was sterilized under validated conditions. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35855 | 1220648-2023-01900 | 16437222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16437222&pc=OZD | 08/23/2022 | 02/24/2023 |  | 02/24/2023 |  |  | 01/26/2023 | 02/24/2023 |  |  | 11/30/2023 |  | 02/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC |  |  | US |  | 2023161285 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/26/2023 |  |  |  |  | The complainant reported a 35-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. following impella explant, the patient was noted to have a lower extremity clot embolization. a bilateral femoral embolectomy and angioplasty were performed to treat the condition. no further treatment was reported following the intervention. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35856 | 1220648-2023-01901 | 16438564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16438564&pc=OZD | 01/27/2023 | 02/24/2023 |  | 02/24/2023 |  |  | 01/27/2023 | 02/24/2023 |  |  | 07/31/2024 |  | 02/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023269005 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/27/2023 |  |  |  |  | The complainant reported a 76 year old female patient presenting for impella support. ischemia developed in the impella inserted limb. an antegrade sheath was attempted to resolve the issue, however, the device was ultimately removed and the issue resolved. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35857 | 1220648-2023-01895 | 16426813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16426813&pc=OZD | 01/24/2023 | 04/24/2023 |  | 02/23/2023 |  |  | 01/24/2023 | 02/23/2023 |  |  | 07/31/2024 |  | 02/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023237817 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an 84 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impela support. during support, multiple hematomas developed at the impella access site. the patient received several blood products and, as a result, the pump was removed. a covered stent was then placed. | [5 empty] | No product was returned for investigation. no data logs were returned for investigation. the cause of the access site bleeding was not determined since product was not returned and there is insufficient clinical evidence. | The device has not been received to date. a final mdr will be filed upon completion of our analysis of the event.
35858 | 1220648-2023-01896 | 16433279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16433279&pc=OZD | 01/07/2023 | 02/23/2023 |  | 02/23/2023 |  |  | 02/10/2023 | 02/23/2023 |  |  | 07/31/2024 |  | 02/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023247123 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/10/2023 |  |  |  |  | The complainant reported a 76-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. several hours following impella explant, it was discovered that the patient had experienced an iliac dissection. surgical intervention was performed and a stent was placed to treat the dissection. a definitive cause for the dissection could not be provided. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35859 | 1220648-2023-01897 | 16433289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16433289&pc=OZD | 02/10/2023 | 09/19/2024 |  | 02/23/2023 |  |  | 02/10/2023 | 02/23/2023 |  |  | 10/31/2024 |  | 02/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023277247 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/10/2023 |  |  |  |  | The complainant reported a 76-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient developed hemolysis. the patient was provided an unknown quantity of blood products and the pump was subsequently explanted as a result of the condition. patient support continued with ecmo following the events. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | Following the initial report, the pump was returned for further evaluation. a visual inspection was performed and it was discovered that there was biomaterial around the impella. there was no noted visible damage on the returned component. upon conclusion of the evaluation the cause was determined to be the presence of biomaterial from the patient around the pump. should additional relevant information become available, a supplemental report will be submitted. as noted in the impella ifu, hemolysis is a known potential adverse event associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The pump was returned and an evaluation of the device is pending. upon completion of the investigation, a supplemental report will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-01766 was provided.
35860 | 1220648-2023-01890 | 16416325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16416325&pc=OZD | 02/08/2023 | 07/05/2023 |  | 02/22/2023 |  |  | 02/08/2023 | 02/22/2023 |  |  | 08/31/2024 |  | 02/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023260335 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2023 |  |  |  |  | The complainant reported a 44-year-old male patient presenting for cardiomyopathy. during impella support, it was discovered during a ct scan that a thrombus had adhered to the impella shaft from the distal arch to the celiac artery. a thrombectomy was performed to treat the thrombus and the pump was subsequently replaced with an impella 5.5 device. heparin was then administered to treat any remaining occlusions in the right renal artery and right internal iliac artery. patient support continued following the event. | [5 empty] | The investigation into the thrombus has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis; only the clinical report was evaluated. based on the clinical details, the cause of the issue was most likely patient condition related to thrombus. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35861 | 1220648-2023-01891 | 16416345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16416345&pc=OZD | 01/23/2023 | 04/21/2023 |  | 02/22/2023 |  |  | 01/23/2023 | 02/22/2023 |  |  | 09/30/2024 |  | 02/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2023263265 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A male patient of unknown age received hemodynamic support from an impella 5.5 heart pump as bridge to lvad. suspicion of thrombus in pump due to hemolysis, red urine and elevated ldh. the attending physician decided to change the pump. a new impella pump has been successfully inserted. the patient has been stable afterwards with no more hemolysis issues. | [5 empty] | The returned pump was visually inspected and biomaterial was observed in the outflow. the pump was run in a basin of water and pump flow was 3.6 l/min with suction.the cause of the issue was most likely biomaterial. | Benchtop investigation of the returned pump is ongoing. when the investigation is complete, a supplemental report will be filed.
35862 | 1220648-2023-01892 | 16416781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16416781&pc=OZD | 01/04/2023 | 08/16/2023 |  | 02/22/2023 |  |  | 01/23/2023 | 02/22/2023 |  |  | 02/29/2024 |  | 02/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 2023199650 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 26 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaint reported a 26 year old female patient presenting with acute myocardial infarction and cardiogenic shock. during support, the impella was explanted due to right leg ischemia with a planned limb amputation. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was most likely patient condition related. | The device was not received, therefore, a physical analysis could not be performed. should we receive the device, or any new information, we will file a supplemental report.
35863 | 1220648-2023-01893 | 16418731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16418731&pc=OZD | 01/24/2023 | 06/07/2023 |  | 02/22/2023 |  |  | 01/24/2023 | 02/22/2023 |  |  | 09/30/2024 |  | 02/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023276671 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/24/2023 |  |  |  |  | The complainant reported a 70 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. limb ischemia developed, resulting from14fr insertion. an antegrade sheath was placed and support continued. after removal of the device, vessel damage was found and required a covered stent to achieve complete hemostasis. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35864 | 1220648-2023-01894 | 16419090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16419090&pc=OZD | 01/17/2023 | 06/15/2023 |  | 02/22/2023 |  |  | 01/17/2023 | 02/22/2023 |  |  | 09/30/2024 |  | 02/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023281417 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2023 |  |  |  |  | The medical center had an impella cp that was placed via the femoral artery access to support a patient admitted with cardiogenic shock and an aicd that was "firing" repeatedly for pacing of the heart. during the pump support the team observed signs of limb ischemia and so the decision was made to have surgery do a cutdown and prelature explant of the impella . the team made decision to place an axillary access for the impella 5.5 instead and escalation of the support. after explant the limb ischemia improved. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center discarded the impella product at the time of explant and care escalation. no benchtop analysis of the root cause of the ischemia was possible; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of the limb ischemia was most likely patient condition related. the failure mode will be monitored and trended. | The medical center discarded the impella pump at the time of explant. no benchtop analysis to root cause is possible. the manufacturer will close the investigation, but should new information be shared that leads to root cause, a final report will follow. the failure mode will be monitored and trended.
35865 | 16408411 | 16408411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16408411&pc=OZD | 11/16/2022 | 11/16/2022 | 11/16/2022 | 02/21/2023 | 11/16/2022 | 02/21/2023 |  | 02/21/2023 |  |  |  |  | 02/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No answer provided |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | No | Risk Manager | [6 empty] | 24090 DA |  | Male |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | Pressure Problem | [5 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Increase of impella purge pressure and decrease in impella purge flow. impella reservoir and filter bypass inspected, no issues. impella rep present at patient's bedside at that time. concerns for abrupt cessation of impella device. note: impella reservoir and filter previously bypassed following impella's guidelines due to purge fluid leakage.
35866 | 16408751 | 16408751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16408751&pc=OZD | 11/11/2022 | 11/22/2022 | 11/22/2022 | 02/21/2023 | 11/22/2022 | 02/21/2023 |  | 02/21/2023 |  |  |  |  | 02/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 370312 |  |  |  | Yes |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 24090 DA |  | Male | 83 KG |  | White | Malfunction | No | Yes | Material Erosion | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Nurse noted a leak in the impella purge connection site. on inspection of the device the purge cassette was intact and not leaking but the luer connection to the purge reservoir and filter was leaking. the impella rep was called to alert them of the issue and was told to call the impella clinical support line. they also suggested changing the pressure line the purge cassette was attached to. the pressure line was changed by the bedside nurse but the leak was still present and purge lows had increased while the pressures decreased. at this time critical care was notified of the issue. imeplla clinical support gave two potential options to the issue: since the leak was at a component connected to the device itself, it could not be replaced. the whole pump would need to be replaced or the purge reservoir/filter would need to be bypassed. the surgeon was contacted about the situation and decided that bypass of the reservoir was the best course of action. the purge reservoir/filter that malfunctioned was saved for the impella rep to pick up. the impella rep was notified as well. during discussion with the imeplla rep, it was made clear that the failure of the female luer connection of the purge reservoir/filter is known to the company. they believe it to be attributed to the new bicarbonate purge solution interacting with the composition of the luer plastic causing it to become brittle. the temporary recommendation (until the plastic composition is updated) is that an intermediate line be attached between the reservoir/filter and purge line so that when the purge line is replaced the connection point getting repeated use is the intermediate line and not the reservoir.
35867 | 1220648-2023-01886 | 16410630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16410630&pc=OZD | 01/23/2023 | 06/07/2023 |  | 02/21/2023 |  |  | 01/23/2023 | 02/21/2023 |  |  | 11/30/2024 |  | 02/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023278454 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2023 |  |  |  |  | The complainant reported a 51 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. persistent ventricular tachycardia was endured during support requiring multiple direct cardioversions required. the cause was believed to be the impella making contact with the septum. as a result, the pump was removed and replaced with an iabp. | [5 empty] | The investigation into the ventricular tachycardia (vt) has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that use issue was the most likely cause of the vt. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35868 | 1220648-2023-01887 | 16410644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16410644&pc=PYX | 01/22/2023 | 06/15/2023 |  | 02/21/2023 |  |  | 01/22/2023 | 02/21/2023 |  |  | 04/30/2023 |  | 02/21/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2022094497 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/22/2023 |  |  |  |  | The complainant reported a 62 male year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. the patient exhibited hemolysis during use of the device. the pump was repositioned, 2 units of blood were delivered, and dialysis was administered. | [5 empty] | The investigation into the hemolysis has been completed since the original report was filed. the impella device was returned for evaluation and the benchtop analysis to root cause revealed that improper positioning of the device was the cause of the hemolysis. the failure mode will be monitored and trended. | The investigation into the hemolysis has been completed since the original report was filed. the medical center discarded the impella product at the time of explant and care escalation. no benchtop analysis of the root cause of the hemolysis was possible; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of the hemolysis was improper positioning of the device. the failure mode will be monitored and trended. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
35869 | 1220648-2023-01888 | 16414226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16414226&pc=OZD | 03/03/2022 | 02/21/2023 |  | 02/21/2023 |  |  | 01/24/2023 | 02/21/2023 |  |  | 12/31/2023 |  | 02/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023165880 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/24/2023 |  |  |  |  | The complainant in japan had an impella cp support that was ongoing til wean and explant in (b)(6) 2022. the support was successful for 8 days. later in (b)(6) 2023 the jpvad registry noted hemolysis treated by haptoglobin infusion and an impella access site bleed treated by skin suture and blood products. | [5 empty] | The medical team in japan discarded the impella pump in 2022 at the time of explant. the jpvad registry team shared no further details that could lead to root cause of the presumed hemolysis or access site bleed. should more information be shared, and a root cause determined, a final report will follow. the failure mode will be monitored and trended.
35870 | 1220648-2023-01889 | 16414247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16414247&pc=OZD | 01/24/2023 | 02/21/2023 |  | 02/21/2023 |  |  | 01/24/2023 | 02/21/2023 |  |  | 09/30/2024 |  | 02/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023278560 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/24/2023 |  |  |  |  | The medical team had an impella cp supporting a 76yo male for a high risk coronary angioplasty. after just over 24 hours of support the pump was explanted. at the explant the pump became stuck and the cannula uncoiled. the cannula's inlet broke and became lodged within the leg. the patient was sent to surgery for removal of the portion of the pump. of note the patient had calcification in the native anatomy the pump was explanted through. | [5 empty] | The medical center has not returned any portion of the impella pump for analysis and benchtop investigation. should the product be returned, and root cause determined, a final report will follow. the failure mode will be monitored and trended.
35871 | 1220648-2023-01881 | 16398124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16398124&pc=OZD | 01/18/2023 | 02/17/2023 |  | 02/17/2023 |  |  | 01/18/2023 | 02/17/2023 |  |  | 10/31/2024 |  | 02/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023283812 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/18/2023 |  |  |  |  | The complainant reported an 80-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. following impella implant, an angiogram was performed and mild to moderate bilateral peripheral vascular disease was identified in the iliac arteries. a pulse check noted that no distal flow to the right leg was present. a fem-fem bypass was subsequently performed to treat the observed ischemia. the patient was stable following the intervention. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35872 | 1220648-2023-01882 | 16399848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16399848&pc=OZD | 01/21/2023 | 02/17/2023 |  | 02/17/2023 |  |  | 01/21/2023 | 02/17/2023 |  |  | 09/30/2024 |  | 02/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023280869 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Abdominal Pain | Dyspnea | Low Blood Pressure/ Hypotension | Loss of Vision | Constipation | Retroperitoneal Hemorrhage | [7 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/21/2023 |  |  |  |  | The complainant reported a 69-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was documented that the patient developed a retroperitoneal bleed and subsequently passed away. the bleed was identified during an abdominal ct scan after the patient had complaints of abdominal pain and constipation. the patient began to lose their vision after the bleed was identified, and had reports of shortness of breath and a drop in blood pressure. the patient underwent a mass transfusion, however a return of spontaneous circulation could not be achieved and the patient passed away. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35873 | 1220648-2023-01883 | 16399898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16399898&pc=OZD | 10/07/2020 | 02/17/2023 |  | 02/17/2023 |  |  | 02/06/2023 | 02/17/2023 |  |  |  |  | 02/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC |  |  | US |  | 1489032 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/06/2023 |  |  |  |  | The complainant reported a 40-year-old male patient presenting for cardiomyopathy. during impella support, it was noted that the patient developed hemolysis and subsequent renal failure. the patient underwent dialysis to manage the condition. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35874 | 1220648-2023-01884 | 16399903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16399903&pc=OZD | 01/18/2023 | 02/17/2023 |  | 02/17/2023 |  |  | 02/15/2023 | 02/17/2023 |  |  | 09/30/2024 |  | 02/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023275785 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/15/2023 |  |  |  |  | The medical center had an impella cp support for a preoperative heart valve patient that was ongoing when there was plasma free hg drawn and seen to be rising. the pump was also seen to be in less than optimal positioning within the left ventricle. in due to hemolysis concerns the team explanted the cp after just 93 hours of support time, and placed va ecmo instead. after explant of the cp, the patient's labs began to trend back down. | [5 empty] | The medical center discarded all impella pump product at the time of explant. no benchtop analysis or hemolysis testing to root cause is possible. the patient was enrolled in the loqi registry, and the registry furnished no further details that lead to root cause. should new information be shared that leads to root cause of the hemolysis a final report will follow. the failure mode will be monitored and trended.
35875 | 1220648-2023-01885 | 16400762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16400762&pc=OZD | 04/07/2022 | 02/17/2023 |  | 02/17/2023 |  |  | 02/06/2023 | 02/17/2023 |  |  | 01/31/2024 |  | 02/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC |  |  | US |  | 2023179444 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hernia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/06/2023 |  |  |  |  | The complainant reported a 52-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that a patient experienced a cerebral infarction and later passed away. upon impella insertion, a thrombus was not initially seen by the cardiology doctor. it was unknown when or why the thrombus developed. roughly ten days after insertion, the patient experienced a cerebral infarction. the patient's condition was believed to be improving following the event and the impella was explanted four days after the infarction. however, following explant, the patient experienced a cerebral hemorrhage that required a craniotomy for hematoma removal and ventricular drainage. the patient passed away as a result of the noted conditions. it was unknown if the impella caused or contributed to the infarction and subsequent passing of the patient. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35876 | 1220648-2023-01879 | 16387813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16387813&pc=OZD | 02/03/2023 | 02/16/2023 |  | 02/16/2023 |  |  | 02/03/2023 | 02/16/2023 |  |  | 05/31/2024 |  | 02/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023224963 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/03/2023 |  |  |  |  | The complainant reported a 66-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. following impella support and explant, it was documented that a doppler pulse in the lower left leg could not be identified. an intravascular ultrasound was performed and it was discovered that a dissection had occurred. endovascular therapy was subsequently completed to treat the dissection. it was noted by the medical examiner following the events that difficulty was experienced during both insertion and explant of the pump. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35877 | 1220648-2023-01880 | 16393393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16393393&pc=OZD | 01/17/2023 | 02/16/2023 |  | 02/16/2023 |  |  | 01/17/2023 | 02/16/2023 |  |  | 08/31/2024 |  | 02/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023269573 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/17/2023 |  |  |  |  | The complainant reported a 67-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. following advancement of the repositioning sheath for impella support, no flow was observed down the impella leg. a fem-fem bypass was subsequently performed to treat the noted ischemia. support continued following the event. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35878 | 1220648-2023-01870 | 16380597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16380597&pc=OZD | 01/22/2023 | 02/15/2023 |  | 02/15/2023 |  |  | 01/22/2023 | 02/15/2023 |  |  | 09/30/2024 |  | 02/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023279104 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/22/2023 |  |  |  |  | The complainant reported a 90-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was documented that the patient developed a hematoma and an access site bleed. femostop was applied to achieve hemostasis and the patient was provided an unknown quantity of blood products to counteract the bleed. the pump was subsequently removed as a result of the condition. the patient was stable following the events. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35879 | 1220648-2023-01871 | 16382305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16382305&pc=OZD | 01/18/2023 | 02/15/2023 |  | 02/15/2023 |  |  | 01/18/2023 | 02/15/2023 |  |  | 10/31/2024 |  | 02/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023284180 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Unspecified Infection | Ulcer | Vascular Dissection | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/18/2023 |  |  |  |  | The complainant reported a 72-year-old female patient presenting for high risk percutaneous coronary intervention. following impella explant, it was documented that the patient developed a hematoma. manual compression and a femostop were applied, however the patient's hgb dropped and a transfusion consisting of 3 units rbc and 1 unit of platelets was performed. upon further consultation, it was identified that a surgical repair was necessary as the femoral artery had a tear. the patient was subsequently taken to the or for surgical repair of the femoral artery and common femoral vein. upon follow up, it was further disclosed that the patient developed an infection at the intervention site. fibrinous tissue was identified on the staple line with ulceration. antibiotics were given to treat the infection and the patient was scheduled to debride and washout the site. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35880 | 1220648-2023-01872 | 16382361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16382361&pc=OZD | 01/12/2023 | 02/15/2023 |  | 02/15/2023 |  |  | 01/12/2023 | 02/15/2023 |  |  | 09/30/2024 |  | 02/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023274480 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [18 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/12/2023 |  |  |  |  | The medical center had placed an impella cp and ecmo for support of a patient who arrested and was down in the field for an extended duration of time after arresting and crashing his car. the team also had placed a lucas device. the critical condition and impella support lasted for 2 days til the family made choice to withdraw care with a "dead gut" and ischemic limb accessed by the impella. the impella caused ischemia was to be treated by external reperfusion catheter, but care was withdrawn and patient expired. | [5 empty] | The medical center did not retain the impella pump or sheath product for analysis. the product was discarded a the time of the patient's death. a root cause of the limb ischemia can not be determined, though should more information be shared that leads to root cause, a final report will follow. the failure mode will be monitored and trended.
35881 | 1220648-2023-01873 | 16382362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16382362&pc=OZD | 01/16/2023 | 02/15/2023 |  | 02/15/2023 |  |  | 01/16/2023 | 02/15/2023 |  |  | 06/30/2024 |  | 02/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023238260 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/16/2023 |  |  |  |  | The complainant reported a 75 year old male patient presenting for acute myocardial infarction and cardiogenic shock. upon explant of the pump, an angiogram was performed, and it was noted that there was a clot in the right femoral artery. a thrombectomy was subsequently performed to treat the observed clot. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35882 | 1220648-2023-01874 | 16382369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16382369&pc=OZD | 01/20/2023 | 05/16/2023 |  | 02/15/2023 |  |  | 01/20/2023 | 02/15/2023 |  |  | 09/30/2024 |  | 02/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023274609 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/20/2023 |  |  |  |  | The medical center placed an impella cp for support of a patient with multiple cardiac and systemic comorbidities. the patient was admitted for a mitral valve replacement and when under anaesthesia she deteriorated and was placed on the cp. after 3 hours of impellla support the patient was diagnosed with an embolic cva and facial asymmetry. the thought was the embolic event was due to the impella support. the patient has survived the event. no intervention was done for the small stroke. | [5 empty] | The medical center did not return the impella pump for any analysis. no imaging was shared. no root cause of the cva can be determined. should new information be shared that leads to a root cause, a final report will be filed. the failure mode will be monitored and trended. | Since the original report was filed, the investigation into the cva has concluded. the product was not returned for any benchtop analysis, however the data logs were able to be reviewed. the logs revealed no evidence of biomaterial ingestion. there was no evidence of the embolic event being due to the use of impella. the causal relationship between the cva and the use of impella could not be determined. the failure mode will be monitored and trended of note the ifu states: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35883 | 1220648-2023-01876 | 16382382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16382382&pc=OZD | 01/17/2023 | 02/15/2023 |  | 02/15/2023 |  |  | 01/18/2023 | 02/15/2023 |  |  | 09/30/2024 |  | 02/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023280201 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/18/2023 |  |  |  |  | The complainant reported an 80 year old female patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was documented that the patient developed an access site bleed. the patient was provided with two units of prbc to counteract the blood loss. the pump was later removed during a planned explant unrelated to the bleed. the patient was considered to be stable following the event. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35884 | 1220648-2023-01875 | 16382384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16382384&pc=OZD | 01/20/2023 | 02/15/2023 |  | 02/15/2023 |  |  | 01/20/2023 | 02/15/2023 |  |  | 08/31/2024 |  | 02/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023265293 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/20/2023 |  |  |  |  | The medial center had an impella cp placed for support of a patient admitted in cardiogenic shock, in which there was an impella access site bleed. the bleed was not remedied by manual hold or femostop. the team had to infuse 2 units of blood back to the patient. | [5 empty] | The medical center discarded the impella product at the time of explant. no benchtop analysis to root cause is possible. should new information be shared that leads to root cause, a final report will be filed. the failure mode will be monitored and trended.
35885 | 1220648-2023-01877 | 16382385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16382385&pc=OZD | 01/20/2023 | 02/15/2023 |  | 02/15/2023 |  |  | 01/20/2023 | 02/15/2023 |  |  | 09/30/2024 |  | 02/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023282908 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/20/2023 |  |  |  |  | The medical center placed an impella cp for support and developed and access site bleed at the femoral site. the bleed developed into a hematoma. the team attempted to troubleshoot by sheath removal and placement of pump's repositioning sheath. the hematoma still grew and so the team explanted and placed a cook 14fr sheath to achieve hemostasis. the team states the patient lost over 2 units of blood and was infused back more than 2 units. | [5 empty] | The medical center discarded the impella sheath and pump at the time of explant. no benchtop analysis to root cause is possible. should new information be shared that leads to cause, a final report will follow. the failure mode will be monitored and trended.
35886 | 1220648-2023-01878 | 16382411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16382411&pc=OZD | 02/01/2023 | 02/15/2023 |  | 02/15/2023 |  |  | 02/01/2023 | 02/15/2023 |  |  | 06/30/2024 |  | 02/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023233938 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/01/2023 |  |  |  |  | The complainant reported a 67-year-old female patient presenting for cardiomyopathy. following escalation from an impella cp to an impella 5.5, the patient was noted to have stiffness in the leg from which the impella cp was removed. a hematoma was observed at the site and bleeding was identified via ct scan. the patient underwent a blood transfusion consisting of more than ten units of blood and an endovascular aortic repair was completed to achieve hemostasis. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35887 | 1220648-2023-01866 | 16375494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16375494&pc=OZD | 01/15/2023 | 06/02/2023 |  | 02/14/2023 |  |  | 01/15/2023 | 02/14/2023 |  |  | 11/30/2024 |  | 02/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023278445 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/15/2023 |  |  |  |  | The complainant reported a patient presenting with acute myocardial infarction and cardiogenic shock for impella support. after approximately 40 hours into support, the patient yanked on the device and the pump stopped and could not be restarted. the pump was removed. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the reported pump stop has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that an open electrical line due to excessive force by the patient was the cause of the pump stop. the failure modes will be monitored and trended. pump stop ifu quotes: impella 5.5, & cp ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
35888 | 1220648-2023-01867 | 16375649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16375649&pc=OZD | 01/18/2023 | 02/14/2023 |  | 02/14/2023 |  |  | 01/18/2023 | 02/14/2023 |  |  | 09/30/2024 |  | 02/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023278068 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/18/2023 |  |  |  |  | The medical center had an impella cp inserted to support a patient admitted suffering from an acute myocardial infarction. the support was for just over 1 hour to allow the medical team to perform coronary angioplasty. the patient had a known perforation of the femoral artery which complicated the impella access and closure by the perclose device. after closure the team ballooned the impella access to achieve hemostasis. this caused the limb to become ischemic and vascular surgery was called in for intervention and repair. | [5 empty] | The medical center discarded the impella product at the time of explant. no benchtop analysis is possible to determine root cause of the ischemia, though the patient condition and known peripheral arterial disease could have contributed. the failure mode will be monitored and trended.
35889 | 1220648-2023-01868 | 16375656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16375656&pc=OZD | 01/19/2023 | 05/17/2023 |  | 02/14/2023 |  |  | 01/19/2023 | 02/14/2023 |  |  | 10/31/2024 |  | 02/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023283768 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/19/2023 |  |  |  |  | The medical center had an impella cp optical placed above the access of the ecmo, which was placed in the right femoral artery. the day prior there was robotic heart valve surgery performed, also accessed in the same right femoral. the impella access caused an access site injury to the vasculature that was remedied by 5 units of blood and surgical cutdown and hemostasis. | [5 empty] | The medical center discarded the impella pump product at the time of explant. no benchtop analysis to root cause is possible. should new information be shared and a root cause determined, a final report will follow. the failure mode will be monitored and trended. | Since the original report was filed, the investigation has concluded. no product was returned for analysis, but the clinical report was investigated. the root cause was determined to be anticoagulation management. the patient had high act values, with heparin dosed continuously at thousands units/hour. the failure mode will be monitored and trended. the ifu notes the following: ¿assess access site for bleeding and hematoma.¿
35890 | 1220648-2023-01869 | 16375787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16375787&pc=OZD | 01/14/2023 | 02/14/2023 |  | 02/14/2023 |  |  | 01/20/2023 | 02/14/2023 |  |  | 10/31/2024 |  | 02/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023278357 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/20/2023 |  |  |  |  | The complainant reported a patient of unknown age or gender presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, an access site bleed developed. unspecified surgical intervention was required to resolve. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35891 | 1220648-2023-01865 | 16369387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16369387&pc=OZD | 01/16/2023 | 10/24/2024 |  | 02/13/2023 |  |  | 01/16/2023 | 02/13/2023 |  |  | 07/31/2024 |  | 02/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023246667 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Vascular Dissection | [12 empty] | Death | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2023 |  |  |  |  | The complainant in japan had an impella cp that failed delivery despite all troubleshooting and measures taken. the pump did not deliver and the team instead placed ecmo. after the delivery failure the team observed a dissection that may have been due to the insertion attempts. the dissection cascaded up from the impella access through the aorta. surgical intervention was considered, however the patient expired from underlying heart failure. | [5 empty] | The medical team in japan discarded the impella product. no benchtop analysis is possible to root cause. should more information be shared that leads to root cause, a final report will be filed. the failure mode will be monitored and trended. | The investigation for the reported delivery issue, vascular damage, and death has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were sufficient to determine a root cause. the cause of the reported issues is patient condition. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
35892 | 16350743 | 16350743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16350743&pc=OZD | 08/06/2022 | 08/24/2022 | 02/03/2023 | 02/10/2023 | 02/03/2023 | 02/10/2023 |  | 02/10/2023 |  |  |  |  | 02/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP PUMP SET |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003.0814 |  | No answer provided |  |  | Not Available | [6 empty] | User Facility report | Yes | Other Health Care Professional | [6 empty] | 32485 DA |  | Male |  |  |  | Malfunction | No | Yes | Inaccurate Flow Rate | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y |  |  |  | Impella placed for ischemic cardiomyopathy. approximately 3 days later, observed increased flow rate and fluid leaking from impella purge cassette, tubing changed. following day flow rate increased again, cassette checked, moisture noted inside. tubing changed again, flow rate reduced below 10 cc/hr, and no additional leakage detected.
35893 | 1220648-2023-01860 | 16356052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16356052&pc=OZD | 01/10/2023 | 09/26/2024 |  | 02/10/2023 |  |  | 01/10/2023 | 02/10/2023 |  |  | 09/30/2024 |  | 02/10/2023 |  | TEMPORARY NON-ROLLERT TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023273040 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Material Twisted/Bent | Adverse Event Without Identified Device or Use Problem | [5 empty] | Sepsis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | Death |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2023 |  |  |  |  | The patient additionally was support by ecmo and experienced runs of vt and arrested. the pump positioning was found to be causing the pump to kink. it was found that the patient underwent septic shock. the procedure outcome was that patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The medical center had an impella cp support, for a patient admitted in cardiogenic shock, that was complicated by poor positioning. the patient additionally was support by ecmo and experienced runs of vt and arrested during the support. the pump positioning was found to be causing the pump to kink. to rectify the kinking and remove the pump the team needed to snare the pump via access at the radial artery. the use of a snare allowed for the pump to be straightened and removed from the patient's vasculature. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-01860 was provided in sections b1, b2, b5, d6, h1, and h6. | The medical center discarded the impella pump at the time of explant. no imaging was shared for analysis. there is no definitive root cause for the kink and necessary snaring of the impella pump. the failure mode will be monitored and trended.
35894 | 1220648-2023-01861 | 16356304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16356304&pc=OZD | 09/30/2020 | 02/10/2023 |  | 02/10/2023 |  |  | 02/06/2023 | 02/10/2023 |  |  | 04/30/2022 |  | 02/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1479524 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/06/2023 |  |  |  |  | The medical team in japan had a successful impella cp support in 2020 for 58 hours, when explanted. the jpvad registry team later in 2/2023 reported upon hemolysis treated by the infusion of blood products. the jpvad team furnished no further details. | [5 empty] | The medical center in japan discarded all product in 2020. no benchtop analysis and hemolysis testing is possible. no root cause can be determined. the jpvad registry furnished no further clinical details. the failure mode will be monitored and trended.
35895 | 1220648-2023-01863 | 16357517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16357517&pc=OZD | 12/17/2022 | 02/10/2023 |  | 02/10/2023 |  |  | 01/13/2023 | 02/10/2023 |  |  | 04/30/2024 |  | 02/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023215778 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 01/13/2023 |  |  |  |  | The complainant reported a 39 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. the patient developed ischemia in the impella inserted limb. a distal perfusion catheter was placed as treatment. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35896 | 1220648-2023-01862 | 16357525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16357525&pc=OZD | 01/11/2023 | 09/19/2024 |  | 02/10/2023 |  |  | 01/11/2023 | 02/10/2023 |  |  | 08/31/2024 |  | 02/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2023267295 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/11/2023 |  |  |  |  | It was reported the patient expired after being successfully explanted. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a 59 year old male patient presenting for impella support. during the course of support, the pump became dislodged and perforated the ventricle. surgical intervention was required. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-01862 was provided. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35897 | 1220648-2023-01864 | 16357848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16357848&pc=OZD | 01/12/2023 | 09/20/2024 |  | 02/10/2023 |  |  | 01/12/2023 | 02/10/2023 |  |  | 04/30/2024 |  | 02/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022147437 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/12/2023 |  |  |  |  | It was reported that the external femoral to femoral bypass that was placed did not result in pedal pulses. the patient was on 5 drips and was started on dobutamine. the medical team was advised to escalate the patient to an impella 5.5, but this suggestion was refused as the patient was too sick to transfer. the patient continued to decline, and went into pulseless electrical activity and expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a 62 year old male patient presenting for impella support. the patient developed limb ischemia in the impella inserted extremity. an external fem-fem bypass was placed in response to the occlusive blood flow. |  |  |  |  | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-01864 was provided in sections b2, b5, d6a, d6b, h1, and h6. | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35898 | MW5114866 | 16363097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16363097&pc=OZD | 10/28/2022 | 02/08/2023 |  | 02/10/2023 |  |  |  | 02/10/2023 |  |  |  |  | 02/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED INC. | [8 empty] | No | [5 empty] | No | Health Professional |  |  | Voluntary report | Not Available | Other Health Care Professional | [6 empty] | 48 YR |  | Female | 78 KG | Non Hispanic | White | Death | Yes | Yes | Break | Inadequate Instructions for Healthcare Professional | Protective Measures Problem |  |  |  |  | Insufficient Information | [12 empty] | Death | [28 empty] | Initial submission | [9 empty] | The cartridge on the purge line is plastic and on the impella cp this cartridge does not have any "protection". the purge bypass line broke at the cartridge rendering the impella inoperable. there is a purge bypass; however, company does not make customers aware of this until specifically asked about it.
35899 | 1220648-2023-01853 | 16333797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16333797&pc=OZD | 01/10/2023 | 06/06/2023 |  | 02/09/2023 |  |  | 01/10/2023 | 02/09/2023 |  |  | 10/31/2024 |  | 02/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023283339 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2023 |  |  |  |  | The complainant reported a 71 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. the patient developed ischemia in the impella extremity, therefore, a fem-fem bypass was configured. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical reviewer noted small patient vessel due to vasopressors. the evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35900 | 1220648-2023-01854 | 16337523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16337523&pc=OZD | 11/22/2022 | 02/09/2023 |  | 02/09/2023 |  |  | 01/10/2023 | 02/09/2023 |  |  | 06/30/2024 |  | 02/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023236615 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/10/2023 |  |  |  |  | The complainant reported a 77-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was documented that there were persistent suction issues and that pump flows could not be maintained despite the addition of a second pump. the initial pump was subsequently replaced with another impella cp pump and the flow issues resolved. the pump was returned for evaluation and it was discovered that the flow issues were caused by broken impeller blades. there was no harm to the patient as a result of the issue. | [5 empty] | The pump was returned by the user for further evaluation. a visual inspection was performed, and it was discovered that the impeller blades were fractured. there was no biomaterial observed within the pump and a definitive cause for the broken blades could not be determined. should additional relevant information be obtained, a supplemental report will be filed.
35901 | 1220648-2023-01855 | 16338570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16338570&pc=OZD | 01/13/2023 | 02/09/2023 |  | 02/09/2023 |  |  | 01/13/2023 | 02/09/2023 |  |  | 10/31/2024 |  | 02/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023272235 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/13/2023 |  |  |  |  | The complainant in canada had an impella 5.5 support that was ongoing for over 22 days, which is beyond indication, when the pump was explanted from the axillary insertion site. the patient had been transferred to the canadian facility from the medical center in the united states on day 9 of the impella support. the pump explant took place after the patient experienced vision issues and numbness in the arm. the pump was explanted with clot burden observed upon it, and as such the team believes the pump caused an embolic stroke to the patient. the stroke was not noted to be treated and the patient survived. | [5 empty] | The medical center in canada discarded all impella product at the time of explant. as such, no benchtop analysis was possible. photos of the impella pump waith adhered biomaterial was shared by the medical team. a definitive root cause of the cerebral vascular event could not be determined.
35902 | 1220648-2023-01856 | 16338625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16338625&pc=OZD | 01/10/2023 | 10/14/2024 |  | 02/09/2023 |  |  | 01/10/2023 | 02/09/2023 |  |  | 01/31/2023 |  | 02/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2022038540 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2023 |  |  |  |  | The complainant reported a 67-year-old male patient presenting for an impella implant. it was noted that the patient began experiencing ischemia with an implanted intra-aortic balloon pump (iabp) prior to insertion of the impella cp pump. the pump replaced the iabp, however, the ischemia continued. one day following implant, a clot was found in the reperfusion sheath prompting placement of a retrograde contralateral sheath. this was connected to the already in place antegrade femoral sheath. at this point in time, the patient's left foot was cool and had a mottled appearance. as a result of the ongoing ischemia, the physician was advised to explant the pump, or transfer the patient to a facility that could remove the pump and place an axillary 5.5 pump. the patient was weaned on the impella, however, it wasn¿t until four days later that the pump was explanted. a follow up with the nursing unit reported that the patient had to have the affected leg amputated several days after the explant. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information for the initial mfr 1220648-2023-01856 was provided in sections b2, d6a, d6b, g2, h1, h6, and h10. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
35903 | 1220648-2023-01857 | 16338994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16338994&pc=OZD | 01/11/2023 | 02/09/2023 |  | 02/09/2023 |  |  | 01/11/2023 | 02/09/2023 |  |  | 03/31/2024 |  | 02/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023209898 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/11/2023 |  |  |  |  | The medical center had an impella cp delivered for support of a patient undergoing coronary angioplasty after being refused surgery due to underlying conditions. the delivery was complicated by bilateral peripheral arterial disease at the lower limbs. the team treated the impella accessed limb by shockwave to allow for delivery of the impella. the team had the cp support for a total of 92 minutes. after explant the team observed the vasculature to be dissected. the team stented the femoral artery and hemostasis was achieved. | [5 empty] | The medical team discarded the impella product at the time of explant. no benchtop analysis to root cause is possible, however the known peripheral arterial disease most likely contributed to the issue and vessel damage. patient condition contributed to the femoral artery damage and necessary intervention.
35904 | 1220648-2023-01858 | 16340034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16340034&pc=OZD | 01/10/2023 | 02/09/2023 |  | 02/09/2023 |  |  | 01/10/2023 | 02/09/2023 |  |  | 08/31/2024 |  | 02/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED INC. |  |  | US |  | 2023251199 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/10/2023 |  |  |  |  | The complainant in the eu had placed an impella cp for support despite known peripheral arterial disease in the lower limbs. the gender and age is not known to the abiomed field representative in the netherlands. the small vasculature and cold foot prompted the team to place fem-fem bypass. the flow was not fully restored. the impella accessed limb remained cold and pale. the team left the pump within the femoral until the family made the choice to withdraw care. | [5 empty] | The medical center in the (b)(6) discarded the impella pump product at the time that care was withdrawn. the pump benchtop analysis to root cause is not possible. the native condition of peripheral arterial disease and small vasculature most likely contributed to the ischemia. the failure mode will be monitored and trended.
35905 | 1220648-2023-01859 | 16340196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16340196&pc=OZD | 01/10/2023 | 02/09/2023 |  | 02/09/2023 |  |  | 01/10/2023 | 02/09/2023 |  |  | 09/30/2024 |  | 02/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023273897 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/10/2023 |  |  |  |  | The complainant reported a female patient presenting for impella implant. during impella support, it was noted that the patient developed lower limb ischemia. as a result of the condition, the decision was made to explant the pump. flow returned to the leg following explant. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35906 | 1220648-2023-01842 | 16327071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16327071&pc=OZD | 01/11/2023 | 02/08/2023 |  | 02/08/2023 |  |  | 01/11/2023 | 02/08/2023 |  |  | 03/31/2023 |  | 02/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2022059082 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/11/2023 |  |  |  |  | The japanese complainant had an impella 5.0 support ongoing when there were signs of hemolysis. the hemolysis was noted by hematuria alone, no lab value changes. the hemolysis was treated with dialysis initiation. the patient had also been supported by the ecmo circuit as well. | [5 empty] | The medical center in japan returned the impella pump product for analysis and benchtop hemolysis testing. the investigation is on-going at this time. upon completion of the investigation a final report with root cause will be filed.
35907 | 1220648-2023-01844 | 16327121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16327121&pc=OZD | 12/28/2022 | 10/25/2023 |  | 02/08/2023 |  |  | 01/09/2023 | 02/08/2023 |  |  | 11/30/2024 |  | 02/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023275546 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/09/2023 |  |  |  |  | Upon further clarification from the field rep, it was noted that the patient did not require medical intervention for the leaflet damage. based on this provided information, the severity of the case has been reduced and it is no longer considered a reportable serious injury. | The complainant reported a 54-year-old male patient presenting for post cardiotomy cardiogenic shock. following pump explant, it was discovered that the aortic valve leaflet had been damaged during impella support. intervention for the condition was not established. as a result of the damage, the patient had moderate aortic insufficiency with an ejection fraction of thirty five percent. the patient was considered stable at the time of case creation. |  |  |  |  | Additional information provided: sections b5, h1, h2, and h6 have been updated based on follow up information provided from the field rep. this case is no longer being considered a reportable serious injury as the noted adverse event did not prompt medical intervention. | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | Correction provided: upon reassessment it was determined that this is a reportable serious injury. the investigation into the reported aortic valve damage was completed. based on the available information, the root cause of the aortic valve damage and its relation to the device were not determined as the data logs and sufficient clinical information were not provided. impella 55.5 ifu ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
35908 | 1220648-2023-01845 | 16327174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16327174&pc=OZD | 01/25/2023 | 02/08/2023 |  | 02/08/2023 |  |  | 01/26/2023 | 02/08/2023 |  |  | 08/31/2024 |  | 02/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC |  |  | US |  | 2023267289 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Difficult to Remove | Detachment of Device or Device Component | [5 empty] | Foreign Body In Patient | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/26/2023 |  |  |  |  | The complainant reported a 77-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during removal, it was noted that the impella got stuck at the anastomosis graft. the physician subsequently pulled on the device with force and the pump separated between the shaft and the motor. the cannula was subsequently left within the graft and had to be removed separately by the physician. there was no additional harm following removal of the broken component. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35909 | 1220648-2023-01846 | 16327215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16327215&pc=OZD | 01/12/2023 | 02/08/2023 |  | 02/08/2023 |  |  | 01/12/2023 | 02/08/2023 |  |  |  |  | 02/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | [5 empty] | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/12/2023 |  |  |  |  | During a periodic literature review a publication was found inclusive of multiple impella pump patients. the us medical center and physicians published upon a retrospective review of 8 impella supported patients that were escalated from impella cp to the impella 5.0 or 5.5 pump due to signs/symptoms of hemolysis. signs were inclusive of ldh levels and pfhg levels rising. additionally as noted in the publication 1/2 of the patients had dialysis applied to treat the renal issues. | [5 empty] | The medical center and authors did not return any impella pump product for analysis or benchtop testing. no root cause can be determined as data points are lacking. the only information for analysis is what is published. should more information be shared, a final report will follow. the failure mode will be monitored and trended.
35910 | 1220648-2023-01847 | 16329039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16329039&pc=OZD | 01/10/2023 | 06/12/2023 |  | 02/08/2023 |  |  | 01/10/2023 | 02/08/2023 |  |  | 10/31/2024 |  | 02/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023272736 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2023 |  |  |  |  | The complainant reported a 74 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. it was noted that malperfusion in the impella-inserted extremity developed during support. as a result, the device was removed prematurely. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the patient condition was the cause of the limb ischemia. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35911 | 1220648-2023-01848 | 16329048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16329048&pc=OZD | 01/10/2023 | 08/17/2023 |  | 02/08/2023 |  |  | 01/10/2023 | 02/08/2023 |  |  | 08/31/2024 |  | 02/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023270480 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2023 |  |  |  |  | The complainant reported a 60 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. a bleed from the introducer site developed prompting introducer removal and 7 units of blood products delivered. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical information provided, the cause of the access site bleeding through the membrane of the introducer was unable to be determined as no product was returned for evaluation. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35912 | 1220648-2023-01849 | 16329064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16329064&pc=OZD | 01/09/2023 | 06/06/2023 |  | 02/08/2023 |  |  | 01/09/2023 | 02/08/2023 |  |  | 06/30/2024 |  | 02/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023276818 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/09/2023 |  |  |  |  | The complainant reported an 89 year old female patient presenting for high risk percutaneous coronary intervention for impella support. during support, while the 14 fr introducer was in the access site, the patient exhibited reduced blood flow proximal to the impella. intervention was required to regain flow via a percutaneous transluminal angioplasty and the introducer was removed. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition related to the pad/pvd was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. limb ischemia ifu quote: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35913 | 1220648-2023-01850 | 16329108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16329108&pc=OZD | 01/07/2023 | 02/08/2023 |  | 02/08/2023 |  |  | 01/09/2023 | 02/08/2023 |  |  | 08/31/2024 |  | 02/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023256748 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/09/2023 |  |  |  |  | The complainant reported a 78 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed an access site bleed. a blood transfusion was required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35914 | 1220648-2023-01851 | 16329118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16329118&pc=OZD | 01/07/2023 | 06/13/2023 |  | 02/08/2023 |  |  | 01/07/2023 | 02/08/2023 |  |  | 08/31/2024 |  | 02/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023269683 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/07/2023 |  |  |  |  | The complainant reported a 42 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed leg ischemia in the impella inserted leg. a fem-fem bypass was configured to address lack of perfusion. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the patient condition was the cause of the limb ischemia. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35915 | 1220648-2023-01852 | 16329128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16329128&pc=OZD | 01/05/2023 | 06/16/2023 |  | 02/08/2023 |  |  | 01/05/2023 | 02/08/2023 |  |  | 09/30/2024 |  | 02/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023280216 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2023 |  |  |  |  | The complainant reported a 77 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. the device required premature removal due to suspected hemolysis. | [5 empty] | The investigation into the reported hemolysis has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis. the data logs showed recurrent suction alarms throughout support. no other abnormal metrics or trends were observed. the root cause for hemolysis remains undetermined since neither the product nor sufficient clinical information were provided. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35916 | 1220648-2023-01843 | 16329748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16329748&pc=OZD | 01/10/2023 | 06/13/2023 |  | 02/08/2023 |  |  | 01/12/2023 | 02/08/2023 |  |  | 10/31/2024 |  | 02/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023276304 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2023 |  |  |  |  | The medical team in japan had an impella cp support ongoing when pulses were lost, at the limb with the impella access site, and the team made choice to place antegrade sheath for perfusion to the limb. the pump remained on for support despite the harm for a total of just over 6 days. the pump was then weaned and explanted. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center in japan discarded the impella pump product at the time of explant. no benchtop analysis of root cause is possible; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. | The medical center in japan discarded the impella pump product at the time of explant. no benchtop analysis to root cause is possible. should any new information be shared, a final report will be filed with root cause of the ischemia. the failure mode will be monitored and trended.
35917 | 1220648-2023-01839 | 16315220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16315220&pc=OZD | 01/13/2023 | 09/22/2024 |  | 02/07/2023 |  |  | 01/13/2023 | 02/07/2023 |  |  | 10/31/2024 |  | 02/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED INC |  |  | US |  | 2023275450 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/13/2023 |  |  |  |  | The medical center in the eu had an impella cp placed for hemodynamic support of a patient admitted in cardiogenic shock. during the support the alarms and pump values prompted the team to believe there was hemolysis caused by the impella. the team explanted the pump to remedy the issue after just over 56 hours of support. at the time of pump explant there was biomaterial adhered to the pump. | Medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the presumed hemolysis is on-going at this time. upon completion of the investigation a final report will be filed. | The returned device was inspected and a piece of biomaterial was found wrapped around the impeller. based on the device inspection and the clinical account of sudden increase in placement signal and motor current, following suspected hemolysis, the root cause of the hemolysis is most likely due to biomaterial. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6.
35918 | 1220648-2023-01840 | 16315225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16315225&pc=OZD | 01/08/2023 | 02/07/2023 |  | 02/07/2023 |  |  | 01/08/2023 | 02/07/2023 |  |  | 11/30/2024 |  | 02/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED INC. |  |  | US |  | 2023278014 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/08/2023 |  |  |  |  | The medical center had an impella 5.5 pump support ongoing for a patient admitted with cardiomyopathy, via an insertion site at the axillary. after 3 days there was a hematoma that developed at the impella graft site. the hematoma prompted the team to infuse 2 units of blood products. additionally on the subsequent day they did a washout to the site. the pump remained on for support for a total of 12 days despite the access site bleed. | [5 empty] | The medical center discarded the impella product at the time of explant. there can be no benchtop analysis to root cause. should more information be shared that leads to cause, a final report will be filed. the failure mode will be monitored and trended.
35919 | 1220648-2023-01841 | 16315230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16315230&pc=OZD | 01/10/2023 | 03/30/2023 |  | 02/07/2023 |  |  | 01/10/2023 | 02/07/2023 |  |  | 12/31/2023 |  | 02/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023161952 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2023 |  |  |  |  | The medical team in the eu had an impella cp placed for support. after placing the patient in the icu for monitoring the patient moved and caused a hematoma at the impella insertion site and acute limb ischemia. the family was consulted and due to a desire to not escalate the care and support, the team made choice to explant the pump and allow for the patient to expire. the hematoma and ischemia were not treated beyond pump explant. | [5 empty] | The complainant in germany discarded the impella pump product at the time of explant. no benchtop analysis to root cause is possible. the manufacturer will close the investigation, though should new information be shared that leads to cause, a final report will be filed. the failure mode will be monitored and trended. | Since filing the initial report the investigation into to limb ischemia and bleeding has concluded. no product was returned, however the clinical report was reviewed. the root cause of the access site bleeding was determined to be due to patient movement and the limb ischemia was most likely patient condition. the failure modes will be monitored and trended.
35920 | 1220648-2023-01838 | 16306641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16306641&pc=OZD | 01/07/2023 | 02/06/2023 |  | 02/06/2023 |  |  | 01/07/2023 | 02/06/2023 |  |  | 09/30/2024 |  | 02/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023273521 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/07/2023 |  |  |  |  | The complainant reported a 72-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient developed a hematoma around the access site. hemostasis was achieved via manual pressure and the patient was provided an unknown quantity of blood to offset the volume loss. patient support continued following the intervention. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35921 | 1220648-2023-01837 | 16301456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16301456&pc=OZD | 12/29/2022 | 07/05/2023 |  | 02/04/2023 |  |  | 01/04/2023 | 02/04/2023 |  |  | 06/30/2024 |  | 02/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023232069 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2023 |  |  |  |  | The complainant reported a 69 year old female patient presenting with cardiomyopathy for impella support. during support, the impella perforated the ventricle while the physician was attempting to reposition the device. the repair was completed expeditiously and support continued. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the ventricle perforation has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis; only the clinical report was evaluated. based on the clinical details, the ventricle perforation was caused by manipulation of the heart during a cabg procedure. the failure mode will be monitored and trended.
35922 | 1220648-2023-01832 | 16295517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16295517&pc=OZD | 01/05/2023 | 02/03/2023 |  | 02/03/2023 |  |  | 01/06/2023 | 02/03/2023 |  |  | 08/31/2024 |  | 02/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023251161 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/06/2023 |  |  |  |  | The medical center in (b)(6) had an impella cp support ongoing when there was an access site bleed from the impella limb. the bleed was treated by blood product infusion and other intervention of a "gizzard" to achieve hemostasis. the patient additionally had ecmo cannula in the limb with the bleed. | [5 empty] | The medical center in (b)(6) discarded the impella pump product. as such, no product analysis was done and no root cause was determined. should more information be shared that leads to root cause, a final report will be filed. the failure mode will be monitored and trended.
35923 | 1220648-2023-01833 | 16295531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16295531&pc=OZD | 01/06/2023 | 06/13/2023 |  | 02/03/2023 |  |  | 01/06/2023 | 02/03/2023 |  |  | 06/30/2024 |  | 02/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023248513 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2023 |  |  |  |  | The medical team had an impella cp placed for support, for a patient undergoing a high risk coronary intervention, via the femoral artery. the limb became ischemic and had fem-fem bypass done to restore flow to the leg. the pump was later explanted after 28 hours of support. the team instead placed escalated axillary support with an impella 5.5 pump. the pulses returned to the leg after explant. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that the patient condition was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. | The medical center discarded all product at the time of impella explant. no root cause can be determined. should more be shared that leads to root cause, a final report will follow.
35924 | 1220648-2023-01834 | 16298130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16298130&pc=OZD | 01/06/2023 | 06/05/2023 |  | 02/03/2023 |  |  | 01/06/2023 | 02/03/2023 |  |  | 10/31/2024 |  | 02/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023274505 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2023 |  |  |  |  | The complainant reported a 45 year old male patient presenting with cardiomyopathy for impella support. the patient developed a hematoma resulting from an unsuccessful impella delivery attempt. it was discovered that the guidewire was improperly loaded through the inlet cage and not through the nose of the catheter. vascular repair was needed. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the cause of the delivery issues. the failure mode will be monitored and trended.
35925 | 1220648-2023-01835 | 16298248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16298248&pc=OZD | 01/05/2023 | 06/06/2023 |  | 02/03/2023 |  |  | 01/05/2023 | 02/03/2023 |  |  | 09/30/2024 |  | 02/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023277401 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2023 |  |  |  |  | The complainant reported a 32 year old male patient presenting with cardiomyopathy. during support, the impella pump continued to interact with the mitral valve, for which repositioning was completed to resolve the issue each time. despite efforts to reposition, the patient's mitral valve was damaged. the pump was removed and repair to the valve was completed. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the heart valve damage has been completed since the original report was filed. the impella device was returned for benchtop analysis and the clinical evaluation revealed that the improper positioning of the device caused the unique patient¿s anatomy was the most likely cause of the heart valve damage. the failure mode will be monitored and trended. heart valve damage ifu quotes: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
35926 | 1220648-2023-01836 | 16299983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16299983&pc=OZD | 01/06/2023 | 06/13/2023 |  | 02/03/2023 |  |  | 01/06/2023 | 02/03/2023 |  |  | 10/31/2024 |  | 02/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023274571 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  |  |  | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Perforation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2023 |  |  |  |  | A 28-year-old male patient in the us has been treated for ami/cgs. after an initially placed iabp the patient has successfully been escalated to an impella 5.5 heart pump. one week after pump insertion the patient had a severe coughing spell during the night. the impella device perforated the left ventricle, causing bleeding into the chest. the patient expired. contributing factors: asthma, suspected rv involvement. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information received from the field revealed that the clinicians determined that the graft retraction post-cpr caused the reported bleeding. the clinicians confirmed that there was no ventricle perforation. the cause of the patient injury was determined to be unrelated to the impella.
35927 | 1220648-2023-01826 | 16282638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16282638&pc=OZD | 01/03/2023 | 08/15/2023 |  | 02/02/2023 |  |  | 01/03/2023 | 02/02/2023 |  |  | 11/30/2024 |  | 02/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023275541 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Patient was being transported to another facility. while being loaded into the elevator the impella catheter that connects to the controller was hit by either the bed or the elevator doors. once the hit occurred the pump stopped and would not restart. the aic was switched out and the new aic would also not recognize the pump. the patient was taken to the or and the pump was replaced successfully. patient stable afterwards. | [5 empty] | The investigation into the pump stop has been completed since the original report was filed. the user facility has returned the impella device, and the benchtop analysis revealed that the root cause of the pump stop was an electrical open caused by the pump plug hitting the elevator doors during transport. | Investigation of the returned pump is ongoing. once the investigation is complete, a supplemental report will be filed.
35928 | 16282948 | 16282948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16282948&pc=OZD | 01/13/2023 | 01/16/2023 | 01/16/2023 | 02/02/2023 | 01/16/2023 | 02/02/2023 |  | 02/02/2023 |  |  |  |  | 02/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0043-0002 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0043-0002 |  | No |  |  | Not Available |  |  | No | Other |  |  | User Facility report | No | Risk Manager | [8 empty] | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella cp leaking at clip-in site on purge fluid tubing.
35929 | 1220648-2023-01827 | 16283316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16283316&pc=OZD | 01/05/2023 | 02/02/2023 |  | 02/02/2023 |  |  | 01/06/2023 | 02/02/2023 |  |  | 08/31/2024 |  | 02/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023268506 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An 82-year-old female patient presented with stemi and underwent a protected pci procedure with impella cp support. femoral access has been achieved using ultrasound guidance and micro puncture. two perclose devices have been used for pre close. single access site technique used for the pci procedure. best practices followed and pci successfully completed. after removal of the impella device both pre-placed perclose devices were tied, and manual pressure has been applied. minor bleeding at the access site was noted. manual pressure has been applied again and a femostop device applied. the access site remained soft and dry prior to transferring the patient to icu. several hours later the patient suffered from access site bleeding. a covered stent has been placed to achieve hemostasis. no further complications occurred. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35930 | 1220648-2023-01829 | 16285050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16285050&pc=OZD | 01/04/2023 | 10/25/2023 |  | 02/02/2023 |  |  | 01/04/2023 | 02/02/2023 |  |  | 06/30/2024 |  | 02/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023234415 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia Stroke | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/04/2023 |  |  |  |  | This device also endured a purge side arm leak during use. as a result, the pump was replaced. | The medical team had an impella 5.5 support ongoing for over 59 days when the team observed purge alarms. the purge was leaking and the team had a purge extension tubing in place. the pump had been used well beyond indication. the team made decision they would replace the pump. at the time of pump explant there was clot observed. the patient suffered an ischemic stroke on the same day, which may or may not have been due to the use of the impella. the left sided effects from the cva are ongoing now on the 2nd impella 5.5 placed for this patient. |  |  |  |  | (h3) the investigation into the reported "stroke/cva and purge leak - pump" has been completed since the original report was filed. product was returned for investigation. the motor of the pump was returned, but the cannula was not returned. a clot was found on the returned catheter. data logs of the final ten days of support were reviewed and no indication of biomaterial interaction with the pump was seen and there were no motor current spikes or saturated flows. due to this, the relationship between stroke and impella is unknown. the cause of the stroke could not be determined. during analysis a leak was reproduced at the purge tubing to filter neck bond. with the sidearm retainer in place, this bond absorbs any stress on the side arm. therefore, the cause of the leak at the purge tubing to filter neck bond was likely due to excessive force. this report is being filed as part of a retrospective review of historical records. | The investigation into the cva has concluded. the impella pump was analyzed, however not all of the pump was returned. the medical center cut the pump at the motor. the data logs were analyzed as well. the analysis did not reveal any biomaterial interaction with the pump that could have lead to the cva. the failure mode will be monitored and trended. | The investigation into the returned impella product is on-going at this time. upon the completion of the investigation, a final report will be filed.
35931 | 1220648-2023-01828 | 16287543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16287543&pc=OZD | 12/28/2022 | 02/02/2023 |  | 02/02/2023 |  |  | 12/28/2022 | 02/02/2023 |  |  | 09/30/2024 |  | 02/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | UNKNOWN |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/28/2022 |  |  |  |  | The medical center had an impella cp placed via the femoral artery to give hemodynamic support to a patient suffering from cardiogenic shock. the cp supported for just over 4 hours when due to limb ischemia the pump was explanted. the team instead placed an impella 5.5 via the axillary artery insertion. after explant the limb was noted to have normal perfusion. no other intervention was needed. | [5 empty] | The medical center did not return the impella pump for benchtop analysis. the clinical report did state the patient was on vasopressors at the time of pump insertion that may have contributed to the ischemia. patient condition may have contributed to the ischemia. the failure mode will be monitored and trended.
35932 | 1220648-2023-01830 | 16288099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16288099&pc=OZD | 01/03/2023 | 09/08/2023 |  | 02/02/2023 |  |  | 02/03/2023 | 02/02/2023 |  |  | 09/30/2024 |  | 02/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023273168 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/03/2023 |  |  |  |  | Initial report also indicated that the patient was administered a fem-fem bypass in response to the reported ischemia, prior to explant. | The complainant reported a 45 year old male patient presenting with acute myocardial infraction and cardiogenic shock for impella support. the physician noted concern for distal ischemia during impella support. as such, the decision was made to remove the device the following morning. |  |  |  |  | This is a duplicate mdr from an erroneous duplicated complaint. please see 1220648-2023-01831 for original mdr. | B(5): initial report also indicated that the patient was administered a fem-fem bypass in response to the reported ischemia, prior to explant. h(6): included code for unexpected medical intervention (4641). | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35933 | 1220648-2023-01831 | 16288100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16288100&pc=OZD | 01/03/2023 | 02/02/2023 |  | 02/02/2023 |  |  | 01/03/2023 | 02/02/2023 |  |  | 09/30/2024 |  | 02/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023273168 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/03/2023 |  |  |  |  | The complainant reported a 45-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient developed ischemia represented by an intermittent doppler pulse in the leg. to treat the condition, a fem-fem bypass was performed and the pump was explanted. there were no further complications noted following the event. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35934 | 1220648-2023-01823 | 16278689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16278689&pc=OZD | 01/03/2023 | 02/01/2023 |  | 02/01/2023 |  |  | 01/03/2023 | 02/01/2023 |  |  | 09/30/2024 |  | 02/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023275603 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Fracture | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/03/2023 |  |  |  |  | The complainant reported a 63-year-old male patient presenting for high risk percutaneous coronary intervention. during impella support, it was noted that the patient developed an access bleed due to a crack in the 14fr sheath at the hub. the patient lost 1 unit of blood as a result of the noted issue. to counteract the blood loss, the 14fr sheath was replaced and the patient was infused with two units of blood and ns. hemostasis was subsequently achieved and the groin was closed via perclose. the patient was deemed stable following the event. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35935 | 1220648-2023-01824 | 16279328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16279328&pc=OZD | 01/03/2023 | 02/01/2023 |  | 02/01/2023 |  |  | 01/03/2023 | 02/01/2023 |  |  | 07/31/2024 |  | 02/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023259523 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/03/2023 |  |  |  |  | The complainant reported a 56-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient developed hemolysis. as a result of the condition, the surgeon opted to remove the pump. once removed, an intra-aortic balloon pump was inserted for continued support. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35936 | 1220648-2023-01825 | 16280387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16280387&pc=OZD | 01/02/2023 | 08/21/2023 |  | 02/01/2023 |  |  | 01/02/2023 | 02/01/2023 |  |  | 09/30/2024 |  | 02/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023278566 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 52-year-old male patient received hemodynamic support per an impella cp heart pump to perform an vt ablation procedure due to vt episodes. after removal of the impella device leg ischemia has been noted below impella access point. other leg also lost its pulse. distal aortic occlusion was found with a thrombus at the bifurcation. during a long and complicated procedure, a thrombus could be removed from the iliac artery, but blood flow was still not reestablished below the common femoral artery. ekos cathether and tpa drip used for thrombolysis. angiomax drip due to suspected hit. patient stable afterwards. | [5 empty] | The cause of the limb ischemia was biomaterial. the ischemia was unknown relation to impella because ischemia occurred in both legs and thrombus was observed at the femoral bifurcation leading to both legs. | The impella device has been discarded and investigation is not completed yet. a supplemental will be filed upon conclusion of our investigation.
35937 | 1220648-2023-01820 | 16270872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16270872&pc=OZD | 01/02/2023 | 06/12/2023 |  | 01/31/2023 |  |  | 01/02/2023 | 01/31/2023 |  |  | 08/31/2024 |  | 01/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023259588 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/02/2023 |  |  |  |  | The complainant reported a 77-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient developed distal limb ischemia in the impella leg. a fem-fem bypass was subsequently completed to treat the ischemia. patient support continued following the intervention. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the limb ischemia has been completed since the original report was filed. the medical center discarded the impella pump product. no benchtop analysis of root cause is possible; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure mode will be monitored and trended.
35938 | 1220648-2023-01821 | 16270885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16270885&pc=OZD | 12/04/2022 | 01/31/2023 |  | 01/31/2023 |  |  | 12/04/2022 | 01/31/2023 |  |  | 08/31/2024 |  | 01/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023270117 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/04/2022 |  |  |  |  | The complainant reported a 78-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient was getting continuous volume related suction alarms. while attempting to gain access to troubleshoot the issue, the physician perforated the right femoral artery. surgical intervention was performed and the pump was subsequently explanted as a result of the condition. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35939 | 1220648-2023-01822 | 16270899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16270899&pc=OZD | 11/28/2022 | 01/31/2023 |  | 01/31/2023 |  |  | 11/28/2022 | 01/31/2023 |  |  | 02/29/2024 |  | 01/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023202729 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/29/2022 |  |  |  |  | The complainant reported a 45-year-old male patient presenting for cardiomyopathy. during impella support, it was noted that the patient developed an access site bleed. additional sutures were placed and stat seal, along with other sealants, were implemented to counteract the bleed. as a result of the condition, the patient was administered an unknown quantity of blood products. patient support continued following the event. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35940 | 1220648-2023-01813 | 16263842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16263842&pc=OZD | 01/01/2023 | 01/30/2023 |  | 01/30/2023 |  |  | 01/01/2023 | 01/30/2023 |  |  | 09/30/2024 |  | 01/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023276141 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/01/2023 |  |  |  |  | The medical center had an impella cp that was supporting a patient for cardiomyopathy. after just under 2 days there was hemolysis. the hemolysis prompted pump explant, when other troubleshooting measures had failed. they had given volume and adjusted pump positioning. the team placed a new impella for continued support. | [5 empty] | The medical center discarded the impella at the time of explant. no benchtop analysis or hemolysis testing is possible. root cause of the hemolysis can not be determined. the investigation will be closed. should new information be shared that leads to root cause, a final report will be filed.
35941 | 1220648-2023-01814 | 16263852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16263852&pc=OZD | 01/02/2023 | 10/06/2025 |  | 01/30/2023 |  |  | 01/02/2023 | 01/30/2023 |  |  | 04/30/2024 |  | 01/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPOR BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 2023217052 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 01/02/2023 |  |  |  |  | The medical center placed an impella 5.5 pump to support a younger patient with multiple cardiac and systemic comorbidities and microvascular disease. the upper limb became ischemic due to the axillary placement of the impella 5.5. the team explanted the pump due to the ischemia after 28 hours of support. | The patient was heparin-induced thrombocytopenia (hit) positive. there was no heparin in purge or prime and bival was running systemically. | The us complainant placed a 5.5 pump to support a younger patient with multiple cardiac and systemic comorbidities and microvascular disease. the 5.5 would not pass. switched to cp implant in r axillary. pt is presumed hit positive. no heparin in purge or prime. bival running systemically. the upper limb became ischemic due to the axillary placement. the team explanted the pump due to the ischemia after 28 hours of support. |  |  |  | The medical center discarded the impella product at the time of pump explant. no benchtop analysis is possible. should more clinical details or imaging be shared the lead to root cause of the ischemia, a final report will be filed. the failure mode will be monitored and trended. | Corrected data in section d1, d4, and h4, original report filed in error against the impella 5.5 when the ischemia was noted after insertion of the impella cp. | Additional information for the initial mfr 1220648-2023-01814 was provided in sections b2, b5, h6 and h10. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. after reviewing the clinical information, it was determined that the cause of the limb ischemia was patient condition related, specifically the patient's microarterial disease.
35942 | 1220648-2023-01815 | 16265952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16265952&pc=OZD | 08/01/2022 | 01/30/2023 |  | 01/30/2023 |  |  | 01/16/2023 | 01/30/2023 |  |  |  |  | 01/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 27 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [5 empty] | V-A ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/16/2023 |  |  |  |  | The complainant reported a 27-year-old male patient presenting for lymphocytic fulminant myocarditis. during insertion of the impella for ecpella support, the patient¿s aortic valve was damaged. the patient's resulting aortic insufficiency was considered moderate and an aortic valve replacement was subsequently performed. additional information pertaining to the event was not provided. it should be noted that the event date is not exactly known, therefore, we estimated the event date as (b)(6) 2022 based on the context of the publication. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35943 | 1220648-2023-01816 | 16267186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16267186&pc=OZD | 12/31/2022 | 06/05/2023 |  | 01/30/2023 |  |  | 12/31/2022 | 01/30/2023 |  |  | 08/31/2024 |  | 01/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023270035 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/31/2022 |  |  |  |  | The complainant reported a 92 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient exhibited limb ischemia in the impella inserted leg. an external fem-fem bypass was configured, however, the impella was ultimately removed and replaced with an axillary impella 5.5. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35944 | 1220648-2023-01817 | 16267229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16267229&pc=OZD | 12/26/2022 | 06/05/2023 |  | 01/30/2023 |  |  | 12/26/2022 | 01/30/2023 |  |  | 06/30/2024 |  | 01/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023247621 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/26/2022 |  |  |  |  | The complainant reported a 69 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the developed ischemia in the impella inserted leg. the surgical team was activated and a fem-fem bypass was configured. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. limb ischemia ifu quote: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35945 | 1220648-2023-01818 | 16267412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16267412&pc=OZD | 01/01/2023 | 06/06/2023 |  | 01/30/2023 |  |  | 01/01/2023 | 01/30/2023 |  |  | 08/31/2024 |  | 01/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023268806 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/01/2023 |  |  |  |  | The complainant reported a 72 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. minimal perfusion was noted in the impella inserted let, post insertion of repositioning sheath. the decision was made to remove the device. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure mode will be monitored and trended. limb ischemia ifu quotes: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
35946 | 1220648-2023-01819 | 16269621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16269621&pc=OZD | 11/22/2022 | 01/30/2023 |  | 01/30/2023 |  |  | 11/22/2022 | 01/30/2023 |  |  | 01/31/2024 |  | 01/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023185200 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/22/2022 |  |  |  |  | The complainant reported a 68-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. upon removal of the introducer during impella support, the patient began to experience blood loss. it was estimated that the patient lost 500cc of blood as a result of the event. the patient was provided two units of prbc and 250cc albumin to counteract the blood loss. patient support continued following the treatment. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35947 | 1220648-2023-01809 | 16252989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16252989&pc=OZD | 12/30/2022 | 05/25/2023 |  | 01/27/2023 |  |  | 12/30/2022 | 01/27/2023 |  |  | 06/30/2024 |  | 01/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023270518 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 76-year-old male patient with known cad has been treated for ami / cgs. hemodynamic support has been provided by an impella cp heart pump. at the impella access site a hematoma formed. contributing factor has been patient¿s obesity. manual pressure successfully reduced the hematoma and ct scan showed no damage to vasculature. two units of red blood cells have been given during the procedure. | [5 empty] | Based on the clinical account, the root cause of the access site bleeding was most likely a use issue, as the staff is unsure if angle matching was used due to the patient's high bmi. | The impella device has not been received by the manufacturer for investigation. should any new information be returned, a supplemental mdr will be filed.
35948 | 1220648-2023-01810 | 16253190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16253190&pc=OZD | 12/30/2022 | 01/27/2023 |  | 01/27/2023 |  |  | 12/30/2022 | 01/27/2023 |  |  | 11/30/2024 |  | 01/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023282209 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/30/2022 |  |  |  |  | The medical center had placed an impella 5.5 via the axillary insertion site after escalating care from an impella cp placed via the femoral artery. the patient had suffered a large myocardial infarction and was in need of hemodynamic support. the patient had extensive blood loss from the axillary site. the team consulted with vascular surgery and gave 8 units of red blood cells and 2 units of ffp. the blood loss did not prompt pump explant. the pump remained on for another 6 days until weaned and explanted. | [5 empty] | The medical center discarded the impella pump product at the time of explant. no investigation to root cause is possible upon the bench top. further clinical details have not been forthcoming. no determination of root cause is made. the failure mode will be monitored and trended. should new information be shared, that leads to cause, a final report will be filed.
35949 | 1220648-2023-01811 | 16258646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16258646&pc=OZD | 12/12/2022 | 01/27/2023 |  | 01/27/2023 |  |  | 12/12/2022 | 01/27/2023 |  |  | 07/31/2024 |  | 01/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023249874 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/12/2022 |  |  |  |  | The complainant reported a 66 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed hemolysis. ischemia was also noted in the device-inserted leg. the decision was made to remove the pump and replace with an axillary impella 5.5 device. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35950 | 1220648-2023-01812 | 16260442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16260442&pc=OZD | 12/30/2022 | 06/06/2023 |  | 01/27/2023 |  |  | 12/30/2022 | 01/27/2023 |  |  | 08/31/2024 |  | 01/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023269654 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 0 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2022 |  |  |  |  | The complainant reported a male patient of undisclosed age presented for impella support. after insertion, an aortic dissection was found. the pump was removed. no further information could be obtained. | [5 empty] | The investigation into the reported vascular damage has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the root cause for vascular damage remains undetermined since neither the product nor sufficient clinical information were provided. it was reported that there was no allegation against the impella device. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35951 | 1220648-2023-01803 | 16245663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16245663&pc=OZD | 12/28/2022 | 04/26/2023 |  | 01/26/2023 |  |  | 12/28/2022 | 01/26/2023 |  |  | 07/31/2024 |  | 01/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023237737 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Female |  |  |  | Malfunction | Yes | Yes | Entrapment of Device | Separation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A female patient of unknown age with known cad was treated for acute myocardial infarction and cardiogenic shock. for hemodynamic support an impella cp cardiac pump has been inserted. attending physician stated that that it was unusually hard to peel the 14-fr peel-away introducer sheath. while peeling with extended force the impella cp heart pump was pushed deep into the ventricle and subsequently got stuck there. the pump was surgically removed. the patient was then stable and received veno-arterial extracorporeal membrane oxygenation (va-ecmo) support. | [5 empty] | Benchtop analysis of returned device is ongoing. a supplemental report will be filed after the investigation is complete. | Field echo imaging revealed that the cannula had folded on itself as the pump was too deep in the ventricle. a field image of the aic confirmed the occurrence of the "impella position in ventricle" alarm. the pump was not returned. no damage or abnormalities were found on the returned introducer (which was already peeled away upon return), including no evidence of peeling issues or scoreline malfunctions. therefore, it was determined that the cause of the introducer issue was difficulty peeling away the sheath due to improper peeling technique. after reviewing the clinical information, it was determined that the cause of the removal issue was most likely use related, stemming from the physician inadvertently pushing the pump deeper into the ventricle during introducer removal.
35952 | 1220648-2023-01804 | 16245747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16245747&pc=OZD | 12/29/2022 | 04/25/2023 |  | 01/26/2023 |  |  | 12/29/2022 | 01/26/2023 |  |  | 09/30/2024 |  | 01/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023277510 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 56-year-old male patient has been treated for cardiomyopathy / cardiogenic shock and received hemodynamic support from an impella cp heart pump. hemolysis has been detected by red urinary output and supporting lab results. replacement of the impella cp for an impella 5.5 resolved the issue. patient recovered and required no other intervention. patient stable afterwards and on bridge to durable lvad. | [5 empty] | The cause of the issue was not determined since product was not returned. | The impella device has been discarded and investigation is not completed yet. a supplemental will be filed upon conclusion of our investigation.
35953 | 1220648-2023-01805 | 16246981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16246981&pc=OZD | 12/27/2022 | 06/06/2023 |  | 01/26/2023 |  |  | 12/27/2022 | 01/26/2023 |  |  | 09/30/2024 |  | 01/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023274415 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2022 |  |  |  |  | The complainant reported a 67 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. a thrombus formed distal to where the impella was inserted, post impella removal. a balloon intervention. a pulse check was successful, thereafter. | [5 empty] | The investigation into the reported thrombus has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the thrombus. there was no report of the device performing out of specification during support, therefore it is likely the blood clot was picked up on explant. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35954 | 1220648-2023-01806 | 16247517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16247517&pc=OZD | 12/31/2022 | 04/19/2023 |  | 01/26/2023 |  |  | 12/31/2022 | 01/26/2023 |  |  | 03/31/2024 |  | 01/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023212661 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/31/2022 |  |  |  |  | The medical center had an impella cp placed for support of a patient with multiple spontaneous coronary artery dissections, when there was a large hematoma at the impella access site. the hematoma was treated by the advancement of the reposheath and when this did not resolve the hematoma, the team infused blood products and placed pressure with a femostop. additionally vascular surgery was called to remedy the superficial femoral artery damage and achieve hemostasis. the team is unsure if the impella accessed artery dissection was due the same cause as the spontaneous coronary artery dissections also. underlying patient condition may have caused the issue, or the placement of the impella. | [5 empty] | Since the original report was filed the investigation has concluded. the root cause of the vascular damage was found to be user related. the clinical review revealed the introducer was peeled away inside the patient's body/vasculature. the failure mode will be monitored and trended. of note the ifu states the following: "potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ | The medical center discarded all impella product at the time of explant. no benchtop analysis to root cause is possible. further clinical details have been asked, but not yet answered. the manufacturer will file a final report if a definite root cause is determined. the failure mode will be monitored and trended.
35955 | 1220648-2023-01807 | 16247603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16247603&pc=OZD | 12/27/2022 | 06/05/2023 |  | 01/26/2023 |  |  | 12/27/2022 | 01/26/2023 |  |  | 08/31/2024 |  | 01/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023267185 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Retroperitoneal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2022 |  |  |  |  | The complainant reported a 74 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. a drop in hemoglobin and hematocrit was noted, however, no visual impella access site bleed or hematoma was identified. the physician suspected a possible retroperitoneal bleed had occurred, however, a ct scan was never performed to confirm. a blood transfusion was completed. | [5 empty] | The investigation into the reported patient injury has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that patient condition was the most likely cause of the drop in the patient's hematocrit/hemoglobin levels. there was no bleeding observed from the impella access site. the failure modes will be monitored and trended. | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35956 | 1220648-2023-01808 | 16247731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16247731&pc=OZD | 12/27/2022 | 01/26/2023 |  | 01/26/2023 |  |  | 12/28/2022 | 01/26/2023 |  |  | 07/31/2024 |  | 01/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023246657 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/28/2022 |  |  |  |  | The complainant reported a 62 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, a cerebral infarct occurred that lead to a cerebral hemorrhage. the patient ultimately expired. the physician alleged the impella device was a factor in the patient's death. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35957 | 1220648-2023-01798 | 16240131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16240131&pc=OZD | 01/10/2023 | 01/25/2023 |  | 01/25/2023 |  |  | 01/11/2023 | 01/25/2023 |  |  | 08/31/2024 |  | 01/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC |  |  | US |  | 2023262865 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/11/2023 |  |  |  |  | The complainant reported a 44-year-old female patient presenting for cardiomyopathy. during impella support, it was noted that the patient developed an access site bleed. a takoseal adhesive sheet was utilized to stop the bleed after the administration of protamine failed to improve the condition. the patient was subsequently provided an unknown quantity of blood to counteract the blood loss. upon signs of cardiac recovery, the decision was made to remove the pump due to the bleed. the patient was deemed stable following the events. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35958 | 1220648-2023-01799 | 16240855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16240855&pc=OZD | 12/24/2022 | 01/25/2023 |  | 01/25/2023 |  |  | 12/26/2022 | 01/25/2023 |  |  | 08/31/2024 |  | 01/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023268451 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/26/2022 |  |  |  |  | The complainant reported an 83-year-old female patient presenting for cardiomyopathy. during impella support, it was noted that the patient developed an access site bleed. the patient was provided three units of prbc to counteract the blood loss. it was noted that the impella and dressing had been taped flat to the groin at the time of the bleed. to achieve hemostasis, this was corrected with an angle match dressing. patient support continued following the correction. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35959 | 1220648-2023-01800 | 16241301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16241301&pc=OZD | 12/26/2022 | 01/25/2023 |  | 01/25/2023 |  |  | 12/26/2022 | 01/25/2023 |  |  | 04/30/2024 |  | 01/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023223263 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/26/2022 |  |  |  |  | The complainant reported a 60-year-old male patient presenting for impella support. during support, it was noted that the patient developed hemolysis. the pump was explanted as a result of the condition and the patient was subsequently treated with iv fluids. the patient's condition was reported to be improving upon discharge. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35960 | 1220648-2023-01801 | 16241891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16241891&pc=OZD | 12/29/2022 | 01/25/2023 |  | 01/25/2023 |  |  | 12/29/2022 | 01/25/2023 |  |  | 09/30/2024 |  | 01/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023273591 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/29/2022 |  |  |  |  | The medical center had an impella cp support that was successful for 4 days when explanted. the explant revealed clot burden which prompted the team to perform penumbra thrombectomy to the impella limb. of note, the patient was off of appropriate anticoagulation for the postoperative time after the bypass surgery (cabg). | [5 empty] | The medical center discarded all product at the time of explant. there can be no benchtop analysis to root cause. the manufacturer will file another report should more clinical details, imaging, etc be shared that leads to a root cause. failure mode will be monitored and trended.
35961 | 1220648-2023-01802 | 16242323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16242323&pc=OZD | 12/28/2022 | 01/25/2023 |  | 01/25/2023 |  |  | 12/28/2022 | 01/25/2023 |  |  | 09/30/2024 |  | 01/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023275604 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/28/2022 |  |  |  |  | The complainant reported an 86-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. following explant of an impella cp pump, it was documented that the patient required a right femoral cutdown and femoral thromboendarterectomy due to ischemia in the right leg caused by a thrombus. the patient was considered stable following the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35962 | 1220648-2023-01789 | 16217937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16217937&pc=OZD | 12/21/2022 | 06/15/2023 |  | 01/20/2023 |  |  | 12/21/2022 | 01/20/2023 |  |  | 07/31/2024 |  | 01/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023252345 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2022 |  |  |  |  | The complainant reported a 48 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. after removal of the device, a clot traveled to the brachial artery. a vascular consult was called and the patient required a cutdown to extract the clot. | [5 empty] | The investigation into the thrombus has been completed since the original report was filed. the medical center discarded the impella product at the time of explant and care escalation. no benchtop analysis of the root cause of the thrombus was possible; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of the thrombus was most likely the patient condition. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35963 | 1220648-2023-01790 | 16218092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16218092&pc=OZD | 12/22/2022 | 08/17/2023 |  | 01/20/2023 |  |  | 12/22/2022 | 01/20/2023 |  |  | 12/31/2023 |  | 01/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023172114 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2022 |  |  |  |  | The complainant reported a 23 year old female patient presenting with cardiomyopathy for impella support. during support, the patient endured limb ischemia that resulted in premature removal of the device. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible to determine the root cause; only the clinical report was evaluated. based on the clinical information provided, the root cause of the limb ischemia was patient condition related. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35964 | 1220648-2023-01791 | 16218314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16218314&pc=OZD | 12/22/2022 | 06/06/2023 |  | 01/20/2023 |  |  | 12/22/2022 | 01/20/2023 |  |  | 02/29/2024 |  | 01/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023198115 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/22/2022 |  |  |  |  | The complainant reported a 23 year old female patient presenting with cardiomyopathy for impella support. during support, the patient endured limb ischemia that resulted in premature removal of the device. | [5 empty] | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that small femoral arteries patient condition was the cause of the limb ischemia. the failure mode will be monitored and trended. limb ischemia ifu quotes: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35965 | 1220648-2023-01792 | 16218610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16218610&pc=OZD | 12/23/2022 | 06/02/2023 |  | 01/20/2023 |  |  | 12/23/2022 | 01/20/2023 |  |  | 07/31/2024 |  | 01/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023259194 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/23/2022 |  |  |  |  | The complainant reported an 85 year old male patient presenting for high risk percutaneous coronary intervention for impella support. during an attempted repositioning, the cannula became kinked and required vascular intervention to remove. | [5 empty] | The investigation into the reported positioning issues has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that wireless re-access was the cause of the positioning issues. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35966 | 1220648-2023-01793 | 16219485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16219485&pc=OZD | 12/23/2022 | 01/20/2023 |  | 01/20/2023 |  |  | 12/23/2022 | 01/20/2023 |  |  | 09/30/2024 |  | 01/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023276700 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/23/2022 |  |  |  |  | The complainant reported an 81-year-old male patient presenting acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient developed hemolysis and the pump was explanted the following day. the patient was deemed stable following the events. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35967 | 1220648-2023-01794 | 16219488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16219488&pc=OZD | 12/21/2022 | 01/20/2023 |  | 01/20/2023 |  |  | 12/21/2022 | 01/20/2023 |  |  | 06/30/2024 |  | 01/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023256552 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/21/2022 |  |  |  |  | The complainant reported a 71 year old male patient presenting for high risk percutaneous coronary intervention using the impella device for hemodynamic support. after insertion, the patient developed ischemia in the impella inserted limb. the 14fr introducer was left in the access site and the patient was known to have poor blood vessels prior to impella use. as treatment, an antegrade sheath was configured. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35968 | 1220648-2023-01795 | 16219493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16219493&pc=OZD | 12/23/2022 | 03/25/2024 |  | 01/20/2023 |  |  | 12/23/2022 | 01/20/2023 |  |  | 06/30/2024 |  | 01/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023229399 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 68 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/23/2022 |  |  |  |  | The complainant reported a 68 year old male patient presenting with post cardiotomy cardiogenic shock for impella support. during an attempted insertion of the device, the physician perforated the patient's left ventricle. the patient ultimately expired. | [5 empty] | The investigation into the reported ventricle perforation has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that improper positioning was the most likely cause of the ventricle perforation. the failure modes will be monitored and trended. ventricle perforation ifu quotes: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The checkbox for death was inadvertently omitted in b2 despite a date of death being provided in the initial mdr. furthermore, h1 was inadvertently marked as a serious injury in the follow up, despite death being selected in the initial. both fields have been updated to coincide with the reportable death being addressed in this case. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35969 | 1220648-2023-01796 | 16219494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16219494&pc=OZD | 06/13/2022 | 01/20/2023 |  | 01/20/2023 |  |  | 12/23/2022 | 01/20/2023 |  |  | 11/30/2023 |  | 01/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023160852 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/23/2022 |  |  |  |  | The complainant reported a 77 year old male patient presenting with post cardiotomy cardiogenic shock for impella support. new information forwarded via jpvad registry on 2022-12-23 indicates the patient endured thrombocytopenia prompting platelet delivery. the patient also endured an access site bleed prompting a blood transfusion. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35970 | 1220648-2023-01797 | 16219495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16219495&pc=OZD | 12/23/2022 | 01/20/2023 |  | 01/20/2023 |  |  | 12/23/2022 | 01/20/2023 |  |  | 02/29/2024 |  | 01/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023195897 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/23/2022 |  |  |  |  | The complainant reported a 68-year-old male patient presenting acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient experienced runs of ventricular tachycardia. as a result of the noted condition, the decision was made to explant the impella pump. an intra-aortic balloon pump was subsequently inserted to continue support. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35971 | 1220648-2023-01788 | 16212315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16212315&pc=OZD | 12/19/2022 | 01/19/2023 |  | 01/19/2023 |  |  | 12/20/2022 | 01/19/2023 |  |  | 06/30/2024 |  | 01/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023231730 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/20/2022 |  |  |  |  | The complainant reported an 87 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. after removal of the device, the patient endured ischemia. a thrombus was surgically removed from the impella inserted leg, as treatment. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35972 | 1220648-2023-01782 | 16195999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16195999&pc=OZD | 01/01/2023 | 01/18/2023 |  | 01/18/2023 |  |  | 01/01/2023 | 01/18/2023 |  |  | 09/30/2024 |  | 01/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023279907 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/01/2023 |  |  |  |  | The complainant reported a 61-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient experienced a right femoral hemorrhage. as a result of the bleed, the patient was administered 10 units of blood. patient support continued throughout the event. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35973 | 1220648-2023-01783 | 16196687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16196687&pc=OZD | 12/29/2022 | 07/26/2023 |  | 01/18/2023 |  |  | 12/29/2022 | 01/18/2023 |  |  | 10/31/2024 |  | 01/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC |  |  | US |  | 2023271972 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Difficult or Delayed Positioning | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/29/2022 |  |  |  |  | The complainant reported a 75-year-old female patient presenting for impella support. following insertion, it was noted that the physician had difficulty sliding the blue hub into the end of the graft. as a result of the complication, the patient experienced blood loss which prompted the transfusion of two units of blood. hemostasis was eventually achieved, and patient support continued with the implanted pump. | [5 empty] | The investigation into the access site bleeding has been completed since the original report was filed. the user facility has returned the impella device, and the benchtop analysis of the root cause revealed that the cause of the access site bleeding was use related, specifically improper technique when using the repositioning sheath. the failure mode will be monitored and trended. | The impella device was returned and is pending evaluation. upon completion of the investigation, a supplemental mdr will be filed with the result of the evaluation.
35974 | 1220648-2023-01785 | 16200331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16200331&pc=OZD | 12/27/2022 | 01/18/2023 |  | 01/18/2023 |  |  | 12/27/2022 | 01/18/2023 |  |  | 08/31/2024 |  | 01/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023267036 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/27/2022 |  |  |  |  | The complainant reported a 87-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient experienced an access site bleed coinciding with a hematoma. manual pressure was applied to the site, however, no pulse was noted in the rfa and the lower limb was cool to the touch. the impella was subsequently removed to resolve the noted issues. the patient was reportedly stable following pump explant. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35975 | 1220648-2023-01784 | 16200333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16200333&pc=OZD | 12/20/2022 | 09/21/2024 |  | 01/18/2023 |  |  | 12/20/2022 | 01/18/2023 |  |  | 08/31/2024 |  | 01/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023270279 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | A 57-year-old male with a known cad has been treated for ami/cgs. after the procedure the patient continued to have oxygenation issues and suboptimal end organ perfusion, so a va-ecmo has been placed. after ecmo has been placed, leg ischemia below the impella access site has been detected. bypass to sfa has successfully been placed to re-perfuse the lower extremity. | Additional information noted that the family decided to withdraw care, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received by the manufacturer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
35976 | 1220648-2023-01786 | 16200334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16200334&pc=OZD | 12/29/2022 | 05/18/2023 |  | 01/18/2023 |  |  | 12/29/2022 | 01/18/2023 |  |  | 06/30/2024 |  | 01/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023235002 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/29/2022 |  |  |  |  | The complainant reported a 56-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. in the course of impella support, it was noted that the patient developed hemolysis. as the pump preload and position were assessed to be correct, the decision was made to remove the device. however, upon removal, the distal pulses were found to be absent and no flow was observed in the arterial duplex. an occlusive thrombus was found in the right femoral artery and vascular surgery was performed to alleviate the condition. | [5 empty] | Upon conclusion of the investigation, it was determined that the ischemia was related to the thrombus present in the right femoral artery. the origin and cause for the thrombus could not be determined and a definitive cause for the hemolysis could not be established. should additional information become available, a supplemental report will be filed. | Following the initial report, the pump was returned for further evaluation. a visual inspection was performed and no abnormalities were noted that could have caused or contributed to the reported issue. the pump was found to be fully function and a definitive cause for the reported allegation could not be established. | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35977 | 1220648-2023-01787 | 16200335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16200335&pc=OZD | 01/05/2023 | 10/20/2025 |  | 01/18/2023 |  |  | 01/05/2023 | 01/18/2023 |  |  | 06/30/2023 |  | 01/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC |  |  | US |  | 2022136430 |  | 005048-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/05/2023 |  |  |  |  | The complainant reported a 79-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient experienced an access site bleed resulting in a hematoma. due to the condition, the patient was administered an unknown quantity of blood products and the pump was subsequently removed. | The patient also experienced thrombocytopenia, however, a direct correlation to the reported bleed could not be confirmed. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information for the initial 1220648-2023-01787 was provided in sections b5, e1, and g1. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g1: updated manf. email address.
35978 | 1220648-2023-01781 | 16185737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16185737&pc=OZD | 12/21/2022 | 01/17/2023 |  | 01/17/2023 |  |  | 12/21/2022 | 01/17/2023 |  |  | 07/31/2024 |  | 01/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023249620 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/21/2022 |  |  |  |  | The medical team had an impella cp support ongoing for 1 week for a patient admitted in heart failure. the impella was inserted via the right femoral artery. during the support the right leg was watched and evaluated due to difficulty finding distal pulses and limb pain. at the time of explant vascular surgery sutured the site, performed thrombectomy to the iliac and placed a drain. at that time pulses were again present. | [5 empty] | The medical center discarded the pump at the time of explant. no benchtop analysis is possible to root cause of the ischemia. no imaging was shared of the native anatomy. should more information be shared, the leads to the definitive root cause, a final report will be filed. the failure mode will be monitored and trended.
35979 | 1220648-2023-01776 | 16164178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16164178&pc=OZD | 12/14/2022 | 06/02/2023 |  | 01/13/2023 |  |  | 12/14/2022 | 01/13/2023 |  |  | 09/30/2024 |  | 01/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023273026 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2022 |  |  |  |  | The complainant reported a 79 year old male presenting with acute myocardial infarction and cardiogenic shock for impella support. the patient developed limb ischemia post explant. a clot was noted within the patient's vasculature that required intervention to remove. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that biomaterial was the cause of the limb ischemia. the failure modes will be monitored and trended. limb ischemia ifu quotes: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35980 | 1220648-2023-01777 | 16164270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16164270&pc=OZD | 12/14/2022 | 06/27/2023 |  | 01/13/2023 |  |  | 12/14/2022 | 01/13/2023 |  |  | 08/31/2024 |  | 01/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023267914 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2022 |  |  |  |  | The complainant reported a 63 year old female patient developing an access site bleed and monophasic pulses in the inserted limb. the pump was removed prematurely and multiple units of blood products were required. hemodynamic support continued via an intra-aortic balloon pump. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the access site bleeding has been completed since the original report was filed. the medical center did not return the impella product for benchtop analysis of the access site bleeding; only the clinical report was evaluated. the clinical evaluation revealed that the root cause of access site bleeding was most likely related to patient condition, stemming from the significant swelling in the leg. the failure mode will be monitored and trended.
35981 | 1220648-2023-01778 | 16164478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16164478&pc=OZD | 12/15/2022 | 06/02/2023 |  | 01/13/2023 |  |  | 12/15/2022 | 01/13/2023 |  |  | 07/31/2024 |  | 01/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023257291 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2022 |  |  |  |  | The complainant reported a 55 year old male patient presenting with cardiomyopathy for impella support. during support, the patient exhibited signs of hemolysis (red urine) that prompted premature removal of the device. an iabp was placed to continue support. | [5 empty] | The investigation into the reported hemolysis has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that improper positioning was the most likely cause of the hemolysis. the failure modes will be monitored and trended. ventricle perforation ifu quotes: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35982 | 1220648-2023-01779 | 16164905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16164905&pc=OZD | 12/20/2022 | 01/13/2023 |  | 01/13/2023 |  |  | 12/20/2022 | 01/13/2023 |  |  | 06/30/2024 |  | 01/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023247258 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/20/2022 |  |  |  |  | The medical team had an imeplla cp inserted along with extracorporeal membrane oxygenation (ecmo) after escalation from an balloon pump (iabp) for a surgical patient admitted for coronary artery bypass. the team explanted the impella after 1 day due to bilateral limb ischemia. the impella was explanted and ecmo cannulas moved to axillary access instead. | [5 empty] | The medical center discarded the impella pump product at the time of explant. there can be no benchtop analysis to root cause for the reported limb ischemia. the manufacturer will close the investigation, though should new information be shared that leads to a definative root cause, a final report will be filed. the failure mode will be monitored and trended.
35983 | 1220648-2023-01780 | 16165064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16165064&pc=OZD | 12/23/2022 | 01/13/2023 |  | 01/13/2023 |  |  | 12/23/2022 | 01/13/2023 |  |  | 10/31/2024 |  | 01/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023274365 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/23/2022 |  |  |  |  | The complainant reported a 56-year-old male patient presenting for pre-op pump placement. during impella support, it was noted that the pump had flipped in the left ventricle prompting the need for device repositioning. the medical director torqued the pump, however, the position remained unchanged. the pump was subsequently pulled back to relieve any stored energy and the pump came across the ao valve. upon attempting to re-cross the valve, the pump buckled and became wedged. a thoracotomy was then needed to remove the pump. the patient was deemed to be stable following the event. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35984 | 1220648-2023-01774 | 16149988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16149988&pc=OZD | 12/13/2022 | 01/12/2023 |  | 01/12/2023 |  |  | 12/13/2022 | 01/12/2023 |  |  | 03/31/2024 |  | 01/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023204590 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/13/2022 |  |  |  |  | The complainant reported an 84 year old female presenting for high risk percutaneous coronary intervention. the impella device was inserted for hemodynamic support. during support, an access site bleed developed along with a drop in the patient's hemoglobin. a surgical cutdown was performed to remove the device and 2 units of blood products were delivered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35985 | 1220648-2023-01775 | 16152509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16152509&pc=OZD | 12/02/2022 | 01/12/2023 |  | 01/12/2023 |  |  | 12/02/2022 | 01/12/2023 |  |  | 04/30/2024 |  | 01/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023225824 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/02/2022 |  |  |  |  | The complainant reported a 67 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. after the procedure, the physician was unable to doppler pedal pulses in the impella inserted limb. the decision was made to wean and explant the device. pulses returned after explant. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35986 | 1220648-2023-01772 | 16141277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16141277&pc=OZD | 12/12/2022 | 01/11/2023 |  | 01/11/2023 |  |  | 12/12/2022 | 01/11/2023 |  |  | 02/29/2024 |  | 01/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP SET JP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/12/2022 |  |  |  |  | The complainant reported a 78 year old male patient presenting with cardiomyopathy for impella support. during support, a bleed formed around the anastomosis/impella insertion site. a blood transfusion was required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35987 | 1220648-2023-01773 | 16141355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16141355&pc=OZD | 12/12/2022 | 06/21/2023 |  | 01/11/2023 |  |  | 12/12/2022 | 01/11/2023 |  |  | 07/31/2024 |  | 01/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023257978 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2022 |  |  |  |  | The complainant reported a 65 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient had a loss of pulses/discoloration in the impella inserted leg. a distal perfusion catheter was applied and the issue resolved. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return the impella product for manufacturer investigation. no benchtop analysis of the root cause of the limb ischemia was possible; only the clinical report was evaluated. based on the clinical data, the cause of the limb ischemia was most likely related to the patient condition. the failure mode will be monitored and trended. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35988 | 1220648-2023-01770 | 16131751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16131751&pc=OZD | 02/02/2022 | 01/10/2023 |  | 01/10/2023 |  |  | 12/27/2022 | 01/10/2023 |  |  | 10/31/2023 |  | 01/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABOIOMED, INC |  |  | US |  | 2022147667 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/27/2022 |  |  |  |  | The complainant reported an 83-year-old male patient presenting for high risk percutaneous coronary intervention. during impella support, it was documented that the patient experienced an access site bleed. a blood transfusion of unknown quantity was provided to counteract the bleed and the pump was subsequently removed. hemostasis was obtained at the site via perclose and manual pressure. the patient was deemed stable following the event. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35989 | 1220648-2023-01771 | 16134743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16134743&pc=OZD | 11/11/2022 | 01/10/2023 |  | 01/10/2023 |  |  | 12/14/2022 | 01/10/2023 |  |  | 06/30/2024 |  | 01/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023235866 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/14/2022 |  |  |  |  | The complainant reported a 34-year-old male patient presenting for cardiomyopathy. during impella support, it was documented that the patient experienced an access site bleed. manual pressure was applied and an additional suture was placed to treat the bleed. ultimately, four units of prbc had to be provided to the patient to counteract the blood loss. patient support continued throughout the event. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
35990 | 1220648-2023-01769 | 16126211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16126211&pc=OZD | 12/10/2022 | 01/09/2023 |  | 01/09/2023 |  |  | 12/10/2022 | 01/09/2023 |  |  | 08/31/2024 |  | 01/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023266837 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/10/2022 |  |  |  |  | The complainant reported an 82 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed limb ischemia in the impella inserted limb. as a result, the impella was removed. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35991 | 1220648-2023-01767 | 16117382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16117382&pc=OZD | 12/08/2022 | 01/06/2023 |  | 01/06/2023 |  |  | 12/08/2022 | 01/06/2023 |  |  | 08/31/2024 |  | 01/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023272227 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/08/2022 |  |  |  |  | The complainant reported a 72 year old male patient presenting for impella support during an ep/ablation procedure. after the procedure, the patient developed a hematoma at the impella access site. hemostasis was achieved, however, multiple blood products were delivered. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35992 | 1220648-2023-01768 | 16117390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16117390&pc=OZD | 12/08/2022 | 01/06/2023 |  | 01/06/2023 |  |  | 12/08/2022 | 01/06/2023 |  |  | 07/31/2024 |  | 01/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023256282 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/08/2022 |  |  |  |  | The complainant reported a 39 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, concern for hemolysis prompted removal of the device. the patient was then administered dialysis. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
35993 | 1220648-2023-01761 | 16103995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16103995&pc=OZD | 12/07/2022 | 10/13/2025 |  | 01/05/2023 |  |  | 12/07/2022 | 01/05/2023 |  |  | 08/31/2024 |  | 01/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023271076 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2022 |  |  |  |  | The medical center had an impella cp support ongoing for a patient, admitted with an acute myocardial infarct, in which the placment was suboptimal. the team attempted to reposition and kinked the pump. the team used a snare to readvance to the lv. the malfunction and necessary snaring caused no last harm or injury, though the patient did expire on support with the 2nd pump that replaced this pump. | The complainant reported that the impella pump placement was shallow. an attempt to reposition the device under transthoracic echocardiogram was ineffective. it was stated that the patient¿s left ventricle was small, and the pigtail consistently became trapped under the papillary muscle when advancing. it was reported that the catheter came out of the aortic valve leaving the impella pump in an arch. the medical professional tried to push the device back across the aortic valve however, a kink occurred above the outflow cage. the kink was stated to have remained until the device proceeded into the left ventricle. the initial impella catheter was removed, and a new device was placed. the procedural outcome was the patient survived explant. |  |  |  |  | The medical center discarded the impella pump product at the time of explant. no benchtop analysis is possible to determine the root cause of the pump positioning issues and subsequent snaring. the failure mode will be monitored and trended. | After reviewing the clinical information, it was determined that the cause of the positioning issues was wireless re-access. the impella device is not indicated for wireless re-access per ifu 0048-9007. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g1: corrected mfg. email address.
35994 | 1220648-2023-01762 | 16104100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16104100&pc=OZD | 12/14/2022 | 01/05/2023 |  | 01/05/2023 |  |  | 12/14/2022 | 01/05/2023 |  |  | 04/30/2024 |  | 01/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED INC. |  |  | US |  | 2023210066 |  | 0048-0014 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/14/2022 |  |  |  |  | The complainant in the netherlands had inserted an impella cp for support in a patient with known peripheral arterial disease (pad). the pad was known to be "very severe". the patient's age, race, and gender were not shared with the abiomed field representative. the team made the choice to explant the femoral cp due to the occlusion of the extremity and place an axillary 5.0 instead. the limb was becoming ischemic after just over 7 hours of support. | [5 empty] | The medical center in europe discarded the impella pump at the time of explant and replacement to the impella 5.0. the pump benchtop analysis is not possible. root cause determination is not possible. the native condition surely contributed to the ischemia. the failure mode will be monitored and trended.
35995 | 1220648-2023-01763 | 16104230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16104230&pc=OZD | 12/13/2022 | 01/05/2023 |  | 01/05/2023 |  |  | 12/13/2022 | 01/05/2023 |  |  | 04/30/2024 |  | 01/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023223708 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/13/2022 |  |  |  |  | The medical team had an impella cp inserted for hemodynamic support of a patient admitted with multiple cardiac comorbidities. the pump was inserted successfully and ran til the pump was replaced due to low flows after just over 2 hours of support. later the pump was run through benchtop testing at abiomed and an observation was made of damage to the impeller blade. | [5 empty] | The medical center returned the impella pump for analysis and during the analysis the impeller damage was noted. because the fluoroscopic image of the pump was not returned, a definitive root cause of the fracture could not be determined. the failure mode will be monitored and trended.
35996 | 1220648-2023-01764 | 16109889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16109889&pc=OZD | 08/03/2022 | 05/31/2023 |  | 01/05/2023 |  |  | 12/06/2022 | 01/05/2023 |  |  | 03/31/2024 |  | 01/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023207552 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia Stroke | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2022 |  |  |  |  | The complainant reported a patient presenting with acute myocardial infarction and cardiogenic shock for impella support. on (b)(6) 2022, an access site bleed developed which resulted in the need for blood products. on (b)(6) 2022, the patient developed a cerebral infarction. the patient ultimately expired on (b)(6) 2022. the cause of death was attributed to cerebral infarction. although there was no specific device malfunction reported, the impella could not be denied as a causal/contributory factor to the cerebral infarction. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the reported stroke/cardiovascular accident (cva), thrombocytopenia, and access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the cva and thrombocytopenia due to the patient's history of hemorrhages a month prior to impella insertion. the patient suffered an ischemic stroke the day after the impella was removed. the root cause for access site bleeding remains undetermined since neither the product nor sufficient clinical information were provided. the failure modes will be monitored and trended. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35997 | 1220648-2023-01765 | 16109938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16109938&pc=OZD | 12/06/2022 | 06/02/2023 |  | 01/05/2023 |  |  | 12/06/2022 | 01/05/2023 |  |  | 02/29/2024 |  | 01/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023195583 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2022 |  |  |  |  | The complainant reported a 69 year old male patient presenting for high risk percutaneous coronary intervention for impella support. the patient developed leg ischemia during support which resulted in premature removal of the impella device. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the reported limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the most likely cause of the limb ischemia. the failure modes will be monitored and trended. limb ischemia ifu quotes: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
35998 | 1220648-2023-01766 | 16110005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16110005&pc=OZD | 12/07/2022 | 09/19/2024 |  | 01/05/2023 |  |  | 12/07/2022 | 01/05/2023 |  |  | 04/30/2024 |  | 01/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023213828 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/07/2022 |  |  |  |  | The complainant reported an 82 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient endured limb ischemia. a reperfusion sheath was configured, however, the patient expired, thereafter. the placement of the reperfusion catheter was alleged as a factor in the patient's death. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the limb ischemia has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the clinical evaluation revealed that patient condition was the cause of the limb ischemia. the failure mode will be monitored and trended. limb ischemia ifu quotes: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-0176 was provided in sections b2, b5, h1, and h6.
35999 | 1220648-2023-01759 | 16094518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16094518&pc=OZD | 12/09/2022 | 01/04/2023 |  | 01/04/2023 |  |  | 12/09/2022 | 01/04/2023 |  |  | 01/31/2024 |  | 01/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023186090 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/09/2022 |  |  |  |  | The medical team had placed an impella 5.5 via direct access for a surgical patient with multiple cardiac comorbidities and congenital disease. the pump stopped after just over 7 hours. the pump was removed but not replaced by the medical team. the patient was supported by inotropes alone. | [5 empty] | The medical center has retained the impella pump for their own risk management assessment. when the pump is returned to abiomed and an analysis completed, with root cause determined, the final report will follow. should no product be returned, no final report will be submitted as root cause can not be determined. the failure mode will be monitored and trended.
36000 | 1220648-2023-01760 | 16095014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16095014&pc=OZD | 08/26/2022 | 01/04/2023 |  | 01/04/2023 |  |  | 12/21/2022 | 01/04/2023 |  |  | 03/31/2024 |  | 01/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023202177 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/21/2022 |  |  |  |  | The medical team in japan had a successful impella cp support in (b)(6) 2022 for 10 days for a patient of unknown age and gender. the jpvad registry team later in 12/2022 reported upon limb ischemia in the impella accessed limb that was treated by perfusion catheter via an antegrade stick. no further clinical details were shared. | [5 empty] | The medical team in japan discarded the impella product at the time of explant months prior. the jpvad registry furnished no further clinical details that lead to root cause of the ischemia. should more information be shared and a root cause be determined, a final report will follow. the failure mode will be monitored and trended.
36001 | 1220648-2023-01753 | 16085544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16085544&pc=OZD | 12/02/2022 | 09/20/2024 |  | 01/03/2023 |  |  | 12/05/2022 | 01/03/2023 |  |  | 07/31/2024 |  | 01/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023237669 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  |  | Injury | Yes | Yes | Detachment of Device or Device Component | Adverse Event Without Identified Device or Use Problem | Separation Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | A 68-year-old female patient was treated for acute myocardial infarction and cardiogenic shock. an impella cp heart pump was inserted for hemodynamic support. the pump was implanted during an ongoing cardiopulmonary resuscitation. during this procedure, the impella folded up in the aortic root and could not be straightened again. it was decided to remove the impella cp. when the pump was pulled out, resistance was noticed, and the pigtail remained in the patient. the treating physician suspected that the pigtail remained distal to the puncture site, as the new impella cp could be placed without issues. with hemodynamic support from the impella cp, the coronary vessel affected by in-stent thrombosis was reopened. after the successful procedure, the patient suffered from continuous pulseless electrical activity (pea). care has been withdrawn and the patient died. | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | Device has not been received for benchtop analysis yet. a supplemental report will be filed when the investigation is complete. | The returned product was visually inspected and the inflow cage was detached from the cannula. cannula coil unravelling was observed and damaged polyurethane. the inflow cage was not returned. no catheter kinks were observed. the cause of the cannula detachment was most likely use related since impella flipped in the aortic root during cpr and then resulted in resistance when passing back through the sheath during removal. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. note: corrected information provided in h6 is an addition to previous coding.
36002 | 1220648-2023-01755 | 16086803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16086803&pc=OZD | 12/05/2022 | 01/03/2023 |  | 01/03/2023 |  |  | 12/06/2022 | 01/03/2023 |  |  | 07/31/2024 |  | 01/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, CANADA |  |  | ABIOMED, INC |  |  | US |  | 2023256973 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/06/2022 |  |  |  |  | The complainant reported a 56-year-old female patient presenting for impella support. during insertion, it was noted that the patient had a very tortuous vasculature that created a tight pathway for the pump to navigate. once the impella cp was properly inserted, it was reported that the patient's left arm was cool to the touch and that both the brachial and radial pulses were faint. an existing condition from a prior procedure attributed to swelling in the vessel and further reduced the flow to the limb. patient support was maintained until the patient was deemed more stable, at which point the pump was removed due to the ischemia. upon removal, flow to the limb improved. further intervention was not required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36003 | 1220648-2023-01754 | 16086949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16086949&pc=OZD | 12/06/2022 | 01/03/2023 |  | 01/03/2023 |  |  | 12/06/2022 | 01/03/2023 |  |  | 08/31/2024 |  | 01/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023270070 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/06/2022 |  |  |  |  | The medical center had an impella cp support ongoing when there was an observed hematoma. the hematoma prompted treatment with covered stenting to the vasculature and blood products. the pump was explanted and the patient expired when family withdrew care. | [5 empty] | The medical center has discarded the impella pump product. no benchtop analysis to root cause is possible. the case will be closed, but should more information be shared that leads to a root cause, a final report will follow. the failure mode will be monitored and trended.
36004 | 1220648-2023-01756 | 16088249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16088249&pc=OZD | 12/04/2022 | 01/03/2023 |  | 01/03/2023 |  |  | 12/04/2022 | 01/03/2023 |  |  | 02/28/2024 |  | 01/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | CANADA 5.0 AIC |  |  | ABIOMED, INC |  |  | US |  | 2023193083 |  | 005064-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Inflammation | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/04/2022 |  |  |  |  | The complainant reported a 56-year-old female patient presenting for impella support. during insertion, it was noted that the patient had a tortuous vasculature that created a tight angle for the pump to navigate. multiple attempts were made to place the pump, however the patient¿s anatomy caused the delivery attempts to fail. as a result of the repeated attempts, the access site became swollen and the decision was made to place a different pump. upon the subsequent completion of support, the access site remained swollen and a surgical repair of the site/vessel was completed. the patient was administered two units of blood and deemed to be stable following support. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36005 | 1220648-2023-01757 | 16089397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16089397&pc=OZD | 10/20/2022 | 01/03/2023 |  | 01/03/2023 |  |  | 12/06/2022 | 01/03/2023 |  |  | 05/31/2024 |  | 01/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023222658 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Information from study documentation: a patient of unknown age and gender in the eu was treated for coronary artery disease with percutaneous coronary intervention (pci). hemodynamic support was provided by an impella cp heart pump. following the procedure, a pseudoaneurysm occurred at the access site and was treated with a pressure bandage and thrombin injection, prolonging hospitalization. the patient was stable afterwards. | [5 empty] | Device has been discarded. if new information emerges, a supplemental mdr will be filed.
36006 | 1220648-2023-01758 | 16092367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16092367&pc=OZD | 12/04/2022 | 01/03/2023 |  | 01/03/2023 |  |  | 12/04/2022 | 01/03/2023 |  |  | 09/30/2024 |  | 01/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023263521 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/04/2022 |  |  |  |  | The medical center had an impella 5.5 support ongoing when there were positioning issues for the pump. the malpositioning issues were observed by suction alarms being present. the team attempted to adjust the pump position but, in doing so the hemolysis persisted and the patient had crrt initiated. the team made the choice to prematurely explant the pump and place a new impella 5.5 to continue the hemodynamic support. when explanted there was visible biomaterial adhered to the pump. the pump had supported the patient admitted in cardiogenic shock for 7 days. | [5 empty] | The medical center discarded the impella pump at the time of explant. as such, no benchtop analysis and hemolysis testing can be performed. root cause can not be determined. further clinical details have been asked for, but to date no other information has been furnished. should new information come to manufacturer that leads to root cause determination of the hemolysis, a final report will follow. the failure mode will be monitored and trended.
36007 | 16077314 | 16077314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16077314&pc=OZD | 12/07/2022 | 12/08/2022 | 12/08/2022 | 12/30/2022 | 12/08/2022 | 12/30/2022 |  | 12/30/2022 |  |  |  |  | 12/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Nurse | [8 empty] | Unknown |  |  |  | Malfunction | No | Yes | Break | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Leaking medication, break in sterile tubing.
36008 | 1220648-2022-01748 | 16079312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16079312&pc=OZD | 12/04/2022 | 12/30/2022 |  | 12/30/2022 |  |  | 12/04/2022 | 12/30/2022 |  |  | 07/31/2024 |  | 12/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023259412 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/04/2022 |  |  |  |  | The complainant reported a 62 year old female presenting with acute myocardial infarction and cardiogenic shock for impella support. the pump was prematurely removed in response to loss of pedal pulse in the impella inserted limb. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36009 | 1220648-2022-01749 | 16079357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16079357&pc=OZD | 12/02/2022 | 12/30/2022 |  | 12/30/2022 |  |  | 12/02/2022 | 12/30/2022 |  |  | 07/31/2024 |  | 12/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023259246 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/02/2022 |  |  |  |  | The complainant reported a 72-year-old male patient presenting for high risk percutaneous coronary intervention. immediately following impella support, it was noted that the patient developed limb ischemia. it was described that there was no pulse in the impella leg after manual pressure had been applied to the site post removal. a thrombectomy was performed which resulted in the removal of a thrombus in the right leg and the placement of the stent. the patient was stable following the intervention. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36010 | 1220648-2022-01750 | 16079378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16079378&pc=OZD | 11/29/2022 | 10/25/2023 |  | 12/30/2022 |  |  | 12/02/2022 | 12/30/2022 |  |  | 03/31/2024 |  | 12/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023206975 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 15 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Material Puncture/Hole | Adverse Event Without Identified Device or Use Problem | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/02/2022 |  |  |  |  | The complainant reported a 15-year-old male patient presenting for cardiomyopathy. during insertion of the impella pump, it was noted that some stretching of the vessel occurred. as a result, a surgical repair was required to the small axillary artery. patient support was able to continue following the intervention. | Upon surgical repair of the previously reported vessel damage, the white cable was punctured and blood was subsequently seen leaking from the sidearm of the red plug. the puncture was sealed using bone wax and derma bond and patient support continued with the implanted pump. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | Updated sections b1 and h6 to correspond with the report of a hole in the catheter: h10: the investigation into the reported "vascular damage - peripheral vessel and product damage (hole in catheter)" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. after reviewing the clinical information, it was determined that the cause of the vascular damage was most likely patient condition related, specifically the interaction between the impella and the patient's small axillary artery. the cause of the red handle leak was most likely blood ingress due to a hole in the catheter resulting from improper use. this supplemental report is being filed as part of a retrospective review of historical records.
36011 | 1220648-2022-01751 | 16080729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16080729&pc=OZD | 12/02/2022 | 09/22/2024 |  | 12/30/2022 |  |  | 12/02/2022 | 12/30/2022 |  |  | 05/31/2023 |  | 12/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022091036 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/02/2022 |  |  |  |  | The complainant reported a 64 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, a device related hematoma was discovered. multiple blood products were delivered as a result. | It was reported a large late presenting hematoma was observed at the impella access site. the patient required a massive transfusion, and despite blood products the patient had worsening hemodynamics. an investigational ultrasound of the left groin revealed no issues. the patient continued to decompensate, the complication could not be identified, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01751 was provided in sections b2, b5, d6a, d6b, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
36012 | 1220648-2022-01752 | 16080845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16080845&pc=OZD | 12/03/2022 | 12/30/2022 |  | 12/30/2022 |  |  | 12/03/2022 | 12/30/2022 |  |  | 05/31/2024 |  | 12/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023228682 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/03/2022 |  |  |  |  | The complainant reported a 60 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. the pump was ultimately removed due to hemolysis. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36013 | 1220648-2022-01746 | 16067460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16067460&pc=OZD | 11/30/2022 | 12/29/2022 |  | 12/29/2022 |  |  | 11/30/2022 | 12/29/2022 |  |  | 08/31/2024 |  | 12/29/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023266097 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/30/2022 |  |  |  |  | The complainant reported a 72 year old male patient presenting with post cardiotomy cardiogenic shock for impella support. during insertion of the device, a ventricular perforation occurred. the injury was repaired and support proceeded. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36014 | 1220648-2022-01747 | 16073663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16073663&pc=OZD | 11/29/2022 | 12/29/2022 |  | 12/29/2022 |  |  | 11/29/2022 | 12/29/2022 |  |  | 07/31/2024 |  | 12/29/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023249534 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/29/2022 |  |  |  |  | The complainant reported a 74 year old female presenting with acute myocardial infarction and cardiogenic shock for impella support. the pump was weaned and removed successfully, however, the physician was unable to doppler pulses in the limb. a cutdown was performed, a clot was removed, and the artery was surgically repaired. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36015 | 1220648-2022-01744 | 16064196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16064196&pc=OZD | 11/29/2022 | 12/28/2022 |  | 12/28/2022 |  |  | 11/29/2022 | 12/28/2022 |  |  | 07/31/2022 |  | 12/28/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  |  |  | 2023231478 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/08/2022 |  |  |  |  | The complainant reported a 67 year old female presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient exhibited ischemia in the impella inserted limb. it was noted that the 14fr peel away introducer was left in place. an external bypass was attempted, but to no avail. the pump was weaned and removed and the issue resolved. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36016 | 1220648-2022-01745 | 16064697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16064697&pc=OZD | 12/17/2022 | 12/28/2022 |  | 12/28/2022 |  |  | 12/22/2022 | 12/28/2022 |  |  | 08/31/2024 |  | 12/28/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023256187 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | Company representation |  |  | 76 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/22/2022 |  |  |  |  | The complainant reported a 76-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was identified that there were suction alarms coinciding with the identification of a thrombus. it was noted that suction alarms had sounded during support and that lowering the p level failed to resolve the issue. flows were lower than expected and an echocardiogram showed a thrombus in the left ventricle. the patient's condition stabilized without intervention and the impella was explanted once the need for support was no longer required. although some biomaterial was noted on the pump following removal, no action was initially taken regarding the thrombus. however, a follow up report noted that the patient experienced a cerebral infarction upon removal of the pump. the stroke was considered embolic at the time of reporting. additional information pertaining to the stroke was not provided. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36017 | 1220648-2022-01741 | 16056721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16056721&pc=OZD | 11/27/2022 | 12/27/2022 |  | 12/27/2022 |  |  | 11/27/2022 | 12/27/2022 |  |  | 03/31/2024 |  | 12/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023209819 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/27/2022 |  |  |  |  | The complainant reported a 62 year old male patient presenting with post cardiotomy cardiogenic shock for impella support. during insertion of the device, the patient's ventricle was perforated. the pump was removed, the perforation was patched, and support continued after re-inserting the original pump. it was noted that the patient's tissue was severely damaged from his pre-existing myocardial infarction. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36018 | 1220648-2022-01742 | 16056955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16056955&pc=OZD | 11/29/2022 | 12/27/2022 |  | 12/27/2022 |  |  | 11/29/2022 | 12/27/2022 |  |  | 08/31/2024 |  | 12/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023266558 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | PHYSICIAN ASSISTANT | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/29/2022 |  |  |  |  | The complainant reported an 44-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient developed hemolysis. as a result of the condition, the cp pump was removed and the patient was escalated to an impella 5.5 pump. the condition was reported to have resolved following the escalation. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36019 | 1220648-2022-01743 | 16056969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16056969&pc=OZD | 10/12/2022 | 10/14/2024 |  | 12/27/2022 |  |  | 12/16/2022 | 12/27/2022 |  |  | 05/31/2024 |  | 12/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023219801 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 18 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2022 |  |  |  |  | The complainant reported an 18-year-old male patient presenting for cardiomyopathy. during impella support, it was noted that the patient experienced an access site bleed from both the impella and swan-ganz insertion sites. an estimated amount of blood loss from each site could not be provided. as a result of the bleed and coinciding thrombocytopenia, a platelet transfusion was performed involving 28 units of prbc and 8 units of ffp. a definitive cause for the thrombocytopenia could not be determined due to the bleeding from multiple sites and existing liver damage. the patient's condition improved following the transfusion. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
36020 | 1220648-2022-01736 | 16043163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16043163&pc=OZD | 11/21/2022 | 10/13/2023 |  | 12/23/2022 |  |  | 11/24/2022 | 12/23/2022 |  |  | 04/30/2024 |  | 12/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP SET JP |  |  | ABIOMED, INC. |  |  | US |  | 2023226873 |  | 000000000001000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | Aortic Valve Insufficiency/ Regurgitation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2022 |  |  |  |  | The complainant reported an 81 year old male patient presenting for impella support. the physician had difficulty removing the device and required surgical intervention via aortic incision for full device removal. residual thrombotic material in the patient's artery required intervention to resolve via a thrombectomy. the patient exhibited aortic regurgitation that prompted an aortic valve replacement. finally, a cerebral infarction had occurred. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the reported "heart valve damage, stroke/cva, thrombus and removal issues" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the data logs from the case were unremarkable. after reviewing the clinical information, it was determined that the cause of the removal issues was most likely the presence of biomaterial in the patient's vessels. the cause of the thrombus was patient condition related. the cause of the aortic valve damage and its relation to the impella were not determined since the patient had a worn out prosthetic aortic valve. the cause of the cva and its relation to the impella were not determined since it occurred after the thoracotomy procedure to remove the impella and replace the patient's valves. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
36021 | 1220648-2022-01737 | 16043221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16043221&pc=OZD | 05/07/2022 | 12/23/2022 |  | 12/23/2022 |  |  | 11/25/2022 | 12/23/2022 |  |  | 10/31/2023 |  | 12/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022147704 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/25/2022 |  |  |  |  | The complainant reported an impella patient suspected as enduring hemolysis during support. as treatment, haptoglobin was delivered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36022 | 1220648-2022-01738 | 16043258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16043258&pc=OZD | 11/25/2022 | 10/25/2023 |  | 12/23/2022 |  |  | 11/25/2022 | 12/23/2022 |  |  | 03/31/2024 |  | 12/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP SET JP |  |  | ABIOMED, INC. |  |  | US |  | 2023195070 |  | 000000000001000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/25/2022 |  |  |  |  | The complainant reported a 55 year old male patient presenting with acute myocardial infarction and cardiogenic shock. during support, the pump fell out of the ventricle and the patient went into cardiopulmonary arrest. a new impella 5.5 was inserted as treatment. the account noted there was no issue with the pump, rather, this was a "management problem." | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Updated sections g2, g3, g6, h2, h3 & h6: (h3) the investigation into the reported "positioning issues" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the cause of the positioning issue was most likely patient movement during rehabilitation. this report is being filed as part of a retrospective review of historical records section d4: unique identifier (udi) #.
36023 | 1220648-2022-01739 | 16043283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16043283&pc=OZD | 06/15/2022 | 12/23/2022 |  | 12/23/2022 |  |  | 11/25/2022 | 12/23/2022 |  |  | 12/31/2023 |  | 12/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023167849 |  | 0048-0024-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/25/2022 |  |  |  |  | The complainant suspected an impella patient enduring hemolysis during support. as treatment, haptoglobin was delivered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36024 | 1220648-2022-01740 | 16043308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16043308&pc=OZD | 11/27/2022 | 12/23/2022 |  | 12/23/2022 |  |  | 11/28/2022 | 12/23/2022 |  |  | 06/30/2024 |  | 12/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023231970 |  | 0048-0024-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/28/2022 |  |  |  |  | The complainant reported a 74 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed an access site bleed that prompted a blood transfusion. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36025 | 1220648-2022-01735 | 16033584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16033584&pc=OZD | 11/28/2022 | 12/06/2023 |  | 12/22/2022 |  |  | 11/28/2022 | 12/22/2022 |  |  | 06/30/2024 |  | 12/22/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023236935 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 30-year-old female (past medical history: obesity, gastric sleeve, cholecystectomy) in the us with peri-partum cardiomyopathy has been treated for cardiogenic shock. initially, an iabp has been used, which has been escalated to an impella cp heart pump. patient required escalation to an impella 5.5 heart pump. impella cp has been explanted per cfa cutdown. limb ischemia has been noted. two clots have been successfully removed with a fogarty balloon. limb perfusion confirmed via doppler ultrasonography afterwards. patient survived and has been sent to icu with impella 5.5 haemodynamic support. | [5 empty] | The impella device has been discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation for the reported the limb ischemia has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the limb ischemia was most likely biomaterial in the patient's vessel. the failure mode will be monitored and trended.
36026 | 1220648-2022-01732 | 16022305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16022305&pc=OZD | 11/20/2022 | 10/13/2025 |  | 12/21/2022 |  |  | 11/24/2022 | 12/21/2022 |  |  | 08/31/2024 |  | 12/21/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023257653 |  | 0048-0014 |  | No answer provided |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age and sex was to receive hemodynamic support from an impella cp heart pump due to severe cgs post cpr and ecmo implant. attempts to insert the pump were unsuccessful. an attempt to insert a second pump was also unsuccessful. the guidewire and pigtail catheter could not cross the aortic valve because it wasn't opening due to ecmo therapy (cpr was applied in the meanwhile). ct later showed dissection of aorta / iliac artery. md not completely sure, if dissection was already there or caused by guidewire/pigtail catheter, but states that the advancing of said went without any resistance. percutaneous coronary intervention was performed with the support of an already connected extracorporeal membrane oxygenation (ecmo). patient died of cgs later, did not undergo vascular surgery. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36027 | 1220648-2022-01726 | 16022552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16022552&pc=OZD | 12/06/2022 | 12/20/2022 |  | 12/21/2022 |  |  | 12/07/2022 | 12/21/2022 |  |  | 04/30/2024 |  | 12/21/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023211811 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/07/2022 |  |  |  |  | The complainant reported a 59-year-old male patient presenting for cardiomyopathy. during impella support with an impella cp pump, it was noted that the patient developed hemolysis. as a result of the condition, and of the need for continued support, the patient was escalated from the impella cp to an impella 5.5 pump. the patient was stable following the escalation. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36028 | 1220648-2022-01725 | 16022553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16022553&pc=OZD | 12/05/2022 | 12/20/2022 |  | 12/21/2022 |  |  | 12/07/2022 | 12/21/2022 |  |  | 02/29/2024 |  | 12/21/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023191802 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 27 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/07/2022 |  |  |  |  | The complainant reported a 27-year-old male patient presenting for cardiomyopathy. during impella support, it was noted that the patient experienced a blood leak from the impella access site. attempts were made to stop the bleeding with gauze, compression and suturing but the condition remained. the patient was transfused an unknown amount of blood to counteract the ongoing condition and the decision was made to prematurely remove the pump and escalate to an impella 5.5. the patient was stable following the event. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36029 | 1220648-2022-01733 | 16029047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16029047&pc=OZD | 11/23/2022 | 08/03/2023 |  | 12/21/2022 |  |  | 11/24/2022 | 12/21/2022 |  |  | 06/30/2024 |  | 12/21/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023233447 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Malfunction | Yes | Yes | Detachment of Device or Device Component | Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 67-year-old male patient underwent cardiac surgery. due to post cardiotomy cardiogenic shock / low coronary output syndrome he received hemodynamic support by an impella 5.0, introduced via an aortic arch side to side graft anastomosis. during the explantation process the impella pump separated from the catheter shaft and stayed in the graft. the surgeons performed an unscheduled resternotomy to explant the remaining part of the impella device from graft and aorta. the pump has been successfully removed. the patient suffered no further complication. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the cannula detachment has been completed since the original report was filed. the device was returned and an evaluation was performed. a visual inspection verified that the pump was found to be detached between below the motor and just before the catheter. evidence of excessive force was noted and the root cause of the cannula detachment was subsequently noted to be most likely use related. the failure mode will be monitored and trended.
36030 | 1220648-2022-01734 | 16029071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16029071&pc=OZD | 10/06/2022 | 10/01/2025 |  | 12/21/2022 |  |  | 12/20/2022 | 12/21/2022 |  |  | 08/31/2023 |  | 12/21/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022119994 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 26 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2022 |  |  |  |  | The complainant in japan hada cp support in 10/2022 that was successful for 4 days when weaned and explanted. later in 12/2022 the jpvad registry noted thrombosis at the time of explant that lead to limb ischemia. the ischemia prompted thrombectomy. | [5 empty] | The medical center discarded all product at the time of explant. the jpvad registry furnished no further details on the cause of the thombosis. the manufacturer will close the investigation without root cause being determined. the failure mode will be monitored and trended.
36031 | 1220648-2022-01727 | 16016871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16016871&pc=OZD | 10/17/2022 | 12/20/2022 |  | 12/20/2022 |  |  | 12/16/2022 | 12/20/2022 |  |  | 05/31/2024 |  | 12/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023222817 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/16/2022 |  |  |  |  | The medical team in japan had an impella cp support in 10/2022 that was successful til the patient expired from underlying unrelated causes. additionally the family made the patient a dnr. the pump had supported the patient beyond indication, as the pump had run for over 21 days. later in 12/2022 the jpvad registry noted an access site bleed and hemolysis treated by blood product and haptoglobin infusion during the time of impella support. | [5 empty] | All product was discarded months ago at the time of explant. no benchtop analysis leading to root cause is possible. the jpvad registry furnished no further clinical details. the manufacturer will close the investigation. should new information be shared, a final report will follow.
36032 | 1220648-2022-01728 | 16016889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16016889&pc=OZD | 12/16/2022 | 10/09/2025 |  | 12/20/2022 |  |  | 12/16/2022 | 12/20/2022 |  |  | 08/31/2024 |  | 12/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023270569 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2022 |  |  |  |  | The medical team placed an impella cp to support a patient admitted in shock. the team then decided to transfer the patient for more advanced support to the (b)(6). b)(6), at admission, found the patient was suffering from hemolysis and began dialysis for the renal issues. | [5 empty] | The medical center discarded the impella pump at the time of explant. no benchtop analysis and testing to root cause is possible. should the manufacturer be furnished with more details that lead to root cause, a final report will follow.
36033 | 1220648-2022-01729 | 16016941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16016941&pc=OZD | 12/12/2022 | 09/14/2024 |  | 12/20/2022 |  |  | 12/13/2022 | 12/20/2022 |  |  | 12/31/2023 |  | 12/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023165713 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | Death |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2022 |  |  |  |  | The medical team in japan had an impella cp support ongoing for an elderly patient admitted in ami. the patient was additionally supported by ecmo. on day 6 of the support the patient was diagnosed as suffering from a cerebral infarct. the infarct took place during the impella support. the family made the choice to withdraw care on following day and the patient expired. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The medical team discarded the impella pump. no analysis to root cause is possible. the medical team in japan furnished no further clinical details. should more information be shared, that lead to root cause, a final report will follow. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01729 was provided.
36034 | 1220648-2022-01730 | 16017002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16017002&pc=OZD | 12/05/2022 | 12/20/2022 |  | 12/20/2022 |  |  | 12/05/2022 | 12/20/2022 |  |  | 08/31/2024 |  | 12/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023262907 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/05/2022 |  |  |  |  | The medical team had an impella cp inserted femorally for the support of a patient admitted and found to have coronary artery disease. the impella placement caused flow issues in the right leg. the vascular surgery team placed 5fr external bypass to relieve the ischemia. the pump remained on for support until care was withdrawn and the patient expired. | [5 empty] | The medical center discarded the pump and introducer at the time of explant when care was withdrawn. no benchtop analysis to root cause of the ischemia is possible. should the medical team furnish more clinical details that lead to root cause of the ischemia, a final report will be filed. the failure mode will be monitored and trended.
36035 | 1220648-2022-01731 | 16018452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16018452&pc=OZD | 09/28/2022 | 12/20/2022 |  | 12/20/2022 |  |  | 11/21/2022 | 12/20/2022 |  |  | 07/31/2024 |  | 12/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023238778 |  | 0048-0014 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Information about adverse event from protect iv study: a patient in the eu received impella cp support during a high risk pci procedure. 1.5 hours after the procedure, the patient suffered from severe access site bleeding. hemoglobin drop of 3 g/l. contributing factor: pump removed with activated clotting time (act) of 256 seconds instead of <150 seconds. balloon tamponade used to stop bleeding. ct showed a hematoma but no active bleeding afterwards. | [5 empty] | The root cause of the access site bleeding was unable to be determined as no product was returned and limited clinical details were provided. 12 months number of usage: 31,078; number of occurrences: 47; occurrence rate (%): 0.15.
36036 | 1220648-2022-01722 | 16004129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16004129&pc=OZD | 07/22/2022 | 12/19/2022 |  | 12/19/2022 |  |  | 12/06/2022 | 12/19/2022 |  |  | 03/31/2024 |  | 12/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023200392 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/06/2022 |  |  |  |  | The complainant reported a patient presenting for impella support. upon conclusion of support and subsequent explant of the pump, it was discovered that the patient had experienced an ischemic stroke. additional information pertaining to the event was not provided. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36037 | 1220648-2022-01723 | 16006013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16006013&pc=OZD | 12/02/2022 | 12/19/2022 |  | 12/19/2022 |  |  | 12/02/2022 | 12/19/2022 |  |  | 08/31/2024 |  | 12/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023269625 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/02/2022 |  |  |  |  | The complainant reported a 54-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient developed limb ischemia. a thrombectomy was performed in an attempt to improve the condition, however, the issue remained. as a result of the ongoing condition, the decision was made to wean the patient and explant the pump. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36038 | 1220648-2022-01724 | 16009702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=16009702&pc=OZD | 11/30/2022 | 12/19/2022 |  | 12/19/2022 |  |  | 11/30/2022 | 12/19/2022 |  |  | 07/31/2024 |  | 12/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023256302 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/30/2022 |  |  |  |  | The complainant reported a 55-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. following implant, reduced flows were noted distal to the 14 fr sheath. the pump was explanted as a result of the condition and the patient¿s ischemia resolved. the physician noted that despite the condition, the pump successfully supported the patient throughout the scheduled procedure. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36039 | 1220648-2022-01720 | 15997752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15997752&pc=OZD | 08/05/2022 | 05/23/2025 |  | 12/16/2022 |  |  | 11/18/2022 | 12/16/2022 |  |  | 02/29/2024 |  | 12/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023191696 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/18/2022 |  |  |  |  | The complainant reported an impella patient developing an access site hematoma during support. a blood transfusion was required. | The complainant reported that the patient had an infection from insertion to withdrawal. action taken to resolve the issue are unknown. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The cause of the issue was not determined since product was not returned and there is insufficient clinical evidence. this report is being filed as part of a retrospective review of historical records. additional information for the initial mfr 1220648-2022-01720 was provided in sections a5, b5, d6a, d6b, g1, h1, and h6. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
36040 | 1220648-2022-01721 | 15998356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15998356&pc=OZD | 11/18/2022 | 12/16/2022 |  | 12/16/2022 |  |  | 11/18/2022 | 12/16/2022 |  |  | 08/31/2024 |  | 12/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023264032 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/18/2022 |  |  |  |  | The complainant reported a 71 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during the 14fr introducer insertion process, the patient's iliac artery endured a dissection. a balloon tamponade was deployed and the patient's opposing femoral artery was used for insertion of new impella. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36041 | 1220648-2022-01716 | 15987094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15987094&pc=OZD | 11/16/2022 | 12/15/2022 |  | 12/15/2022 |  |  | 11/16/2022 | 12/15/2022 |  |  | 08/31/2023 |  | 12/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC |  |  | US |  | 2022146969 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/16/2022 |  |  |  |  | The complainant reported a 56-year-old male patient presenting for cardiomyopathy. upon a planned explant of the impella pump, it was noted that there was a clot on the tip of the pigtail. no visual clots in the patient were initially noted, however, upon full explant, the patient's cerebral oximetry dropped. as a result, the patient's aorta was opened and an organized clot was discovered prompting surgical removal. the patient¿s condition improved following the intervention. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36042 | 1220648-2022-01717 | 15991728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15991728&pc=OZD | 11/16/2022 | 12/15/2022 |  | 12/15/2022 |  |  | 11/16/2022 | 12/15/2022 |  |  | 03/31/2024 |  | 12/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023209833 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/16/2022 |  |  |  |  | The complainant reported a 58 year old female patient presenting with post cardiotomy cardiogenic shock for impella support. during an attempted wireless insertion of the impella device, the left ventricle became perforated. the perforation was repaired surgically and support continued. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36043 | 1220648-2022-01718 | 15991882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15991882&pc=OZD | 08/14/2022 | 12/15/2022 |  | 12/15/2022 |  |  | 11/25/2022 | 12/15/2022 |  |  | 02/29/2024 |  | 12/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023191670 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/25/2022 |  |  |  |  | The medical team in japan had an impella cp support ongoing in 8/2022 when there were signs/symptoms of hemolysis that prompted the team to troubleshoot and intervene. this was noted in the jpvad registry, in 11/2022, that haptoglobin was infused. | [5 empty] | The medical center discarded all impella product months prior. the jpvad registry furnished no further details that lead to root cause of the presumed hemolysis. there can be no benchtop testing of the impella pump. the manufacturer will close the analysis without root cause. the failure mode will be monitored and trended.
36044 | 1220648-2022-01719 | 15991969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15991969&pc=OZD | 11/28/2022 | 12/15/2022 |  | 12/15/2022 |  |  | 11/28/2022 | 12/15/2022 |  |  | 06/30/2024 |  | 12/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023247165 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 70 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/28/2022 |  |  |  |  | The medical center had an impella cp patient enrolled in the stemi dtu trial. the pump support was successful but, at the time of explant there was thrombosis noted and the team performed penumbra aspiration. the flow was partially resolved. | [5 empty] | The medical center discarded all product at the time of explant. no benchtop analysis of the pump or introducer sheath is possible. the stemi dtu documentation furnished no further details regarding the causality of the thrombosis. the case will be closed without root cause determination, however should more information be shared that leads to root cause, a final report will follow.
36045 | 1220648-2022-01715 | 15977684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15977684&pc=OZD | 11/21/2022 | 12/14/2022 |  | 12/14/2022 |  |  | 11/21/2022 | 12/14/2022 |  |  | 03/31/2023 |  | 12/14/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022067463 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Valvular Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/21/2022 |  |  |  |  | The medical center placed an impella 5.5 via direct aortic access when in the operating room for mitral valve and bypass (cabg) surgery. the support was for just over 7 days when explanted. at the time of pump explant, there was new damage to the aortic valve. the patient was suffering from new regurgitation, and found to have damage to the non-coronary cusp. the team had aortic valve replacement to remedy the issue. | [5 empty] | The medical center discarded the impella pump product. no benchtop analysis is possible to lead to root cause. the team did not furnish any images of the valve anatomy or surgical report. should more information be shared that leads to root cause, a final report will follow. the failure mode will be monitored and trended.
36046 | 1220648-2022-01714 | 15970738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15970738&pc=OZD | 08/18/2022 | 10/23/2024 |  | 12/13/2022 |  |  | 11/29/2022 | 12/13/2022 |  |  | 01/31/2024 |  | 12/13/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023190908 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia Stroke | [11 empty] | Required Intervention | Death | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2022 |  |  |  |  | The complainant reported a 56-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was reported that the patient experienced an ischemic stroke. impella support continued following the stroke. additional information regarding the stroke could not be obtained. | Additional information that the patient had a blood transfusion for bleeding including hematoma. approximately eleven (11) days following the explant of the pump, the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | As noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2022-01714 was provided.
36047 | 1220648-2022-01713 | 15963486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15963486&pc=OZD | 11/14/2022 | 08/21/2023 |  | 12/12/2022 |  |  | 11/14/2022 | 12/12/2022 |  |  | 06/30/2024 |  | 12/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023247751 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/14/2022 |  |  |  |  | The medical team had inserted an impella cp for support, via the femoral artery, for a patient admitted in cardiogenic shock. the team found that there was blood loss from the impella access site. the site was treated by the vascular surgery team. the impella was removed, the artery was opened and repaired. of note the patient was obese and on multiple anticoagulant drugs which may have contributed to the bleed and need for surgical repair. after the repair the patient remained on impella support, but via the axillary artery and an impella 5.5 placed via that access. | [5 empty] | The investigation is on-going at this time. the medical center has returned for analysis the impella pump. upon conclusion of the investigation, a supplemental report will follow. | The investigation into the vascular damage has been completed since the original report was filed. the medical center has returned for analysis the impella pump and the benchtop analysis revealed that the root cause of the vascular damage was most likely patient condition related.
36048 | 1220648-2022-01709 | 15952505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15952505&pc=OZD | 11/10/2022 | 12/09/2022 |  | 12/09/2022 |  |  | 11/10/2022 | 12/09/2022 |  |  | 08/31/2024 |  | 12/09/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023265587 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/10/2022 |  |  |  |  | The complainant reported a 51 year old male patient presenting for high risk percutaneous coronary intervention using the impella cp optical for hemodynamic support. the physician noted a concern of limb ischemia distal to the device's repositioning sheath, therefore, a fem-fem bypass was performed. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36049 | 1220648-2022-01710 | 15953472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15953472&pc=OZD | 05/07/2021 | 12/09/2022 |  | 12/09/2022 |  |  | 05/07/2021 | 12/09/2022 |  |  | 02/28/2023 |  | 12/09/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022047435 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/07/2021 |  |  |  |  | The complainant reported a 69 year old male patient presenting with cardiomyopathy for impella support. during support, an access site hematoma developed. a washout via surgical incision was required. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36050 | 1220648-2022-01711 | 15954534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15954534&pc=OZD | 11/18/2022 | 12/09/2022 |  | 12/09/2022 |  |  | 11/18/2022 | 12/09/2022 |  |  | 07/31/2024 |  | 12/09/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2023236831 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/21/2022 |  |  |  |  | The complainant reported a 67 year old female patient presenting with post cardiotomy cardiogenic shock for impella support. upon insertion of the device, the tip of the catheter perforated the patient's left ventricle. the injury was repaired and support proceeded successfully. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36051 | 1220648-2022-01712 | 15955455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15955455&pc=OZD | 11/06/2022 | 12/23/2022 |  | 12/09/2022 |  |  | 11/16/2022 | 12/09/2022 |  |  | 02/29/2024 |  | 12/09/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP SET JP |  |  | ABIOMED INC. |  |  | US |  | 2023183296 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/16/2022 |  |  |  |  | The medical center in japan had an impella 5.5 support ongoing when the pump stopped on day 8 of support. the team had changed the purge solution and the pump stopped after changes in purge pressure and purge flow. the pump was explanted and not replaced. | [5 empty] | The medical center in japan did return the pump for analysis. the analysis is on-going at this time. upon completion of the investigation, and when root cause is determined, a final report will be filed. | Since the original report was filed, the investigation into the pump stop has concluded. the medical team in japan returned the pump and data logs for analysis. the analysis on the benchtop revealed native patient biomaterial had been ingested which caused the pump to stop. the failure mode will be monitored and trended.
36052 | 1220648-2022-01705 | 15939233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15939233&pc=OZD | 12/05/2020 | 12/08/2022 |  | 12/08/2022 |  |  | 11/09/2022 | 12/08/2022 |  |  | 02/28/2022 |  | 12/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 1466638 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A patient in the eu suffered from access site bleeding two days after implant, which required temporary suspension of anticoagulation and blood products transfusion. after 10 days of support, the patient was weaned from the right- ventricular support and on the day after from impella 5.0 as complete recovery of cardiac function and end-organ damage was observed. input from a publication: "full myocardial recovery following covid-19 fulminant myocarditis after biventricular mechanical support with bipella: a case report" | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36053 | 1220648-2022-01706 | 15940341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15940341&pc=PYX | 10/05/2022 | 12/04/2024 |  | 12/08/2022 |  |  | 11/08/2022 | 12/08/2022 |  |  | 09/30/2023 |  | 12/08/2022 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023186633 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Pulmonary Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2022 |  |  |  |  | The complainant reported a 54 year old male patient presenting with acute myocardial infarction and cardiogenic shock. the pump supported the patient until explant without any reported problems. approximately 27 days later, a bacterial growth (vegetation) was found on the patient's pulmonic valve. the physician feels the impella may have damaged the valve's leaflet which allowed for bacterial growth to occur. the bacterial growth was noted as a result of using an infected dialysis catheter, not the impella device. an angiovac procedure was completed to remove the growth. the patient continues to recover and was noted as doing well. | Additional information reported that after the impella rp was implanted the peel away sheath was removed and the repositioning sheath was advanced. the previously deployed mattress suture was locked down, however the groin site continued to bleed. an additional mattress suture was deployed and manual pressure was applied to resolve the bleeding. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information provided in b5 ad h6 for additional bleeding reported. d6a/d6b included for implant and explant dates. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp with smart assist system with automated impella controller instructions for use for the related event are as follows: section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
36054 | 1220648-2022-01707 | 15945771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15945771&pc=OZD | 10/03/2022 | 08/21/2023 |  | 12/08/2022 |  |  | 11/08/2022 | 12/08/2022 |  |  | 07/31/2024 |  | 12/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023240451 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 27 YR |  | Male |  |  |  | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Other | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Input from a study case adverse event reporting: a patient in the eu suffered from severe access site bleeding with shock after impella removal. patient has been under full anticoagulation, contributing to the bleeding. bleeding classification barc type 3b. prolonged existing hospitalization. bleeding treated with three units of blood products. bleeding stopped with femostopp. patient recovered, no residual effects. | [5 empty] | The impella device was received and an evaluation is ongoing. a supplemental report will be filed upon conclusion of our investigation. | The investigation into the pump stop and the access site bleeding has been completed since the original report was filed. the benchtop analysis revealed that the root cause of the pump stop was bearing wear; however, the cause of the access site bleeding was not determined since there are insufficient clinical details.
36055 | 1220648-2022-01708 | 15946784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15946784&pc=OZD | 11/22/2022 | 12/08/2022 |  | 12/08/2022 |  |  | 11/25/2022 | 12/08/2022 |  |  | 05/31/2024 |  | 12/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023219805 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/25/2022 |  |  |  |  | The complainant reported a 58-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support it was noted that the patient developed lower limb ischemia. a bypass sheath was placed in an attempt to improve the patient's flow but the condition remained. an endovascular thrombectomy was also completed but the patient's ischemia continued to worsen and the decision was subsequently made to remove the pump. the patient recovered from the condition following removal of the pump. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36056 | 1220648-2022-01700 | 15924128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15924128&pc=OZD | 11/25/2022 | 09/22/2024 |  | 12/06/2022 |  |  | 11/25/2022 | 12/06/2022 |  |  | 07/31/2024 |  | 12/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023246570 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | [12 empty] | Death | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/25/2022 |  |  |  |  | The medical team in japan had an impella cp support initiated after a balloon pump (iabp) for care escalation. after 16 hours the patient expired on the support. the patient had suffered a tamponade but, the causality was not known. the unknown relationship of the tamponade and the impella causing the tamponade can not be denied. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The medical center in japan discarded all product at explant. no benchtop analysis to root cause is possible. a request for more clinical information and imaging has not been met. the manufacturer will close the investigation but, will file a final report should more be shared that leads to a root cause of the cardiac tamponade. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
36057 | 1220648-2022-01702 | 15929274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15929274&pc=OZD | 10/18/2022 | 12/06/2022 |  | 12/06/2022 |  |  | 12/06/2022 | 12/06/2022 |  |  | 05/31/2024 |  | 12/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023225157 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/06/2022 |  |  |  |  | The medical team in japan had an impella cp support in 10/2022. the support was for 7 days when explanted and patient survived. later in 12/2022 the jpvad registry noted an ischemic/embolic cva had occurred on day 6 of the support. | [5 empty] | The jpvad registry team furnished no clinical details on the reported cva. the causality between impella use and the cva can not be determined with known information. no root cause is possible. the failure mode will be monitored and trended.
36058 | 1220648-2022-01701 | 15929307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15929307&pc=OZD | 08/24/2022 | 10/25/2023 |  | 12/06/2022 |  |  | 10/24/2022 | 12/06/2022 |  |  | 05/31/2024 |  | 12/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023223986 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2022 |  |  |  |  | The medical center had an impella 5.5 support initiated on a surgical patient. the pump was placed via direct access to the ascending aorta. the 5.5 pump stopped after 3 hours. the pump was explanted and replaced. | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 support initiated on a surgical patient. the pump was placed via direct access. the 5.5 pump stopped after 3 hours. the pump was explanted and replaced. |  |  |  |  | The medical center discarded the impella pump at explant. no benchtop analysis to root cause of the malfunction is possible. should a root cause be determined via new information being shared, a final report will follow. | Although the pump was not returned, data was reviewed. data logs show the pump running to specification until (b)(6) 2022 at around 21:10, when the mc loses pulsatility, snr drops to around 500, and the raw ps becomes erratic and begins to "barcode". as this occurs without a change in motor speed, it is likely the cause of the placement signal issue is due to biomaterial. the flow decreased as a result of the impact to the placement signal, however the pump did not stop. following the flow issue, it appears as though the team removed the pump from the body while running, and it was then unplugged. in conclusion, the root cause of the position signal issue is most likely due to biomaterial, and the change in flow was likely a result of the position signal issue. the report of pump stop could not be confirmed, however biomaterial was suspected in restricting flow. the impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
36059 | 1220648-2022-01703 | 15929344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15929344&pc=OZD | 11/07/2022 | 12/06/2022 |  | 12/06/2022 |  |  | 11/07/2022 | 12/06/2022 |  |  | 09/30/2024 |  | 12/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023256991 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/07/2022 |  |  |  |  | The complainant reported an 87 year old female patient presenting for pre-operative impella support. attempts to deliver the impella device were unsuccessful due to patient's tortuous vasculature and caused damage to the patient's anastomosis. surgical repair was required as a result. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36060 | 1220648-2022-01704 | 15929429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15929429&pc=PYX | 11/08/2022 | 12/06/2022 |  | 12/06/2022 |  |  | 11/08/2022 | 12/06/2022 |  |  | 07/31/2023 |  | 12/06/2022 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022131574 |  | 0046-0035 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/08/2022 |  |  |  |  | The complainant reported a 52 year old male patient presenting post-operative requiring right heart support via the impella rp. during support, the patient was suspected as enduring hemolysis as evident by rising ldc values. as a result, the pump was removed and the patient was placed on alternative therapy. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36061 | 1220648-2022-01698 | 15914946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15914946&pc=OZD | 09/15/2022 | 12/05/2022 |  | 12/05/2022 |  |  | 11/30/2022 | 12/05/2022 |  |  | 03/31/2024 |  | 12/05/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023207554 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/30/2022 |  |  |  |  | The medical center in japan had a successful impella cp support of 6 days in (b)(6) 2022. later in (b)(6) 2022 the jpvad registry team noted hemolysis. the hemolysis was treated by haptoglobin infusion. | [5 empty] | The medical center in japan discarded all product at explant. no benchtop analysis or hemolysis testing of the pump is possible. root cause can not be determined. the manufacturer will monitor and trend the failure mode.
36062 | 1220648-2022-01699 | 15914995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15914995&pc=OZD | 04/19/2022 | 12/05/2022 |  | 12/05/2022 |  |  | 11/25/2022 | 12/05/2022 |  |  | 08/31/2023 |  | 12/05/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022127403 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/25/2022 |  |  |  |  | In 4/2022 there was a successful 5 day impella cp support in japan. the age and gender was unknown and not documented. in 11/2022 the jpvad registry noted hemolysis to have occurred during the support. it was noted on day 4 of the support. the pump remained on despite this and the hemolysis was treated by haptoglobin infusion. | [5 empty] | The medical center discarded all product in 4/2022 and as such no benchtop analysis to root cause is possible. no hemolysis testing is possible. the investigation will be closed, but should more information be shared, by the jpvad registry, a final report can be filed with root cause.
36063 | 1220648-2022-01697 | 15897579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15897579&pc=OZD | 11/02/2022 | 12/30/2022 |  | 12/01/2022 |  |  | 11/02/2022 | 12/01/2022 |  |  | 09/30/2023 |  | 12/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2022145594 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Necrosis | Thrombosis/Thrombus | [9 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/04/2022 |  |  |  |  | The complainant reported a 42-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, the patient experienced an arterial occlusion leading to lower limb ischemia. upon initial observation, the patient had oozing at the access site that coincided with the presence of a hematoma. the patient had no distal pulse in the impella leg and an arterial ultrasound showed a total left leg arterial occlusion. the patient's condition prompted a fem to fem external bypass to restore blood flow after foley placement proved to be unsuccessful. femostop was used to reach hemostasis, however, following completion of the bypass, the patient moved and pulled the repo sheath out of position. the sheath was repositioned and re-sutured to resolve the bleed. during the noted interventions, three units of prbc were transfused to maintain support. a follow up echo was completed and a clot was shown to be developing on the pigtail and inlet cage of the pump. as a result, the pump was removed and replaced with a new impella cp for continued support. | Upon additional follow up from the field rep, it was reported that the patient underwent a thrombectomy at the time of pump removal to treat the occlusion noted in the left leg. seven days later, the limb was amputated as the leg became necrotic following the thrombectomy. per the vascular surgeon, it was believed that the late addition of the fem to fem bypass resulted in compartment syndrome and subsequently led to the muscle tissue from the mid thigh down to die. support was maintained throughout the intervention. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information was provided by the field rep regarding the reported event. please see field b2, b5 and h6 for the updated information pertaining the patient's condition and treatment.
36064 | 1220648-2022-01695 | 15888023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15888023&pc=OZD | 11/08/2022 | 11/30/2022 |  | 11/30/2022 |  |  | 11/08/2022 | 11/30/2022 |  |  | 07/31/2024 |  | 11/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023253844 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/08/2022 |  |  |  |  | The complainant reported a 17-year-old female patient presenting for impella support due to a potential heart transplant rejection. upon attempting to remove the pump, complications were noted that resulted in a vascular dissection. it was documented that the physician was unable to remove the pump out of the small vessel despite attempts to cut down the graft and clear the pathway. eventually, the impella 5.5 was pulled with force and a hole formed in the artery which eventually caused the distal axillary artery to detach. the artery was pulled back and three covered stents coinciding with a bare-metal stent were placed to repair the artery. the patient was recovering following the events and further intervention was not required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36065 | 1220648-2022-01696 | 15888053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15888053&pc=OZD | 11/11/2022 | 10/25/2023 |  | 11/30/2022 |  |  | 11/11/2022 | 11/30/2022 |  |  | 06/30/2024 |  | 11/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023232712 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Material Twisted/Bent | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2022 |  |  |  |  | The complainant reported a 63-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during implantation attempts for the impella, it was noted that the patient was bleeding from the access site. floseal was subsequently placed around the anastamosis to achieve hemostasis. additional intervention was not required for the bleed and the patient remained stable following the event. | In addition to the original filing for access site bleed, it should be known that the pump had become bent. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | (h3) the investigation into the reported "access site bleeding and delivery issues" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. based on the clinical account, the root cause of the delivery issues was due to patient condition. the cause of the bleeding could not be determined due to lack of clinical details and that the product was not returned. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
36066 | 15864662 | 15864662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15864662&pc=OZD | 11/15/2022 | 11/16/2022 | 11/16/2022 | 11/25/2022 | 11/16/2022 | 11/25/2022 |  | 11/25/2022 |  |  |  |  | 11/25/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Nurse | [6 empty] | 18250 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | Failure to Pump | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Crack at side arm connection of pressure reservoir purge cassette. sodium bicarb used as purge fluid- known to cause breakage at side arm- extension tubing is current workaround- extension tubing already in place at time of device failure- currently if side arm breaks there are 2 options: cut purge cassette and replace connection with 21g needle (temp fix) or replace entire impella device (not an option as patient is not candidate and it would require another surgery).
36067 | 1220648-2022-01691 | 15857791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15857791&pc=OZD | 10/26/2022 | 10/25/2023 |  | 11/23/2022 |  |  | 10/26/2022 | 11/23/2022 |  |  | 02/29/2024 |  | 11/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST |  |  | ABIOMED, INC. |  |  | US |  | 2023188544 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Injury | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 10/26/2022 |  |  |  |  | A 62-year-old male patient has been treated for ami / cgs. for hemodynamic support, an impella 5.5 device has been inserted. insertion has been difficult due to small vessel size (6 mm) and tortuosity. no vasodilators used. during that process, the pump dissected the aortic arch. aortic arch stent placement to repair the dissection. biv pacer placed at that time as well. new impella 5.5 pump inserted. patient hemodynamically stable afterwards. | It was also reported during support, there was low flows on the pump despite being in a good position. the decision was made to remove the pump. a new pump was placed. |  |  |  |  | The returned pump was visually inspected and dried blood was observed in the outflow. the pump was run in a basin of water and pump flow was within specification. the pump was removed from the basin, visually inspected and two kinks were observed in the cannula. the kinks did not result from the lab run. | This supplemental is being filed to correct some erroneous information listed in initial mdr. see below for corrected data: h(3): the device has been returned, but investigation is still ongoing. h(6): please disregard codes: investigation findings 213 - no device problem found. investigation conclusions 50- adverse event related to patient condition. | The investigation into the reported low pump flow and vascular damage has been completed since the original report was filed. the pump was returned for evaluation. the evaluation revealed that user error was the most likely cause of the low pump flow: "the cause of the low pump flow was most likely left axillary tortuosity necessitating excessive force resulting in cannula kinks during delivery. the 6mm measured vessel diameter was below the 7mm minimum vessel diameter per ifu 0550-9002." the root cause for vascular damage is user error: the cause of the vascular damage was most likely left axillary tortuosity patient condition related. the 6mm measured vessel diameter was below the 7mm minimum vessel diameter per ifu 0550-9002. | H6-health impact code, medical device codes, component codes, investigation codes, investigation conclusion code.
36068 | 1220648-2022-01693 | 15858598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15858598&pc=OZD | 11/08/2022 | 11/23/2022 |  | 11/23/2022 |  |  | 11/08/2022 | 11/23/2022 |  |  | 08/31/2024 |  | 11/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023261265 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/08/2022 |  |  |  |  | The complainant reported a 71-year-old female patient presenting for high risk percutaneous coronary intervention. following explant of the impella pump, it was discovered that the patient had a retroperitoneal bleed originating from the distal descending aorta. a cause for the vascular damage was unknown. a surgical repair was completed to resolve the bleed and the patient was reported to be recovering from the condition. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36069 | 1220648-2022-01692 | 15858798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15858798&pc=OZD | 10/26/2022 | 11/23/2022 |  | 11/23/2022 |  |  | 10/26/2022 | 11/23/2022 |  |  | 07/31/2024 |  | 11/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023254842 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  | White | Malfunction | No | Yes | Detachment of Device or Device Component | Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/26/2022 |  |  |  |  | A 53-year-old white male presented with ami/cgs and received hemodynamic support with an impella cp heart pump, introduced via the left femoral artery. during the procedure the physician tried to reposition the device and pulled it out of the left ventricle. subsequently, the impella device ended up folding up on itself. the physician tried to straighten it back out by pulling it into the introducer sheath. while trying to pull it through the sheath, the distal end of the cannula right above the inlet cage separated and remained in the femoral artery attached by the nitinol coil. the cardiovascular surgeon did a cutdown on the femoral artery and removed the piece without difficulty. the surgeon used the left femoral access site to place a 19fr arterial ecmo cannula. the venous cannula was placed in the right femoral vein and a new impella cp was placed through the right femoral artery without issue. | [5 empty] | A visual inspection of the pump confirmed that the cannula detachment occurred near the inflow cage. no other abnormalities were observed. the peel away sheath remained on the catheter and no kinks on the sheath were observed. product did not fail to meet specification. root cause verified to be use issue related to excessive force.
36070 | 1220648-2022-01694 | 15858805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15858805&pc=OZD | 11/05/2022 | 11/23/2022 |  | 11/23/2022 |  |  | 11/05/2022 | 11/23/2022 |  |  | 04/30/2024 |  | 11/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023226425 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/05/2022 |  |  |  |  | The complainant reported a 17-year-old female patient presenting for cardiomyopathy. during impella support, it was discovered that the patient had vascular damage resulting in the development of bleeding complications and hematomas. a cutdown was initially attempted with the pump in place, however the pump ultimately had to be removed so that proper access to the femoral artery could be obtained for the repair. coinciding with the artery repair, the patient was provided two units of blood to counteract the bleed. a definitive cause for the damage was unknown. the patient was reported to be stable following the intervention. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36071 | 1220648-2022-01690 | 15858955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15858955&pc=OZD | 10/12/2022 | 10/13/2025 |  | 11/23/2022 |  |  | 11/21/2022 | 11/23/2022 |  |  | 05/31/2024 |  | 11/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023225054 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Renal Failure | Tachycardia | Thrombocytopenia | [8 empty] | Death | Required Intervention |  |  |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/21/2022 |  |  |  |  | The medical team in japan had an impella cp support that was successful for 9 days when it was weaned and explanted in (b)(6) 2022. later in (b)(6) 2022 the jpvad registry noted an unrelated death due to sepsis and limb ischemia. the ischemia was thought to be due to clot embolization from the impella at explant. the limb was treated by thrombectomy. | Additional information noted that, the patient had access site bleeding requiring blood transfusion, thrombocytopenia, renal failure and ventricular tachycardia during support. the procedure outcome was that patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The medical center discarded all product at explant. no benchtop analysis and testing to root cause is possible. the jpvad registry furnished no further clinical details. the failure mode will be monitored and trended. should the medical team in japan furnish more details that lead to root cause, a final report will follow. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. as noted in the impella ifu: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information was provided in sections b2, b5, d6, h1, and h6.
36072 | 1220648-2022-01684 | 15843212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15843212&pc=OZD | 10/31/2022 | 11/22/2022 |  | 11/22/2022 |  |  | 11/08/2022 | 11/22/2022 |  |  | 02/29/2024 |  | 11/22/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023195315 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Paralysis | Ischemia Stroke | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/08/2022 |  |  |  |  | The complainant reported a 57-year-old male patient presenting for cardiomyopathy. following removal of the impella pump, it was observed that the patient experienced right upper and lower extremity paralysis. a ct scan was performed and it was discovered that the patient had experienced an ischemic stroke. additional information pertaining to the condition was not available and a cause could not be determined. the patient was reported to be recovering in rehab following the event. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36073 | 1220648-2022-01685 | 15847802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15847802&pc=OZD | 08/27/2022 | 10/14/2024 |  | 11/22/2022 |  |  | 11/18/2022 | 11/22/2022 |  |  | 10/31/2023 |  | 11/22/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022147710 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility |  |  | 35 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Renal Failure | [10 empty] | Required Intervention | [5 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/18/2022 |  |  |  |  | The medical team in japan had an access site bleed and renal failure reported, in 11/2022 in the jpvad registry, from an impella cp support of 12 days in 8/2022. the registry reported upon blood product delivery for the access bleed, of an unknown amount of units transfused. no other intervention was needed. the renal failure also occurred during the support and was treated by haptoglobin infusion. the jpvad registry furnished no further details. | Subsequently, additional information was received indicating that the patient also experienced an access site infection, in (b)(6) 2022 in the jpvad registry from an impella cp support of 12 days in (b)(6) 2022. the access site infection occurred during support and unknown additional measures were taken; however, the additional measures resolved the issue. no further details were provided within the jpvad registry. |  |  |  |  | The medical center in japan discarded all product in 8/2022. as such, no investigation to root cause is possible. no further information has been shared by the jpvad registry team that leads to a root cause for either the bleeding nor the renal failure. the failure modes will be monitored and trended. | Additional information for the initial mfr 1220648-2022-01685 was provided in sections b5, d6a, d6b and h6. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
36074 | 1220648-2022-01686 | 15847944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15847944&pc=OZD | 11/04/2022 | 11/22/2022 |  | 11/22/2022 |  |  | 11/04/2022 | 11/22/2022 |  |  | 04/30/2024 |  | 11/22/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023220627 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/04/2022 |  |  |  |  | The complainant reported a 57-year-old male patient presenting for cardiomyopathy. during impella support, it was discovered that the pump had gotten stuck under the mitral valve and, as a result, the patient started hemolyzing. repositioning of the pump was attempted but ultimately failed to resolve the issue. as a result of the ongoing condition, the pump was removed and a new impella pump was implanted for continued support. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36075 | 1220648-2022-01689 | 15847991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15847991&pc=OZD | 10/24/2022 | 10/25/2023 |  | 11/22/2022 |  |  | 10/24/2022 | 11/22/2022 |  |  | 08/31/2024 |  | 11/22/2022 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMART ASSIST |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2023252888 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2022 |  |  |  |  | Patient in the eu had a impella 5.5 pump stop. pump unable to restart. pump successfully explanted. patient survived. no further measures required. patient has already been hemodynamically stable again. | In addition to the original filing of pump stop, it should be known the high purge pressure was also exhibited. |  |  |  |  | A visual inspection of the pump revealed large amounts of ingested biomaterial. after removal of the biomaterial the pump ran within specification. | (h3) the investigation into the reported "high purge pressure and pump stop" has been completed since the original report was filed. product was returned for investigation. the device was inspected and found the lt log confirms high purge pressure and a pump stop. during the case the purge pressure recovered and was within specification at the time of the sudden pump stop. a visual inspection of the pump revealed large amounts of biomaterial on and around the impeller. the inspection also revealed that the patient cable was loose in the red handle. the cannula was removed to image the biomaterial present. the pump was then run and purged and the clot was removed from impeller. the high purge pressure was not reproduced and the pump ran within specification. the cause of the high purge pressure was most likely biomaterial. the cause of the pump stop was biomaterial. the cause of the product damage was bond failure at the patient cable to kink protector.
36076 | 1220648-2022-01687 | 15848626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15848626&pc=OZD | 10/26/2022 | 11/22/2022 |  | 11/22/2022 |  |  | 10/26/2022 | 11/22/2022 |  |  | 07/31/2024 |  | 11/22/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023246521 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/26/2022 |  |  |  |  | The complainant reported a 53 year old male patient presenting for impella support. during support, an impella access site bleed developed. as treatment, blood products were delivered. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36077 | 1220648-2022-01688 | 15849753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15849753&pc=OZD | 10/24/2022 | 11/22/2022 |  | 11/22/2022 |  |  | 10/24/2022 | 11/22/2022 |  |  | 07/31/2024 |  | 11/22/2022 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2023236844 |  | 0550-0002 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Hospitalization | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/24/2022 |  |  |  |  | A male patient of unknown age was treated for cardiomyopathy. an impella 5.5 smart assist cardiac pump was used for hemodynamic support, which perforated the left ventricle. after explantation of the impella heart pump, veno-arterial extracorporeal membrane oxygenation (va-ecmo) was implanted. subsequently, the patient was awake and stable. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36078 | 1220648-2022-01683 | 15829706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15829706&pc=OZD | 10/27/2022 | 03/25/2024 |  | 11/19/2022 |  |  | 10/27/2022 | 11/19/2022 |  |  | 07/31/2024 |  | 11/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023254399 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2022 |  |  |  |  | The medical team had an impella 5.5 pump being inserted post bypass surgery (cabg) in the operating suite. the 5.5 pump caused a left ventricle (lv) perforation. as with delivery the pump punctured the weakened tissue at the site of the previous myocardial infarct (mi). the team patched the lv and support proceeded. the support was for just under 11 hours when the patient expired on support. | Additional information was received regarding the patient's passing. it was noted by the physician that the death was primarily attributed to the patient's three day long untreated myocardial infarction leading up to support. however, it was additionally expressed that although the perforation was considered to be controlled, it was believed that the perforation worsened the already poor health condition of the patient. |  |  |  |  | The medical center has not returned the impella pump nor any imaging for analysis. no benchtop testing to root cause is possible. should the product be returned and root cause be determined, a final report will follow. the failure mode will be monitored and trended. | Additional information provided for sections b2, b5, h1, h2, and h6 to address the reported death. should additional relevant information become available, a supplemental report will be submitted.
36079 | 1220648-2022-01679 | 15822681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15822681&pc=OZD | 07/05/2022 | 10/23/2024 |  | 11/18/2022 |  |  | 11/14/2022 | 11/18/2022 |  |  | 03/31/2024 |  | 11/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023202314 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Hospitalization | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/14/2022 |  |  |  |  | The complainant in japan had an impella cp support of 5 days in 7/2022 until wean and explanted. later in 11/2022 the jpvad registry updated that the patient had suffered from an ischemic stroke (cva) during the support. the causal link between impella use and the cva could not be determined or denied. | Additional information reported that the patient expired, and the cause of death was unknown. there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The medical center in japan discarded all impella pump products months ago at the time of explant. no benchtop analysis to root cause is possible. the jpvad registry team furnished no further details on the cva and as such the investigation has concluded. the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01679 was provided in sections b2, b5, d6, h1, and h6.
36080 | 1220648-2022-01681 | 15825258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15825258&pc=OZD | 11/03/2022 | 09/19/2024 |  | 11/18/2022 |  |  | 11/03/2022 | 11/18/2022 |  |  | 07/31/2024 |  | 11/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023255170 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/03/2022 |  |  |  |  | The complainant reported a 69-year-old male patient presenting for cardiomyopathy. during impella support, an increase in the motor current was noted and blood was seen in the hub of the purge bypass needle. the pump subsequently stopped and the decision was made to remove and replace the pump. there was no harm to the patient as a result of the noted issue. | Additional information noted that patient coded and did not survive the procedure. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. |  |  |  |  | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | Following the initial report, the pump was returned for further evaluation. the yellow luer was inspected and it was noted that a crack had formed in the luer from which a fluid leak was confirmed. it was further identified that a thrombus had gotten into the inflow cage. based on the conclusion of the investigation, it was determined that the leak caused low purge pressure which allowed blood the opportunity to ingress into the motor. this led to an increase in the motor current and the pump subsequently stopped. should additional relevant information become available, a supplemental report will be submitted. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01681was provided in sections b1, b2, b5, h1, and h6.
36081 | 1220648-2022-01682 | 15827338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15827338&pc=OZD | 10/12/2022 | 10/11/2023 |  | 11/18/2022 |  |  | 11/07/2022 | 11/18/2022 |  |  | 06/30/2024 |  | 11/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023240269 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemolysis | Ischemia | [11 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2022 |  |  |  |  | The us complainant had an impella 5.5 support ongoing when on day 38 of support the team observed signs of hemolysis and pump in aorta positional alarms. this prompted the team to explant and replace the pump. the team observed the loss of the upper limb pulse due to presumed clot embolization from the impella. the team observed clot adhered to the pump at the surgical explant. the team had vascular surgery intervene to the limb. | It was also reported that during support, there was a leaking purge sidearm. the purge solution was running sodium bicarbonate at the time. retainer was in use. |  |  |  |  | The investigation into the hemolysis that prompted pump explant and the subsequent limb ischemia has concluded. the hemolysis cause was not determined as the pump was not returned for benchtop analysis and hemolysis testing. the cause of the limb ischemia is due to patient condition. the clot developed due to the patient being off of heparin for the underlying bleed from the gastrointestinal tract. the failure modes will be monitored and trended. | The investigation that included that positioning issues and the purge sidearm leak has been completed. the cause of the positioning issue was most likely a false position in aorta alarm due to patient condition since the pump position was confirmed under echo. the cause of the purge sidearm leak was not determined since product was not returned. patient was on sodium bicarbonate. the cause of the limb ischemia was most likely a clot noted in cannula due to patient being off heparin several days earlier due to gi bleed. the cause of the hemolysis was not determined since product was not returned. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
36082 | 1220648-2022-01680 | 15817206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15817206&pc=OZD | 11/04/2022 | 09/20/2024 |  | 11/17/2022 |  |  | 11/04/2022 | 11/17/2022 |  |  | 04/30/2024 |  | 11/17/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023213445 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Vascular Dissection | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/04/2022 |  |  |  |  | The complainant reported an 81-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during removal of the peel-away sheath for impella support, it was noted that the patient developed vascular damage and a hematoma. a stent was placed to achieve hemostasis and a new impella pump was placed to continue with patient support. | The procedural outcome noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01680 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
36083 | 1220648-2022-01676 | 15805876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15805876&pc=OZD | 10/19/2022 | 11/16/2022 |  | 11/16/2022 |  |  | 10/19/2022 | 11/16/2022 |  |  | 05/31/2024 |  | 11/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023217512 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/19/2022 |  |  |  |  | The complainant reported a 70 year old male patient presenting with cardiomyopathy for impella support. just before planned explant on day of support, the pump stopped. the pump was removed and no new impella was used. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36084 | 1220648-2022-01677 | 15808761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15808761&pc=OZD | 11/04/2022 | 11/16/2022 |  | 11/16/2022 |  |  | 11/04/2022 | 11/16/2022 |  |  | 07/31/2024 |  | 11/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023240725 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/04/2022 |  |  |  |  | The complainant reported a 71-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient experienced hemolysis and ventricular tachycardia when the level of support was set to p-9. the patient was subsequently administered haptoglobin to counteract the hemolysis and the level of support was lowered to p-4. support continued with the impella pump following the intervention. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36085 | 1220648-2022-01678 | 15810721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15810721&pc=OZD | 10/22/2022 | 10/26/2024 |  | 11/16/2022 |  |  | 10/22/2022 | 11/16/2022 |  |  | 07/31/2024 |  | 11/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC, |  |  | US |  | 2023254109 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Material Twisted/Bent | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/22/2022 |  |  |  |  | The medical team had an impella cp pump inserted for support of an elderly patient suffering from cgs. the team placed the pump, and had the pump became kinked within the left ventricle. the team had to snare the pump. the pump was snared successfully via the radial artery and the team made choice to explant and place a new pump. | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella cannula bent inside the ventricle and was unable to unbend. the impella pulled into the aorta and was still unable to unbend. it was necessary for a snare to be used to straighten, but at this point physician was concerned about the integrity of the pump. the physician decided to explant the impella and exchanged for another second impella device with no complications. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The medical center has not yet returned the impella pump product for analysis. no benchtop testing is possible to determine why the pump kinked and was snared out of the patient. should the product be returned and root cause determined, a final report will be filed. | Since the original report was filed, the medical center returned the impella pump for analysis and benchtop testing to root cause. the pump was returned and found to have a kink in the cannula and pigtail. theis was due to positioning by the user. the failure mode will be monitored and trended. | B2: the date of the patient's death is unknown. this report is being filed as part of a retrospective review of historical records.
36086 | 1220648-2022-01674 | 15797333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15797333&pc=OZD | 11/02/2022 | 11/15/2022 |  | 11/15/2022 |  |  | 11/02/2022 | 11/15/2022 |  |  | 12/31/2023 |  | 11/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023171662 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/02/2022 |  |  |  |  | The complainant reported a 65 year old male patient presenting for impella support. one day after implant, it was noted that the patient experienced hemolysis due to poor positioning of the pump. haptoglobin was administered to treat the hemolysis and patient support continued with no replacement parts required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36087 | 1220648-2022-01675 | 15797659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15797659&pc=OZD | 10/18/2022 | 11/15/2022 |  | 11/15/2022 |  |  | 10/18/2022 | 11/15/2022 |  |  | 08/31/2024 |  | 11/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2023247375 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  |  |  | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/18/2022 |  |  |  |  | A male patient of unknown age in the eu presented with cardiomyopathy and received support by an impella 5.5 heart pump as bridge to durable lvad. during transportation in operation room the cable between pump and automated impella controller (aic) has been accidentially stuck and torn off from aic, followed by pump stop. a new impella 5.5 pump has been successfully implanted. also, a new aic has been used, as the use of excessive force destroyed the plug that could not be removed from the aic. patient stable afterwards. | [5 empty] | Pump has not been returned for investigation. automated impella controller has been returned for repair / removal of plug fragments. from the information currently available, it can be concluded that the user unintentionally applied excessive force to the medical device. should any new information be returned, a supplemental mdr will be filed.
36088 | 1220648-2022-01673 | 15791203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15791203&pc=OZD | 10/17/2022 | 12/06/2022 |  | 11/14/2022 |  |  | 10/17/2022 | 11/14/2022 |  |  | 05/31/2024 |  | 11/14/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023225562 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2022 |  |  |  |  | The medical team had placed an impella 5.5 support via a direct access to the left ventricle for hemodynamic support. the pump placement signal was unreliable and in an effort to troubleshoot the white cable was unplugged and torque relieved. upon plugging the 5.5 back in the pump failed to restart. the team attempted to troubleshoot but when aic would not restart the pump the team removed and replaced the 5.5. | [5 empty] | The investigation into the impella 5.5 pump stopping, and unable to be restarted, is ongoing. at the conclusion of the investigation a final report with root cause will be submitted. | Since the original report was filed, the investigation into the pump stop has concluded. the product was returned and the benchtop analysis found an open electrical line was the cause of the pump stop. the failure mode will be moniotred and trended.
36089 | 1220648-2022-01671 | 15785693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15785693&pc=OZD | 10/15/2022 | 10/25/2023 |  | 11/12/2022 |  |  | 10/15/2022 | 11/12/2022 |  |  | 06/30/2024 |  | 11/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023232796 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/15/2022 |  |  |  |  | The medical team escalated care from an impella cp to the 5.5 pump for hemodynamic support. the patient suffered an embolic stroke (cva) during the 5.5 support. there is no noted intervention to rectify or treat the harm to the patient. | In addition to the originally reported stroke, thrombus was found on the tip of the device during explant. |  |  |  |  | The investigation into the cva is ongoing at this time. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation into the stoke (cva) has concluded. the pump was returned, as were the data logs. the root cause of the cva could not be determined, nor could the relationship between the stroke and impella use be confirmed. the data logs did not reveal any ingestion of clot that could have caused the ischemic cva. the failure mode will be monitored and trended. | Included additional narrative in b(5) speaking to thrombus found upon explant. the investigation into the reported stroke and thrombus was completed. the device was returned for evaluation. the device was returned cut into pieces and a clot was returned alongside the pump. analysis of the data logs revealed suction alarms and a change in waveform indicating a potential blockage of the cannula or change in volume causing a decrease in flows on 10/14. the cause for this is unknown, as no clot was visualized under transesophageal echocardiogram (tee) just prior to the stroke. the pump operated within specification for the duration of support. based on the available information, the returned evidence cannot confirm relation between the stroke and the device. the root cause of the reported thrombus was most likely due to patient condition. the exact origin of the thrombus is unknown. impella 5.5 ifu ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
36090 | 1220648-2022-01672 | 15785694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15785694&pc=OZD | 10/10/2022 | 10/24/2024 |  | 11/12/2022 |  |  | 10/15/2022 | 11/12/2022 |  |  | 06/30/2024 |  | 11/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023229221 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/15/2022 |  |  |  |  | The medical team in japan had an impella cp support ongoing when there was thrombosis that developed at the pump explant. the team had to intervene and perform peripheral angioplasty/thrombectomy. | Additional information noted that after impella was removed patient experienced lower limb ischemia in both legs and the thrombectomy was performed for both legs. because of thrombectomy, resulted in reperfusion injury and then the patient was died due to this reperfusion. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The medical center in japan discarded all impella product at explant. no benchtop analysis to root cause is possible. the case investigation will be closed without root cause. should more information be shared, that leads to root cause, a final report will be filed. the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information was provided in sections b2, b5, d6, h1, and h6.
36091 | 1220648-2022-01663 | 15777783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15777783&pc=OZD | 10/24/2022 | 11/09/2022 |  | 11/11/2022 |  |  | 10/24/2022 | 11/11/2022 |  |  | 07/31/2024 |  | 11/11/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023257032 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/24/2022 |  |  |  |  | The complainant reported an 80-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, the patient experienced an access site bleed for which a femostop was applied. however, the bleeding continued and the patient was subsequently administered a blood transfusion amounting to 4 units of prbc. impella support continued throughout the reported condition. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36092 | 1220648-2022-01668 | 15773073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15773073&pc=OZD | 10/13/2022 | 11/10/2022 |  | 11/10/2022 |  |  | 10/13/2022 | 11/10/2022 |  |  | 07/31/2024 |  | 11/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023249644 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/13/2022 |  |  |  |  | The complainant reported an 82 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. red urine was noted during support that prompted an attempted repositioning of the device. the issue did not resolve, therefore, the pump was removed and replaced. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36093 | 1220648-2022-01669 | 15775913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15775913&pc=OZD | 10/15/2022 | 11/10/2022 |  | 11/10/2022 |  |  | 10/15/2022 | 11/10/2022 |  |  | 06/30/2024 |  | 11/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023234983 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A 64-year-old male presented with anterior stemi and total left anterior descending (lad) artery occlusion. during revascularization of lad artery per balloon, the patient¿s condition worsened to require impella support. patient unrestful during insertion but otherwise without issue. pci continued to restore flow in lad artery. after pci procedure, a large hematoma at the impella access site was noted. angiogram showed access at the bifurcation of the common femoral artery and the epigastric caused extravasation. impella cp pump had to be removed and balloon tamponade used to achieve hemostasis. patient transferred to another hospital, where pink / red color of urine was observed, which was attributed to hemolysis. escalated to impella 5.5. hemolysis was resolved and patient stable. | [5 empty] | The impella device has not yet been returned to the manufacturer for investigation. should any new information be returned, a supplemental mdr will be filed.
36094 | 1220648-2022-01670 | 15777007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15777007&pc=OZD | 10/14/2022 | 11/10/2022 |  | 11/10/2022 |  |  | 10/14/2022 | 11/10/2022 |  |  | 04/30/2024 |  | 11/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMART ASSIST |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2023215900 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Impella 5.5 has been in the patient for 42,5 days together with a v-v ecmo of an unknown manufacturer. due to increasing hemolysis parameters a new impella 5.5 has been inserted. patient stable. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36095 | 1220648-2022-01664 | 15761604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15761604&pc=OZD | 05/21/2022 | 10/25/2024 |  | 11/09/2022 |  |  | 10/11/2022 | 11/09/2022 |  |  | 12/31/2023 |  | 11/09/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023164175 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Thrombocytopenia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The complainant reported a 73 year old male patient presenting for acute myocardial infarction and cardiogenic shock for impella support. during support, the an access site bleed and thrombocytopenia developed. as treatment, blood products were infused. the patient also exhibited renal failure. as treatment, dialysis was administered. | Additional information noted that patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The cause of the hemolysis was most likely patient condition related, specifically patient volume issues. the cause of the renal failure and its relation to the impella were not determined since sufficient clinical information was not provided; additionally, the hemolysis resolved several days prior to the issue. the cause of the access site bleeding and its relation to the impella were not determined since the product and sufficient clinical information were not provided. the cause of the thrombocytopenia and its relation to the impella were not determined since sufficient clinical information was not provided. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information for the was provided in sections b2, b5, d6, h1, and h6.
36096 | 1220648-2022-01665 | 15763197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15763197&pc=OZD | 01/06/2022 | 05/16/2025 |  | 11/09/2022 |  |  | 10/11/2022 | 11/09/2022 |  |  | 10/31/2023 |  | 11/09/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022138493 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | No | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Distributor |  |  |  |  | 69 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Sepsis | Tachycardia | Thrombocytopenia | [8 empty] | Required Intervention | Death | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 10/11/2022 |  |  |  |  | The complainant reported a 69 year old male patient presenting for impella support. during support, an access site bleed developed which prompted a blood transfusion. the patient also developed thrombocytopenia which was treated with a platelet transfusion. lastly, the patient exhibited renal failure that prompted dialysis as intervention. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | Additional information received from the complainant reported the patient was also septic and had ventricular tachycardia with unknown relationship to the device. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  | The impella device was discarded by the customer and therefore; investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | As noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | No product returned, there fore no investigation conducted.
36097 | 1220648-2022-01666 | 15763830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15763830&pc=OZD | 09/25/2022 | 11/09/2022 |  | 11/09/2022 |  |  | 10/12/2022 | 11/09/2022 |  |  | 02/29/2024 |  | 11/09/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP SET JP |  |  | ABIOMED, INC. |  |  | US |  | 2023191131 |  | 000000000001000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/12/2022 |  |  |  |  | The complainant reported a 60 year old male patient presenting with cardiomyopathy for impella support. during support, a hematoma formed around the access site/graft region. as treatment, a blood transfusion was administered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36098 | 1220648-2022-01667 | 15764038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15764038&pc=OZD | 07/13/2021 | 11/09/2022 |  | 11/09/2022 |  |  | 10/15/2022 | 11/09/2022 |  |  | 01/31/2023 |  | 11/09/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022045961 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/15/2022 |  |  |  |  | The complainant reported a 68 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient exhibited hemolysis. as treatment, haptoglobin was delivered. the patient also developed an access site bleed prompting a blood transfusion. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36099 | 1220648-2022-01658 | 15742879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15742879&pc=OZD | 10/09/2022 | 10/25/2023 |  | 11/07/2022 |  |  | 10/09/2022 | 11/07/2022 |  |  | 06/30/2024 |  | 11/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023234453 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2022 |  |  |  |  | A 52-year-old patient presented with severe multiple mvd / cad and severely depressed left ventricular function. an ecmo and an impella 5.5 heart pump have been inserted to support the patient until a heart transplant or a durable lvad. the impella heart pump stopped and couldn¿t be restarted. a new impella 5.5 pump has successfully been delivered. no harm has been inflicted to the patient. | In addition to the initial report of pump stop, it should be noted that the device's purge side arm was found leaking from the yellow luer connection. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Updated sections b5, d9, g2, g3, g6, h2, h3 & h6: (h3) the investigation into the reported pump stop and purge side arm leak has been completed since the original report was filed. product was returned for investigation. the device was visually inspected and found the lt log shows the purge pressure drop to 0mmhg on october 9th around 2:30am. the purge pressure then increases to 1100mmhg around 11:00am. around this time the motor current increases until a pump stop occurred at 21:23 on 10/9. however, only the purge sidearm was returned. upon inspection of the sidearm, damage to the yellow luer was observed. the cause of the purge leak was sidearm yellow luer damage. following the sidearm yellow luer damage during the time in which the purge flow was close to zero, the purge fluid was unable to protect the motor allowing for blood ingress to occur. a pump stop then occurred. the cause of the pump stop was most likely blood ingress. the cause of the blood ingress was the broken yellow luer. this report is being filed as part of a retrospective review of historical records
36100 | 1220648-2022-01659 | 15744314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15744314&pc=OZD | 09/20/2022 | 11/07/2022 |  | 11/07/2022 |  |  | 10/26/2022 | 11/07/2022 |  |  | 04/30/2024 |  | 11/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 PUMP SET JP |  |  | ABIOMED, INC |  |  | US |  | 2023226875 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/26/2022 |  |  |  |  | The complainant reported a 75-year-old female patient presenting for impella support. during support, the patient experienced an access site bleed for which more than two units of blood were transfused. the bleeding eventually subsided without further intervention and support continued with the implanted impella. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36101 | 1220648-2022-01661 | 15747714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15747714&pc=OZD | 10/12/2022 | 11/07/2022 |  | 11/07/2022 |  |  | 10/12/2022 | 11/07/2022 |  |  | 05/31/2024 |  | 11/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023229087 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/12/2022 |  |  |  |  | The medical center had a surgical patient with an impella cp support. the surgical team cross clamped the aorta and observed the cp to have perforated the lv. the team patched the perforation and the surgical case proceeded, along with the cp for hemodynamic support. | [5 empty] | The medical center has not yet returned the impella pump product for the investigation. should the product be returned and a root cause determined, a final report will be filed. the failure mode will be monitored and trended.
36102 | 1220648-2022-01662 | 15748112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15748112&pc=OZD | 08/24/2022 | 11/07/2022 |  | 11/07/2022 |  |  | 10/13/2022 | 11/07/2022 |  |  | 04/30/2024 |  | 11/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023214687 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Hospitalization | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/13/2022 |  |  |  |  | The medical team had an impella 5.5 support for 6 days when the patient recieved a heart transplant. the transplant and explant occurred in (b)(6) 2022. later in (b)(6) 2022 the local abiomed field representative was notified the patient suffered an ischemic stroke attributed to the use of the impella and it's explant. | [5 empty] | The team discarded all the product at explant and as such there was no benchtop analysis. the root cause has not been determined. the failure mode will be monitored and trended.
36103 | 1220648-2022-01660 | 15748408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15748408&pc=OZD | 07/20/2022 | 11/07/2022 |  | 11/07/2022 |  |  | 10/13/2022 | 11/07/2022 |  |  | 03/31/2024 |  | 11/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023206302 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/13/2022 |  |  |  |  | The medical center was utilizing an impella cp in a stemi dtu ide trial when post explant the patient had ischemia that prompted intervention. the injury was treated by thrombectomy, endarterectomy, and femoral artery patch. | [5 empty] | The medical center discarded all product at the point of explant. no benchtop analysis to root cause is possible. the study team furnished no further clinical details. the failure mode will be monitored and trended.
36104 | 1220648-2022-01657 | 15741076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15741076&pc=OZD | 12/31/2021 | 05/22/2025 |  | 11/05/2022 |  |  | 10/11/2022 | 11/05/2022 |  |  | 07/31/2023 |  | 11/05/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022109131 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Lay User/Patient |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Renal Failure | Sepsis | Thrombocytopenia | [9 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/11/2022 |  |  |  |  | The complainant reported a 56 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient exhibited signs of hemolysis and renal failure. as treatment, a blood transfusion was performed. | Additional information from the complainant reported a pump stop that was successfully restarted. in addition, the patient experienced thrombocytopenia and sepsis. the patient hemodynamically collapsed, resulting in multiple organ failure and ultimately the patient declined and expired while on support. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information for the initial mfr 1220648-2022-01657 was provided. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm wheth er a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ sepsis section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ renal failure and thrombocytopenia section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the pump pause was unable to be determined as no product or logs were returned for analysis. the pump stop occurred on day 10 of support which is beyond the design life of 8 days per design trace matrix 0046-9910. this report is being filed as part of a retrospective review of historical records.
36105 | 1220648-2022-01654 | 15729846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15729846&pc=OZD | 10/06/2022 | 11/03/2022 |  | 11/04/2022 |  |  | 10/06/2022 | 11/04/2022 |  |  | 04/30/2024 |  | 11/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP SET JP |  |  | ABIOMED, INC. |  |  | US |  | 2023200765 |  | 000000000001000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/06/2022 |  |  |  |  | The complainant reported a 71 year old male patient presenting for high risk percutaneous coronary intervention using the impella for hemodynamic support. during support, the patient developed an access site hematoma. as treatment, multiple blood products were delivered. | [5 empty] | The impella device was not returned from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36106 | 1220648-2022-01653 | 15731040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15731040&pc=OZD | 10/01/2022 | 11/03/2022 |  | 11/04/2022 |  |  | 10/27/2022 | 11/04/2022 |  |  | 01/31/2024 |  | 11/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023181685 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/27/2022 |  |  |  |  | The medical team in japan had an impella cp support in 9/2022 of 8 days . the month after the support was completed, the jpvad registry team notified abiomed that there was hemolysis treated by the infusion of blood products. | [5 empty] | The medical center in japan discarded all pump product at explant. no benchtop analysis and hemolysis testing is possible. the jpvad registry did not furnish further clinical details. the root cause can not be determined. the failure mode will be monitored and trended.
36107 | 1220648-2022-01656 | 15738820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15738820&pc=OZD | 09/24/2022 | 10/23/2024 |  | 11/04/2022 |  |  | 10/25/2022 | 11/04/2022 |  |  | 04/30/2024 |  | 11/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023211941 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | Other | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2022 |  |  |  |  | The complainant reported a 46-year-old male patient presenting for impella support. during support, the patient experienced an ischemic stroke, for which treatment information was not provided. the pump was subsequently removed two days following the stroke. | It was reported 10 days after the removal of the impella the patient expired. the patient experienced a cerebral hemorrhage while on impella cp support. the patient was heparinized for impella use at the time and this may have contributed to the cerebrovascular event. the patient later died due to the complications from the cerebral hemorrhage. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | Impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. this report is being filed as part of a retrospective review of historical records. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01656 was provided in sections b2, b5, d6a, d6b, g1, h1, and h6.
36108 | 1220648-2022-01655 | 15721836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15721836&pc=OZD | 05/29/2021 | 11/03/2022 |  | 11/03/2022 |  |  | 11/01/2022 | 11/03/2022 |  |  | 04/30/2023 |  | 11/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022072582 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/01/2022 |  |  |  |  | The medical team had a stemi dtu study patient impella cp support in may of 2021. in (b)(6) 2022 the stemi dtu documentation was reviewed. anemia was attributed to the pump and patient movement. the team infused 1 unit of rbcs. later in (b)(6) 2022 the stemi dtu team further documented a vascular damage of the femoral artery that had been accessed for the impella. the damage had been observed after explant and patient movement. the artery was treated by an atrium stent. | [5 empty] | The medical center discarded all impella product in (b)(6) 2021. no benchtop investigation and analysis to root cause is possible. the investigation will close without root cause. should more information be shared, a final report will follow.
36109 | 1220648-2022-01651 | 15714901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15714901&pc=OZD | 02/03/2022 | 11/02/2022 |  | 11/02/2022 |  |  | 02/03/2022 | 11/02/2022 |  |  | 09/30/2023 |  | 11/02/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022141328 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/03/2022 |  |  |  |  | The complainant reported a 54 year old female patient presenting with acute myocardial infarction and cardiogenic shock. the patient had been transferred from an outside hospital when the impella inserted leg was found ischemic. the pump was removed as a result. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36110 | 1035166-2022-00111 | 15715090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15715090&pc=DYB | 10/01/2022 | 02/24/2023 |  | 11/02/2022 |  |  | 10/05/2022 | 11/02/2022 |  |  | 07/01/2024 |  | 11/02/2022 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | DP-13298 |  | 0052-3021 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 72 YR |  | Female |  |  |  | Injury | Yes | No | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup | [8 empty] | It was reported that the customer was notified of a complaint involving an oscor 23 fr introducer on (b)(6) 2022 during a post-op cardiothoracic surgery. there was a report of a cracked sheath in the right femoral vein upon insertion on behalf of (b)(6) healthcare - billings. a second peel-away sheath was inserted to replace the damaged component. the event was reported under (b)(4) and salesforce case # (b)(4). patient presented with rv failure overnight/through the day. pump insertion for patient support was successful following the peel-away sheath replacement. report of access site bleed, skin sutures and two units of packed red blood cells were transfused. impella support continued throughout the event. the customer discarded the product. reference report # 1035166-2022-00116. | Additional information received 11/01/2022: the cause for the blood leak was unknown. it was not attributed to the damaged sheath. a new sheath was placed and skin sutures were applied to counter the blood leak. due to the pre-existing condition of the patient, the family opted to withdraw impella support when the patient became unstable and the patient passed away. customer unable to provide any additional information regarding patient's death. 1035166-2022-00111 is associated with lot number dp-13298. |  |  |  |  | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Reporter was not sure which lot the cracked device belonged to; therefore, two mdrs were submitted. b5: additional information received. patient death. the device was not returned for analysis, as it was discarded. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. the device was not returned for analysis, therefore, the exact cause of the device issue cannot be determined and the clinical observation could not be confirmed. however, the following controls are in place to mitigate the reported product issue. per procedure (adelante s2s introducer sheath in-process and final inspection): final assembly inspect 100% per procedure visual inspection: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. foreign matter (fm) fiber is limited to 2 particles < 0.10 mm2 or 1 particle < 0.2 mm2 with naked eye at a distance of 12" to 18", verify the sheath hub and split caps are free of excess silicone oil applied between the seals during assembly. verify split caps are properly placed and secured onto sheath hub and free of cracks/damages or excessive adhesive. vacuum leak test. perform pressure and vacuum leak test per procedures. the instructions for use (ifu) informs the user: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36111 | 1035166-2022-00116 | 15715357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15715357&pc=DYB | 10/01/2022 | 02/24/2023 |  | 11/02/2022 |  |  | 10/05/2022 | 11/02/2022 |  |  | 07/01/2024 |  | 11/02/2022 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | DP-13404 |  | 0052-3021 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 72 YR |  | Female |  |  |  | Injury | Yes | No | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup | [8 empty] | It was reported that the customer was notified of a complaint involving an oscor 23 fr introducer on 10/01/2022 during a post-op cardiothoracic surgery. there was a report of a cracked sheath in the right femoral vein upon insertion on behalf of st. vincent healthcare - billings. a second peel-away sheath was inserted to replace the damaged component. the event was reported under (b)(4) and salesforce case #(b)(4). patient presented with rv failure overnight/through the day. pump insertion for patient support was successful following the peel-away sheath replacement. report of access site bleed, skin sutures and two units of packed red blood cells were transfused. impella support continued throughout the event. the customer discarded the product. reference report # 1035166-2022-00111. | Additional information received 11/01/2022: the cause for the blood leak was unknown. it was not attributed to the damaged sheath. a new sheath was placed and skin sutures were applied to counter the blood leak. due to the pre-existing condition of the patient, the family opted to withdraw impella support when the patient became unstable and the patient passed away. customer unable to provide any additional information regarding patient's death. 1035166-2022-00111 is associated with lot number dp-13298. |  |  |  |  | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Reporter was not sure which lot the cracked device belonged to; therefore, two mdrs were submitted. b5: additional information received. patient death. the device was not returned for analysis, as it was discarded. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. the device was not returned for analysis, therefore, the exact cause of the device issue cannot be determined and the clinical observation could not be confirmed. however, the following controls are in place to mitigate the reported product issue. per procedure (adelante s2s introducer sheath in-process and final inspection): final assembly inspect 100% per procedure visual inspection: - with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. foreign matter (fm) fiber is limited to 2 particles < 0.10 mm2 or 1 particle < 0.2 mm2 - with naked eye at a distance of 12" to 18", verify the sheath hub and split caps are free of excess silicone oil applied between the seals during assembly. - verify split caps are properly placed and secured onto sheath hub and free of cracks/damages or excessive adhesive. vacuum leak test perform pressure and vacuum leak test per procedures. the instructions for use (ifu) informs the user: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36112 | 1220648-2022-01652 | 15716002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15716002&pc=OZD | 10/03/2022 | 10/23/2023 |  | 11/02/2022 |  |  | 10/03/2022 | 11/02/2022 |  |  | 05/31/2024 |  | 11/02/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023225670 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/03/2022 |  |  |  |  | The complainant reported a 69 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. towards the end of support, the patient was found to have endured a stroke. the neurological intensivist believes this injury was caused by a clot becoming dislodged from the impella. | The user facility reported that a 69-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced high purge pressure, stroke, and low pump flows. the issue arose while coming off bypass and onto impella support. the motor current (mc) was also reported to be higher than normal for the current p-level. at p-4 and above, the device stopped and alarmed for retrograde flows; however, the surgeon did not want to replace pump and decided instead to initiate tissue plasminogen activator (tpa) protocol via the purge system. the patient decompensated while awaiting the tpa, and an intra-aortic balloon pump (iabp) was subsequently emergently implanted via the left groin. once the tpa was administered through the purge system, pump flows and pressures normalized. upon chest closure, however, the patient began to bleed significantly and the tpa was stopped after approximately 15cc of the 2mg/ 50cc bag had been administered. pump pressures were able to increase to p-8. the left iabp was left in place as a precaution. the patient's activated clotting time (act) was 300 sec. the patient also suffered a stroke which was believed to be caused by the dislodging of a clot from the catheter per the neurological intensivist. no further information was provided. |  |  |  |  | The impella device was withheld by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the reported stroke, thrombus, and low pump flows was completed. the device was not returned for evaluation. based on the available information, the root cause of the high purge pressure, low pump flow, and stroke remain undetermined since device was not returned. impella 5.5 ifu ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
36113 | 1220648-2022-01645 | 15709675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15709675&pc=OZD | 04/05/2022 | 11/01/2022 |  | 11/01/2022 |  |  | 10/03/2022 | 11/01/2022 |  |  | 10/31/2023 |  | 11/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022137404 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/03/2022 |  |  |  |  | The complainant reported a 77 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, an access site bleed developed. as treatment, a blood transfusion was required and the vessel was sutured. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36114 | 1220648-2022-01646 | 15710018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15710018&pc=OZD | 10/03/2022 | 11/01/2022 |  | 11/01/2022 |  |  | 10/03/2022 | 11/01/2022 |  |  | 05/31/2024 |  | 11/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023231645 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/03/2022 |  |  |  |  | The complainant reported a 70 year old male patient presenting for high risk percutaneous coronary intervention. the impella cp optical was selected for hemodynamic support. the physician was unable to advance the supplied 14fr introducer past the patient's existing iliac stents. a vessel perforation had developed and 2 covered stents were required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36115 | 1220648-2022-01647 | 15710143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15710143&pc=OZD | 09/28/2022 | 11/01/2022 |  | 11/01/2022 |  |  | 10/04/2022 | 11/01/2022 |  |  | 04/30/2024 |  | 11/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023218707 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/04/2022 |  |  |  |  | The complainant reported an impella patient developing an access site infection. as a result the pump was removed and antibiotics were given. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36116 | 1220648-2022-01648 | 15710213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15710213&pc=OZD | 07/28/2022 | 11/01/2022 |  | 11/01/2022 |  |  | 10/05/2022 | 11/01/2022 |  |  | 03/31/2024 |  | 11/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023209901 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/05/2022 |  |  |  |  | The complainant reported a 58 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient was suspected as enduring hemolysis. as treatment, a blood transfusion was performed. the issue resolved, thereafter. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36117 | 1220648-2022-01649 | 15712730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15712730&pc=OZD | 09/26/2022 | 10/25/2023 |  | 11/01/2022 |  |  | 09/26/2022 | 11/01/2022 |  |  | 01/31/2024 |  | 11/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023183577 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Pressure Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/26/2022 |  |  |  |  | The medical team had an impella 5.5 support ongoing for a patient suffering from cardiomyopathy. the impella was placed via axillary grafted insertion. there was an access site hematoma that was surgically evacuated. the hematoma developed after a week of support. the pump remains on for support on day 45 at time of the report. the pump has now been used beyond indication. | It was further reported that tissue plasminogen activator was administered to prevent lower purge flow and increased purge pressures. no purge alarms were triggered and all pressures and flow were within normal limits. |  |  |  |  | The impella pump remains on for support and as such no investigation upon the benchtop has yet begun. upon completion of the investigation, a final report will be filed. | The investigation for the high purge pressures has been completed. the data logs were not returned for review. additionally, clinical details do not suggest a cause for the reported high purge pressures. therefore, the root cause of the high purge pressures could not be determined. this report is being filed as part of a retrospective review of historical records. | Since the original report was filed, some of the impella product was returned, but not the introducer or pump. only the purge cassette was returned from the medical center. as such, no benchtop analysis is possible to lead to a root cause of the access site bleeding. the root cause is not determined. the failure mode will be monitored and trended.
36118 | 1220648-2022-01650 | 15712731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15712731&pc=OZD | 09/29/2022 | 11/01/2022 |  | 11/01/2022 |  |  | 10/07/2022 | 11/01/2022 |  |  | 03/31/2024 |  | 11/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023198230 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/07/2022 |  |  |  |  | The medical team in japan had an impella cp inserted via the femoral in an obese patient. the patient's anatomy did not allow the reposheath to fully deploy down the tissue track to prevent bleeding. there was blood loss. the team explanted and ecmo was placed. more than 2 units of rbcs were given. | [5 empty] | The medical center in japan discarded the pump product at explant. no benchtop analysis is possible. the medical team did relay the clinical update of the patient being obese with a bmi of 30-35 which could contribute to the inability to achieve hemostasis and cause the bleed at the femoral access site. the failure mode will be monitored and trended.
36119 | 1220648-2022-01644 | 15698905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15698905&pc=OZD | 10/03/2022 | 11/16/2022 |  | 10/31/2022 |  |  | 10/03/2022 | 10/31/2022 |  |  | 05/31/2024 |  | 10/31/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023217508 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/03/2022 |  |  |  |  | The medical team had an impella 5.5 supporting a cardiac surgery patient, which on day 13 stopped repeatedly. the team attempted to troubleshoot but, when issue did not resolve and pump restart appropriately, the pump was explanted. the pump has not been replaced. | [5 empty] | The investigation into the pump stop is on-going at this time. upon the completion of the investigation, a final report will be filed. | Since the original report was filed the investigation into the pump stop has completed. the pump and data logs were analyzed. blood ingress was found to have occurred due to misuse. the misuse, thought to be have occurred when sutures were applied caused a hole which allowed for the blood ingress and following pump stop. the failure mode will be monitored and trended.
36120 | 1220648-2022-01642 | 15691006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15691006&pc=OZD | 07/04/2022 | 10/28/2022 |  | 10/28/2022 |  |  | 07/05/2022 | 10/28/2022 |  |  | 02/29/2024 |  | 10/28/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023202739 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Use of Device Problem | [6 empty] | Low Blood Pressure/ Hypotension | Decreased Respiratory Rate | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/05/2022 |  |  |  |  | A 72 year old male patient presented with acute myocardial infarction and cardiogenic shock for impella support. during support, an alarm sounded which prompted attention from the medical staff. the patient's blood pressure had decreased, lips were blue, and respiration was noted as "agonal." the physician found the position of the catheter was more shallow than when originally placed. further examination found the patient without a pulse and pronounced dead. the touhy borst was noted as loose, the impella was not sutured to the thigh, and a knee immobilizer was not being used. the account reported no malfunction of the touhy borst valve as they deemed it functioning properly when testing, post-mortem. the abiomed representative was notified the following day and training was conducted. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. the user stated that the touhy borst valve was functioning as intended, however, was not tightened. tightening the touhy borst prevents unwanted movement. the following labeling details proper technique to ensure the catheter does not migrate during use: section 3, pg. 25 (use of suture pads): "the statlock compatible suture pads help secure the repositioning sheath to the patient¿s skin." section 5, pg 79 (use of repositioning sheath): "5. secure the repositioning unit to the patient with the blue suture pads or a statlock stabilization device." "7. attach the anti contamination sleeve to the blue section of the repositioning sheath. lock the anchoring ring in place by turning it clockwise. secure the catheter shaft in place by tightening the connected anchoring ring." section 7, pg. 179 (use of knee immobilizer): "use knee immobilizer as needed to maintain access site straight." section 7, pg. 183 (correct placement of catheter): "for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. "
36121 | 1220648-2022-01641 | 15691112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15691112&pc=OZD | 09/15/2022 | 09/23/2025 |  | 10/28/2022 |  |  | 09/29/2022 | 10/28/2022 |  |  | 05/31/2024 |  | 10/28/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224954 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | Death | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/29/2022 |  |  |  |  | The complainant reported a 66 year old male patient presenting for impella support. during support, the patient was suspected as developing hemolysis and subsequent renal injury. as treatment, haptoglobin and dialysis were administered. | Additional information noted that the procedure outcome was that patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01641 was provided in sections b2, b5, h1, and h6.
36122 | 1220648-2022-01643 | 15693287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15693287&pc=OZD | 09/29/2022 | 10/28/2022 |  | 10/28/2022 |  |  | 10/01/2022 | 10/28/2022 |  |  | 05/31/2024 |  | 10/28/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023222453 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/01/2022 |  |  |  |  | The complainant reported a 71 year old patient presenting for high risk percutaneous coronary intervention, using the impella cp smart assist for hemodynamic support. during support, the patient was suspected as enduring hemolysis. the device was prematurely removed and patient was escalated to an impella 5.5, resolving the issue. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36123 | 1220648-2022-01635 | 15680093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15680093&pc=OZD | 10/15/2022 | 10/27/2022 |  | 10/27/2022 |  |  | 10/15/2022 | 10/27/2022 |  |  | 05/31/2023 |  | 10/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2022089490 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/15/2022 |  |  |  |  | The complainant reported a 72-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient experienced an access site bleed resulting in a hematoma. to treat the bleed, the patient received a blood transfusion of unknown quantity and a skin suture was placed. impella support continued throughout the event. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36124 | 1220648-2022-01636 | 15686786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15686786&pc=OZD | 09/27/2022 | 12/01/2022 |  | 10/27/2022 |  |  | 09/27/2022 | 10/27/2022 |  |  | 02/29/2024 |  | 10/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023187497 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2022 |  |  |  |  | The complainant reported a 48 year old male patient presenting for impella support. at the conclusion of support, the device separated upon explant. the cannula "came apart at the motor" and required use of hemostats to retrieve from the graft. | [5 empty] | The device has been received and an investigation is underway. upon completion, a supplemental mdr will be filed. | The device was returned and visual inspection found kinks in the cannula and a horizontal crease. evidence suggests the detachment was use related (excessive force used).
36125 | 1220648-2022-01637 | 15686790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15686790&pc=OZD | 10/19/2022 | 10/27/2022 |  | 10/27/2022 |  |  | 10/19/2022 | 10/27/2022 |  |  | 03/31/2024 |  | 10/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023202267 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/19/2022 |  |  |  |  | The medical team in japan had a successful impella cp support in 9/2022. the jpvad registry team later noted in 10/2022 that there was hemolysis during the support. the hemolysis was observed on day 2 of the support and was treated by blood product infusion. the pump remained on for 5 days beyond that without any other treatment given. | [5 empty] | The medical center discarded all product at the time of explant. no benchtop analysis and hemolysis testing is possible. the jpvad registry furnished no further clinical details. no root cause has been determined. the failure mode will be monitored and trended. a final report will be filed should new information be shared.
36126 | 1220648-2022-01638 | 15686791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15686791&pc=PYX | 10/01/2022 | 09/25/2024 |  | 10/27/2022 |  |  | 10/01/2022 | 10/27/2022 |  |  | 05/31/2023 |  | 10/27/2022 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP PUMP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC |  |  | US |  | 2022103186 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  |  | White | Death | Yes | Yes | Defective Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/01/2022 |  |  |  |  | The complainant reported a 72-year-old female patient presenting for post-op cardiothoracic surgery. during impella support, it was noted that the patient experienced an access site bleed. to treat the bleed, additional skin sutures were placed and two units of packed red blood cells were transfused. impella support continued throughout the event. | Additional information provided indicated the 23 fr peel away sheath was found to be cracked inside the right femoral vein. in addition, it was noted the family withdrew care and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01638 was provided.
36127 | 1220648-2022-01639 | 15686793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15686793&pc=OZD | 10/17/2022 | 10/27/2022 |  | 10/27/2022 |  |  | 10/17/2022 | 10/27/2022 |  |  | 04/30/2023 |  | 10/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2022079811 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/17/2022 |  |  |  |  | The complainant reported a 51-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient experienced an access site bleed resulting in a hematoma. to treat the bleed, the patient received a blood transfusion of unknown quantity and a vascular suture was placed. impella support continued throughout the event. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36128 | 1220648-2022-01640 | 15686795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15686795&pc=OZD | 10/17/2022 | 10/27/2022 |  | 10/27/2022 |  |  | 10/17/2022 | 10/27/2022 |  |  | 03/31/2024 |  | 10/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023212644 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 10/17/2022 |  |  |  |  | The complainant reported an 87-year-old female patient presenting for high risk percutaneous coronary intervention. upon insertion of a 14 fr peel-away sheath for impella implant, the patient experienced total limb occlusion. the sheath was removed as a result of the condition and blood flow subsequently returned to the patient¿s limb. no further attempts to implant the patient were performed as a result of the patient¿s anatomy. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36129 | 1220648-2022-01599 | 15672640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15672640&pc=OZD | 09/27/2021 | 09/30/2022 |  | 10/26/2022 |  |  | 09/27/2022 | 10/26/2022 |  |  | 08/31/2023 |  | 10/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2022118860 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/27/2022 |  |  |  |  | The complainant reported an 80-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient experienced a vascular injury that required unspecified treatment. additional information pertaining to the injury and subsequent intervention was not available. the patient was considered stable following treatment. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36130 | 1220648-2022-01631 | 15673026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15673026&pc=OZD | 10/04/2021 | 10/26/2022 |  | 10/26/2022 |  |  | 10/15/2022 | 10/26/2022 |  |  | 05/31/2023 |  | 10/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022089495 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/15/2022 |  |  |  |  | The medical team in japan had an impella cp that was successfully supporting the patient in 8/2021. the team then filed the jpvad registry document for limb ischemia in 2/2022. there was not any noted intervention at that time. later in 10/2022 the jpvad registry further noted an antegrade sheath was placed for perfusion for the impella accessed limb. | [5 empty] | The medical center discarded the pump last year. there can be no benchtop analysis. the jpvad team furnished no further details regarding the ischemia. no root cause can be determined. should more information be shared, a final report will be filed.
36131 | 1220648-2022-01633 | 15673063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15673063&pc=OZD | 08/14/2022 | 10/25/2024 |  | 10/26/2022 |  |  | 10/11/2022 | 10/26/2022 |  |  | 01/31/2024 |  | 10/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023190998 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | Thrombocytopenia | [10 empty] | Required Intervention | Hospitalization | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The medical center in japan in 8/2022 had a successful impella cp support of 14 days. the pump was then explanted as the support was no longer needed. later in 10/2022 the jpvad registry noted hemolysis and renal failure occurring. to treat the renal failure the team initiated dialysis. later the patient expired from underlying condition. | Additional information noted that the patient experienced thrombocytopenia and the procedure outcome was that patient expired, medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The medical team in japan discarded all impella product at explant. no benchtop analysis is possible. the jpvad registry team shared no further details. no root cause can be determined. should more information be shared that leads to root cause, a final report will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information was provided in sections b2, b5,d6, h1, and h6.
36132 | 1220648-2022-01634 | 15673082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15673082&pc=OZD | 04/23/2022 | 10/26/2022 |  | 10/26/2022 |  |  | 10/11/2022 | 10/26/2022 |  |  | 12/31/2023 |  | 10/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023164177 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/11/2022 |  |  |  |  | The medical team in japan had an impella cp support in 4/2022. the support was successful for 5 days until weaned and explanted. later in 10/2022 the jpvad registry team noted an access site bleed and thrombocytopenia attributed to the use of impella. the access site bleed was treated by compression and blood products, and the thrombocytopenia was treated by blood product infused. | [5 empty] | The medical center in japan had discarded all product. no benchtop analysis is possible. the jpvad registry shared no further details that could lead to root cause. the investigation will close without root cause. should information be shared later, that leads to root cause, a supplemental report will be filed.
36133 | 1220648-2022-01632 | 15673208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15673208&pc=OZD | 09/29/2022 | 10/26/2022 |  | 10/26/2022 |  |  | 09/29/2022 | 10/26/2022 |  |  | 12/31/2023 |  | 10/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023176793 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/29/2022 |  |  |  |  | The medical team was placing an impella 5.5 via the graft tunneled sub-xiphoid wirelessly. the pump was delivered across the valve and caused an lv perforation when the team was repositioning the pump. the perforation was surgically closed and the pump was explanted. | [5 empty] | The medical center discarded all impella pump product at the explant in the surgical suite. no imaging or other clinical details have been shared. should a root cause be determined, a final report will be filed at the conclusion of the investigation. the failure mode will be monitored and trended.
36134 | 1220648-2022-01627 | 15668168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15668168&pc=OZD | 09/14/2022 | 12/05/2022 |  | 10/25/2022 |  |  | 09/28/2022 | 10/25/2022 |  |  | 04/30/2024 |  | 10/25/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023211819 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2022 |  |  |  |  | The complainant reported a patient of unknown age/gender presenting with acute myocardial infarction and cardiogenic shock. the patient was suspected as enduring hemolysis during support, therefore, the pump was removed prematurely and the patient was placed on an intra-aortic balloon pump. | [5 empty] | The device has been received and an investigation is underway. upon completion, a supplemental mdr will be filed. | The pump was received and an investigation was performed. while no noted damage or abnormalities were confirmed regarding the returned product, we are unable to definitively determine a root cause as no data logs were provided.
36135 | 1220648-2022-01629 | 15668181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15668181&pc=OZD | 07/01/2022 | 10/25/2022 |  | 10/25/2022 |  |  | 09/26/2022 | 10/25/2022 |  |  | 12/31/2023 |  | 10/25/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023171157 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 09/26/2022 |  |  |  |  | The complainant reported a 77 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient endured an access site bleed. as treatment, a blood transfusion was given. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36136 | 1220648-2022-01628 | 15668426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15668426&pc=OZD | 07/09/2022 | 10/25/2022 |  | 10/25/2022 |  |  | 09/26/2022 | 10/25/2022 |  |  | 02/29/2024 |  | 10/25/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023193779 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/26/2022 |  |  |  |  | The complainant reported a 51 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, an access site bleed formed. as treatment, a blood transfusion was required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36137 | 1220648-2022-01630 | 15670361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15670361&pc=OZD | 10/11/2022 | 12/02/2022 |  | 10/25/2022 |  |  | 10/11/2022 | 10/25/2022 |  |  | 02/29/2024 |  | 10/25/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC |  |  | US |  | 2023191194 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The complainant reported a 79-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. following direct cardioversion during impella support, a rise in motor current was observed and the pump subsequently stopped. attempts to restart the device with a new aic failed and the decision was made to remove the pump. patient support was maintained via ecmo and an intra-aortic balloon pump. there was no harm to the patient as a result of the pump stop. | [5 empty] | The impella device was returned and is pending evaluation. upon completion of the evaluation, a supplemental mdr will be filed. | Following the initial report, the pump was returned for further evaluation. a review of the data logs verified that the motor current spiked resulting in a pump stop. a visual inspection of the returned pump discovered that a thrombus was wrapped around the impeller. upon conclusion of the evaluation, the pump stop was confirmed and the cause was traced to biomaterial in the impeller. should additional relevant information become available, a follow up report will be submitted.
36138 | 15645265 | 15645265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15645265&pc=OZD | 07/30/2022 | 08/12/2022 | 09/29/2022 | 10/21/2022 | 09/29/2022 | 10/20/2022 |  | 10/21/2022 |  |  |  |  | 10/21/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No answer provided |  |  | Not Available | [6 empty] | User Facility report | No | Risk Manager | [8 empty] | Female |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella allegedly unable to purge properly.
36139 | 15645301 | 15645301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15645301&pc=OZD | 01/02/2022 | 09/30/2022 | 09/30/2022 | 10/21/2022 | 09/30/2022 | 10/20/2022 |  | 10/21/2022 |  |  |  |  | 10/21/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 18980 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Infuse | [6 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Pt with an impella 5.5, awaiting a heart transplant, the impella alarmed that the purge flow was blocked. please refer to nursing notes that delineate the measures that were taken to resolve this problem but were unsuccessful and ultimately led to a complete pump failure.
36140 | 1220648-2022-01625 | 15654204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15654204&pc=PYX | 09/20/2022 | 09/24/2024 |  | 10/21/2022 |  |  | 09/20/2022 | 10/21/2022 |  |  | 07/31/2023 |  | 10/21/2022 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP PUMP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC |  |  | US |  | 2022122212 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/20/2022 |  |  |  |  | The complainant reported a 62-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was documented that the patient had an unknown source of bleeding for which four units of packed red blood cells were required. at the time of the bleed, the patient's hemoglobin dropped from 9.4 to 6.5. the cause for the bleed could not be determined and impella support continued. | It was reported care was withdrawn, and the patient expired within 10 minutes of withdrawal of care. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | Upon further review of the provided information, it has been determined that there is currently no evidence or allegation of a causal relationship between the impella pump and the report of an unknown source of bleeding for which a blood transfusion was required. it was further documented that the access site was clean, dry, and intact, at the time of the noted event. as such, this case will be deemed not reportable. should additional information become available to support an allegation or malfunction of the impella pump in lines with a reportable event, a supplemental report will be submitted. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01625 was provided in sections b2, b5, d6a, d6b, h1, and h6.
36141 | 1220648-2022-01626 | 15654721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15654721&pc=OZD | 09/22/2022 | 11/22/2022 |  | 10/21/2022 |  |  | 09/22/2022 | 10/21/2022 |  |  | 06/30/2024 |  | 10/21/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023240250 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fracture | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2022 |  |  |  |  | The complainant reported a 44 year old male patient presenting with cardiomyopathy for impella support. the patient developed signs of hemolysis, shortly after pump initiation. the pump was removed as a result. the product was returned and a visual inspection of the pump revealed damage to the impeller. | [5 empty] | The impella device was received and an investigation is underway. once our analysis is complete, a supplemental mdr will be filed. | The device was returned and an analysis was performed. fractured impeller blades were identified. review of the clinical account found the physician used a 0.035" guidewire, which is not recommended for use with the impella device. we believe the root cause of the reported hemolysis was related to the fractured impeller blades.
36142 | 1220648-2022-01622 | 15627875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15627875&pc=OZD | 09/21/2022 | 08/17/2023 |  | 10/18/2022 |  |  | 09/21/2022 | 10/18/2022 |  |  | 07/31/2024 |  | 10/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023238319 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Detachment of Device or Device Component | Separation Problem | [5 empty] | Perforation of Vessels | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient of unknown age/gender presenting for impella support. during explantation of the device, at the conclusion of support, the pump broke and components remained in the patient. intervention was attempted to remove the separated components on the following day, however, the patient expired due to vessel rupture. | [5 empty] | The device was sequestered by law enforcement for investigation. the defective pump has been requested for return so that we can complete an investigation. a supplemental mdr shall be filed should the device be returned by law enforcement. | The investigation into the pump damage/detachment has been completed since the original report was filed. the benchtop analysis revealed that the root cause of the product damage was a detached inflow cage most likely as a result of a bond failure.
36143 | 1220648-2022-01624 | 15629497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15629497&pc=OZD | 10/14/2022 | 10/18/2022 |  | 10/18/2022 |  |  | 10/14/2022 | 10/18/2022 |  |  | 04/30/2024 |  | 10/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023218567 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/14/2022 |  |  |  |  | The complainant reported a 72-year-old male patient presenting for high risk percutaneous coronary intervention. during impella support, there was a spike in the motor current and the pump subsequently stopped. the pump was removed and replaced with a new impella device. the explanted pump was returned for evaluation, and it was discovered that the impeller blades were broken. | [5 empty] | Following the report of a pump stop, the impella device was returned by the customer for further evaluation. upon analysis of the event logs, it was discovered that there was a spike in motor current which was followed by a low flow. a visual inspection was performed and it was noted that the impeller blades were broken. a cause for the damage to the blades could not be definitively established as fluoroscopic imaging and/or additional clinical details were not provided. should any new information be returned, a supplemental mdr will be filed.
36144 | 1220648-2022-01623 | 15629518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15629518&pc=OZD | 09/20/2022 | 10/21/2022 |  | 10/18/2022 |  |  | 09/20/2022 | 10/18/2022 |  |  | 02/29/2024 |  | 10/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023191573 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2022 |  |  |  |  | The medical center was placing an impella 5.5 for support of a surgical patient. the patient was undergoing coronary artery bypass surgery and was unable to be weaned off of bypass. the impella pump stopped repeatedly. the team made decision to replace the pump and have support proceed on 2nd pump. no lasting harm or injury from delay. | [5 empty] | The investigation into the pump stop is on-going at this time. upon completion, a final report will be filed. | Since the original report was filed the investigation into the pump failing has concluded. the pump was run upon the benchtop and the failure was not able to be reproduced. the pump operated to specification. the cause was determined to be technique/use related. the failure mode will be monitored and trended.
36145 | MW5112733 | 15633053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15633053&pc=DYB | 10/13/2022 | 10/17/2022 |  | 10/18/2022 |  |  |  | 10/18/2022 |  |  |  |  | 10/18/2022 |  | INTRODUCER, CATHETER |  |  | IMPELLA CP |  |  | DYB | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 72 YR |  | Male | 163 KG | Non Hispanic | White | Injury | No | Yes | Therapeutic or Diagnostic Output Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | Impella inserted the optical which gives an ao pressure malfunctioned. a new impella replaced it. vender confirmed this was a catheter issue and not a console issue. when they used a second catheter and plugged into console, no problems identified. fda safety report id# (b)(4).
36146 | 1220648-2022-01621 | 15606481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15606481&pc=OZD | 09/14/2022 | 10/14/2022 |  | 10/14/2022 |  |  | 09/14/2022 | 10/14/2022 |  |  | 04/30/2024 |  | 10/14/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023212015 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Difficult to Insert | Difficult to Advance | [5 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/14/2022 |  |  |  |  | The complainant reported a 61-year-old male patient presenting for cardiomyopathy. during the initiation of impella support, a delivery issue was documented involving an impella 5.5 pump. after the pump failed to advance across the anastomosis site, the decision was made to load an rp sheath over a .18 wire. the pump insertion was subsequently successful, however, the anastomosis graft site tore upon insertion. blood products were provided and the anastomosis site was repaired. following the repair, it was noted that the patient's left hand was cold with no pulse. to treat the resulting ischemia, a portion of the anastomosis graft was excised. it was further identified that the multiple insertion attempts had damaged the vessel attached to the graft. a patch was placed over the anastomosis site and a slit was created for the impella to sit in. the site was sutured and patient support continued. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36147 | 1220648-2022-01620 | 15600561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15600561&pc=OZD | 09/14/2022 | 11/14/2022 |  | 10/13/2022 |  |  | 09/14/2022 | 10/13/2022 |  |  | 03/31/2024 |  | 10/13/2022 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023205564 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 09/14/2022 |  |  |  |  | The medical center in germany placed an impella cp for hemodynamic support and found it's access site to have a bleed that lead to a hematoma in the groin. the physician suspected patient obesity being a reason for the repositioning sheath being not deep enough in the vessel. to remedy the issue, a hemostop was applied, 4 units of blood products transfused, protamin was given, and surgical clearing out of the hematoma was performed. additionally the pump was explanted. at the conclusion of the interventions, the patient was noted to be stable. | [5 empty] | The impella device was discarded by the customer and therefore; investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The cause of the bleeding was due to the patient's condition. best practices were followed and no use-related problem was reported. review of the manufacturing documentation did not reveal any abnormalities. the product was in compliance with the specification. incidence rate 0.05% in the last 12 months.
36148 | 15592548 | 15592548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15592548&pc=OZD | 09/11/2022 | 09/26/2022 | 09/26/2022 | 10/12/2022 | 09/26/2022 | 10/04/2022 |  | 10/12/2022 |  |  |  |  | 10/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Nurse | [6 empty] | 10220 DA |  | Male |  |  |  | Malfunction | No | Yes | Therapeutic or Diagnostic Output Failure | [6 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y |  |  |  | Impella waveform to indicate placement went flat--sensor malfunction, but an echo showed that the device was in the correct place. and, there was good flow. patient moved to cardiothoracic icu (cticu) for further monitoring. confirmed sensor failed.
36149 | 1220648-2022-01616 | 15581049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15581049&pc=OZD | 10/07/2022 | 12/02/2022 |  | 10/11/2022 |  |  | 10/08/2022 | 10/11/2022 |  |  | 04/30/2024 |  | 10/11/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC |  |  | US |  | 2023243443 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | Company representation |  |  | 53 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/08/2022 |  |  |  |  | The complainant reported a 53-year-old male patient presenting for off-pump coronary artery bypass surgery. one day after implant, it was noted that the purge pressure increased and the pump subsequently stopped. the pump was removed and the patient was placed on ecmo for continued support. there was no harm to the patient as a result of the pump stop. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | Following the initial report, the pump was returned for further evaluation. testing of the device successfully replicated the reported failure. a visual inspection found that the catheter to kink protector bond had been compromised due to an inadequate amount of adhesive. this inadequacy allowed the catheter to twist freely. as a result, the purge line was found kinked/coiled in the red handle and subsequently caused the high purge pressure. as the purge flow was then inhibited, blood was able to enter the purge line and stopped the pump. should additional relevant information become available, a follow up report will be submitted.
36150 | 1220648-2022-01615 | 15581050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15581050&pc=OZD | 02/08/2022 | 09/25/2025 |  | 10/11/2022 |  |  | 09/27/2022 | 10/11/2022 |  |  | 10/31/2023 |  | 10/11/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2022144074 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | Aortic Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/27/2022 |  |  |  |  | The complainant reported an 81-year-old male patient presenting for myocarditis. upon being placed on impella support, it was noted that the patient developed worsening aortic regurgitation. after repositioning of the pump failed to relieve the issue, ecmo was initiated and the condition stabilized. | During support, the patient also experienced thrombus, the resolution for this issue is unknown. additionally, it was reported that the patient died from intestinal necrosis. no further information is available. there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | This report is being filed as part of a retrospective review of historical records.
36151 | 1220648-2022-01617 | 15581087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15581087&pc=OZD | 09/14/2022 | 10/26/2022 |  | 10/11/2022 |  |  | 09/14/2022 | 10/11/2022 |  |  | 03/31/2024 |  | 10/11/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023203557 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2022 |  |  |  |  | The medical center had an impella 5.5 pump support ongoing when there was hemolysis. the labs were hemolyzing but the pump was in appropriate position. dialysis was initiated as an intervention. the pump remained on until lvad placed. | [5 empty] | The investigation is on-going at this time. upon completion of the investigation and hemolysis testing of the pump, a followup report will be filed. | Since the original report was filed, the investigation into the hemolysis has concluded. the pump was returned and investigated and run on the benchtop. the pump passed hemolysis testing. there was no damage or malfunction noted on the pump. the hemolysis root cause was determined to be patient condition. the failure mode will be monitored and trended.
36152 | 1220648-2022-01619 | 15581141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15581141&pc=OZD | 10/07/2022 | 10/11/2022 |  | 10/11/2022 |  |  | 10/07/2022 | 10/11/2022 |  |  | 05/31/2024 |  | 10/11/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023225055 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/07/2022 |  |  |  |  | The complainant reported a 74-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, hemolysis was identified in the patient. to treat the condition, the patient was placed on ecmo and was administered haptoglobin. furthermore, the p-level of the pump was reduced from p8 to p2. the hemolysis subsequently resolved and patient support continued. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36153 | 1220648-2022-01618 | 15581156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15581156&pc=OZD | 09/16/2022 | 10/13/2023 |  | 10/11/2022 |  |  | 09/16/2022 | 10/11/2022 |  |  | 03/31/2024 |  | 10/11/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023209666 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 12 YR |  | Female |  |  |  | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Vascular Dissection | [12 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/16/2022 |  |  |  |  | The medical center was utilizing an impella 5.5 for support of a pediatric patient with multiple cardiac comorbidities. there was a purge leak that was bypassed. the patient had been mobile and walking for physical therapy. the team made choice to replace the pump and found at the removal that the pump broke to 2 pieces. at the removal of the pump the artery was dissected and was treated by covered stenting. | [5 empty] | The medical center returned the impella pump product for analysis and investigation to root cause of the detachment. the catheter was found to have multiple kinks. the cause of the detachment was determined to be excessive force by the operator. the failure mode will be monitored and trended. | The investigation into the reported 'purge leak - pump, cannula/pigtail detachment and vascular damage - peripheral vessel" has been completed since the original report was filed. product was returned for investigation. the device was visually inspected and found that the catheter was observed to be detached from the motor with kinks in the catheter and the sidearm was detached from the red handle tubing. the sidearm was not returned for review. biomaterial observed in the outflow. the cause of the purge leak was most likely damage to the filter to tubing joint due to patient movement. the cause of the vascular damage was most likely excessive force leading to the catheter detachment since kinks in the catheter were observed.
36154 | 1220648-2022-01613 | 15562259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15562259&pc=OZD | 09/27/2022 | 07/09/2025 |  | 10/07/2022 |  |  | 09/27/2022 | 10/07/2022 |  |  | 05/31/2024 |  | 10/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023224958 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/27/2022 |  |  |  |  | The complainant reported a 60-year-old male patient presenting for cardiomyopathy. during impella support, the patient experienced an access site bleed leading to a hematoma. adjusting the angle of the repositioning sheath slowed the bleed, but the patient required more than two unit of prbc to counteract the blood loss. the following day, it was noted that hemolysis was also observed in the patient. two additional units of prbc were transfused to treat the hemolysis. the device was not removed as a result of the noted conditions and patient support continued. | It was reported the patient expired from multiple organ failure while on impella support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01613 was provided in sections b2, b5, d6a, d6b, h1, and h6. | The root cause of the hemolysis was determined to be most likely patient condition as there was insufficient volume in the left ventricle. the root cause of the access site bleeding was most likely use related as adjusting the angle of the repo resolved the bleeding. the root cause of the patient death was determined to be patient condition and unrelated to the use of impella. instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction . ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ b2 required intervention was inadvertently omitted on manufacturer device report number 1220648-2022-01613-1. d1 brand name was erroneously entered on the initial manufacturer device report number 1220648-2022-01613. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2022-01613. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2022-01613. g1 reporting contact first name, last name, and email revised on manufacturer device report 1220648-2022-01613 in accordance with updated procedures. h6 investigation findings and investigation conclusions was erroneously entered on the initial manufacturer device report number 1220648-2022-01613. h10 investigation results was inadvertently omitted on manufacturer device report number 1220648-2022-01613. h10 instructions for use were inadvertently omitted on manufacturer device report number 1220648-2022-01613.
36155 | 1220648-2022-01614 | 15567685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15567685&pc=OZD | 09/10/2022 | 10/14/2024 |  | 10/07/2022 |  |  | 09/10/2022 | 10/07/2022 |  |  | 05/31/2024 |  | 10/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023229069 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/10/2022 |  |  |  |  | The complainant reported a 67 year old female patient presenting for impella support. during support, the artery was found occluded under angiogram. the pump was removed in response. | It was also reported the patient experienced pump stop during support. several attempts to restart the device occurred which were unsuccessful. the power cable was removed from aic and plugged back in and the pump was eventually restarted with suction alarms and placement signal low alarms. the device position was confirmed; however, the decision was made to remove the pump. |  |  |  |  | The impella device was received and an investigation is underway into the reported issue. a supplemental mdr will be filed at the conclusion of our analysis. | The pump was returned and an investigation was conducted. visual inspection of the device found no abnormalities observed. we feel the root cause of the ischemia was related to the patient's condition. | Additional information for the initial mfr was provided in sections b1, b5, d6a, d6b and h6. the impella device was returned for evaluation. upon visual inspection, no abnormalities were observed. the cause of the pump stop was use related since the pump stop was unable to be reproduced and resolved by unplugging and plugging cable back in. the cause of the optical signal issue was most likely patient condition related since the issue did not persist and was unable to be reproduced. the cause of the ischemia was most likely patient condition related. as noted in the impella ifu: impella cp with smartassist system (pump stop) section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ limb ischemia section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
36156 | 1220648-2022-01612 | 15556949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15556949&pc=OZD | 10/10/2022 | 11/27/2024 |  | 10/06/2022 |  |  | 09/08/2022 | 10/06/2022 |  |  | 05/31/2023 |  | 10/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFTHEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 IMC PUMPEN SET EU |  |  | ABIOMED, INC. |  |  | US |  | 2022085043 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 24 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2022 |  |  |  |  | During an impella 5.0 support, according to the surgeon an injury likely related to 0.018" guidewire occurred during difficult implant in challenging anatomy, and imaging issues/failure during implantation. ifu requires fluoroscopy for placement, with the guidewire reliably observed at all times. small dissection in brachiocephalic. the impella was planned for removal/replacement with impella 5.5, however due to dissection, the pump will remain in longer than planned in order to monitor the injury. during explant, additional intervention is expected. | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support as a bridge to transplant. additional information reported that there was a purge system open alarm and leaking at the clear side arm, proximal to the bacterial filter. the team was advised of the warning signs of possible pump failure if the motor current increased beyond the motor current of 840/760, which at the time had not changed. the heart team was consulted, and the decision was made to continue support with the impella 5.0 with plans to emergently exchange the pump under venoarterial extracorporeal membrane oxygenation if the pump began to fail or stop. the patient remained on impella 5.0 support and was successfully weaned. |  |  |  |  | The impella device is still inside the patient and therefore could not be returned yet. investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the purge leak and vascular damage has been completed. the impella 5.0 was not received from the customer and therefore, an evaluation of the device by the manufacturer was not possible. the root cause of the vascular damage was determined to be patient condition as the patient had a challenging anatomy making it difficult to place the guidewire. the root cause of the purge leak was unable to be determined as no product was returned for investigation. the leak occurred on day 78 of support which is beyond the 28 day design life per design trace matrix 0046-9717. the cause of the vascular damage issue was patient condition related. the cause of the purge leak issue was not determined since the product was not returned. additional information: a5 ethnicity and race. b5,b6, b7: additional malfunction information and relevant history. d6a, d6b: implant date and explant dates. e1: reporting facility phone number. g1: current manufacturer contact email address. g3: date received. this report is being filed as part of a retrospective review of historical records.
36157 | 1220648-2022-01607 | 15545123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15545123&pc=OZD | 10/04/2022 | 10/05/2022 |  | 10/05/2022 |  |  | 10/04/2022 | 10/05/2022 |  |  | 03/31/2024 |  | 10/05/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023200212 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/04/2022 |  |  |  |  | The complainant reported a 72-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during ecpella support, the patient showed evidence of hemolysis. to treat the hemolysis, the patient was administered haptoglobin and transfused 2 units of blood product daily for a total of 24 units. patient support was maintained throughout the event without further issue. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36158 | 1220648-2022-01608 | 15545169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15545169&pc=OZD | 09/23/2022 | 11/28/2022 |  | 10/05/2022 |  |  | 09/23/2022 | 10/05/2022 |  |  | 04/30/2024 |  | 10/05/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023223936 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/23/2022 |  |  |  |  | The complainant reported a 71-year-old female patient presenting for post cardiotomy cardiogenic shock. during impella support, a blood leak was observed from the repositioning sheath of the impella pump. as treatment for the leak, two units of prbc were transfused and the pump was replaced with a new impella pump. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | The pump was returned for further evaluation. a visual inspection was performed and no damage was observed on the returned component. furthermore, leak testing was also completed with no noted issues that could have caused or contributed to the original allegation. upon conclusion of the evaluation the reported failure could not be replicated. the pump was found to be fully function and a definitive cause for the reported allegation could not be established.
36159 | 1220648-2022-01609 | 15545214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15545214&pc=OZD | 09/26/2022 | 10/05/2022 |  | 10/05/2022 |  |  | 09/26/2022 | 10/05/2022 |  |  | 04/30/2024 |  | 10/05/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 PUMP SET JP |  |  | ABIOMED, INC |  |  | US |  | 2023223248 |  | 000000000001000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 32 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/26/2022 |  |  |  |  | The complainant reported a 32-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. during escalation from an impella cpsa to an impella 5.5, the patient experienced an access site bleed. it was estimated that roughly 400ml of blood were lost during the procedure to connect the graft for the new implant. to counteract the blood loss, over 2 units of prbc and ffp were provided to the patient. additional intervention was not required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36160 | 1220648-2022-01610 | 15547637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15547637&pc=OZD | 09/20/2022 | 10/11/2023 |  | 10/05/2022 |  |  | 09/20/2022 | 10/05/2022 |  |  | 03/31/2024 |  | 10/05/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023209953 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2022 |  |  |  |  | The complainant reported a 51-year-old male patient presenting for cardiomyopathy. during impella support, the patient experienced hemolysis. despite repositioning attempts, the hemolysis did not improve and the pump was subsequently explanted. the patient stabilized and additional support was not required. | [5 empty] | The impella device has not been returned by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed. | Since the original report was filed the investigations have concluded. product was not returned for analysis, and as such only the clinical report and data logs were evaluated. the cause of the hemolysis was ingested biomaterial.the cause of the positioning issues was most likely use related. the failure modes will be monitored and trended.
36161 | 1220648-2022-01611 | 15547654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15547654&pc=OZD | 09/11/2022 | 10/25/2023 |  | 10/05/2022 |  |  | 09/11/2022 | 10/05/2022 |  |  | 05/31/2024 |  | 10/05/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023218327 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 28 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/11/2022 |  |  |  |  | The medical team had an impella 5.5 support ongoing for a patient bridging to transplant. the team observed signs of hemolysis that prompted treatment. the team initiated dialysis and infused blood products. the pump remained on for support for over 12 days while awaiting heart transplant. | It was also reported that during support, the pump had lows flows. troubleshooting was able to resolve this issue. there was no patient harm reported. |  |  |  |  | The investigation into the hemolysis is on-going at this time. no product nor data logs have yet been returned from the medical facility for analysis and/or benchtop testing. should product be returned, and a root cause be determined, a final report will be filed. the failure mode will be monitored and trended. | Since the original report was filed, the investigation has concluded. the medical center returned the pump and benchtop analysis determined the cause of the hemolysis was biomaterial. the failure mode will be monitored and trended. | The investigation for the reported low flow and hemolysis. the data log shows a pump flow of 5.0 l/min on 9/1 when low pump flow was reported. on 9/11 the pump flow drops to around 1.0 l/min. there is also a motor current spike observed in the data log on 9/11. a visual inspection of the returned pump revealed a clot on both the inflow and the outflow. no other abnormalities were observed. the cause of the low pump flow and hemolysis was biomaterial. this report is being filed as part of a retrospective review of historical records.
36162 | MW5112454 | 15550471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15550471&pc=OZD | 09/24/2022 | 10/03/2022 |  | 10/05/2022 |  |  |  | 10/05/2022 |  |  |  |  | 10/05/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED INC. |  |  | US | [5 empty] | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 79 YR |  | Male |  |  |  | Injury | Yes | Yes | Obstruction of Flow | Therapeutic or Diagnostic Output Failure | [5 empty] | Failure of Implant | Thrombosis/Thrombus | [11 empty] | Other | [28 empty] | Initial submission | [9 empty] | A 79 year old male admitted with cardiogenic shock post myocardial infarction. an impella was placed on (b)(6) 2022. the right axillary impella was found to be dysfunctional on the morning of (b)(6) 2022. it was felt to be likely clotted, as evidenced by normal position on echo and x-ray but pulseless current flow. the impella was removed and a complete thrombectomy was performed on the graft. a new right axillary impella was placed successfully and the patient was returned to the icu for further care. fda safety report id# (b)(4).
36163 | 1220648-2022-01604 | 15537103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15537103&pc=OZD | 08/31/2022 | 10/14/2024 |  | 10/04/2022 |  |  | 09/05/2022 | 10/04/2022 |  |  | 01/31/2024 |  | 10/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023182362 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2022 |  |  |  |  | The complainant reported a 70 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. after insertion, the patient was alleged as enduring hemolysis, as evident by hemolytic urine. haptoglobin was delivered and the patient's urine color improved. | The complainant reported that the impella device purge housing experienced a fluid leak since insertion. it was reported that the purge house leaked and was replaced twice. |  |  |  |  | The impella pump was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information for the initial mfr 1220648-2022-01604 was provided in sections b1, b3, b5, d6a, d6b, h1, h6. the investigation has been completed for the reported issue of purge leak. the device was received for evaluation. the cause of the purge leak was a piston failure. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
36164 | 1220648-2022-01605 | 15537112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15537112&pc=OZD | 05/18/2022 | 10/04/2022 |  | 10/04/2022 |  |  | 09/06/2022 | 10/04/2022 |  |  | 11/30/2023 |  | 10/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022151421 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/06/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed an access site bleed that was treated with blood product delivery. the patient was also suspected as enduring hemolysis, which was treated with haptoglobin. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36165 | 1220648-2022-01606 | 15538123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15538123&pc=PYX | 09/19/2022 | 10/04/2022 |  | 10/04/2022 |  |  | 09/19/2022 | 10/04/2022 |  |  | 02/28/2024 |  | 10/04/2022 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC |  |  | US |  | 2023194542 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/19/2022 |  |  |  |  | The complainant reported a patient presenting for impella support. while implanted, the patient experienced hemolysis. despite repositioning attempts, the hemolysis did not improve and the pump was subsequently explanted. the patient was placed on ecmo for continued support following explant. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36166 | 1220648-2022-01603 | 15532996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15532996&pc=OZD | 09/15/2022 | 10/03/2022 |  | 10/03/2022 |  |  | 09/19/2022 | 10/03/2022 |  |  | 05/31/2024 |  | 10/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023222784 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/19/2022 |  |  |  |  | The medical team in japan had an impella cp support for 4 days. the cp was inserted after cpr and rib fracture. due to the fracture the heparin was reduced. there was thrombosis observed at explant and so the team used a fogarty embolectomy. additionally during the ecmo-cp support there was hemolysis. the team made the choice to infuse blood products back to the patient. | [5 empty] | All product was discarded by the user in japan. no benchtop analysis to root cause is possible. the failure modes will be monitored and trended.
36167 | 1220648-2022-01602 | 15533109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15533109&pc=OZD | 03/31/2021 | 10/03/2022 |  | 10/03/2022 |  |  | 09/27/2022 | 10/03/2022 |  |  | 01/31/2023 |  | 10/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2022038566 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/27/2022 |  |  |  |  | The medical team in japan had an impella cp support in 3/2021. the support was successful til wean and explant. later in 9/2022 the jpvad registry noted an ischemic stroke during the support. the contribution of the impella use to the stroke could not be ruled out by the medical team in japan. | [5 empty] | The medical center discarded all impella product and as such no benchtop analysis and testing to root cause is possible. the jpvad registry team shared no further clinical details. the failure mode will be monitored and trended.
36168 | 1220648-2022-01601 | 15533365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15533365&pc=OZD | 06/01/2022 | 10/25/2024 |  | 10/03/2022 |  |  | 09/27/2022 | 10/03/2022 |  |  | 12/31/2023 |  | 10/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023163711 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Other |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2022 |  |  |  |  | The medical team in japan had a cp support of 7 days. this occurred in 5/2022. the jpvad registry team documented a stroke, in 9/2022, to have occurred during the support. the stroke was determined to be ischemic. the causality of the stroke was unknown, thought to be caused or contributed to by the use of the impella. | [5 empty] | All impella pump product was discarded in 5/2022 and as such no investigation to root cause is possible. the jpvad registry team shared no further clinical details. the failure mode will be monitored and trended. | B2: date of patient death is unknown. the investigation for the reported stroke has been completed. the clinical details provided were not sufficient to determine a root cause. the cause of the stroke and its relation to the impella were not determined since the product, data logs, and sufficient clinical information were not provided. the patient death was determined to be unrelated to the impella. this report is being filed as part of a retrospective review of historical records.
36169 | 1220648-2022-01600 | 15523043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15523043&pc=OZD | 09/21/2022 | 09/30/2022 |  | 10/01/2022 |  |  | 09/21/2022 | 10/01/2022 |  |  | 08/31/2022 |  | 10/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC |  |  | US |  | 1498090 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/21/2022 |  |  |  |  | The complainant reported an 81-year-old male patient presenting for post cardiotomy cardiogenic shock. during impella support, it was noted that the patient experienced an ischemic stroke. the pump was subsequently explanted two days following the stroke. additional details pertaining to the event or treatment were not available. | [5 empty] | The impella device was discarded by the customer and therefore; investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36170 | 1220648-2022-01597 | 15524949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15524949&pc=OZD | 12/27/2021 | 09/30/2022 |  | 10/01/2022 |  |  | 12/27/2021 | 10/01/2022 |  |  | 06/30/2023 |  | 10/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC |  |  | US |  | 2022109036 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/27/2021 |  |  |  |  | The complainant reported a 60-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. while on impella support, the patient experienced hemolysis. as treatment, the impella 2.5 pump was prematurely removed and the patient was escalated to an impella 5.5 pump. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36171 | 1220648-2022-01595 | 15512204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15512204&pc=OZD | 09/01/2022 | 11/10/2022 |  | 09/30/2022 |  |  | 09/01/2022 | 09/30/2022 |  |  | 04/30/2024 |  | 09/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023226021 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2022 |  |  |  |  | The complainant reported a 68 year old male patient presenting nstemi for impella support. patient was taken to the operating room approximately 4 days later for planned impella removal and ventricular septal defect repair procedure, when an aortic thrombus was noted encased around the impella device. intervention was required via surgical removal of the cannula and an embolectomy was required on the right iliac and femoral arteries. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | In the initial mdr, we erroneously stated the device had been discarded. the device was not discarded and received by our facility on 2022-10-04 for investigation. the device was returned severed, so simulated use testing could not be performed. no other abnormalities were identified. with the information provided, we believe the root cause of the reported issue was related to the patient's condition.
36172 | 1220648-2022-01596 | 15520415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15520415&pc=OZD | 09/27/2022 | 09/30/2022 |  | 09/30/2022 |  |  | 09/27/2022 | 09/30/2022 |  |  | 12/31/2022 |  | 09/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC |  |  | US |  | 2022038490 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 89 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/27/2022 |  |  |  |  | The complainant reported an 89-year-old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient experienced an access site bleed during pump explant and developed a hematoma. the patient underwent a blood transfusion of unknown quantity to counteract the blood loss. additional information pertaining to the bleed and subsequent intervention was not available. the patient was considered stable following treatment. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36173 | 1220648-2022-01598 | 15521901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15521901&pc=OZD | 09/14/2022 | 09/30/2022 |  | 09/30/2022 |  |  | 09/19/2022 | 09/30/2022 |  |  | 03/31/2024 |  | 09/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2023212174 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  |  | White | Injury | Yes | Yes | Material Twisted/Bent | Migration | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/19/2022 |  |  |  |  | The complainant reported a 37-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, the pump shifted out of position and subsequently prolapsed in the middle of the cannula. the decision was subsequently made to remove the pump using a snare. however, during removal, an ascending aortic dissection occurred prompting an aortic valve replacement. after removal of the pump, the access site was successfully closed and the patient remained on ecmo support. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36174 | 1220648-2022-01594 | 15505542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15505542&pc=OZD | 08/29/2022 | 09/28/2022 |  | 09/29/2022 |  |  | 08/29/2022 | 09/29/2022 |  |  | 03/31/2024 |  | 09/29/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023204889 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/29/2022 |  |  |  |  | The complainant reported a 75 year old female patient presenting for impella support prior to procedure. after the procedure, severe bleeding was discovered within the aorta immediately after removing the cross clamp. the patient's tissues were not amenable to repair/salvage and the patient expired. the impella was found within the aortic wall, therefore, the physician alleged the device had dissected the artery and led to the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36175 | 1220648-2022-01593 | 15499452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15499452&pc=OZD | 09/12/2022 | 09/27/2022 |  | 09/28/2022 |  |  | 09/12/2022 | 09/28/2022 |  |  | 04/30/2024 |  | 09/28/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023210747 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/12/2022 |  |  |  |  | The complainant reported a 73-year-old female patient presenting for post cardiotomy cardiogenic shock. during impella support, the patient experienced hemolysis for which 11 units of prbc were provided. a transesophageal echocardiogram (tee) showed that the impella had flipped positioning and was pointed towards the septum. as a result of the hemolysis, the decision was made to explant the pump prematurely. | [5 empty] | The impella device was not returned by the customer and therefore; investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36176 | 1220648-2022-01592 | 15488118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15488118&pc=OZD | 09/15/2022 | 09/26/2022 |  | 09/26/2022 |  |  | 09/15/2022 | 09/26/2022 |  |  | 04/30/2024 |  | 09/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023215936 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/15/2022 |  |  |  |  | The complainant reported a 68-year-old female patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, the patient developed lower limb ischemia of both extremities. in an attempt to treat the reported condition, the pump flow was initially increased but the decision was later made to explant the pump. upon explant, a thrombus was discovered and surgically removed. an intra-aortic balloon pump (iabp) was then implanted to continue with patient support. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36177 | 1220648-2022-01591 | 15480239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15480239&pc=OZD | 12/15/2021 | 09/23/2022 |  | 09/24/2022 |  |  | 09/22/2022 | 09/24/2022 |  |  | 07/31/2023 |  | 09/24/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022108809 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/22/2022 |  |  |  |  | The medical team in japan had an impella cp support in (b)(6) 2021. later in (b)(6) 2022 the jpvad registry noted an access site bleed. the bleed was treated by the infusion of blood products. the team did note the heparin concentration was adjusted in preparation for a coronary artery bypass surgery, and this could have contributed to the access site bleed. | [5 empty] | The medical center discarded all impella product in 2021. no investigation and benchtop analysis to root cause is possible. only the jpvad registry information is available. should more be shared which leads to root cause, another report will be filed.
36178 | 1220648-2022-01583 | 15470108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15470108&pc=OZD | 03/31/2022 | 09/22/2022 |  | 09/23/2022 |  |  | 03/31/2022 | 09/23/2022 |  |  | 12/31/2023 |  | 09/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023180841 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/31/2022 |  |  |  |  | The medical center had a patient in the loqi study supported by an impella cp. the cp was explanted when the team observed the pump to be occlusive to the leg/perfusion down the leg. this prompted the team to explant the pump and close the groin site by a manta. the patient was later placed on an iabp for hemodynamic support. the limb ischemia resolved after pump removal. | [5 empty] | The medical center discarded all product at the time of explant. the loqi study gave no further details on the cause of the ischemia attributed to the use of impella. the manufacturer will close the investigation without root cause. should more information be shared, a final report will be filed with definitive root cause.
36179 | 1220648-2022-01584 | 15470643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15470643&pc=OZD | 08/25/2022 | 09/22/2022 |  | 09/23/2022 |  |  | 08/25/2022 | 09/23/2022 |  |  | 05/31/2024 |  | 09/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023218344 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/25/2022 |  |  |  |  | The complainant reported an 84 year old male patient presenting with post cardiotomy cardiogenic shock for impella support. after approximately 45 hours of support, the pump cable became disconnected while moving the patient which resulted in a pump stop. the pump was unable to be restarted. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36180 | 1220648-2022-01586 | 15471434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15471434&pc=OZD | 01/29/2022 | 09/22/2022 |  | 09/23/2022 |  |  | 08/25/2022 | 09/23/2022 |  |  | 10/31/2023 |  | 09/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022148664 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/25/2022 |  |  |  |  | The complainant reported a 28 year old patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient endured an impella related access site bleed. blood products were given as treatment. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36181 | 1220648-2022-01588 | 15471944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15471944&pc=OZD | 08/06/2022 | 10/14/2024 |  | 09/23/2022 |  |  | 08/26/2022 | 09/23/2022 |  |  | 03/31/2024 |  | 09/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023200262 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  |  | Asian | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2022 |  |  |  |  | The complainant reported a 57 year old female patient presenting with acute myocardial infarction and cardioigenic shock for impella support. during support, the patient had developed an access site bleed and thrombocytopenia. a blood transfusion was delivered for the access site bleed and platelets were delivered in response to thrombocytopenia. | The complainant reported a pump stop during use of the device. troubleshooting was able to restart the device with no harm to the patient. |  |  |  |  | The impella impella pump was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The cause of the pump stop was not determined since product was not returned. the pump likely interacted with the patient since the pump stop was temporary and support continued. additional information for the initial mfr was provided in sections b5, d6a, d6b, g1, h1, h6. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
36182 | 1220648-2022-01589 | 15472365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15472365&pc=OZD | 08/29/2022 | 09/22/2022 |  | 09/23/2022 |  |  | 08/29/2022 | 09/23/2022 |  |  | 02/29/2024 |  | 09/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023191683 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/29/2022 |  |  |  |  | The complainant reported a 78 year old female patient presenting with post cardiotomy cardiogenic shock for impella support. during support, the patient developed ischemia in the impella inserted limb. an antegrade sheath was administered as intervention and the issue resolved. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36183 | 1220648-2022-01590 | 15472503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15472503&pc=OZD | 11/23/2021 | 09/22/2022 |  | 09/23/2022 |  |  | 11/23/2021 | 09/23/2022 |  |  | 09/30/2023 |  | 09/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022136160 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/26/2021 |  |  |  |  | The complainant reported a 57 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient's impella inserted extremity was cold to the touch without pulses. the pump was removed and the decision was made to escalate care to an impella 5.5. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36184 | 1220648-2022-01587 | 15473841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15473841&pc=OZD | 08/11/2022 | 09/22/2022 |  | 09/23/2022 |  |  | 08/26/2022 | 09/23/2022 |  |  | 03/31/2024 |  | 09/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023205513 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/26/2022 |  |  |  |  | The complainant reported a 73 year old female patient presenting for impella support. during support, hemolysis was suspected due to patient's hemolytic urine. haptoglobin was delivered as treatment. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36185 | 1220648-2022-01585 | 15473854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15473854&pc=OZD | 03/18/2022 | 09/22/2022 |  | 09/23/2022 |  |  | 08/25/2022 | 09/23/2022 |  |  | 05/31/2021 |  | 09/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1428037 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/25/2022 |  |  |  |  | The complainant reported a 78 year old male patient presenting with cardiomyopathy for impella support. during support, the patient developed an access site hematoma that prompted a blood transfusion. additionally, the patient endured hemolysis which lead to pump removal. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36186 | 15462266 | 15462266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15462266&pc=OZD | 06/16/2022 | 09/09/2022 | 09/09/2022 | 09/22/2022 | 09/09/2022 | 09/22/2022 |  | 09/22/2022 |  |  |  |  | 09/22/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA PUMP |  |  | OZD | [5 empty] | ABIOMED INC. |  |  | US |  | 375094 |  | 00048-0003 |  | Yes |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 31025 DA |  | Female |  |  |  | Malfunction | No | Yes | Suction Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | After case completed, impella device started to have error/suction issue. doctor tried to reposition device, had to remove and put in brand new impella device. representative present during case, new device was needed. new impella device inserted. manufacturer response for impella device, (brand not provided) company rep was present during case.
36187 | 1220648-2022-01582 | 15457579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15457579&pc=OZD | 08/23/2022 | 09/20/2022 |  | 09/20/2022 |  |  | 08/23/2022 | 09/20/2022 |  |  | 04/30/2024 |  | 09/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023220615 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2022 |  |  |  |  | The complainant reported an 82 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. while the patient was being taken off the table in the cardiac catheterization lab, the patient patient kicked their leg. subsequently, the patient endured a vascular dissection. vascular surgery completed a cutdown to confirm the dissection, repaired the injury, and placed a new impella. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36188 | 1220648-2022-01579 | 15449337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15449337&pc=OZD | 08/25/2022 | 10/21/2022 |  | 09/19/2022 |  |  | 08/25/2022 | 09/19/2022 |  |  | 04/30/2024 |  | 09/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 IMC PUMPEN SET EU |  |  | ABIOMED, INC. |  |  | US |  | 2023197310 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Patient in the eu had a impella 5.0 pump stop. pump unable to restart. pump successfully exchanged for a new one. patient survived. no new mechanical cardiovascular support devices used. | [5 empty] | The investigation is ongoing at this time. a final report will be filed upon completion. | The cause of the pump stop was ingested biomaterial. product did not fail to meet specification.
36189 | 1220648-2022-01581 | 15452259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15452259&pc=OZD | 06/27/2022 | 09/19/2022 |  | 09/19/2022 |  |  | 09/06/2022 | 09/19/2022 |  |  | 12/31/2023 |  | 09/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023171127 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility |  |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/06/2022 |  |  |  |  | The medical center in japan had a successful impella cp support in 6/2022. later in 9/2022 the jpvad registry noted limb ischemia that prompted explant. there was no other treatment noted and the limb resolved after pump explant. | [5 empty] | All product was discarded month prior to investigation opening. there can be no benchtop investigation to root cause. the jpvad registry team provided no further clinical details. the failure mode will be monitored and trended.
36190 | 1220648-2022-01580 | 15452260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15452260&pc=OZD | 08/14/2022 | 09/19/2022 |  | 09/19/2022 |  |  | 09/08/2022 | 09/19/2022 |  |  | 04/30/2024 |  | 09/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023216078 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Hospitalization | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/08/2022 |  |  |  |  | In japan a covid + patient admitted in cardiogenic shock and a history of cardiomyopathy, was placed on an impella cp and ecmo for support in (b)(6) 2022. later in (b)(6) 2022 the jpvad registry noted there had been an impella access site bleed and hemolysis. both were treated with intervention, in the form of blood product and haptoglobin infusion. | [5 empty] | The medical team in japan discarded all product at the explant. no benchtop analysis to root cause is possible. the information provided by the jpvad registry alone did not lead to root cause. the failure mode will be monitored and trended.
36191 | MW5112134 | 15455579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15455579&pc=OZD | 08/11/2022 | 09/15/2022 |  | 09/19/2022 |  |  |  | 09/19/2022 |  |  | 04/30/2024 |  | 09/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA CP |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED INC. |  |  | US |  |  |  | SET IMPELLA CP 9.3IN 10-40C |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 37 YR |  | Male | 72 KG | Hispanic | White | Injury | Yes | Yes | Therapeutic or Diagnostic Output Failure | [6 empty] | Failure of Implant | [12 empty] | Life Threatening | [28 empty] | Initial submission | [9 empty] | Impella cp inserted on (b)(6) 2022 spontaneously malfunctioned required return to surgery for impella cp exchange on (b)(6) 2022. fda safety report id #(b)(4).
36192 | MW5112136 | 15455634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15455634&pc=OZD | 08/19/2022 | 09/15/2022 |  | 09/19/2022 |  |  |  | 09/19/2022 |  |  | 04/30/2024 |  | 09/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA CP |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED INC. |  |  | US |  |  |  | SET IMPELLA CP 9.3IN 10-40C |  | Yes | [5 empty] | No |  |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 37 YR |  | Male | 72 KG | Hispanic | White | Injury | Yes | No | Therapeutic or Diagnostic Output Failure | [6 empty] | Failure of Implant | [12 empty] | Life Threatening | [28 empty] | Initial submission | [9 empty] | Impella cp inserted on (b)(6) 2022 spontaneously malfunctioned required return to surgery for impella cp explantation on (b)(6) 2022. fda safety report id (b)(4).
36193 | 1220648-2022-01575 | 15442240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15442240&pc=OZD | 09/12/2022 | 09/16/2022 |  | 09/16/2022 |  |  | 09/12/2022 | 09/16/2022 |  |  | 12/31/2023 |  | 09/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023171618 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/12/2022 |  |  |  |  | The complainant reported a 78 year old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was noted that the patient experienced hemolysis. as treatment, haptoglobin was administered to the patient and the pump was later removed to relieve the condition. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36194 | 1220648-2022-01576 | 15442603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15442603&pc=OZD | 12/07/2021 | 09/16/2022 |  | 09/16/2022 |  |  | 08/23/2022 | 09/16/2022 |  |  | 10/31/2022 |  | 09/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2021019905 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2022 |  |  |  |  | The medical team in japan had a successful impella cp support in (b)(6) 2021. the support was just under 18 hours when explanted after the coronary angioplasty procedure. later in (b)(6) 2022 the jpvad registry noted a vascular injury that required treatment. the registry noted the team performed vessel suture. | [5 empty] | The medical center discarded the impella pump in 2021. only information shared by the jpvad registry could be reviewed for analysis. the root cause can not be determined with known information. the failure mode will be monitored and trended.
36195 | 1220648-2022-01577 | 15442613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15442613&pc=OZD | 05/10/2022 | 09/16/2022 |  | 09/16/2022 |  |  | 08/23/2022 | 09/16/2022 |  |  | 10/31/2023 |  | 09/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED IN. |  |  | US |  | 2022147721 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2022 |  |  |  |  | The medical team in japan had an 11 day support with an impella cp in 5/2022. the jpvad registry team later reported in 8/2022 an access site bleed. the bleed was treated by blood product infusion. the bleed was attributed to the impella access and combination with anticoagulation. | [5 empty] | The medical center discarded all impella product. no investigation to root cause is possible. the failure mode will be monitored and trended.
36196 | 1220648-2022-01578 | 15442623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15442623&pc=OZD | 08/20/2022 | 09/16/2022 |  | 09/16/2022 |  |  | 08/25/2022 | 09/16/2022 |  |  | 01/31/2024 |  | 09/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023182405 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/25/2022 |  |  |  |  | The medical team in japan had an impella cp support of 5 days. during the support time there was an observation made of an access site bleed that was treated by surgical suture and blood product transfusion. the number of units infused is not known. | [5 empty] | The medical center in japan had discarded all the product at pump explant. no investigation to root cause is possible. the failure mode will be monitored and trended.
36197 | 1220648-2022-01573 | 15419316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15419316&pc=OZD | 08/20/2022 | 09/14/2022 |  | 09/14/2022 |  |  | 08/20/2022 | 09/14/2022 |  |  | 05/31/2024 |  | 09/14/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023222119 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | Mitral Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/20/2022 |  |  |  |  | The complainant reported a 73 year old female patient presenting with mild aortic insufficiency (ai) and moderate/severe mitral regurgitation (mr). after the physician removed the pump, at the end of support, the patient's mr and ai were found in worse condition. the patient's aortic valve was replaced and a new impella 5.5 was inserted. the site was concerned removal of the initial pump caused further damage to the valve. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36198 | 1220648-2022-01574 | 15419325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15419325&pc=OZD | 08/16/2022 | 10/24/2022 |  | 09/14/2022 |  |  | 08/16/2022 | 09/14/2022 |  |  | 05/31/2024 |  | 09/14/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023228822 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/16/2022 |  |  |  |  | The complainant reported a 73 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, a dissection in the impella inserted artery occurred, as evident by a large hematoma. the pump was removed and intervention was required via a covered stent to resolve the injury. | [5 empty] | The impella device was received from the customer and an evaluation is underway. a supplemental mdr will be filed at the conclusion of our analysis. | The pump was returned with no damage observed. the introducer was not returned. we are unable to definitively determine the root cause of the reported injury.
36199 | 1220648-2022-01570 | 15411818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15411818&pc=OZD | 09/03/2022 | 09/13/2022 |  | 09/13/2022 |  |  | 09/03/2022 | 09/13/2022 |  |  | 12/31/2022 |  | 09/13/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2022029191 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/03/2022 |  |  |  |  | The complainant reported a 68 year old male patient presenting for acute myocardial infarction and cardiogenic shock. during impella support, it was discovered that the patient experienced hemolysis. at the time of the noted issue, the pump was found to have migrated into the aorta and repositioning attempts subsequently failed. as such, the pump was prematurely removed and replaced with a new pump to relieve the patient condition. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36200 | 1220648-2022-01571 | 15412560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15412560&pc=OZD | 08/19/2022 | 09/23/2022 |  | 09/13/2022 |  |  | 08/20/2022 | 09/13/2022 |  |  | 09/30/2022 |  | 09/13/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1501589 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/20/2022 |  |  |  |  | The medical team had made the choice to escalate hemodynamic care from an impella cp to the impella 5.5. the team was performing a bypass surgery (cabg) and placing the 5.5. the 5.5 was placed and turned on with wire still present and as such the pump stopped and was replaced. there was no harm or injury from delay. a new 5.5 was successfully placed. | [5 empty] | The medical center has returned the pump for evaluation. the evaluation and benchtop analysis is on-going at this time. upon completion of the investigation, a final report will be drafted. | Since the original report was filed, the investigation into the pump stop has completed. the pump was returned and analyzed upon the benchtop. the pump was run without any issues. the cause of the pump stop occurring during the patient use was use related, as the medical team operated the pump with the guidewire still retained. the failure mode will be monitored and trended.
36201 | 1220648-2022-01572 | 15412617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15412617&pc=OZD | 12/15/2021 | 09/16/2022 |  | 09/13/2022 |  |  | 12/15/2021 | 09/13/2022 |  |  | 09/30/2023 |  | 09/13/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022135361 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2021 |  |  |  |  | The medical center was utilizing an impella cp for support of a high risk coronary angioplasty patient in 12/2021. the patient was additionally supported by ecmo. after 1 day of support there was bilateral limb ischemia that prompted pump explant and replacement by an axillary 5.5 pump. the surpass registry review found in 9/2022 adverse events/harms and a complaint was created for the ischemia and early explant. blood products additionally were infused for the access site hematoma of the impella. the hematoma was treated by multiple units of blood products, platelets, and ffp. | [5 empty] | The medical center discarded the impella pump product at explant in 2021. no investigation to root cause is possible. the failure mode will be monitored and trended. should new information be shared, another report will follow. | Since the original report was filed the clinical review and investigation has completed. the access site bleed was found to be due to improper technique by the user. the introducer sheath was peeled while still in the vessel resulting in bleeding. the limb ischemia was determined to be due to patient condition. both failure modes will be monitored and trended.
36202 | 1035166-2022-00100 | 15404092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15404092&pc=DYB | 08/05/2022 | 01/10/2023 |  | 09/12/2022 |  |  | 08/17/2022 | 09/12/2022 |  |  | 11/01/2024 |  | 09/12/2022 |  | INTRODUCER, CATHETER |  |  | BIOMED INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | DP-14289 |  | 0052-3006 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [7 empty] | Unknown |  |  |  | Malfunction | No | Yes | Break | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a 73 year old male was presented to the physician with shortness of breath and underwent an acute myocardial infarction/cardiogenic shock procedure on (b)(6) 2022 involving an oscor 23 fr introducer. after inserting the introducer through the surgical graft, blood was noted leaking around the sheath, despite tightening the grafts locks. the bleed resolved after the introducer was removed and a new 23 fr oscor introducer was inserted. a split was noted; however, no further information was provided. a additional device, vascutek gelweave graft was used during the procedure. there was a delay in inserting the impella due to bleeding. no intervention reported, beyond changing out the introducer. there was no harm to the patient reported. | Additional information received 09/15/2022: patient had normal anatomy. it was a problem introducing the sheath into the graft. no excessive force was observed. |  |  |  |  | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | One 23f abiomed axillary introducer sheath was received back from the customer without the dilator. there were no other accessories. traces of blood were found on and inside the sheath. according to the event description summary, blood was leaking from the sheath and a split was found. it was not clear from the event description exactly where the leak in the sheath occurred. the valve was viewed under a 10x microscope and looked normal with no anomalies. the sheath's hub and shaft looked normal and no cracks or splits were found. the sheath was manually leak tested with the syringe connected to the stopcock and the tip occluded. the sheath did not exhibit leakage when tested using this method. the sheath was then tested using the iso 11070 for liquid leakage test method. the device was occluded at the distal tip and pressurized to 38kpa and held at this pressure for 30 seconds and no leakage was observed from the hemostatic valve. the sheath was further pressurized beyond 300kpa and no leakage was observed from the stopcock, sideport tubing or the hub. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. returned device analysis revealed the sheath was within manufacturing specification. the sheath did not leak when tested manually. the sheath also met the iso leakage test method requirement and did not leak through the hemostatic valve, hub or shaft. the sheath's hub and shaft looked normal and no cracks or splits were found. according to the dhr, the sheath passed all in-process and final inspection steps including visual, mechanical, dimensional and leak testing. no manufacturing defects were found. per qa procedure (abiomed introducer sheath in-process and final inspection) sample size 100%. with naked the eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. verify split caps are properly placed and secured onto sheath hub and free of cracks/ damages or excessive adhesive. perform pressure and vacuum leak test per procedures. the instructions for use (ifu) informs the user: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36203 | 1220648-2022-01569 | 15406815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15406815&pc=OZD | 08/15/2022 | 10/14/2024 |  | 09/12/2022 |  |  | 08/18/2022 | 09/12/2022 |  |  | 01/31/2024 |  | 09/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023187101 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Hospitalization | Required Intervention | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2022 |  |  |  |  | The medical team had an impella cp supporting a patient admitted with a myocardial infarction, a nstemi. the cp was placed while coronary bypass surgery was discussed for the patient. on day 3 of support the patient had limb ischemia, and as such the cp was explanted. the support was replaced by an impella 5.5 via axillary insertion. the limb ischemia prompted fasciotomy to the impella accessed leg and amputation followed later. the amputation was above the knee. | It was also reported the patient experienced hemolysis which the device was manipulated, due to improper positioning of the pump. |  |  |  |  | The medical center discarded the impella cp at explant. the patient is still supported by the impella 5.5. the investigation into the limb ischemia and amputation is ongoing at this time. no root cause has yet been determined. further clinical details have been asked, but not yet furnished. should a root cause be determined, another report will be filed. the failure mode will be monitored and trended. | Additional information for the initial mfr 1220648-2022-01569 was provided in sections b5, d6a, d6b, g1, h1, and h6. the investigation for the reported hemolysis has been completed. the impella device nor data logs were received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the hemolysis was most likely improper positioning of the device. the failure modes will be monitored and trended. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
36204 | 1220648-2022-01568 | 15396231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15396231&pc=OZD | 08/11/2022 | 09/09/2022 |  | 09/09/2022 |  |  | 08/11/2022 | 09/09/2022 |  |  | 03/31/2024 |  | 09/09/2022 |  | (01)813502011876(10)2023208255(21)380726(17)240331 |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023208255 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/11/2022 |  |  |  |  | The complainant reported a 69 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during placement of the 14 fr introducer, the patient's vessel was perforated. femoral insertion of the impella was aborted and intervention was required to stabilize the injury. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36205 | 1220648-2022-01567 | 15396299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15396299&pc=OZD | 08/11/2022 | 09/09/2022 |  | 09/09/2022 |  |  | 08/11/2022 | 09/09/2022 |  |  | 04/30/2024 |  | 09/09/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023217054 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/11/2022 |  |  |  |  | The complainant reported a 57 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. while delivering the short introducer over the guide wire, the patient's femoral artery was perforated. vascular surgery was required to repair the injury. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36206 | 1220648-2022-01564 | 15372760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15372760&pc=OZD | 08/31/2022 | 09/07/2022 |  | 09/07/2022 |  |  | 08/31/2022 | 09/07/2022 |  |  | 03/31/2024 |  | 09/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED. INC |  |  | US |  | 2023210455 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/31/2022 |  |  |  |  | The complainant reported a 79 year old male patient presenting for impella support. upon placement of the sheath during implantation, there was a decrease in blood pressure. a fluoroscopy with contrast was utilized to identify a vascular dissection with extravasation. the patient was subsequently administered two units of blood and norepinephrine, before a covered stent was placed to treat the dissection. following intervention, the impella pump was successfully implanted for patient support. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36207 | 1220648-2022-01563 | 15373630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15373630&pc=OZD | 08/10/2022 | 09/06/2022 |  | 09/07/2022 |  |  | 08/10/2022 | 09/07/2022 |  |  | 01/31/2024 |  | 09/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023188331 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/10/2022 |  |  |  |  | The complainant reported a 59 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed limb ischemia in the impella inserted leg. the pump was removed and the ischemia resolved. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36208 | 1220648-2022-01565 | 15375403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15375403&pc=OZD | 09/05/2020 | 09/07/2022 |  | 09/07/2022 |  |  | 08/16/2022 | 09/07/2022 |  |  | 10/31/2021 |  | 09/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1459401 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/16/2022 |  |  |  |  | The medical team in japan had an impella cp support in 2020. later in 2022 the jpvad registry noted a vascular injury that required treatment. the patient had moved during support and caused a hematoma. the team noted the treatment was a vessel suture. no additional treatment was noted, nor any lasting harm/injury to the vasculature. | [5 empty] | The medical center discarded all impella product at the explant in 2020. no benchtop analysis to root cause is possible. should the registry furnish more details that lead to a definitive root cause, a final report will follow. the failure mode will be monitored and trended.
36209 | 1220648-2022-01566 | 15375436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15375436&pc=OZD | 08/16/2022 | 09/20/2022 |  | 09/07/2022 |  |  | 08/16/2022 | 09/07/2022 |  |  | 06/23/2024 |  | 09/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023217288 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/16/2022 |  |  |  |  | The medical team had an impella cp support in which there was an access site complication. the patient had a hematoma at the impella site. the hematoma and blood loss were remedied by a 16fr introducer being placed and subsequent surgical repair. the patient had multiple perclose fail, and the patient was on thrombolytics. the perclose failure, impella placement, and thrombolytics all contributed to the patient injury per the medical team. | [5 empty] | The medical center discarded all impella pump and introducer product. no benchtop analysis to root cause is possible. should more clinical information be shared that leads to a definitive root cause, a final report will be filed. the failure mode will be monitored and trended. | The investigation has concluded. the cause of the access site bleeding was most likely due to the high patient act values. this patient condition allowed for the bleeding to occur. the failure mode will be monitored and trended.
36210 | MW5111905 | 15379915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15379915&pc=OJE | 07/14/2022 | 09/04/2022 |  | 09/07/2022 |  |  |  | 09/07/2022 |  |  |  |  | 09/07/2022 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA RP |  |  | OJE | [5 empty] | ABIOMED INC. | [8 empty] | No | [6 empty] | Health Professional |  |  | Voluntary report | Not Available | Patient Family Member or Friend | [6 empty] | 69 YR |  | Male | 118 KG | Non Hispanic | White | Death | Yes | Yes | Difficult to Remove | [6 empty] | Perforation | Retroperitoneal Hemorrhage | [11 empty] | Hospitalization | Death |  |  |  |  | HEART MEDICATIONS FOR BLOOD PRESSURE PRESCRIBED BY HOSPITAL CARDIAC CARE UNIT. | [22 empty] | Initial submission | [9 empty] | I recently read your fda 2018 article by deborah kontz on the impella heart pump. i would like my husband's death as a result of the impella pump to be included in the fda data. my husband was treated with the impella pump during a high risk pci procedure on (b)(6) 2022. removing the pump catheter caused a severe internal bleed (retroperitoneal) botched surgery caused my husband to die seven days later. we were told that there was a risk for death, but had no idea the risk would be from withdrawing the catheter and puncturing his femoral artery. he lost about 4 liters of blood within 30 minutes following the surgery. the pci surgery appeared to be a success but the blood loss put a strain on his already debilitated and revascularized heart. your article helped me process whether we even needed to do this surgery and if this risk was worth it. it seems "either die from the equipment malfunctions" or "die without the surgery"; pick your hard. thank you for addressing the issues with this impella pump. i hope the fda will include my husband's case in their data. his name was (b)(6). he was a patient of dr (b)(6) at (b)(6) hospital in (b)(6). his surgery was (b)(6) 2022 and he died (b)(6) 2022. regards, (b)(6). fda safety report id# (b)(4).
36211 | 15360943 | 15360943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15360943&pc=OZD | 08/06/2022 | 08/12/2022 | 08/30/2022 | 09/06/2022 | 08/30/2022 | 09/06/2022 |  | 09/06/2022 |  |  |  |  | 09/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP PUMP SET |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003.0814 |  | Yes |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 32485 DA |  | Male | 52 KG |  | White | Malfunction | No | Yes | Fluid/Blood Leak | Improper Flow or Infusion | [5 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Patient had impella placement for ischemic cardiomyopathy, with flow rate monitored every hour. approximately 3 days later, noted fluid leaking from impella purge cassette, and low rate at 28- 30cc/hr. abiomed representative notified. purge cassette tubing changed, no more leak and purge flow rate started to come down to 20's. following day purge flow rate increased again, to 26 cc/hr. cassette opened; damp moisture felt inside. abiomed representative notified. purge cassette tubing changed again, and the flow rate reduced to 7cc/hr, with no additional leakage observed.
36212 | 1220648-2022-01559 | 15362220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15362220&pc=OZD | 08/29/2022 | 09/06/2022 |  | 09/06/2022 |  |  | 08/29/2022 | 09/06/2022 |  |  | 04/30/2024 |  | 09/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC |  |  | US |  | 2023200762 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/29/2022 |  |  |  |  | The complainant reported a 73 year old male patient presenting for impella support. during a scheduled extraction of an impella 5.5 pump, a thrombus was discovered within the subclavian artery. the pump was extracted and a fogarty catheter was successfully utilized to treat the thrombus. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36213 | 1220648-2022-01560 | 15366514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15366514&pc=OZD | 08/10/2022 | 09/06/2022 |  | 09/06/2022 |  |  | 08/10/2022 | 09/06/2022 |  |  | 07/31/2023 |  | 09/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022111011 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/10/2022 |  |  |  |  | The complainant reported a 68 year old male patient presenting for high risk percutaneous coronary intervention using the impella cp optical for hemodynamic support. during support, flows were lower than expected. a new pump was used without any harm to the patient. upon investigation of the returned device, broken blades were identified. | [5 empty] | The pump was returned and fractured impeller blades were found to be the root cause of the reported pump flow. however, because the fluoroscopic image of the pump was not returned, a definitive root cause of the fracture could not be determined.
36214 | 1220648-2022-01561 | 15366716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15366716&pc=OZD | 08/10/2022 | 09/06/2022 |  | 09/06/2022 |  |  | 08/10/2022 | 09/06/2022 |  |  | 03/31/2024 |  | 09/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023206522 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/10/2022 |  |  |  |  | The complainant reported a 72 year old female patient presenting for high risk percutaneous coronary intervention using the impella cp optical device for support. during support, blood was found leaking from the 14 fr introducer that prompted multiple units of blood product delivered. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36215 | 1220648-2022-01562 | 15366916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15366916&pc=OZD | 06/27/2022 | 09/06/2022 |  | 09/06/2022 |  |  | 08/10/2022 | 09/06/2022 |  |  | 11/30/2023 |  | 09/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022148713 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/10/2022 |  |  |  |  | The complainant reported a 57 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient endured an ischemic stroke in which the device could not be ruled out as a causal/contributory factor. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36216 | 1220648-2022-01552 | 15349986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15349986&pc=OZD | 08/03/2022 | 10/13/2022 |  | 09/02/2022 |  |  | 08/05/2022 | 09/02/2022 |  |  | 11/30/2023 |  | 09/02/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022157953 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/05/2022 |  |  |  |  | The complainant reported a 51 year old male patient presenting with cardiomyopathy for impella support. during support, signs of hemolysis were exhibited. as a result, the pump was removed and the patient was escalated to an impella 5.5. | [5 empty] | The impella device was returned by the customer and an investigation is underway. a supplemental mdr will be filed at the conclusion of our analysis. | The device was returned for evaluation with no abnormalities noted. evidence of the clinical account suggests the reported hemolysis was due to patient's condition.
36217 | 1220648-2022-01553 | 15353004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15353004&pc=OZD | 08/09/2022 | 09/02/2022 |  | 09/02/2022 |  |  | 08/09/2022 | 09/02/2022 |  |  | 02/29/2024 |  | 09/02/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023195409 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/09/2022 |  |  |  |  | The complainant reported a 74 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, an access site bleed formed. as treatment, multiple blood products were delivered. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36218 | 1220648-2022-01554 | 15353017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15353017&pc=OZD | 07/04/2021 | 11/27/2024 |  | 09/02/2022 |  |  | 08/09/2022 | 09/02/2022 |  |  | 07/31/2021 |  | 09/02/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1432800 |  | 005048-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 52 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia Stroke | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2022 |  |  |  |  | The complainant reported a 52 year old female patient presenting with cardiomyopathy for impella support. during support, the patient endured an ischemic stroke. the relationship to the stroke and the use of impella could not be ruled out. | Additionally, the complainant reported that the patient experienced a gastrointestinal bleed during impella placement and received blood transfusions to manage the complication. it is unknown how hemostasis was achieved. weaning was successful, the device was removed, and the patient survived the explant procedure. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information for the initial mfr 1220648-2022-01554 was provided in sections b5, d6, g2, and h6. this file is being submitted as part of a retrospective review of historical records.
36219 | 1220648-2022-01555 | 15353026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15353026&pc=OZD | 08/09/2022 | 09/02/2022 |  | 09/02/2022 |  |  | 08/09/2022 | 09/02/2022 |  |  | 04/30/2024 |  | 09/02/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023211945 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 08/09/2022 |  |  |  |  | The complainant reported a 48 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. after removing the introducer and advancing the repositioning sheath, the patient developed an access site bleed. as treatment, multiple blood products were delivered and a surgical suture was applied. | [5 empty] | The impella device was not received and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36220 | 1220648-2022-01557 | 15353041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15353041&pc=OZD | 08/06/2022 | 09/02/2022 |  | 09/02/2022 |  |  | 08/06/2022 | 09/02/2022 |  |  | 05/31/2024 |  | 09/02/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023220796 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/06/2022 |  |  |  |  | The complainant reported a 64 year old male patient presenting with cardiomyopathy for impella support. after just under 6 days of support, the pump stopped when the patient stepped on the white cord of the device. the pump could not be restarted, therefore, was removed and replaced. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36221 | 1220648-2022-01556 | 15353042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15353042&pc=OZD | 08/08/2022 | 09/02/2022 |  | 09/02/2022 |  |  | 08/08/2022 | 09/02/2022 |  |  | 03/31/2024 |  | 09/02/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023214686 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/08/2022 |  |  |  |  | The complainant reported a 59 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed an access site hematoma/bleed. as treatment, red blood cell blood products were delivered. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36222 | 1220648-2022-01558 | 15353090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15353090&pc=PYX | 08/05/2022 | 12/18/2024 |  | 09/02/2022 |  |  | 08/05/2022 | 09/02/2022 |  |  | 10/31/2022 |  | 09/02/2022 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2021009630 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/05/2022 |  |  |  |  | The complainant reported an 83 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. after removing the pump, an access site bleed had formed. as treatment, 2 units of blood products were required. | It was noted that prior to impella rp insertion, the patient was admitted to the hospital as a tier 1 trauma following a vehicle accident where he hit a parked car at approximately 30 miles per hour. the patient was found unconscious and unresponsive, and transported by an ambulance to the hospital. during the patient's hospitalization, the patient experienced an acute myocardial infarction and eventually developed acute right heart failure. the patient was referred for extracorporeal membrane oxygenation but was turned down. the decision was made to implant the impella rp. it was reported that after impella rp support was initiated, the impella rp had a blocked purge alarm. the purge cassette was changed, but the alarm continued. the impella rp was then exchanged. during the exchange, after the first impella rp was removed, the patient experienced bleeding at the impella access site. due to the bleeding, the patient was transfused with packed red blood cells, fresh frozen plasma and platelets. two mattress sutures were deployed at the site and bleed was controlled. it was noted that the patient expired while on support. given the limited information surrounding the details the patient's cause of death the impella cannot be excluded as a possible contributing factor. |  |  |  |  | The impella device was returned and an evaluation is underway. upon conclusion of our analysis, a supplemental mdr will be filed. | The device was returned and no abnormalities were observed on the repositioning sheath. while the cause of the access site bleeding could not be definitively determined, evidence suggests this to be related to the patient's condition. | A visual inspection of the pump revealed a kink in the catheter by the blue handle. no other abnormalities were observed. the cause of the high purge pressure was a kinked purge line in the patient cable. this report is being filed as part of a retrospective review of historical records correction to initial mfr 1220648-2022-01558 made to b2 - death, date of death . b5 - additional patient information, and purge block alarms were not previously reported. h1 - death reportability. h6 - additional codes added for replacement, death, and increase in pressure, and investigation for purge block.
36223 | 1220648-2022-01549 | 15334591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15334591&pc=OZD | 08/04/2022 | 09/01/2022 |  | 09/01/2022 |  |  | 08/04/2022 | 09/01/2022 |  |  | 04/30/2024 |  | 09/01/2022 |  | 3)TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023223692 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/04/2022 |  |  |  |  | The complainant reported a 76 year old male patient presenting for high risk percutaneous coronary intervention for impella support. after a failed attempt to deliver the device (due to calcified patient vasculature), the guide wire became kinked at the exit of the introducer. a snare was used via contralateral access, however, became entangled with the pump. vascular surgery intervened to remove the device from the patient. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36224 | 1220648-2022-01550 | 15334593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15334593&pc=OZD | 05/12/2022 | 10/26/2024 |  | 09/01/2022 |  |  | 08/04/2022 | 09/01/2022 |  |  | 01/31/2024 |  | 09/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023181755 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Ventricular Fibrillation | Asystole | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2022 |  |  |  |  | The complainant reported a 74 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, an access site bleed developed and prompted delivery of blood products. hemolysis and renal failure were also suspected, thus, haptoglobin and dialysis were administered. | Additional information was received indicating, the patient had an episode of ventricular fibrillation. during the insertion procedure, the patient developed asystole from ventricular fibrillation. venoarterial extracorporeal membrane oxygenation (va-ecmo) was added. after establishing impella and va-ecmo (ecpella), the percutaneous coronary intervention was completed. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information for the initial mfr 1220648-2022-01550 was provided in sections b5, d6a, d6b and h6. corrected information for the initial mfr 1220648-2022-01550 was provided in sections b3 and g2. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2022-01550 in error.
36225 | 1220648-2022-01551 | 15338767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15338767&pc=OZD | 08/12/2022 | 09/01/2022 |  | 09/01/2022 |  |  | 08/12/2022 | 09/01/2022 |  |  | 04/30/2024 |  | 09/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED. INC |  |  | US |  | 2023220603 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/12/2022 |  |  |  |  | The complainant reported a 57 year old female patient presenting for cardiomyopathy. during escalation from an impella cp to an impella 5.5 pump, an infection was discovered at the femoral access site. antibiotics were administered to the patient and a wound vac was utilized to aid in the treatment. the impella cp was subsequently explanted and the planned escalation to an impella 5.5 pump was completed. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36226 | 1220648-2022-01547 | 15319550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15319550&pc=OZD | 07/29/2022 | 08/29/2022 |  | 08/30/2022 |  |  | 07/29/2022 | 08/30/2022 |  |  | 04/30/2024 |  | 08/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023218242 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/29/2022 |  |  |  |  | The complainant reported an 85 year old white male patient presenting for impella support for acute myocardial infarction and cardiogenic shock. during insertion of the repositioning sheath, the "blue hub" was inserted too far and a bleed developed. the pump was removed and hemostasis was achieved via manual pressure, however, a contralateral balloon insertion was also required as intervention. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36227 | 1220648-2022-01545 | 15319574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15319574&pc=OZD | 08/23/2022 | 08/29/2022 |  | 08/30/2022 |  |  | 08/23/2022 | 08/30/2022 |  |  | 02/29/2024 |  | 08/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023191100 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 32 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2022 |  |  |  |  | The complainant reported a 32 year old male patient presenting for impella support. while on the device, the patient experienced hemolysis. as treatment, haptoglobin was administered to the patient and the condition subsequently improved. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36228 | 15310441 | 15310441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15310441&pc=OZD | 06/12/2022 | 08/12/2022 | 08/12/2022 | 08/29/2022 | 08/12/2022 | 08/29/2022 |  | 08/29/2022 |  |  |  |  | 08/29/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | PURGE CASSETTES |  |  | ABIOMED, INC. |  |  | US |  | 1663047 |  | 0043-0003 |  | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 22630 DA |  | Male |  |  |  | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | Followup |  |  |  |  | Y | HOSPITAL |  |  | Impella cassette leaking, no harm to patient. cassette given to rep for follow-up. | Impella cassette leaking, no harm to patient. cassette given to rep for follow-up.
36229 | 1220648-2022-01544 | 15310641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15310641&pc=OZD | 08/23/2022 | 08/29/2022 |  | 08/29/2022 |  |  | 08/23/2022 | 08/29/2022 |  |  | 03/31/2024 |  | 08/29/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023198343 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2022 |  |  |  |  | The complainant reported a 50 year old male patient presenting for myocarditis. while on impella support, the patient experienced hemolysis. as treatment, a blood transfusion was performed. the patient was later weaned from the device and the impella was explanted after six days of use without issue. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36230 | 1220648-2022-01546 | 15310698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15310698&pc=OZD | 08/23/2022 | 08/29/2022 |  | 08/29/2022 |  |  | 08/23/2022 | 08/29/2022 |  |  | 03/31/2024 |  | 08/29/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2023200209 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2022 |  |  |  |  | The complainant reported a 78 year old male patient presenting for myocarditis. while on impella support, it was noted that the patient experienced hemolysis. haptoglobin was administered to the patient and the performance level was changed from p-9 to p-7 to treat the reported condition. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of the device was not possible. should any new information be returned, a supplemental mdr will be filed.
36231 | 1220648-2022-01548 | 15310726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15310726&pc=OZD | 07/30/2022 | 10/07/2022 |  | 08/29/2022 |  |  | 07/30/2022 | 08/29/2022 |  |  | 03/31/2024 |  | 08/29/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023206779 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2022 |  |  |  |  | The complainant reported a 60 year old male patient presenting for impella support for acute myocardial infarction and cardiogenic shock. during support, the patient developed ischemia in the impella inserted leg. vascular surgery removed the pump and a thrombectomy and fasciotomy were performed. | [5 empty] | The impella device was returned and an evaluation is underway. a supplemental mdr will be filed at the conclusion of our analysis. | The device was received and no damage was identified. based on the clinical account, the cause of the reported ischemia was related to patient's condition.
36232 | 1220648-2022-01542 | 15303554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15303554&pc=OZD | 07/21/2022 | 05/28/2025 |  | 08/26/2022 |  |  | 07/29/2022 | 08/26/2022 |  |  | 04/30/2024 |  | 08/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213554 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Thrombosis/Thrombus | [9 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/29/2022 |  |  |  |  | The complainant a 20 year old male patient presenting with cardiomyopathy for impella support. during support, positive cultures were identified and the device was removed. upon removal, a large thrombus was found on the device with residual clot burden in the access site. manual extraction to remove residual clot was required as intervention. after a new pump was inserted, cultures came back negative. | Additional information was reported that the patient had an onset of hemolysis and pink/red tinged urine) that required one unit of packed red blood cells. urine cleared overnight after pump repositioning. placement suction alarms reported, but alarms resolved after patient received sedation and body repositioning. also, right groin bleeding was present. the patient had popped retention suture with coughing overnight. site was redressed and sutured. the patient was taken to the cath lab for repositioning of cp despite multiple unsuccessful bedside attempts. unable to successfully reposition pump in cath lab, but large amounts of bubbles present whenever pump moved visible on echo (similar to a bubble study). survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2022-01542 was provided in sections b2, b5, d4, d6, and h6. | Data logs showed there were all pump performance metrics were within specification toward the end of support, including the purge pressure and flow. no "air in purge" alarms occurred. since the product was not returned, the cause of the embolism and its relation to the impella were not determined. images received from the field confirmed the presence of thrombus around the cannula and on the pigtail. the impella was used beyond its design life of 8 days, which increases the risk of potential complications. regarding the infection, the impella device was confirmed to have been sterilized under validated conditions and the relation to impella was determined to be unknown. this report is being filed as part of a retrospective review of historical records.
36233 | 1220648-2022-01543 | 15306346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15306346&pc=OZD | 08/01/2022 | 08/26/2022 |  | 08/26/2022 |  |  | 08/01/2022 | 08/26/2022 |  |  | 03/31/2024 |  | 08/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023204601 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  |  | Death | Yes | No | Leak/Splash | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/02/2022 |  |  |  |  | The complainant reported a 65 year old female patient presenting for impella support for acute myocardial infarction and cardiogenic shock. during support, the physician discovered blood leaking from the impella. the leak was addressed with surgical tape. although the pump was subsequently switched out for a new one, the patient ultimately expired after enduring a "significant bleed" from a cannulation site on the patient's aorta. the blood leak from the initial impella was alleged as a contributing factor to the patient's death. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36234 | 1220648-2022-01541 | 15296771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15296771&pc=OZD | 08/22/2022 | 08/25/2022 |  | 08/25/2022 |  |  | 08/22/2022 | 08/25/2022 |  |  | 04/30/2024 |  | 08/25/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023218733 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/22/2022 |  |  |  |  | The complainant reported 73 year old male patient presenting for impella support. the patient experienced hemolysis while on the device. as treatment, the pump was prematurely removed and the patient was escalated to an impella 5.5. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36235 | 1220648-2022-01540 | 15282520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15282520&pc=OZD | 07/18/2022 | 08/24/2022 |  | 08/24/2022 |  |  | 07/25/2022 | 08/24/2022 |  |  | 11/30/2023 |  | 08/24/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023161096 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/25/2022 |  |  |  |  | The complainant reported an 88 year old female patient presenting for high risk percutaneous coronary intervention. the impella cp optical was selected for hemodynamic support. during the procedure, an impella access site bleed developed. as treatment, 2 units of blood products were delivered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36236 | 1220648-2022-01537 | 15268499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15268499&pc=OZD | 07/21/2022 | 08/22/2022 |  | 08/22/2022 |  |  | 07/27/2022 | 08/22/2022 |  |  | 03/31/2024 |  | 08/22/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023214746 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/27/2022 |  |  |  |  | The medical center had an impella cp support that was ongoing, for a patient admitted in shock, with signs of hemolysis despite pump repositioning. the team explanted and escalated to the impella 5.5 and began dialysis. additionally the patient was infused with blood product replacement in the amount of 9 units of blood. | [5 empty] | The medical center discarded the impella pump product at explant and as such no benchtop analysis and hemolysis testing is possible. should more information be shared that does lead to root cause, a final report will be filed. the failure mode will be monitored and trended.
36237 | 1220648-2022-01538 | 15268507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15268507&pc=OZD | 07/28/2022 | 09/21/2024 |  | 08/22/2022 |  |  | 07/28/2022 | 08/22/2022 |  |  | 04/30/2024 |  | 08/22/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023212729 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes |  |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Leak/Splash | Pumping Stopped | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | Hospitalization | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/28/2022 |  |  |  |  | The medical center had an impella 5.5 pump support that was ongoing for a patient with cardiomyopathy, when on day 5 it stopped. the pump was explanted and not replaced due to plan of care discussions with the medical team and the family. | In addition to the originally reported pump stop, the purge side arm was noted as leaking purge fluid. the issue was remedied by configuring a purge system bypass. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  | The medical center did not return the impella pump for benchtop analysis. the team removed the pump as part of care/comfort conversations with the family and then discarded the pump. no investigation to root cause is possible. should more information be shared and a root cause determined, a final report will be filed. the failure mode will be monitored and trended. | (h3) the investigation into the reported "purge leak - pump and pump stop" has been completed since the original report was filed. the medical center did not return any impella products for benchtop an analysis; only the clinical report was evaluated. during data logs analysis it showed the pump stopped immediately upon the sudden occurrence of the "impella stopped - motor current high" alarm. preceding the pump stopped, the motor current gradually increased over several hours and the pump flow rapidly became saturated, losing all pulsatility. the purge pressure rapidly increased and remained elevated. after reviewing the clinical information, it was determined that the cause of the pump stop was most likely blood ingress due to the lack of purge fluid. the cause of the purge leak was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records section d4: unique identifier (udi) #. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6.
36238 | 1220648-2022-01539 | 15268651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15268651&pc=OZD | 07/14/2022 | 08/22/2022 |  | 08/22/2022 |  |  | 07/28/2022 | 08/22/2022 |  |  | 12/31/2023 |  | 08/22/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023167751 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/28/2022 |  |  |  |  | The medical team in japan had a vessel damage at the insertion of the 14fr introducer sheath for the impella cp. the native anatomy was known to be too small for the sheath. the team performed angioplasty and stented the area. addtionally the team observed thrombus as the area of the vessel dissection when the pump was explanted. | [5 empty] | The medical team in japan discarded the impella pump and introducer sheath. no benchtop analysis is possible. the root cause can not be determined, though due to the known peripheral arterial disease and small vasculature, patient condition contributed to the outcome. the failure mode will be monitored and trended.
36239 | MW5111503 | 15241442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15241442&pc=KFM | 08/01/2022 | 08/15/2022 |  | 08/17/2022 |  |  |  | 08/17/2022 |  |  |  |  | 08/17/2022 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA |  |  | KFM | [5 empty] | ABIOMED, INC. |  |  | US |  | 378794 |  |  |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 65 YR |  | Female | 86 KG | Non Hispanic | White | Death | Yes | Yes | Crack | Fluid/Blood Leak | [5 empty] | Exsanguination | [12 empty] | Death | [28 empty] | Initial submission | [9 empty] | Patient undergoing cabg (coronary artery bypass graft) when it was noted the patient was bleeding from impella. blood was dripping under drapes down onto floor. staff member noticed blood around feet and told surgeon. patient hypotensive. patient had significant blood loss before it was recognized impella had crack in it. patient expired. fda safety report id# (b)(4).
36240 | 1220648-2022-01535 | 15224928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15224928&pc=OZD | 07/21/2022 | 08/15/2022 |  | 08/15/2022 |  |  | 07/21/2022 | 08/15/2022 |  |  | 05/31/2024 |  | 08/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023218321 |  | 0550-0008 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/21/2022 |  |  |  |  | The medical center had an impella 5.5 support on-going for 5 days when it was explanted due the patient improvement. after the explant the abiomed field representative was notified of valve damage presumed to be caused by the impella pump. the aortic valve leaflet was damaged as observed by echo. the patient of note had the 5.5 pump repositioned on day 2 of the support, this repositioning may have cause the valve damage. the conclusion could not be confrmed. the patient was monitored by cardiac imaging in the mri suite. | [5 empty] | The medical center did not return the impella pump product or echo or mri imaging for analysis. no root cause could be determined for the impella pump presumed valve damage. should the medical team furnish more details, that lead to root cause, a final report will be filed.
36241 | 1220648-2022-01536 | 15230335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15230335&pc=OZD | 07/18/2022 | 08/15/2022 |  | 08/15/2022 |  |  | 07/18/2022 | 08/15/2022 |  |  | 04/30/2024 |  | 08/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023216131 |  | 0550-0008 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/18/2022 |  |  |  |  | A patient was admitted for the planned surgeries of bypass grafts (cabg), mitral valve replacement, and left atrial appendage closure. the team made the choice to place an impella 5.5 for hemodynamic support of the surgical patient as he was in shock post cardiotomy. the pump stopped after 6 days and was replaced. | [5 empty] | The medical center has not yet returned any impella pump product for investigation. should product be returned and benchtop analysis determine a root cause of the pump stop, a final report will be filed. the failure mode will be monitored and trended.
36242 | 1035166-2022-00094 | 15198412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15198412&pc=DYB | 07/13/2022 | 12/21/2022 |  | 08/10/2022 |  |  | 07/14/2022 | 08/10/2022 |  |  | 08/01/2024 |  | 08/10/2022 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | DP-13777 |  | 0052-3021 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 85 YR |  | Female |  |  |  | Malfunction | No | Yes | Break | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a 85 year old female underwent a procedure for placement of the impella rp. a complaint was entered inclusive of your 23fr sheath. we expect the product to be returned, but to date nothing has been returned. the field rep noted the following: the abiomed 23fx30cm sheath was then inserted without issue. heparin was not given per the request of the dr., the dilator was then removed. upon removal of the dilator, a crack in the 23f abiomed sheath became evident, and bleeding was present. a replacement 23fx30cm sheath was quickly located and the broken sheath was replaced without issue. the patient expired when family chose to withdraw care. | [5 empty] | Our investigation is still in progress, follow-up report will be submitted if we found any further additional information. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Additional information received 08/15/2022: the family withdrew care and patient expired at 0936 on (b)(6) 2022. the family withdrew care as the patient had decompensated after tricuspid valve repair. the patient suffered from atrial fibrillation, hypotension, bradycardia, chf, ef of 55%, copd, dilated rv, and severe tricuspid regurgitation. the product was not a factor in the cause of death. one 23f abiomed introducer sheath was received. there were no other accessories. traces of blood were found on the sheath. crack was observed on the sheath. according to the event description summary the sheath was cracked during removal of the dilator. upon evaluation of the returned product, during the visual inspection with the naked eye it was found that the sheath presented 2 cracks, 10-11mm in length, approximately 20cm from the distal tip. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. returned device analysis revealed the sheath was within manufacturing specification. two 10-11mm cracks were found on the score line of the sheath shaft. according to the event description summary, upon removal of the dilator, a crack in the 23f abiomed sheath became evident. it is possible that the dilator removal was performed on an angle and damaged the score lines of the sheath. according to the device history record, the sheath passed all in-process and qa final inspection steps before shipping to the customer including visual inspection and leak test. per procedure (abiomed introducer sheath in-process and final) sample size: 100% with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. foreign matter (fm)/fiber is limited to 2 particles < 0.10 mm2 or 1 particle < 0.2 mm2 per ifu dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36243 | 1220648-2022-01534 | 15198623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15198623&pc=OZD | 12/23/2021 | 10/29/2024 |  | 08/10/2022 |  |  | 07/11/2022 | 08/10/2022 |  |  | 08/31/2023 |  | 08/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022127524 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  |  | Asian | Death | Yes | Yes | Positioning Failure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2022 |  |  |  |  | The complainant reported an 85 year old male patient presenting with high risk percutaneous coronary intervention for impella support. during support the patient exhibited ventricular fibrillation and required a direct cardioversion. | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, device had positioning issue and patient experienced access site bleeding and ventricular fibrillation. the pcps and impella was removed on (b)(6) 2021. the removal of impella was not planned, but when the position of the impella in the intensive critical unit (icu) was being adjusted, it pulled out of the aortic valve and was removed as is. after removing the impella, it was replaced with an intra-aortic balloon pump (iabp). on (b)(6) 2024, japanese registry for percutaneous ventricular assist device (j-pvad) reported that on (b)(6) 2021, patent experienced bleeding, and blood transfusion was needed and on (b)(6) 2021, the patient had ventricular tachycardia (vt). patient later expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information for the initial mfr 1220648-2022-01534 has been provided. corrected information for the initial mfr 1220648-2022-01534 was provided.
36244 | 15155658 | 15155658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15155658&pc=OZD | 07/11/2022 | 07/19/2022 | 07/29/2022 | 08/03/2022 | 07/29/2022 | 08/03/2022 |  | 08/03/2022 |  |  |  |  | 08/03/2022 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 363274 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 005042 |  | No |  |  | Not Available | [6 empty] | User Facility report | Yes | Other Health Care Professional | [8 empty] | Unknown |  |  |  | Malfunction | No | Yes | Material Separation | Physical Resistance/Sticking | [5 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y |  |  |  | Procedure: impella assisted pci. impella was turned off and removed from heart. upon trying to remove impella completely, impella became lodged and stuck in sheath. patient had received angiomax during procedure. unable to advance or remove impella despite multiple attempts. fluoroscopy showed impella stuck in sheath and not moving. unable to use manta closure device due to being unable to access sheath for wire insertion. decision was made by physician to remove entire sheath with impella device inside and hold manual pressure and then apply femstop device. femstop was applied under patient hips at proper position by the 2 rns in the room. sheath with impella stuck inside was then removed and manual pressure was held by doctor and then by tech, rt. impella device inspected upon removal and was still lodged in sheath intact, however further manipulation outside the body caused impella to come apart. it did not come apart inside the patient.
36245 | 1220648-2022-01531 | 15114095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15114095&pc=OZD | 07/20/2022 | 07/27/2022 |  | 07/27/2022 |  |  | 07/20/2022 | 07/27/2022 |  |  | 01/31/2024 |  | 07/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023181705 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/20/2022 |  |  |  |  | The medical team in japan had an impella cp support for 5 days when the pump was explanted. the team later noted that the explant site had a flow disorder attributed to either thrombus or embolus. the next day thrombectomy was performed to the limb and blood flow improved. | [5 empty] | The medical center in japan discarded all impella pump and introducer sheath product. no benchtop analysis to root cause is possible. the failure mode will be monitored and trended.
36246 | 1220648-2022-01532 | 15114280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15114280&pc=OZD | 07/20/2022 | 07/27/2022 |  | 07/27/2022 |  |  | 07/20/2022 | 07/27/2022 |  |  | 01/31/2024 |  | 07/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023178480 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/20/2022 |  |  |  |  | The medical team in japan had an impella cp support in (b)(6) 2022 that was successful til patient expired from underlying and unrelated condition. the team later in 7/2022 reported upon an access site bleed and limb ischemia via the jpvad registry. the bleed was treated by rbc infusion and the limb ischemia by external bypass. | [5 empty] | The medical center in japan discarded all impella pump product months ago. no benchtop analysis to root cause is possible. the failure mode will be monitored and trended.
36247 | 1220648-2022-01533 | 15114332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15114332&pc=OZD | 07/21/2022 | 07/27/2022 |  | 07/27/2022 |  |  | 07/21/2022 | 07/27/2022 |  |  | 02/29/2024 |  | 07/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023191529 |  | 0048-0024-JP |  | No |  |  | No |  |  |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/21/2022 |  |  |  |  | The medical team in japan had an impella cp support ongoing for a patient admitted in shock. there was a diagnosis made of a left ventricle rupture. the rupture/perforation of the tissue was of either the underlying myocardial damage and/or the placement of the impella cp pigtail. the team could not determine. the surgical intervention resolved the issue and the patient has survived. | [5 empty] | The medical center in japan has not returned the impella pump product for analysis. no imaging was shared, though asked for. a full analysis to root cause is not possible with known information. should more be shared that leads to root cause, a final report will be filed. the failure mode will be monitored and trended.
36248 | 1220648-2022-01524 | 15104538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15104538&pc=OZD | 07/05/2022 | 09/18/2024 |  | 07/26/2022 |  |  | 07/05/2022 | 07/26/2022 |  |  | 04/30/2024 |  | 07/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED INC. |  |  | US |  | 2023215740 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/05/2022 |  |  |  |  | The medical team in (b)(6) had an impella cp support that was ongoing with signs of hemolysis. the plasma free hemoglobin (pfhg) levels were high as was bilirubin levels. the team assessed pump position and began dialysis care for the renal issue. the patient's impella support proceeded til the patient expired from underling colon ischemia. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation is on-going at this time. no product has been returned yet for investigation. upon completion of the investigation, a final report will be filed. | The impella pump was returned for benchtop analysis and hemolysis testing. when the pump was returned there was adhered biomaterial. when the biomaterial was removed the pump performed and passed benchtop hemolysis testing. the root cause of the hemolysis was the patient condtition and ingression of the biomaterial. the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01524 was provided in sections b2, b5, d6a, d6b, h1, and h6.
36249 | 1220648-2022-01525 | 15104834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15104834&pc=OZD | 07/05/2022 | 07/26/2022 |  | 07/26/2022 |  |  | 07/05/2022 | 07/26/2022 |  |  | 02/29/2024 |  | 07/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023195210 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/05/2022 |  |  |  |  | The medical center had an impella cp support which was supporting a patient, in need of hemodynamic support during coronary angioplasty, via an axillary access. the pump was explanted after over an hour of support due to the limb becoming ischemic. the limb ischemia was diagnosed by the lost pulses in the arm. | [5 empty] | The medical team discarded the impella product at explant. no benchtop analysis to root cause is possible. the patient was known to have peripheral arterial disease, and as such underlying native anatomy could have contributed to the limb ischemia. should more information be shared, that leads to definitive cause, a final report will be filed.
36250 | 1220648-2022-01526 | 15104894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15104894&pc=OZD | 07/07/2022 | 07/26/2022 |  | 07/26/2022 |  |  | 07/07/2022 | 07/26/2022 |  |  | 11/30/2023 |  | 07/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023161716 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/07/2022 |  |  |  |  | The medical team had an impella cp support when at the explant, there was a harm to the limb accessed for the placement of the cp. the limb was seen to be thrombosed and needed intervention in the form of angioplasty. the team placed stents to the superficial femoral artery. | [5 empty] | All product was discarded at explant. no benchtop analysis to root cause is possible. should more details be shared that lead to root cause of the thrombosis, a final report will be filed.
36251 | 1220648-2022-01527 | 15105059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15105059&pc=OZD | 07/07/2022 | 08/17/2022 |  | 07/26/2022 |  |  | 07/07/2022 | 07/26/2022 |  |  | 03/31/2024 |  | 07/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED INC. |  |  | US |  | 2023203818 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2022 |  |  |  |  | The medical team in the eu had an impella cp support along with ecmo for a patient in need of hemodynamic support. after over 2 days of support the team removed the pump due to limb ischemia. there was no additional treatment beyond the explant. the team attributed the ischemia to thrombus burden at the implant site. | [5 empty] | The medical center in the eu has not yet returned the impella product for analysis. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation has concluded. the medical team in the eu never returned the impella pump nor introducer sheath for analysis. as such, the root cause of the limb ischemia and the causal relationship to the impella could not be proven. the assumption is that patient condition may have contributed to the limb injury. the failure mode will be monitored and trended.
36252 | 1220648-2022-01528 | 15105237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15105237&pc=OZD | 07/07/2022 | 07/26/2022 |  | 07/26/2022 |  |  | 07/07/2022 | 07/26/2022 |  |  | 06/30/2023 |  | 07/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022096610 |  | 0550-0008 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Valvular Insufficiency/ Regurgitation | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/07/2022 |  |  |  |  | The medical center had an impella support in 2021. the support was successful and no complaints or allegations were made at the time of support. later in 2022 there was a publication made inclusive of the impella support in which there was documented aortic valve damage that was treated by placement of the aortic valve ring. there is a possible correlation of the impella use/placement contributing or causing the valve damage and aortic insufficiency (ai) noted in the publication. | [5 empty] | Root cause of the valve injury has not been determined. followup and additional information has been requested, but not yet furnished. this is inclusive of imaging. should more details be shared that lead to a root cause, a final report will be filed.
36253 | 1220648-2022-01529 | 15105649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15105649&pc=OZD | 07/14/2022 | 08/22/2022 |  | 07/26/2022 |  |  | 07/14/2022 | 07/26/2022 |  |  | 06/30/2023 |  | 07/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022102403 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2022 |  |  |  |  | The medical center had an impella 5.5 support ongoing for a patient with cardiomyopathy. the support ended on day 16 when they had observed the bilateral lower limbs to be clotted. the team was unsure as to why the patient was hypercoagulable but both ecmo and 5.5 were explanted. after the explant a stroke was diagnosed. there was additional neurological surgery to intervene and then the vad was placed. | [5 empty] | The medical center discarded the impella pump product at explant. requests for additional information and imaging have not been fulfilled. a root cause has not been determined. should more be shared, and a definitive root cause be determined, a final report will be filed. | The medical center did not return the impella pump, however the team did return console logs. the team did relay more clinical information. due to the bilateral limb thrombosis, it was determined the ischemic stroke was most likely due to native patient condition. the failure mode will be monitored and trended.
36254 | 1220648-2022-01530 | 15105858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15105858&pc=OZD | 07/14/2022 | 09/20/2024 |  | 07/26/2022 |  |  | 07/14/2022 | 07/26/2022 |  |  | 03/31/2024 |  | 07/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023201417 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Pressure Problem | [5 empty] | Cardiac Perforation | [12 empty] | Hospitalization | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 07/14/2022 |  |  |  |  | The medical center had an impella 5.5 support in the operating suite ongoing for a patient with an aortic dissection. the 5.5 perforated the left ventricle. the posterior aspect had perforated, and due to the chest being open in the or, the suture was applied and patient survived the surgery and was discharged to the unit for continued monitoring and care. | It was also reported during support that there was a high purge pressure alarm. a purge cassette change was performed and the issue resolved. there was no patient harm reported. | Additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  | The medical center has not yet returned the impella pump nor any imaging for evaluation. upon the completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation into the left ventricle perforation has concluded. the impella pump product was not returned. the clinical report alone was analyzed. the root cause was determined to be patient condition related. the pump was known to be placed in less than optimal, as it was due to shallow graft anastamosis post the patient's aortic dissection. the failure mode will be monitored and trended. | The investigation for the high purge pressure has been completed. the clinical reported provides details to determine the root cause. therefore, the cause of the high purge pressure was due to biomaterial, as the purge pressure slowly increased after concentration of bicarbonate increased. the instructions for use provides troubleshooting for high pressure issues. it states unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records. h6-medical device codes, component code, investigation findings, investigation conclusion. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01530 was provided in sections b2, b5, d6a, d6b, h1, and h6.
36255 | 1220648-2022-01523 | 15084553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15084553&pc=OZD | 07/04/2022 | 08/17/2022 |  | 07/22/2022 |  |  | 07/04/2022 | 07/22/2022 |  |  | 04/30/2024 |  | 07/22/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023211959 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/04/2022 |  |  |  |  | The medical team had an impella 5.5 pump stop occurr after just 5 days. this was the patient's 2nd 5.5 pump during hosptialization for cardiomyopathy. there was no noted harm or injury. | [5 empty] | The medical center has not yet returned the impella 5.5 for analysis and testing. upon the return of the product and investigation completion, a final report will be filed. the failure mode will be monitored and trended. | Since the original report was filed, the investigation into the pump stop has concluded. the pump was returned for benchtop analysis. the pump was found to have a kinked purge line that allowed for blood ingress. the blood ingress was due to the technique/improper use. the failure mode will be monitored and trended.
36256 | 15084900 | 15084900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15084900&pc=OZD | 07/06/2022 | 07/20/2022 | 07/20/2022 | 07/22/2022 | 07/20/2022 | 07/22/2022 |  | 07/22/2022 |  |  |  |  | 07/22/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED INC. |  |  | US | [5 empty] | No answer provided |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 30660 DA |  | Male |  |  |  | Malfunction | No | Yes | Loose or Intermittent Connection | Migration or Expulsion of Device | Use of Device Problem |  |  |  |  | Cardiac Arrest | Low Blood Pressure/ Hypotension | Failure of Implant | Respiratory Failure | [34 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Reporter called into the room because the impella was alarming. the patient was a dnr (do-not-resuscitate)/dni (do-not-intubate). blood pressure (bp) rapidly decreased and he began turning blue with what appeared to be agonal respirations. primary rn had already turned the impella down to p level of p2 from previous level of p9. another rn was instructed to go to the call room and bring back a doctor to assess the patient. the impella was originally at 96cm of inserted depth and then migrated to 80cm of depth. it was then advanced back to 79cm of depth. epi was increased to 0.3 mcg/kg/min from 0.08 mcg/kg/min. on further assessment the patient was found pulseless, no heart sounds, and in a paced rhythm. 2nd md confirmed there was no cardiac activity and that the impella had migrated. the patient was pronounced dead. it was later discovered the touhy borst valve was loose and the impella could easily be repositioned without any resistance. the impella was not sutured to the thigh, nor was the knee immobilizer in use. the impella did not have any additional anchors nor tegaderm dressings to secure it. the touhy borst valve lock was turned in a clockwise manner until it stopped and then tested. it functioned as designed and the catheter became unmovable with the valve closed. it was then turned counterclockwise which released the catheter and it once again returned to an adjustable state. there did not appear to be any malfunction with the valve.
36257 | 1035166-2022-00085 | 15076235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15076235&pc=DYB | 06/04/2022 | 12/08/2022 |  | 07/21/2022 |  |  | 06/23/2022 | 07/21/2022 |  |  | 07/01/2024 |  | 07/21/2022 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | DP-13327 |  | 0052-3021 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | [6 empty] | Cardiogenic Shock | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | On (b)(6)2022, the patient a 65 year old, male was presented to the physician with deteriorated hemodynamics. the physician reported kinking of the impella catheter (23fr introducer) and guidewire when attempting to pass the impella device into the pulmonary artery. the introducer broke within the patient's body. the pump and introducer were removed and replaced with a new pump/introducer. no intervention was required, as the introducer was manually removed. the procedure was successfully completed when using a new impella and new introducer. the procedure took 30 minutes due to using new devices. it was reported there was no known harm or injury to the patient. the physician discarded the introducer and guidewire. no additional information is available. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | The following sections were updated in follow-up 1: b4, g3, g6, h2, h6 and h10 the device was not returned for analysis, as it was discarded. the investigation was focused on a review of product documentation. the device used in treatment. the device was not returned for analysis, therefore, the exact cause of the product issue cannot be determined and the clinical observation could not be confirmed. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. however, the following controls are in place to mitigate the reported product issue. as per procedure (abiomed introducer sheath in-process and final inspection), the devices in this lot were inspected as follows: 9.2.2 visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. per section 12.2, visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. the instructions for use (ifu) informs the user: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. abiomed provides the ifu for this product. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36258 | 1220648-2022-01522 | 15065689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15065689&pc=OZD | 06/28/2022 | 07/20/2022 |  | 07/20/2022 |  |  | 06/28/2022 | 07/20/2022 |  |  | 08/31/2022 |  | 07/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1496670 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/28/2022 |  |  |  |  | The medical team in japan had an imeplla cp support ongoing with an access site bleeding. the bleed was treated with blood product infusion when manual pressure and skin suture failed. | [5 empty] | The medical center in (b)(6) discarded all impella product. no investigation with benchtop analysis is possible. should the team furnish more details that lead to root cause, a final report will be filed. failure mode will be monitored and trended.
36259 | 15052115 | 15052115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15052115&pc=OZD | 07/11/2022 | 07/13/2022 | 07/13/2022 | 07/19/2022 | 07/13/2022 | 07/19/2022 |  | 07/19/2022 |  |  |  |  | 07/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Yes |  |  | Not Available | [6 empty] | User Facility report | No | Risk Manager | [8 empty] | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y |  |  |  | Liquid was on the floor, device failure and failure of purge to keep impella functioning.
36260 | 1220648-2022-01518 | 15052990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15052990&pc=OZD | 06/25/2022 | 07/19/2022 |  | 07/19/2022 |  |  | 06/25/2022 | 07/19/2022 |  |  | 03/31/2024 |  | 07/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023202315 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/25/2022 |  |  |  |  | The medical team in (b)(6) had an impella cp inserted for hemodynamic support of a patient admitted and suffering from a myocardial infarction. the impella insertion site had a bleed treated by surgical hemostasis and 6 units of blood. | [5 empty] | The medical team in (b)(6) did not return the impella product for analysis or benchtop testing. no further clinical details have been shared. should more information be shared that leads to root cause of the bleed, a final report will be filed.
36261 | 1220648-2022-01519 | 15053154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15053154&pc=OZD | 06/25/2022 | 09/17/2024 |  | 07/19/2022 |  |  | 06/25/2022 | 07/19/2022 |  |  | 09/30/2024 |  | 07/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED INC. |  |  | US |  | 2022141818 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | Required Intervention | Hospitalization |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/25/2022 |  |  |  |  | The medical team in the eu had an impella 5.5 support ongoing for hemodynamic support along with ecmo for a patient admitted in cardiogenic shock. on day 4 the impella pump stopped. the patient had coughed moved position in the bed. after the pump stop they were able to restart it with high motor current alarms. the team initiated pump replacement. | Additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The medical team in germany has returned the pump for analysis. the investigation is on-going at this time. upon completion, a final report will be submitted. | The pump was returned amd benchtop testing was performed. the pump was found to have biomaterial adhered to it and wrapped around the impella. the cause of the pump stoppage was biomaterial, and as such the pump stop was due to patient condition. the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01519 was provided.
36262 | 1220648-2022-01520 | 15055091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15055091&pc=OZD | 06/25/2022 | 10/14/2024 |  | 07/19/2022 |  |  | 06/25/2022 | 07/19/2022 |  |  | 03/31/2024 |  | 07/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023209599 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Failure | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/25/2022 |  |  |  |  | The medical team had an impella cp placed via single access at the right femoral artery for hemodynamic support. there was an access site perforation that was treated by artery repair and a surgical cut-down. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited suction issues indicating low flow. the staff reduced flows to 1-1.3 l/min on p9, but the motor current was nearly flat. the staff was unable to restore flows to expected range or pulsitility to motor current. the patient also experienced limb ischemia and heparin was added to the purge. due to these issues, the physician decided to explant the pump and replace with another impella cp. during replacement the doctor encountered positioning issues, the orignal impella positioning was too deep into the ventricle and exited the tip of the sheath. the impella could not pass through, so a peelaway was utilized. once the cook sheath was in place, the pigtail catheter was used to cross the aortic valve. the first impella was pulled back into the ascending aorta, while the second impella was inserted over the wire in standard fashion. it was reported that the patient survived the explant and the procedure. |  |  |  |  | The medical center returned for analysis the impella pump, but no introducer sheath. the root cause of the access site damage was presumed to be a use error. the failure mode will be monitored and trended. | Low flow: the investigation for the reported low flow has been completed. the product was returned. per the results of the clinical review and investigation: pump was returned and inspected in the bhl. biomaterial was visualized to be adhered on the impeller. the pump was run at p3 and p7 and the flow was continuously .2 l/min. the ingestion of biomaterial and improper positioning most likely contributed to the low pump flow. limb ischemia and positioning issues the impella device was returned, and an investigation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: limb ischemia: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. corrections: b5: updated to address additional reported issues. g1: corrected manufacturer contact.
36263 | 1220648-2022-01521 | 15055585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15055585&pc=OZD | 06/26/2022 | 07/19/2022 |  | 07/19/2022 |  |  | 06/26/2022 | 07/19/2022 |  |  | 02/29/2024 |  | 07/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023193766 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/26/2022 |  |  |  |  | The medical team in japan had an impella cp support in which they treated hemolysis presumed to be caused by impella. the team infused blood products of unknown amount and attempted to adjust pump position. when adjusting position the pump was mistakenly pulled to the aorta. when patient was seen to be stable without impella, they removed the pump and did not replace it. | [5 empty] | The medical team in japan discarded the impella product at explant. no benchtop analysis is possible to determine root cause of the presumed hemolysis. should the team furnish more details that lead to a root cause. a final report will be filed.
36264 | 1220648-2022-01516 | 15028395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15028395&pc=OZD | 07/05/2022 | 10/25/2024 |  | 07/15/2022 |  |  | 07/05/2022 | 07/15/2022 |  |  | 08/31/2023 |  | 07/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022122120 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional |  |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Hospitalization | Other |  |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/05/2022 |  |  |  |  | The medical team in japan had an ecmo and impella cp support in which the patient expired in (b)(6) 2022. the death was deemed unrelated to the use of impella. later in (b)(6) 2022 the jpvad registry updated that the team had observed a cva during support. the patient had suffered an ischemic stroke followed by a hemorrhagic one. the team noted a hemorrhagic event after impella placed with anticoagulation. the anticoagulation is md responsibility and normal practice. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  | The team in japan discarded all product months ago in (b)(6) 2022. no investigation is possible. should any new information be shared that leads to root cause, a final report will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01516 was provided in sections b2, b5, h1, and h6. | This report is being filed as part of a retrospective review of historical records. the following sections were updated from the original mdr 1220648-2022-01516: b2, b5, d6a, g3, h1, h6, h8.
36265 | 1220648-2022-01517 | 15028636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15028636&pc=OZD | 07/12/2022 | 08/24/2022 |  | 07/15/2022 |  |  | 07/12/2022 | 07/15/2022 |  |  | 02/29/2024 |  | 07/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023202866 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/12/2022 |  |  |  |  | The medical team reported an impella cp pump stop after a coronary angioplasty. the pump was removed and replaced without any harm to the patient. later in the pump investigation, upon the benchtop, the engineering team discovered a broken impeller blade. | [5 empty] | The investigation into the broken impeller blade is on-going at this time. upon completion of the investigation, a final report will be filed. | The pump and data logs involved in this event were returned for analysis. broken blades were found on the device. the data logs confirm a pump stop with successful restart and low flows. evidence suggests a brief interaction with another device was the cause of the pump stop and broken blades, however, fluoroscopic imagery of the pump was not returned to confirm this.
36266 | 1220648-2022-01515 | 15014347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=15014347&pc=OZD | 06/19/2022 | 07/22/2022 |  | 07/14/2022 |  |  | 06/19/2022 | 07/14/2022 |  |  | 03/31/2024 |  | 07/14/2022 |  | TEMPORARY NON-ROLLER TYPE LERFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023206387 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/19/2022 |  |  |  |  | The medical team had a younger patient with underlying condition of muscular dystrophy and coronary artery disease supported by an imeplla cp. the patient was transferred and care escalated from the femorally accessed cp to the axillary impella 5.5. the 5.5 supported for 5 days when it stopped and could not be troubleshot by aic replacement. the 5.5 was replaced. | [5 empty] | The investigation into the impella pump stop is on-going at this time. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation into the pump stop has concluded. the product was returned and the pump analysis was completed on the benchtop. the analysis revealed blood ingress into the pump, via a hole created from improper use, by medical staff. the failure mode will be monitored and trended.
36267 | 1220648-2022-01514 | 14978344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14978344&pc=OZD | 06/10/2022 | 10/14/2024 |  | 07/11/2022 |  |  | 06/16/2022 | 07/11/2022 |  |  | 02/29/2024 |  | 07/11/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023193812 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aneurysm | Hematoma | Ventricular Fibrillation | [10 empty] | Hospitalization | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/16/2022 |  |  |  |  | The medical center had enrolled a patient in the stemi dtu study. the support with an impella cp was a successful one for 5 hours of support for the angioplasty. at the point of sheath/pump explant the team noted no harm. later study documentation noted a pseudoaneurysm treated by thrombin injection and imaging followup. the team attributed the need for the treatment to the impella sheath placement. | It was also reported the patient experienced ventricular fibrillation during the percutaneous coronary intervention (pci) procedure and the patient was defibrillated one time, then the patient transitioned into bradycardia. 2.5mg iv metoprolol was administered and the patient fully transitioned into normal sinus rhythm. |  |  |  |  | No product, data logs, or imaging have been returned for analysis. investigation to the root cause of the pseudoaneurysm is possible. should more information be shared that leads to cause, a final report will be filed. | Since the original report was filed the study team for the stemi-dtu has a new report of another adverse event, of a hematoma, documented. the hematoma was at the impella access site that also had the prior reported upon pseudoaneurysm. the hematoma was only treated by manual pressure. the failure mode will be monitored and trended. no root cause can be determined as no product was returned for analysis. | Additional information for the initial mfr was provided in sections b1, b2, b5, d6a, d6b, g1, h1, and h6. the investigation for the reported ventricular fibrillation has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the ventricular fibrillation was most likely patient condition related. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
36268 | 1220648-2022-01507 | 14959411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14959411&pc=OZD | 06/17/2022 | 07/08/2022 |  | 07/08/2022 |  |  | 06/17/2022 | 07/08/2022 |  |  | 03/31/2024 |  | 07/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023203918 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Retroperitoneal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/17/2022 |  |  |  |  | The complainant reported an 85 year old female patient presenting for high risk percutaneous coronary intervention for impella support. a retroperitoneal bleed had formed after removing the 14 fr introducer and advancing the impella's repositioning sheath. the patient required 8 units of blood products and the impella was removed and replaced. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36269 | 1220648-2022-01508 | 14959492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14959492&pc=OZD | 06/24/2022 | 07/08/2022 |  | 07/08/2022 |  |  | 06/24/2022 | 07/08/2022 |  |  | 12/31/2023 |  | 07/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023174836 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/24/2022 |  |  |  |  | The complainant reported a 62 year old male patient presenting for high risk percutaneous coronary intervention for impella support. during placement of 14 fr introducer, the patient's vessel was dissected. a balloon stent was deployed to control the bleed and a covered stent with viabahn to seal the arteriotomy was put in place, thereafter. an unknown amount of blood products were also delivered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36270 | 1220648-2022-01510 | 14960955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14960955&pc=OZD | 06/12/2022 | 07/18/2022 |  | 07/08/2022 |  |  | 06/12/2022 | 07/08/2022 |  |  | 01/31/2024 |  | 07/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023186126 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2022 |  |  |  |  | The medical team had an impella 5.5 support on-going in which there was a purge leak that lead to the pump stop. the 5.5 pump stop occurred on day 6 of support, which is within the indicated use time. the team then replaced the pump and the patient survived and hemodynamic support continued. | [5 empty] | The medical center has returned the impella 5.5 pump for investigation and analysis. the analysis is on-going at this time. upon completion, the final report with root cause will be drafted. | Since the original report was filed, the investigation into the pump stop has concluded. product was returned and investigation revealed the pump had stopped due to an ingress of blood into the pump's purge line. the failure mode will be monitored and trended.
36271 | 1220648-2022-01509 | 14960957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14960957&pc=OZD | 06/12/2022 | 07/08/2022 |  | 07/08/2022 |  |  | 06/12/2022 | 07/08/2022 |  |  | 11/30/2023 |  | 07/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED INC. |  |  | US |  | 2023161233 |  | 005062 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/12/2022 |  |  |  |  | The medical center had an imeplla cp support escalated to the impella 5.0. after insertion of the 5.0 there was an observation made of a cva. the cva prompted heparin reduction and imaging/analysis. the team observed a cerebral artery (mca) occlusion. an ischemic cva was survived, but three days later the team made the choice to withdraw all care. | [5 empty] | The medical center has not returned for analysis any impella pump product, data logs, or imaging. no root cause of the cva has been determined. should the medical team share any product or additional details that lead to a definitive root cause, a final report will be filed.
36272 | 1220648-2022-01513 | 14961320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14961320&pc=OZD | 06/25/2022 | 10/15/2024 |  | 07/08/2022 |  |  | 06/25/2022 | 07/08/2022 |  |  | 02/29/2024 |  | 07/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203010 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2022 |  |  |  |  | The complainant reported a 71 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed an uncontrollable bleed from the impella's repositioning sheath, despite intervention, including 5 units of blood product delivery. the patient ultimately expired in part, as the cause of death was identified as cardiogenic shock, hemorrhagic shock, and late presentation myocardial infarction. | Additional information was reported that the patient had runs of wide complex ventricular tachycardia with poor perfusion. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2022-01513 was provided in sections b5, d1, d4, d6, and h6.
36273 | 1220648-2022-01512 | 14961398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14961398&pc=OZD | 12/23/2021 | 10/26/2024 |  | 07/08/2022 |  |  | 06/28/2022 | 07/08/2022 |  |  | 09/30/2023 |  | 07/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022135429 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 85 YR |  | Male |  |  | Asian | Malfunction | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | No Clinical Signs, Symptoms or Conditions | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/28/2022 |  |  |  |  | The complainant reported an (b)(6) male patient presenting with high risk percutaneous coronary intervention for impella support. during support, the patient developed an access site bleed that required a blood transfusion. the patient also exhibited ventricular fibrillation and required a direct cardioversion. | The complainant reported an 85 year old male patient presenting with high risk percutaneous coronary intervention for impella support. during support, the patient developed an access site bleed that required a blood transfusion. | During support of the impella cp, there seemed like there was a purge liquid leakage. there was not any obvious damage from the outside. although they wiped the leakage around the pressure reservoir, leakage occurred again so purge liquid leakage was suspected. there was no change of purge pressure and flow. the impella was removed and intra-aortic balloon pump (iabp) was performed to the patient, and was taken back to intensive care unit icu where the impella was again. |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | B5: we have corrected this section to exclude the report of ventricular fibrillation requiring cardioversion as it was listed in error. h6: please disregard health effect - clinical code 2130 and health effect - impact code 4641 as these were made in reference to the erroneous reporting of ventricular fibrillation. | This report is being filed as part of a retrospective review of historical records. impella cp with smartassist for use during cardiogenic shock and high risk cpi instructions for use for the related event are as follows: section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.". ¿clean the connector cable with 70% isopropyl alcohol.¿ corrected information for the follow-up number 1 mfr 1220648-2022-01512 was provided in sections b1, b5, d6, h1, and h6. b(5): we have corrected this section to exclude the report of access site bleed requiring a blood transfusion as it was listed in error. h(6): please disregard health effect - clinical code 1888 and health effect - impact code 4643. medical device problem code 2993 as these were made in reference to the erroneous reporting of access site bleed.
36274 | 1220648-2022-01511 | 14961427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14961427&pc=OZD | 07/07/2022 | 07/08/2022 |  | 07/08/2022 |  |  | 07/07/2022 | 07/08/2022 |  |  | 01/31/2024 |  | 07/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023174023 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/07/2022 |  |  |  |  | The complainant reported a 47 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. shortly after removing the pump, at the conclusion of support, the patient was noted to have limb ischemia. a thrombus was identified in the vascular system that prompted intervention to remove via fogarty catheter. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36275 | 1220648-2022-01498 | 14936729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14936729&pc=OZD | 06/20/2022 | 07/06/2022 |  | 07/06/2022 |  |  | 06/21/2022 | 07/06/2022 |  |  | 02/29/2024 |  | 07/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICA |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023194091 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/21/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with post cardiotomy cardiogenic shock for impella support. during support, the patient developed an access site bleed that prompted 2 units of blood products delivered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36276 | 1220648-2022-01500 | 14936769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14936769&pc=OZD | 06/01/2022 | 07/06/2022 |  | 07/06/2022 |  |  | 06/03/2022 | 07/06/2022 |  |  | 01/31/2024 |  | 07/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023181744 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/03/2022 |  |  |  |  | The complainant reported a 54 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient had hemolytic urine. after removing the impella, the symptoms improved. | [5 empty] | The impella device was not returned from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36277 | 1220648-2022-01499 | 14936796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14936796&pc=OZD | 06/18/2022 | 07/06/2022 |  | 07/06/2022 |  |  | 06/19/2022 | 07/06/2022 |  |  | 03/31/2024 |  | 07/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023205295 |  | 0550-0008 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/19/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient endured hemolysis. the pump was removed prematurely as a result. | [5 empty] | The impella device was not returned from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36278 | 1220648-2022-01502 | 14936868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14936868&pc=OZD | 06/08/2022 | 07/06/2022 |  | 07/06/2022 |  |  | 06/08/2022 | 07/06/2022 |  |  | 03/31/2024 |  | 07/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023204376 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/08/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with post cardiotomy cardiogenic shock for impella support. during support, the patient developed ischemia in the impella inserted leg. as treatment, the pump was removed and a below the knee amputation was required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36279 | 1220648-2022-01503 | 14936943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14936943&pc=OZD | 05/29/2022 | 10/14/2024 |  | 07/06/2022 |  |  | 06/08/2022 | 07/06/2022 |  |  | 11/30/2023 |  | 07/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2023159553 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/08/2022 |  |  |  |  | The complainant reported a 53 year old female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient was suspected as enduring hemolysis. haptoglobin was delivered as a result. | It was also reported the patient experience high purge pressure while on support. heparin in the purge solution was changed and the act had increase, there was no bleeding, and the purge flow improved. however, the issue was still not resolved. further investigation was requested by the physician on the returned device. |  |  |  |  | The impella device was received and an investigation is underway. upon completion of our analysis, a supplemental mdr will be filed. | The device was returned and an investigation was conducted. inspection found no damage or abnormalities. metrics were analyzed and found within specification. we are unable to determine the root cause of the reported complaint. | The impella device was returned for evaluation; however, no data logs were received. upon inspection, the reported metrics were within specification, therefore, no purge issue was found. additionally, the investigation for the hemolysis found: the cause of the hemolysis was patient condition related since the administration of haptoglobin resolved the issue. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records. additional information for the initial mfr 1220648-2022-01503 was provided in sections b1,b2,b5, d6a, d6b, g1, h1, and h6.
36280 | 1220648-2022-01505 | 14937386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14937386&pc=OZD | 06/22/2022 | 07/06/2022 |  | 07/06/2022 |  |  | 06/22/2022 | 07/06/2022 |  |  |  |  | 07/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/22/2022 |  |  |  |  | The complainant reported a 74 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. when being transferred into another hospital, the pump connector broken off inside the aic console after hospital doors were closed on it resulting in a pump stop just after 6 days of use. a new impella was emergently placed and support continued. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36281 | 1220648-2022-01504 | 14940423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14940423&pc=OZD | 06/21/2022 | 07/06/2022 |  | 07/06/2022 |  |  | 06/21/2022 | 07/06/2022 |  |  | 10/31/2023 |  | 07/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022156290 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/21/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed ischemia in the impella inserted leg. the pump was removed as a result. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36282 | 1220648-2022-01501 | 14940617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14940617&pc=OZD | 05/18/2022 | 07/06/2022 |  | 07/06/2022 |  |  | 06/09/2022 | 07/06/2022 |  |  | 12/31/2023 |  | 07/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2023162633 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/09/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient endured an ischemic stroke. | [5 empty] | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36283 | 1220648-2022-01506 | 14940886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14940886&pc=OZD | 06/21/2022 | 07/06/2022 |  | 07/06/2022 |  |  | 06/23/2022 | 07/06/2022 |  |  | 06/30/2023 |  | 07/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022101639 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/23/2022 |  |  |  |  | The complainant reported a 76 year old male patient presenting with post cardiotomy cardiogenic shock for impella support. during support, a cardiac tamponade had formed. the physician alleged the impella was a contributing factor. a pericardiocentesis was required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36284 | 14914459 | 14914459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14914459&pc=OZD | 03/19/2022 | 03/23/2022 | 03/23/2022 | 07/05/2022 | 03/23/2022 | 07/05/2022 |  | 07/05/2022 |  |  |  |  | 07/05/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Yes |  |  | Not Available |  |  | No | Lay User/Patient |  |  | User Facility report | No | Risk Manager | [6 empty] | 19710 DA |  | Male | 148 KG |  | White | Death | Yes | Yes | Failure to Pump | [6 empty] | Chest Pain | Hyperglycemia | Liver Damage/Dysfunction | Myocardial Infarction | Pain | Cardiogenic Shock | Unspecified Kidney or Urinary Problem | [6 empty] | Death | [24 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | At 2332, impella cp alarmed red alarm stating "pump off; motor current high" rn in room and perfusion and cca called to bedside. p level increased to p3 and pump restarted. at 2336, pump stopped again with same red alarm- p level increased again and pump restarted. after 3rd time of pump stopping, at 2339, rn unable to increase p level to restart pump. pt. remained hemodynamically stable at this time- norepinephrine increased and epinephrine increased to 0.06. plan made for patient to go to cath lab. impella pulled back per protocol. at 2350 patient's maps dropped to 30, epinephrine administered and compressions started. patient unable to be resuscitated and expired. prior to patient's transfer to hospital, he presented to a different hospital with chest and left arm pain and was found to have an acute anterolateral stemi. underwent pci (95% occlusion prox lad, 100% mid lad) with prox and mid lad stents placed. right femoral impella cp was placed post-cath for ci 1.2 and elevated pa pressures. the next day he was noted to have a worsening aki, shock liver and hyperglycemia. the next day cardiogenic shock worsened and he was transferred for higher level care.
36285 | 14885134 | 14885134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14885134&pc=QOC | 03/18/2022 | 03/23/2022 | 03/23/2022 | 07/01/2022 | 03/23/2022 | 07/01/2022 |  | 07/01/2022 |  |  |  |  | 07/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | QOC | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Yes |  |  | Not Available |  |  | No | Lay User/Patient |  |  | User Facility report | No | Risk Manager | [6 empty] | 23725 DA |  | Male | 64 KG |  | White | Malfunction | No | Yes | Unintended Power Up | Defective Device | [5 empty] | Insufficient Information | [12 empty] | Other | [24 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Perfusion note- per discussion with abiomed representative, impella pump showing signs of strain/possible impending pump failure (increasing motor current, narrowing min/max flow range). pt now going today for impella cp exchange in or or cath lab due to impending failure of device. abiomed rep aware - is also aware of other event over the weekend where pt had code 4 event following apparent device malfunction. cath lab or where device is exchanged has been asked to save device; abiomed rep will take on behalf of company for internal investigation. able to get impella sn off controller screen as well as controller device number. per abiomed rep, there is not a lot number available. this patient was a transfer from an outside hospital where the device was placed. patient then transferred to this hospital for transplant or lvad consideration.
36286 | 1220648-2022-01495 | 14887376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14887376&pc=OZD | 06/02/2022 | 07/01/2022 |  | 07/01/2022 |  |  | 06/02/2022 | 07/01/2022 |  |  | 10/31/2023 |  | 07/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022143672 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/02/2022 |  |  |  |  | The complainant suspected a patient exhibiting signs of hemolysis during impella support. as treatment, haptoglobin was delivered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36287 | 1220648-2022-01496 | 14887422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14887422&pc=OZD | 03/15/2022 | 10/09/2025 |  | 07/01/2022 |  |  | 06/02/2022 | 07/01/2022 |  |  | 11/30/2023 |  | 07/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CPOPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023159651 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 17 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Sepsis | [11 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 06/02/2022 |  |  |  |  | The complainant reported a 17 year old male patient presenting for impella support. the patient developed ischemia in the impella inserted leg. as treatment, an external bypass was administered to restore blood flow. | [5 empty] | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36288 | 1220648-2022-01497 | 14887448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14887448&pc=OZD | 09/07/2021 | 07/01/2022 |  | 07/01/2022 |  |  | 06/02/2022 | 07/01/2022 |  |  | 07/31/2023 |  | 07/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022105865 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/02/2022 |  |  |  |  | The complainant suspected an 80 year old patient exhibiting signs of hemolysis during use of the device. as treatment, haptoglobin was delivered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36289 | 14842354 | 14842354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14842354&pc=OZD | 04/04/2022 | 06/15/2022 | 06/15/2022 | 06/28/2022 | 06/15/2022 | 06/28/2022 |  | 06/28/2022 |  |  |  |  | 06/28/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | No answer provided |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 24455 DA |  | Male |  |  |  | Malfunction | No | Yes | Crack | Leak/Splash | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella housing was leaking and required a "bypass" to be done to eliminate the cracked housing. circuit had been in use as our customary means with a 6 inch piece of pressure tubing placed between housing and the cassette tubing. pt reported the leak. bypass was performed by dr. per the instructions from the manufacturer and phone call support from impella.
36290 | 1220648-2022-01492 | 14848371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14848371&pc=OZD | 01/01/2022 | 06/28/2022 |  | 06/28/2022 |  |  | 05/31/2022 | 06/28/2022 |  |  |  |  | 06/28/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature |  |  |  |  | 72 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/31/2022 |  |  |  |  | A literature review discovered an impella cp pump report out of the (b)(6) in which post explant there was aortic valve damage observed and attributed to the use of the cp. the valve had noted regurgitation at removal. one year later the valve damage was fully appreciated and the valve was replaced. in the or the valve noncoronary cusp was observed to be perforated. | [5 empty] | The literature piece was the only reference available. at the time of the support no complaint was made and no product was returned for investigation. no further analysis or investigation has been possible. should more details be furnished regarding the impella support, a final report will be filed. the failure mode will be trended and monitored.
36291 | 1220648-2022-01493 | 14848419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14848419&pc=OZD | 05/09/2022 | 06/28/2022 |  | 06/28/2022 |  |  | 06/02/2022 | 06/28/2022 |  |  | 12/31/2023 |  | 06/28/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023166695 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/02/2022 |  |  |  |  | The medical team in japan had an impella cp support in combination with ecmo. the support was successful for over 15 days and then the patient expired due to underlying medical condition. weeks later the team reported that the patient had suffered from bilateral limb ischemia that was treated by external bypass for the impella limb. | [5 empty] | The medical center in japan discarded the impella pump product, and as such no investigation is possible. should more information be shared that leads to root cause, a final report will be filed.
36292 | 1220648-2022-01494 | 14848422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14848422&pc=OZD | 05/30/2022 | 06/28/2022 |  | 06/28/2022 |  |  | 06/03/2022 | 06/28/2022 |  |  | 05/31/2023 |  | 06/28/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022085946 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/03/2022 |  |  |  |  | The medical team in japan had an impella cp support with an allegation of hemolysis. the hemolysis was treated with both rbc infusion and haptoglobin. the pump repositioning didn't resolve the issue. the pump remains on for support despite the hemolysis. the pump support is currently on day 30, which is beyond its indicated use. | [5 empty] | The medical center in japan has not yet returned the imeplla pump for analysis. should product be returned and benchtop hemolysis testing be performed, a final report will be filed with root cause.
36293 | 1220648-2022-01490 | 14786993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14786993&pc=OZD | 05/25/2022 | 06/23/2022 |  | 06/23/2022 |  |  | 05/25/2022 | 06/23/2022 |  |  | 09/30/2023 |  | 06/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022137999 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/25/2022 |  |  |  |  | The complainant reported a 59 year old male cardiomyopathy patient presenting for impella support. during placement of the device in the left femoral artery, an access site hematoma formed. a vascular dissection was discovered during intervention to achieve hemostasis. as such, surgical intervention and blood products were required to resolve the injury. a new impella was prepped and inserted without further issue. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36294 | 1220648-2022-01491 | 14787074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14787074&pc=OZD | 05/28/2022 | 06/23/2022 |  | 06/23/2022 |  |  | 05/28/2022 | 06/23/2022 |  |  | 01/31/2024 |  | 06/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023191142 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/28/2022 |  |  |  |  | The complainant reported a 40 year old male cardiomyopathy patient presenting for impella support. during support, a blood flow occlusion was noted in the patient's brachial artery that prompted premature removal of the device and intervention to remove a thrombus. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36295 | 1220648-2022-01489 | 14753209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14753209&pc=OZD | 05/19/2022 | 06/21/2022 |  | 06/21/2022 |  |  | 05/24/2022 | 06/21/2022 |  |  | 12/31/2023 |  | 06/21/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023170375 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/24/2022 |  |  |  |  | The complainant reported a 72 year old patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient endured a hemorrhagic cerebral infarction. the impella was weaned and removed. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36296 | 14754020 | 14754020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14754020&pc=OZD | 03/28/2022 | 04/04/2022 | 04/07/2022 | 06/21/2022 | 04/07/2022 | 06/21/2022 |  | 06/21/2022 |  |  |  |  | 06/21/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 364273 |  | 0048-0003.0814 |  | Device was returned to manufacturer |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 32485 DA |  | Female |  |  | White | Malfunction | No | Yes | Detachment of Device or Device Component | Physical Resistance/Sticking | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Foreign Body In Patient | [36 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | During an intervention while patient was on impella device the intervention was completed. the patient was weaned off the device and then prepped for removal. the perclose was prepped and almost ready to be deployed when the impella was being pulled out it seemed to have some resistance. the doctor then took over to realize that the devise indeed was stuck, but it was pulled harder which led to some unraveling of the tip of the device and partial detachment (as shown on xray). with the continuous bleeding from the sight pressure was held and the patient had to have the device removed by cut down. the device was removed successfully.
36297 | 1220648-2022-01477 | 14711927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14711927&pc=OZD | 04/19/2019 | 07/01/2022 |  | 06/16/2022 |  |  | 03/25/2022 | 06/16/2022 |  |  | 08/31/2020 |  | 06/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1368961 |  | 0048-3096 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2022 |  |  |  |  | The medical center had an impella cp support in 2019 in which the patient expired after 0.73 hours of support. the patient had experienced deterioration of his health and had a diagnosis of pericardial effusion, vtach, a rv laceration, and code. the patient was unable to be revived and expired with a failed code. impella contribution to the event is unknown. | [5 empty] | The investigation into the event is on-going. upon completion of the investigation, a final report will be filed. | Since the original report was filed the investigation has concluded. there was no product returned and further clinical details were not shared by medical team in japan. the root cause of the injury and death are not determined. the failure mode will be monitored and trended.
36298 | 1220648-2022-01478 | 14711931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14711931&pc=OZD | 05/18/2022 | 06/16/2022 |  | 06/16/2022 |  |  | 05/18/2022 | 06/16/2022 |  |  | 02/29/2024 |  | 06/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023195209 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/18/2022 |  |  |  |  | The medical center had a patient transferred to the tertiary medical center for optimal care and support. at the first center the team had placed an impella cp and had the 14fr sheath left in. the impella sheath caused limb thrombosis and intervention. the team ballooned and had to explant the pump. the medical team had retained the sheath in the artery for 3 days which is against the ifu recommendations. | [5 empty] | The medical center discarded all product. no benchtop analysis is possible. the medical team had not followed ifu recommendations and removed the 14fr sheath from the femoral artery. limb ischemia most likely developed due to the user error. should more details be shared, another report will be filed.
36299 | 1220648-2022-01479 | 14711935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14711935&pc=OZD | 05/19/2022 | 03/10/2025 |  | 06/16/2022 |  |  | 05/19/2022 | 06/16/2022 |  |  | 08/31/2023 |  | 06/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022127624 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/19/2022 |  |  |  |  | The medical team in japan had an access site bleed at the impella cp site. the bleed was treated by vessel suture. the suture and pressure were effective for hemostasis. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced thrombocytopenia and hemolysis, the resolution for these issues is unknown. no further information is available. |  |  |  |  | The medical center in japan had discarded all impella pump product. no benchtop analysis or testing to root cause is possible. should more information be shared, a final supplemental report will be filed. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure for the additional reported issues, a supplemental mdr will be filed. as noted in the impella ifu: thrombocytopenia: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. corrections: b5: updated to reflect additional issues. e1: corrected address and added phone number. g1: corrected contact information. g2: corrected report sources. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
36300 | 1220648-2022-01480 | 14711951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14711951&pc=OZD | 05/20/2022 | 07/01/2022 |  | 06/16/2022 |  |  | 05/20/2022 | 06/16/2022 |  |  | 02/29/2024 |  | 06/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023199406 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | Detachment of Device or Device Component | Adverse Event Without Identified Device or Use Problem | Separation Problem |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2022 |  |  |  |  | The medical center had an impella cp support which was successful. when explanting the pump from the axillary artery there was damage, and the pump came apart when removed and this prompted surgical removal. the patient survived the event. during the pump investigation at the manufacturer the impeller blade was found to be broken. this was not observed by the medical center. | [5 empty] | The investigation into the pump damage and need for surgical explant is on-going at this time. the investigation into the broken impeller blade is also on-going at this time. upon completion of the analysis, a final report will be filed. | Since the original report was filed the product investigation has concluded in to the pump separation as well as the broken impeller blade. the detachment and separation occurred due to a use error. the use error was from the application of the cross clamp and an inadvertent interaction of the pump and clamp. the broken blade cause is not determined. because fluoroscopic imaging of the pump was not returned, a definitive root cause of the fracture could not be determined. the failure mode will be monitored and trended.
36301 | 1220648-2022-01481 | 14715669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14715669&pc=OZD | 05/19/2022 | 06/16/2022 |  | 06/16/2022 |  |  | 05/19/2022 | 06/16/2022 |  |  | 02/28/2023 |  | 06/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 2022045364 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Nausea | Tachycardia | Vomiting | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/19/2022 |  |  |  |  | The complainant reported a patient of unknown age and gender presenting for high risk percutaneous coronary intervention using the impella cp optical for hemodynamic support. after initiating the pump, the patient was reported as having hypotension, nausea, vomiting, and tachycardia. as such, the pump was removed and the patient's symptoms resolved. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36302 | 1220648-2022-01482 | 14715719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14715719&pc=OZD | 05/17/2022 | 06/16/2022 |  | 06/16/2022 |  |  | 05/19/2022 | 06/16/2022 |  |  | 10/31/2023 |  | 06/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 2022148313 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Nausea | Tachycardia | Vomiting | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/19/2022 |  |  |  |  | The complainant reported a patient of unknown age and gender presenting for high risk percutaneous coronary intervention using the impella cp optical for hemodynamic support. after initiating the pump, the patient was reported as having hypotension, nausea, vomiting, and tachycardia. as such, the pump was removed and the patient's symptoms resolved. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36303 | 1220648-2022-01483 | 14715742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14715742&pc=OZD | 04/19/2022 | 06/16/2022 |  | 06/16/2022 |  |  | 05/23/2022 | 06/16/2022 |  |  | 12/31/2023 |  | 06/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023166718 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/23/2022 |  |  |  |  | The complainant reported a 41 year old female patient presenting with myocarditis for impella support. when removing the pump, a thrombus was found adhered to the impella within the thoracic aorta. a fogarty catheter procedure was required to fully remove the thrombus. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36304 | 1220648-2022-01484 | 14715765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14715765&pc=OZD | 05/20/2022 | 06/16/2022 |  | 06/16/2022 |  |  | 06/03/2022 | 06/16/2022 |  |  | 12/31/2023 |  | 06/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023177731 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/03/2022 |  |  |  |  | The us complainant had an impella cp with sensor issues that prompted the pump to be explanted. the sensor and low flows prompted the explant and no pump replacement. the procedure completed without the support. later in the benchtop analysis of the pump product the blade of the impeller was found to be broken. | [5 empty] | Because the fluoroscopic image of the pump was not returned, a definitive root cause of the fracture could not be determined. the failure mode will be monitored and trended.
36305 | 1220648-2022-01485 | 14715832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14715832&pc=OZD | 05/23/2022 | 06/16/2022 |  | 06/16/2022 |  |  | 05/23/2022 | 06/16/2022 |  |  | 01/31/2024 |  | 06/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023174066 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Hospitalization | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/23/2022 |  |  |  |  | The medical team in (b)(6) had an imeplla cp support. at the point of explant of the pump and sheath the team observed thrombotic debris and later there were neurological changes. the team had arrousal issues with the patient. the team performed imaging and found a cerebral infarct that was treated by anticoagulation only. the causal relationship between impella use and the cva could not be determined by the medical team in (b)(6). | [5 empty] | The cause of the cva could not be determined as the team in (b)(6) discarded all impella product. no imaging or further clinical details were shared. the failure mode will be monitored and trended.
36306 | 1220648-2022-01488 | 14715847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14715847&pc=OZD | 05/28/2022 | 07/05/2022 |  | 06/16/2022 |  |  | 05/28/2022 | 06/16/2022 |  |  | 01/31/2024 |  | 06/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023181253 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/28/2022 |  |  |  |  | The medical center had an impella 5.5 inserted via the axillary graft. the patient movement dislodged the pump and position in the lv was lost. the patient had become confused and dislodged his pump despite pump fixation. when taken to the or for pump replacement, the team observed arm pulses to be lost. the team replaced the pump and performed embolectomy. | [5 empty] | The medical center has not yet returned the impella pump product for analysis. upon completion of the investigation, a final report will be filed. | The medical center never returned the impella pump product for an investigation into the limb ischemia. the only analysis possible was of the clinical report. the ischemia was related to the underlying patient condition. the failure mode will be monitored and trended.
36307 | 1220648-2022-01486 | 14715851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14715851&pc=PYX | 05/24/2022 | 12/03/2024 |  | 06/16/2022 |  |  | 05/24/2022 | 06/16/2022 |  |  | 02/28/2023 |  | 06/16/2022 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED INC. |  |  | US |  | 2022043148 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Defective Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | Hospitalization |  |  |  |  | IMPELLA CP. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/24/2022 |  |  |  |  | The medical center had a bipella support, of both right and left sided impella pumps. at the placement of the rp to support the right side of the heart, that team mistakenly placed a 14fr for insertion of the rp. the rp was eventually placed appropriately through a 23 fr. the rp was explanted at day 6. upon explant pump damage was observed that the md is aware may have been caused by the improper sheath at the start. the team also noted access site bleeding/ooze during support. the bleed prompted the team to infuse 5 units of blood products over a few days of the support. | It was additionally reported that upon impella rp explant, tissue was observed in the struts of the outflow cage. |  |  |  |  | The investigation into the access site bleeding and intervention is on-going at this time. the product is expected but not yet returned for testing. upon completion of the investigation, a final report will be filed. | The medical center failed to return the impella pump product. only a photograph was shared. the cause of the damage to the product was a user related error. the cause was attempting to insert the pump via the incorrect sheath, a 14fr rather than a 23fr. as for the bleeding at the access site, it too was caused by user error in allowing the sheath to back out of the anatomy. the failure modes will be monitored and trended. | The investigation for the reported access site bleeding, outflow cage damage, and thrombus has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the issue was a damaged outflow cage from attempting to insert pump through a 14 fr introducer instead of a 23 fr introducer. the cause of the access site bleeding was use related since the sheath backed out. as the necessary clinical information was not provided, the root cause of the thrombus was not determined. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. the following have been reported incorrectly or missing from manufacturer device report 1220648-2022-01486. b5 updated with missing information d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact first name, last name, and email. g1 phone number should have been left blank. g4 pma/510(k)number. h6 health effect - clinical code and health effect - impact code added. h6 component code revised from the initial submission in accordance with updated procedures.
36308 | 1220648-2022-01487 | 14715860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14715860&pc=OZD | 05/26/2022 | 06/16/2022 |  | 06/16/2022 |  |  | 05/26/2022 | 06/16/2022 |  |  | 12/31/2023 |  | 06/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023175368 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/26/2022 |  |  |  |  | The (b)(6) medical center had an impella 5.5 support in which after insertion the pump stopped. the restart failed and so the pump was replaced. the 2nd pump successfully supports the patient without any harm or injury from the delayed support initiation. | [5 empty] | The medical center has not yet returned the impella pump product for analysis and benchtop testing. upon return of the product and investigation conclusion, a final report will be filed.
36309 | 1220648-2022-01474 | 14687249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14687249&pc=OZD | 05/17/2022 | 06/14/2022 |  | 06/14/2022 |  |  | 05/17/2022 | 06/14/2022 |  |  | 01/31/2024 |  | 06/14/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023188097 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/17/2022 |  |  |  |  | The complainant reported a 69 year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. while attempting to wirelessly re-cross the aortic valve, the pump's cannula folded over on itself. a snare was required to straighten and remove the pump. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36310 | 1220648-2022-01475 | 14687762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14687762&pc=OZD | 05/17/2022 | 10/11/2023 |  | 06/14/2022 |  |  | 05/18/2022 | 06/14/2022 |  |  | 01/31/2024 |  | 06/14/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023185220 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2022 |  |  |  |  | The complainant reported a 61 year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella 5.5 device. on the 14th day of usage, the pump stopped and could not be restarted. a new pump was placed. the patient did not sustain any injury. | It was reported that during support ,a purge leak occurred at the yellow luer while sodium bicarbonate was running through the purge solution. a purge bypass was performed. following the bypass, the impella began to exhibit saturated flow waveforms. the decision was then made to explant the pump due to the saturated flows. however, before the clinicians could remove the device, the pump suddenly stopped with high motor current and was unable to be restarted. a new impella was placed afterwards. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation for the reported purge leak and the saturated flows has been completed. no product was returned. data logs showed the "purge pressure low" alarm occurred with a rapid increase in purge flow and decrease in purge pressure. soon after, the pump flow became saturated and the motor current gradually increased. about 90 minutes later, the pump stopped immediately upon the occurrence of the "impella stopped - motor current high" alarm. after reviewing the clinical information, it was concluded that the cause of the pump stop was most likely blood ingress due to the sidearm yellow luer damage from the use of sodium bicarbonate. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records. d4-udi.
36311 | 1220648-2022-01476 | 14689142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14689142&pc=OZD | 05/17/2022 | 06/14/2022 |  | 06/14/2022 |  |  | 05/18/2022 | 06/14/2022 |  |  | 11/30/2023 |  | 06/14/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023168248 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Other | Death | [23 empty] | A |  |  |  | Initial submission |  |  |  |  | 05/18/2022 |  |  |  |  | The complainant reported a male patient of unknown age presenting for pre-operative percutaneous coronary intervention (pci). the impella 5.5 device was selected for support leading up to the pci. the device was inserted without any noted issues. approximately 40 minutes into support, the patient was rolled to the left and right sides in order to place bilateral chest tubes. after placement, suction alarms occurred and the patient's condition declined. effusion was noted under tee imaging with a posterior left ventricle perforation via the impella device. intervention was unsuccessful and the patient was withdrawn from care and expired. the physician attributed the perforation to the size of the left ventricle after 40 minutes of unloading. the physician also felt the injury had occurred when patient was rolled for chest tube placement as the impella device was against the ventricle at this time. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. evaluation of the reported incident suggests the perforation was likely caused by patient movement. should any new information be returned, a supplemental mdr will be filed.
36312 | 1220648-2022-01472 | 14655235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14655235&pc=OZD | 05/09/2022 | 06/10/2022 |  | 06/10/2022 |  |  | 05/11/2022 | 06/10/2022 |  |  | 07/31/2023 |  | 06/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022112962 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/11/2022 |  |  |  |  | The complainant reported a (b)(6)female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella 5.5 device. after repositioning the device, the patient exhibited bleeding from the impella insertion site and blood found pooling around the heart and lungs. further inspection found the graft had become detached from the aorta. the patient expired after failed attempts to intervene. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36313 | 1220648-2022-01473 | 14655284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14655284&pc=OZD | 05/11/2022 | 06/10/2022 |  | 06/10/2022 |  |  | 05/11/2022 | 06/10/2022 |  |  | 02/29/2024 |  | 06/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023194947 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 05/11/2022 |  |  |  |  | The complainant reported a 70 year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical. during support, the physician opted to leave the 14 fr introducer in the patient's access site. the patient's leg lost blood flow, therefore, the introducer was removed. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36314 | 1220648-2022-01469 | 14633025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14633025&pc=OZD | 05/25/2020 | 06/08/2022 |  | 06/08/2022 |  |  | 05/10/2022 | 06/08/2022 |  |  | 09/30/2021 |  | 06/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1446479 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/10/2022 |  |  |  |  | The complainant reported a 55 year old male patient presenting with myocarditis for impella support. during support, the patient endured an ischemic stroke for which the device was alleged to have a relationship. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36315 | 1220648-2022-01470 | 14633061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14633061&pc=OZD | 11/22/2021 | 06/08/2022 |  | 06/08/2022 |  |  | 05/10/2022 | 06/08/2022 |  |  | 05/31/2023 |  | 06/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022085988 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/10/2022 |  |  |  |  | The complainant reported a 57 year old female patient presenting for hemodynamic support using the impella cp optical device. during support, the patient developed an impella related access site bleed that prompted a blood transfusion. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36316 | 1220648-2022-01467 | 14633111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14633111&pc=OZD | 05/11/2022 | 06/08/2022 |  | 06/08/2022 |  |  | 05/11/2022 | 06/08/2022 |  |  | 09/30/2023 |  | 06/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022144939 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/11/2022 |  |  |  |  | The complainant reported a 67 year old white male patient enduring a stroke shortly after removal of the impella device. it was alleged that a clot had traveled through the catheter to the patient's brain. three hours later, the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36317 | 1220648-2022-01468 | 14633123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14633123&pc=OZD | 01/24/2022 | 06/08/2022 |  | 06/08/2022 |  |  | 05/10/2022 | 06/08/2022 |  |  | 07/31/2023 |  | 06/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022114302 |  | 0048-0024-JP |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/10/2022 |  |  |  |  | The complainant reported an (b)(6) year old male patient presenting for an aortic valvuloplasty. the impella cp optical was used for hemodynamic support. on the last day of impella support, the patient endured an ischemic stroke. the device could not be ruled out as a causal/contributory factor. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36318 | 1220648-2022-01471 | 14633598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14633598&pc=OZD | 05/15/2022 | 08/04/2022 |  | 06/08/2022 |  |  | 05/15/2022 | 06/08/2022 |  |  | 01/31/2024 |  | 06/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023185185 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2022 |  |  |  |  | The complainant reported a 57 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. shortly after initiation, the pump stopped and could not be restarted. a new impella was used. | [5 empty] | The impella device was returned and an evaluation is underway. a supplemental mdr will be filed upon completion. | The device was returned and an investigation was conducted. an open circuit was found within the catheter that was concluded as the root cause of the reported pump stop.
36319 | 1220648-2022-01466 | 14605302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14605302&pc=OZD | 05/11/2022 | 06/13/2022 |  | 06/06/2022 |  |  | 05/11/2022 | 06/06/2022 |  |  | 02/28/2024 |  | 06/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD AIC PUMP SET US |  |  | ABIOMED INC. |  |  | US |  | 2023188889 |  | 005082 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2022 |  |  |  |  | The medical center had an impella ld support for escalation from an impella cp. after 7 days the team explanted the ld. at the explant the pump separated due to aggressive removal. the medical team had felt resistance at the pump explant. the pump was surgically removed from the patient. | [5 empty] | The investigation is on-going a this time. upon completion, a final report will be submitted. | Since the original report was filed, the investigation has concluded. the medical center returned the impella pump product, however what was returned was a damaged catheter. the failure mode could not be reproduced. the root cause was determined to be use related. the team had shared that the user had used excessive force. the failure mode will be monitored and trended.
36320 | 1220648-2022-01464 | 14605326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14605326&pc=OZD | 05/08/2022 | 11/29/2024 |  | 06/06/2022 |  |  | 05/10/2022 | 06/06/2022 |  |  | 03/31/2023 |  | 06/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2022059088 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2022 |  |  |  |  | The medical team in (b)(6) had an impella 5.0 support of over 3 weeks. the 5.0 stopped and was able to be restarted. the pump had been utilized beyond indication. after a few days the abiomed (b)(4) team was further updated that at pump explant there had been thrombus burden dislodged and embolized to the elbow. surgical removal and thrombectomy was needed. | It was also reported that the pump stopped during support. the pump was able to be restarted, and the medical team decided to replace the pump due to the pump stop. there was no harm to the patient due to the pump stop. |  |  |  |  | The medical team in (b)(6) discarded all impella pump and sheath product at explant. there can be no investigation to root cause. should the team share more information that leads to a definate root cause, a supplemental report will be filed. the failure mode will be monitored and trended. | The investigation for the reported pump stop and limb ischemia issues has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2022-01464 was provided in sections b1, b5, d6, and h6. this report is being filed as part of a retrospective review of historical records.
36321 | 1220648-2022-01465 | 14605377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14605377&pc=OZD | 12/08/2021 | 06/06/2022 |  | 06/06/2022 |  |  | 05/10/2022 | 06/06/2022 |  |  | 08/31/2023 |  | 06/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022127485 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Hospitalization | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/11/2022 |  |  |  |  | The medical team in (b)(6) in (b)(6) 2021 had an impella cp and ecmo support that was successful til wean and explant. later in (b)(6) 2022 the jpvad registry noted an ischemic cva 2 days before explant. the event can not be investigated fully as all product and data were discarded in 2021. | [5 empty] | The medical center discarded all impella prior in 2021. no investigation to root cause is possible. the jpvad registry did not furnish further details. no root cause is determined. the failure mode will be monitored and trended.
36322 | 1035166-2022-00069 | 14592051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14592051&pc=DYB | 05/25/2022 | 11/10/2022 |  | 06/03/2022 |  |  | 05/26/2022 | 06/03/2022 |  |  |  |  | 06/03/2022 |  | INTRODUCER, CATHETER, PRODUCT CODE: DYB |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  |  |  | 0052-3006 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | On (b)(6) 2022 a (b)(6) male admitted suffering from cardiomyopathy underwent a procedure for access at axillary graft for placement of the impella 5.5. today a complaint was filed inclusive of your introducer packaged within our axillary set. the field rep noted the sheath cracked during placement into the graft and placing graft locks. patient reported as stable. it is unknown if there was a delay in the procedure. no medical intervention was given. patient is still on impella support. the procedure was completed with a new sheath. | Additional information received 06/03/2022. sheath crack was confirmed by physician before placement. sheath cracked on distal tip of sheath along striations. leaking not confirmed as physician did not want to use this sheath in this procedure, new sheath requested. |  |  |  |  | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. | One abiomed axillary 23f sheath was received. there were no other accessories. traces of blood were found on the sheath. according to the event description summary sheath cracked during placement into the graft and placing graft locks. upon evaluation with naked eye a 3-4mm crack was found on the score line of the sheath shaft. the lot number was not provided by the customer; therefore, the device history record could not be reviewed, however, inspection procedures require any oscor product to pass all in-process and qa final inspection before shipping to the customer. returned device analysis revealed the sheath was within manufacturing specification. a 3-4mm crack was found on the score line of the sheath shaft. although the exact cause is unknown, it is possible that resistance was encountered while in use causing the sheath to crack. the sheaths are inspected by qa for cracks and damage and are also leak tested in-process 100%. since the lot number is unknown, the dhr cannot be reviewed to see if there's any related issues or anomalies. per procedure (adelante s2s introducer sheath in-process and final inspections) sample size: 100% with naked eye at a distance of 12" to 18", qa personnel must to verify that the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm. the instructions for use (ifu) informs the user: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the.
36323 | 1035166-2022-00067 | 14592271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14592271&pc=DYB | 04/08/2022 | 10/13/2022 |  | 06/03/2022 |  |  | 05/11/2022 | 06/03/2022 |  |  |  |  | 06/03/2022 |  | INTRODUCER, CATHETER, PRODUCT CODE: DYB |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  |  |  | 0052-3025 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a (b)(6) male presented to the emergency department with an st-elevation myocardial infarction. the patient underwent an acute myocardial infarction/cardiogenic shock procedure when presented to the physician which attempted to deliver a 14 fr introducer device, however, insertion was ultimately aborted after the device became damaged due to patient's calcified vasculature. it was noted that the common femoral artery was dissected while attempting to access. as a result, the introducer was removed and an intra-aortic balloon pump was placed. patient reported as stable. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. | Device was used in treatment. the device was discarded and not returned for analysis. several attempts were made to obtain additional patient and event information. there was no specific performance related failure reported by the user. no further investigation is required. the lot number was not provided by the customer, therefore the device history records could not be reviewed, however, inspection procedures require any oscor product pass all in-process and qa final inspection before shipping to the customer. no further follow-up is required. the device was not returned and there was no specific performance related failure reported by the user. based on the investigation, a capa is not required. the event will be re-evaluated if additional information becomes available. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36324 | 1220648-2022-01457 | 14594276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14594276&pc=OZD | 04/08/2021 | 05/23/2025 |  | 06/03/2022 |  |  | 05/07/2022 | 06/03/2022 |  |  | 12/31/2022 |  | 06/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022029849 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/07/2022 |  |  |  |  | The complainant suspected an (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient was suspected as enduring hemolysis. haptoglobin was delivered as treatment. | Additional information was reported by the physician indicating there was access site bleeding noted by a hematoma near the impella insertion site, resulting in blood transfusion and pressure being held for hemostasis. thrombocytopenia was also found from insertion to withdrawal. the patient remained on impella support and was successfully weaned. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information for the initial mfr was provided in sections b2, b5, d6, g1, g2, h1, and h6. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr was provided in sections b5, d6, g1, g2, h1, and h6. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
36325 | 1220648-2022-01456 | 14594282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14594282&pc=OZD | 05/06/2021 | 06/03/2022 |  | 06/03/2022 |  |  | 05/06/2022 | 06/03/2022 |  |  | 12/31/2022 |  | 06/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022028937 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/06/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, an access site bleed developed that prompted a blood transfusion. after removal of the device, an infection occurred at the impella access site prompting debridement as intervention. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36326 | 1220648-2022-01458 | 14594443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14594443&pc=OZD | 05/08/2022 | 06/03/2022 |  | 06/03/2022 |  |  | 05/08/2022 | 06/03/2022 |  |  | 01/31/2024 |  | 06/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023184778 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/08/2022 |  |  |  |  | The complainant reported a 39 year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient developed limb ischemia distal to the impella device. the pump was removed and a thrombectomy was required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36327 | 1220648-2022-01459 | 14595548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14595548&pc=OZD | 02/27/2021 | 06/03/2022 |  | 06/03/2022 |  |  | 05/09/2022 | 06/03/2022 |  |  | 12/31/2022 |  | 06/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022023291 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Wound Dehiscence | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/09/2022 |  |  |  |  | The complainant reported an (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. after support, the impella site became infected and required vac therapy as treatment. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36328 | 1220648-2022-01460 | 14595631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14595631&pc=OZD | 05/06/2022 | 06/30/2022 |  | 06/03/2022 |  |  | 05/06/2022 | 06/03/2022 |  |  | 02/29/2024 |  | 06/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023196153 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2022 |  |  |  |  | The medical center had a patient with known peripheral arterial disease, and many other cardiac and systemic comorbidities, suffer from a dissection to the femoral after attempted insertion of the impella cp. the vessel was observed to be dissected and the patient had contralateral leg angiography done and was sent for iliac artery repair. three days after the explant of the impella the patient had limb amputation. | [5 empty] | The investigation is on-going at this time. no product has yet been returned for analysis. upon completion of the investigation, a final report will be filed. | Since the original report was filed the investigation has completed. there was no product returned for analysis, only the clinical report was reviewed. the patient was reported to have native peripheral arterial disease. the native patient condition allowed for the formation of the clot and limb ischemia. the failure mode will be monitored and trended.
36329 | 1220648-2022-01461 | 14595632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14595632&pc=OZD | 03/24/2021 | 06/03/2022 |  | 06/03/2022 |  |  | 05/07/2022 | 06/03/2022 |  |  | 12/31/2022 |  | 06/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022028946 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/07/2022 |  |  |  |  | The medical team in (b)(6) had an impella cp support in 2021 which was successful. then later in (b)(6) 2022 the jpvad registry noted an access site bleed in which blood products (rbcs) were given to treat the bleed. | [5 empty] | The medical center in (b)(6) discarded all product in 2021. no investigation to root cause is possible. no further clinical details have been shared. the failure mode will be monitored and trended.
36330 | 1220648-2022-01462 | 14595682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14595682&pc=OZD | 04/29/2022 | 06/03/2022 |  | 06/03/2022 |  |  | 05/09/2022 | 06/03/2022 |  |  | 11/30/2023 |  | 06/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023159639 |  | 0048-0024-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/09/2022 |  |  |  |  | The medical team in (b)(6) had an impella cp-ecmo support of over 5 days. the cp was removed and after explant thrombosis was observed. of note patient had underlying infection and peritonitis that could have contributed to the event. some days later the abiomed japan team was updated that heparin and the thrombolytic uk were administered for the clot as treatment/intervention. | [5 empty] | The medical center in (b)(6) discarded all impella product at explant. further clinical details have not been forthcoming. a root cause of the thrombosis has not been determined. should more information be relayed that leads to a definitive root cause, a final report will be filed. the failure mode will be monitored and trended.
36331 | 1220648-2022-01463 | 14595714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14595714&pc=OZD | 08/27/2020 | 09/24/2024 |  | 06/03/2022 |  |  | 05/11/2022 | 06/03/2022 |  |  | 01/31/2022 |  | 06/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1474555 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Hospitalization | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2022 |  |  |  |  | In 2020 the medical center in japan had an impella cp supporting a patient with atrial fibrillation, that was thought to be the cause of sma thrombosis and other clotting issues systemically. the patient expired due to multi organ failure due to sma thrombosis and predisposition to clotting. two years later, in 3/2022, the jpvad registry noted injuries that occurred during the time of impella support, being a cva, hemolysis, and lower limb ischemia. none of these were directly attributed to the use of impella with any causal relationship. later in 5/2022 the jpvad further updated that blood product/rbcs were given for the hemolysis and limb ischemia was treated by bypass. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The medical center in japan discarded all product in 2020 at the time of explant. further clinical details have not been forthcoming. there can be no investigation to root cause. the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01463 was provided in sections b2, b5, d6a, d6b, h1, and h6.
36332 | 1220648-2022-01455 | 14569273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14569273&pc=OZD | 05/02/2022 | 06/01/2022 |  | 06/01/2022 |  |  | 05/06/2022 | 06/01/2022 |  |  | 11/30/2023 |  | 06/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023159566 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/06/2022 |  |  |  |  | A younger patient was admitted in (b)(6) suffering from cardiomyopathy and had ecmo support placed. the medical center had an impella cp pump added to his care, as escalation from the balloon pump (iabp). the impella access site needed surgical repair due to a bleed. after 94 hours of impella support the pump was explanted as the patient was stable hemodynamically. | [5 empty] | The medical center in (b)(6) discarded the impella product at the explant. no benchtop analysis is possible. no root cause can be determined for the access site bleeding. should the medical center share more details that lead to a definitive root cause a final report will be filed. the failure mode will be monitored and trended.
36333 | 1220648-2022-01453 | 14549022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14549022&pc=OZD | 10/21/2021 | 09/25/2025 |  | 05/31/2022 |  |  | 05/02/2022 | 05/31/2022 |  |  | 06/30/2023 |  | 05/31/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022100431 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | Thrombocytopenia | [10 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/02/2022 |  |  |  |  | The complainant suspected an (b)(6) male patient, presenting with acute myocardial infarction and cardiogenic shock, enduring hemolysis during use of the impella device. the patient's urine was red and ldh values were elevated. haptoglobin was delivered as treatment. | During support, the patient experienced renal failure and thrombocytopenia, the resolution for these issues is unknown. no further information is available. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
36334 | 1220648-2022-01454 | 14549031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14549031&pc=OZD | 11/09/2021 | 05/31/2022 |  | 05/31/2022 |  |  | 05/02/2022 | 05/31/2022 |  |  | 07/31/2023 |  | 05/31/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022114705 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/02/2022 |  |  |  |  | The complainant suspected an (b)(6) male patient, presenting with acute myocardial infarction and cardiogenic shock, enduring hemolysis during use of the impella device. the patient's urine was red and ldh values were elevated. haptoglobin was delivered as treatment. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36335 | 1220648-2022-01451 | 14522459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14522459&pc=OZD | 04/28/2022 | 09/22/2024 |  | 05/27/2022 |  |  | 04/28/2022 | 05/27/2022 |  |  | 11/30/2023 |  | 05/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022154227 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/28/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with cardiomyopathy for hemodynamic support using the impella 5.0. during support, an access site bleed developed around the repositioning sheath. a blood transfusion was administered as treatment. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
36336 | 1220648-2022-01452 | 14526832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14526832&pc=OZD | 04/29/2022 | 05/27/2022 |  | 05/27/2022 |  |  | 04/29/2022 | 05/27/2022 |  |  | 07/31/2023 |  | 05/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022122803 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/29/2022 |  |  |  |  | The complainant reported a 78 year old female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient exhibited ischemia in the impella inserted limb. the pump was removed and the patient was escalated to an axillary pump. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36337 | 1220648-2022-01449 | 14499353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14499353&pc=OZD | 10/20/2021 | 05/25/2022 |  | 05/25/2022 |  |  | 05/02/2022 | 05/25/2022 |  |  | 12/31/2022 |  | 05/25/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2022075859 |  | 005048-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/02/2022 |  |  |  |  | The medical team in (b)(6) in (b)(6) 2021 had a successful impella 2.5 support. in (b)(6) 2022 the team noted in the (b)(6) registry there was hemolysis of unknown causality. later in (b)(6) 2022 the (b)(6) team further updated that haptoglobin was infused to treat the hemolysis, and that the relationship to the use of impella was presumed. | [5 empty] | The medical center discarded the impella pump product in 2021. no benchtop analysis or hemolysis testing can be performed. root cause determination is not possible for the hemolysis. the failure mode will be monitored and trended.
36338 | 1220648-2022-01450 | 14499389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14499389&pc=OZD | 11/26/2021 | 10/24/2024 |  | 05/25/2022 |  |  | 05/02/2022 | 05/25/2022 |  |  | 07/31/2023 |  | 05/25/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022111251 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Sepsis | [9 empty] | Required Intervention | Death |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2022 |  |  |  |  | The medical team in (b)(6) had a successful impella cp support in (b)(6) 2021 along with ecmo. in (b)(6) 2022 the jpvad registry noted hemolysis and death, the patient expired due to causes not related to the use of impella. the death occurred post explant. later in (b)(6) 2022 the jpvad team further updated that in fact the hemolysis was treated by infusion of haptoglobin. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there was noted hemolysis during the time of support. actions taken are unknown. bleeding including hematoma from insertion to removal was also noted. blood transfusion was required. additional information noted the patient experienced cerebrovascular accident and sepsis. survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The medical center discarded all product in 2021 at the time of explant. no benchtop analysis and testing can be performed. the jpvad registry did not furnish further details. no root cause can be determined. the failure mode will be monitored and trended. | The investigation for the reported sepsis, access site bleeding and stroke/cva has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the reported observations could not be determined. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system (infection/sepsis) section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist system (cva). section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist system (access site bleeding): section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information for the initial mfr 1220648-2022-01450 was provided in sections b2, b5, g1, h1 and h6.
36339 | 1220648-2022-01446 | 14488670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14488670&pc=OZD | 04/25/2022 | 06/01/2022 |  | 05/24/2022 |  |  | 04/25/2022 | 05/24/2022 |  |  | 01/31/2024 |  | 05/24/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023181728 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2022 |  |  |  |  | The medical center had an impella cp support on going with poor pump position within the left ventricle. the poor position had caused signs of hemolysis. the team explanted and placed a new pump. additionally the team put the patient on dialysis. | [5 empty] | The medical center did return the impella pump product for testing and benchtop analysis. the analysis is on-going at this time, and upon completion a final report will be filed. | Since the original report was filed, the investigation into the reported hemolysis has concluded. the pump was returned for analysis. the investigation determined the root cause to be a user error of improper pump positioning. the failure mode will be monitored and trended.
36340 | 1220648-2022-01447 | 14488709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14488709&pc=OZD | 04/26/2022 | 05/24/2022 |  | 05/24/2022 |  |  | 04/26/2022 | 05/24/2022 |  |  | 11/30/2023 |  | 05/24/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022151766 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/26/2022 |  |  |  |  | The medical center in (b)(6) had an access site bleed at the impella limb. the bleed was treated by the infusion of 4 units of blood. the team then also prematurely explanted the pump. | [5 empty] | The medical team in (b)(6) discarded the impella product at the explant. there can be no investigation of the product to root cause of the bleed. the failue mode will be monitored and trended.
36341 | 1220648-2022-01448 | 14488890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14488890&pc=OZD | 04/23/2022 | 06/28/2022 |  | 05/24/2022 |  |  | 04/27/2022 | 05/24/2022 |  |  | 12/31/2023 |  | 05/24/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED INC. |  |  | US |  | 2023175759 |  | 0550-00002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2022 |  |  |  |  | The medical team in the eu hada an impella 5.5 pump stop after just 4 hours of support. the pump was explanted and replaced. the second pump successfully supports the patient to date, now a month later. | [5 empty] | The medical center in the eu has returned the impella pump product and investigation of it is on-going at this time. upon completion of the analysis a final report will be submitted. | Since the initial report was filed, the investigation into the impella pump stop has concluded. the pump was returned and benchtop analysis discovered biomaterial was on and around the pump impeller. the ingested material stopped the pump. the failure mode will be monitored and trended.
36342 | 1220648-2022-01445 | 14472921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14472921&pc=OZD | 04/24/2022 | 05/23/2022 |  | 05/23/2022 |  |  | 04/24/2022 | 05/23/2022 |  |  | 12/31/2023 |  | 05/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023180637 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/24/2022 |  |  |  |  | The complainant reported an (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient developed an impella related access site bleed that prompted 2 units of blood products delivered. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36343 | 1220648-2022-01442 | 14464706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14464706&pc=OZD | 04/22/2022 | 06/01/2022 |  | 05/21/2022 |  |  | 04/22/2022 | 05/21/2022 |  |  | 11/30/2023 |  | 05/21/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023168255 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Detachment of Device or Device Component | Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 04/22/2022 |  |  |  |  | The medical center had an impella cp support of just under 16 hours for hemodynamic support of a patient having a high risk coronary angioplasty. the team explanted and when explanting felt resistance. the pump had separated at the cannula. the pump was therefore surgically explanted. the patient was stable afterwards without lasting harm or injury. | [5 empty] | The investigation is on-going at this time. the medical center did return the impella product and analysis is on-going at this time. at the conclusion of the investigation, a final report will be filed. | Since the original report was filed, the investigation into the pump detachment has concluded. the pump was returned and the partial detachment was observed. prior to the detachment a kink was observed. the damage was indicative of a shearing force being applied. the cause was use related. the failure mode will be monitored and trended.
36344 | 1220648-2022-01443 | 14464921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14464921&pc=OZD | 04/22/2022 | 05/20/2022 |  | 05/21/2022 |  |  | 04/22/2022 | 05/21/2022 |  |  | 11/30/2023 |  | 05/21/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023159572 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/23/2022 |  |  |  |  | The medical team in (b)(6) had an impella cp placed for hemodyamic support. the team reported upon an access site bleed that was treated with 2 units of blood replacement products. the team did allude to the fact the patient was bleeding from other sites as well. | [5 empty] | The medical center in (b)(6) did not return the entire impella pump product for analysis. no introducer sheath or pump were returned for benchtop testing. the team only returned the purge cassette. the root cause of the access site bleed can not be determined with known information. should the medical team forward more, and a root cause is definatively determined, another report will be filed.
36345 | 1220648-2022-01444 | 14464969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14464969&pc=OZD | 05/16/2021 | 05/20/2022 |  | 05/21/2022 |  |  | 05/16/2021 | 05/21/2022 |  |  | 03/31/2023 |  | 05/21/2022 |  | TEMPORARY NON ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022059751 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 38 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/16/2021 |  |  |  |  | The medical team in 2021 had an impella cp support. in (b)(6) 2022 the surpass study team reviewed the case and found a hemolysis complaint. when the cp was replaced the team made choice to escalate to the impella 5.5 and ecmo. | [5 empty] | The medical center discarded the impella pump in 2021, and as such no benchtop testing and analysis is possible. only the clinical report and surpass study documents could be reviewed. the root cause of the hemolysis was determined to be improper pump positionining. this is a technique/use error. the failure mode will be monitored and trended.
36346 | 1220648-2022-01435 | 14454578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14454578&pc=OZD | 10/29/2021 | 05/19/2022 |  | 05/20/2022 |  |  | 04/20/2022 | 05/20/2022 |  |  | 05/31/2023 |  | 05/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022085920 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/20/2022 |  |  |  |  | The complainant reported a male patient presenting with acuter myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical. during support, the patient developed an impella related access site bleed that prompted a blood transfusion. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36347 | 1220648-2022-01438 | 14454663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14454663&pc=OZD | 04/23/2022 | 07/06/2022 |  | 05/20/2022 |  |  | 04/23/2022 | 05/20/2022 |  |  | 12/31/2023 |  | 05/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023171572 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2022 |  |  |  |  | The complainant reported a 73 year old male patient presenting with cardiomyopathy. during support, the pump stopped. the device was removed and replaced without any harm to the patient. | [5 empty] | The impella device was returned and an evaluation is underway. a supplemental mdr will be filed at the conclusion of our analysis. | The device was returned and an investigation was performed. biomaterial was found in the outflow cage, however, there was no evidence of a kinked purge line or catheter. analysis of the returned data logs found a sudden spike in motor current, indicative of a kink in the purge line. based on our analysis, we feel the root cause of the pump stop was due to user technique/use resulting from a kink to the purge line.
36348 | 1220648-2022-01432 | 14454710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14454710&pc=OZD | 04/19/2022 | 10/23/2024 |  | 05/20/2022 |  |  | 04/19/2022 | 05/20/2022 |  |  | 01/31/2024 |  | 05/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023174064 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock and a ventricular aneurysm. the impella was selected for hemodynamic support. when advancing the repositioning sheath, during device placement, the tip of the catheter penetrated an existing aneurysm. the patient expired. | The complainant reported a patient presenting with a ventricular aneurysm for impella support. when advancing the repositioning sheath, during device placement, the tip of the catheter penetrated the existing aneurysm. the patient expired. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation for the reported ventricle perforation issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the ventricle perforation was use related since the catheter was not held to maintain position while the repo unit was inserted. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. as noted in the impella ifu ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2022-01432 was provided.
36349 | 1220648-2022-01430 | 14454715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14454715&pc=OZD | 10/14/2021 | 05/19/2022 |  | 05/20/2022 |  |  | 10/14/2021 | 05/20/2022 |  |  | 08/31/2023 |  | 05/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022123494 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/14/2021 |  |  |  |  | The complainant reported a 75 year old male patient presenting with cardiomyopathy for impella support. during support, the patient developed an access site bleed that prompted 2 units of blood product delivered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36350 | 1220648-2022-01434 | 14454718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14454718&pc=OZD | 04/06/2022 | 05/19/2022 |  | 05/20/2022 |  |  | 04/20/2022 | 05/20/2022 |  |  | 01/31/2024 |  | 05/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023171651 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/20/2022 |  |  |  |  | The complainant suspected a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical. during support, the patient developed signs of hemolysis as evident by hematuria output. haptoglobin was delivered as treatment. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36351 | 1220648-2022-01431 | 14454776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14454776&pc=OZD | 04/19/2022 | 09/14/2022 |  | 05/20/2022 |  |  | 04/19/2022 | 05/20/2022 |  |  |  |  | 05/20/2022 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA AIC CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1163151 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiogenic Shock | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/19/2022 |  |  |  |  | The complainant reported a 56 year old male patient presenting with cardiomyopathy for impella support. approximately 7 days into use, the pump stopped and the patient went into cardiogenic shock requiring 25 minutes of cardiopulmonary resuscitation. troubleshooting was successful in restarting the pump with use of a new aic controller, however, the patient's condition deteriorated after the event, resulting in the therapy being discontinued and the patient expiring. | [5 empty] | The device is en-route to our facility for evaluation. upon completion of our analysis, a supplemental mdr will be filed. | The aic console was returned and an investigation concluded the pump stop was related to a defective sub-circuit component within the console (pump motor driver).
36352 | 1220648-2022-01436 | 14454796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14454796&pc=OZD | 06/01/2021 | 05/19/2022 |  | 05/20/2022 |  |  | 04/20/2022 | 05/20/2022 |  |  | 01/31/2023 |  | 05/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022038033 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/20/2022 |  |  |  |  | The complainant reported an (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. a thrombus was found on the tip of the impella upon removal. a fogarty catheter procedure was then required to remove remaining thrombi from the artery. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36353 | 1220648-2022-01437 | 14454799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14454799&pc=OZD | 04/21/2022 | 05/19/2022 |  | 05/20/2022 |  |  | 04/21/2022 | 05/20/2022 |  |  | 01/31/2024 |  | 05/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023186186 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/21/2022 |  |  |  |  | The complainant reported a 64 year old female patient presenting with post cardiotomy cardiogenic shock for impella support. during placement of the device, the patient's ventricle was perforated. surgical repair was completed and impella was removed. the patient remained stable. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36354 | 1220648-2022-01433 | 14454941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14454941&pc=OZD | 04/06/2022 | 05/19/2022 |  | 05/20/2022 |  |  | 04/20/2022 | 05/20/2022 |  |  | 12/31/2023 |  | 05/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023165770 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/20/2022 |  |  |  |  | The complainant reported a 53 year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp smart assist. during support the patient developed an access site bleed. a blood transfusion was completed. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36355 | 1220648-2022-01440 | 14461467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14461467&pc=OZD | 05/22/2021 | 05/20/2022 |  | 05/20/2022 |  |  | 05/02/2022 | 05/20/2022 |  |  | 01/31/2023 |  | 05/20/2022 |  | TEMPORARY NON ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022038051 |  | 0048-0024-JP |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [18 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/02/2022 |  |  |  |  | The medical team in (b)(6) in (b)(6) 2021 had a successful cp support. later in (b)(6) 2022 the (b)(6) registry noted hemolysis occurring during support. no intervention was noted. two months later in (b)(6) 2022 the registry team further updated that haptoglobin was infused to treat the hemolysis. | [5 empty] | The medical center in (b)(6) discarded all product in 2021. no benchtop testing for hemolysis analysis is possible. no root cause determination is possible. should more information be shared that leads to root cause, a final report will be filed.
36356 | 1220648-2022-01439 | 14461531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14461531&pc=OZD | 04/21/2022 | 05/20/2022 |  | 05/20/2022 |  |  | 04/21/2022 | 05/20/2022 |  |  | 01/31/2024 |  | 05/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023174012 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Valvular Insufficiency/ Regurgitation | Mitral Valve Insufficiency/ Regurgitation | [11 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/21/2022 |  |  |  |  | The medical center in (b)(6) had an impella cp placed for hemodyanamic support for a patient suffering from shock. the cp was presumed to contribute to the chordawe tendonae rupture and associated mtiral valve regurgitaton, and need for mitraclip procedure. | [5 empty] | The impella device was returned and an evaluation is underway. a supplemental mdr will be filed at the conclusion of our analysis.
36357 | 1220648-2022-01441 | 14463199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14463199&pc=OZD | 04/22/2022 | 07/06/2022 |  | 05/20/2022 |  |  | 04/23/2022 | 05/20/2022 |  |  | 10/31/2023 |  | 05/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022157459 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support. after just over 30 hours of runtime, the pump stopped. the device was removed and replaced with an impella cp. | [5 empty] | The device was received and an investigation is underway. a supplemental (b)(6) will be filed upon completion of our analysis. | The device and data logs were returned and an investigation was performed. data logs confirm a pump stop had occurred. the device was found to have punctures in the catheter, indicative of a suture needle mistakenly hitting the catheter. blood ingress was found and determined to have caused an electrical short, resulting in the pump stop. the root cause of the pump stop was blood ingress due to user technique resulting from the punctured catheter, allowing for blood ingress.
36358 | 1220648-2022-01429 | 14410784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14410784&pc=OZD | 04/13/2022 | 05/16/2022 |  | 05/16/2022 |  |  | 04/19/2022 | 05/16/2022 |  |  | 07/31/2022 |  | 05/16/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1490580 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 65 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [18 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/19/2022 |  |  |  |  | The medical center in japan had an impella cp support in which there was an episode of hemolysis attributed to the use of impella. the team stated the hemolysis signs changed/improved with pump's p level adjustment and volume delivery. days later the team confirmed the hemolysis was actually treated by 2 units of rbcs and albumin. later the team escalated care to the impella 5.0 and hemolysis resolved. | [5 empty] | The medical center in (b)(6) discarded all impella product at the explant. there can be no benchtop analysis and hemolysis testing, and as such no definative root cause determination. the failure mode will be monitored and trended.
36359 | 1220648-2022-01420 | 14373926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14373926&pc=OZD | 12/28/2021 | 05/12/2022 |  | 05/12/2022 |  |  | 04/13/2022 | 05/12/2022 |  |  | 08/31/2023 |  | 05/12/2022 |  | IMPELLA 5.5 |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022133359 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/13/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for impella support as a bridge to heart transplant. the patient did not endure any adverse events during support, however, developed acute kidney injury approximately 3-5 days post impella explant, requiring dialysis. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36360 | 1220648-2022-01422 | 14374119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14374119&pc=OZD | 02/12/2022 | 05/12/2022 |  | 05/12/2022 |  |  | 04/13/2022 | 05/12/2022 |  |  | 10/31/2023 |  | 05/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  |  |  | 2022155813 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/13/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for impella support as a bridge to heart transplant. the patient did not endure any adverse events during support, however, developed acute kidney injury approximately 3-5 days post impella explant, requiring dialysis. the patient was implanted on (b)(6) 2022 and explanted on (b)(6) 2022. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36361 | 1220648-2022-01421 | 14374163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14374163&pc=OZD | 02/19/2022 | 05/28/2025 |  | 05/12/2022 |  |  | 04/13/2022 | 05/12/2022 |  |  | 10/31/2023 |  | 05/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022152037 |  | 14-DEC-2021 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2022 |  |  |  |  | The complainant reported a (b)(6) female patient presenting for impella support as a bridge to heart transplant. the patient did not endure any adverse events during support, however, developed acute kidney injury approximately 3-5 days post impella explant, requiring dialysis. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The field long term log was reviewed and no non-pulsatile waveforms were observed throughout support. the cause of the renal failure was patient condition and unrelated to impella since it occurred beyond 8 hours after explant. this report is being filed as part of a retrospective review of historical records. g1: corrected ¿reporting contact first name¿, ¿reporting contact last name¿ and ¿reporting contact email¿.
36362 | 1220648-2022-01425 | 14374203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14374203&pc=OZD | 04/08/2022 | 05/12/2022 |  | 05/12/2022 |  |  | 04/13/2022 | 05/12/2022 |  |  | 01/31/2024 |  | 05/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023182465 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/13/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. attemps to place the 14 fr introducer resulted in a dissection of the common femoral artery. the patient was noted to have vessel calcification. introducer insertion and impella support was aborted and an intra-aortic balloon pump was used for support. it's unclear what intervention was administered as a result of the vessel dissection. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36363 | 1220648-2022-01423 | 14374219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14374219&pc=OZD | 04/06/2022 | 05/12/2022 |  | 05/12/2022 |  |  | 04/13/2022 | 05/12/2022 |  |  | 09/30/2023 |  | 05/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022132615 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/14/2022 |  |  |  |  | The complainant suspected a (b)(6) patient presenting with cardiomyopathy for impella support. during support, hemolysis was suspected. as treatment, haptoglobin administration was required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36364 | 1220648-2022-01424 | 14374239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14374239&pc=OZD | 05/19/2021 | 05/12/2022 |  | 05/12/2022 |  |  | 04/14/2022 | 05/12/2022 |  |  | 01/31/2023 |  | 05/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022038048 |  | 17-FEB-2021 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/14/2022 |  |  |  |  | The complainant reported a patient enduring an impella related thrombus in the common iliac artery that prompted intervention via a fogarty catheter to remove. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36365 | 1220648-2022-01427 | 14380079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14380079&pc=OZD | 03/04/2021 | 05/12/2022 |  | 05/12/2022 |  |  | 05/11/2022 | 05/12/2022 |  |  | 07/31/2022 |  | 05/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1490681 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/11/2022 |  |  |  |  | The medical team in jb)(6) had a successful impella cp support in (b)(6) 2021 and later in (b)(6) 2022 they had a (b)(6) registry reported entry for an ischemic cva occurring during the impella support, of unknown causality. later the (b)(6) registry team further noted in (b)(6) 2022 that the patient additionally had limb ischemia. the limb ischemia was treated by bypass and attributed to the use of impella. | [5 empty] | The medical center discarded all impella product in 2021. the (b)(6) registry team gave no further clinical details. root cause of the limb ischemia can not be determined. the failure mode will be monitored and trended.
36366 | 1220648-2022-01428 | 14380105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14380105&pc=OZD | 01/01/2022 | 05/12/2022 |  | 05/12/2022 |  |  | 05/02/2022 | 05/12/2022 |  |  | 07/31/2023 |  | 05/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022114708 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/02/2022 |  |  |  |  | The medical center in (b)(6) in (b)(6) 2022 had a successful support with an impella cp. the team later in (b)(6) 2022 had a (b)(6) registry noted hemolysis. later in (b)(6) 2022 the team followed up that in fact haptoglobin was infused for the hemolysis treatment. | [5 empty] | All product was discarded in (b)(6) in 2021. no investigation to root cause is possible. the failure mode will be monitored and trended.
36367 | 1035166-2022-00063 | 14363604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14363604&pc=DYB | 03/24/2022 | 09/21/2022 |  | 05/11/2022 |  |  | 04/25/2022 | 05/11/2022 |  |  |  |  | 05/11/2022 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  |  |  | 0052-3021 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 74 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Myocardial Infarction | Cardiogenic Shock | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We have been notified of a complaint involving an oscor 23 fr introducer (material #: (b)(4). a (b)(6) male was presented to the physician decompensated after recently receiving a stent. the decision was made to use impella rp for an acute myocardial infarction/cardiogenic shock procedure. the physician reported blood leaking from the introducer's valve after removing the dilator. no blood transfusion was required and the procedure was successfully completed when using a new 23 fr oscor introducer. there was no additional intervention or injury sustained as a result (no harm). the devices were discarded by the physician. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete. | The device was used for treatment. the device was not returned for evaluation, as it was discarded; as a result, the exact cause of the product issue cannot be determined and the clinical observation could not be confirmed. however, the following controls are in place to mitigate the reported product issue. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections before shipping to the customer. per procedure adelante s2s introducer sheath in-process and final inspection: 12. final assembly inspect 100% for steps 12.2 and 12.3 12.2 visual inspection: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. foreign matter (fm)/fiber is limited to 2 particles < 0.10 mm2 or 1 particle < 0.2 mm2 with naked eye at a distance of 12" to 18", verify the sheath hub and split caps are free of excess silicone oil applied between the seals during assembly. verify split caps are properly placed and secured onto sheath hub and free of cracks/damages or excessive adhesive. 12.3 vacuum leak test perform pressure and vacuum leak test per procedure. the instructions for use (ifu) informs the user: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36368 | 1220648-2022-01414 | 14317768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14317768&pc=OZD | 04/09/2022 | 05/06/2022 |  | 05/07/2022 |  |  | 04/09/2022 | 05/07/2022 |  |  | 04/30/2023 |  | 05/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2022082513 |  | 005066-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 79 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/11/2022 |  |  |  |  | The medical team in (b)(6) had difficulty inserting an impella 5.0. the team had pump passing issues via the axillary and a bleed at the surgical graft site that prompted blood products to be infused. once successfully inserted, the pump stopped after just .83 hours of support. the pump stop prompted the team to explant. | [5 empty] | The medical center in (b)(6) discarded the impella pump and introducer product. no analysis to root cause of the bleeding and the pump stop is possible. should the team relay more clinical information that leads to root cause, a final report will be filed.
36369 | 1220648-2022-01415 | 14317793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14317793&pc=OZD | 04/12/2022 | 05/17/2022 |  | 05/07/2022 |  |  | 04/12/2022 | 05/07/2022 |  |  | 01/31/2024 |  | 05/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023186066 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup |  |  |  | 04/12/2022 |  |  |  |  | The medical center had an impella cp patient escalated to an impella 5.5 pump. the 5.5 supported the patient but, due to an alarm, the team made the choice to troubleshoot pump position. when the pump-to-aic cable was adjusted, in troubleshooting attempt, the pump would failed recognition on the aic and so the pump was replaced due to the pump stop/failure to start. the patient expired on the 2nd pump after cpr failed. | [5 empty] | The investigation into the pump stop/failure to start is ongoing at this time. pump product return is expected, but has not yet occurred. the pump is currently being held at the medical center. upon completion of the investigation, a final report will be submitted. | Since the original report was filed, the manufacturer has closed the investigation. no product was returned for benchtop analysis into the pump stop, and as such no root cause has been determined. the failure mode will be monitored and trended.
36370 | 1220648-2022-01413 | 14317853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14317853&pc=OZD | 01/23/2022 | 05/06/2022 |  | 05/07/2022 |  |  | 04/11/2022 | 05/07/2022 |  |  | 09/30/2023 |  | 05/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022132412 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/11/2022 |  |  |  |  | In (b)(6) 2022 the complainant in (b)(6) had a successful cp support of over 8 days when it was explanted. the patient was additionally supported by ecmo. later in (b)(6) 2022 the patient outcomes were updated from the (b)(6) registry that hemolysis, an access site bleed, and renal failure prompted rbc infusion and dialysis. | [5 empty] | The medical center in (b)(6) discarded all impella pump and introducer sheath product. no investigation to root cause is possible. should the (b)(6) registry or medical center share more information, a final report will be filed with root cause.
36371 | 1220648-2022-01416 | 14317888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14317888&pc=OZD | 11/28/2021 | 05/06/2022 |  | 05/07/2022 |  |  | 04/08/2022 | 05/07/2022 |  |  | 09/30/2023 |  | 05/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022128444 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 59 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/08/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for hemodynamic support using the impella cp optical for hemodynamic support. the patient endured ischemia after placing the 14 fr introducer into the access site. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36372 | 1220648-2022-01417 | 14317890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14317890&pc=OZD | 12/05/2021 | 10/11/2024 |  | 05/07/2022 |  |  | 04/08/2022 | 05/07/2022 |  |  | 07/31/2023 |  | 05/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022108771 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2022 |  |  |  |  | The complainant reported an (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device for support. during support, the patient endured an impella related access site bleed. blood product delivery was administered. | During support, vascular injury was reported in the descending aorta due to the effects of cardiac massage. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation has been completed for the reported issue of vascular damage - peripheral vessel. the device was not received. as there is no supporting evidence, the relationship between the vascular injury and impella is unknown. as no details were provided, relationship of bleeding and impella is unknown. additional information for the initial mfr 1220648-2022-01417 was provided in sections b5, e1, g1, h1, and h6. this report is being filed as part of a retrospective review of historical records.
36373 | 1220648-2022-01418 | 14317891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14317891&pc=OZD | 03/24/2022 | 05/06/2022 |  | 05/07/2022 |  |  | 04/11/2022 | 05/07/2022 |  |  | 11/30/2023 |  | 05/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023171638 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/11/2022 |  |  |  |  | The complainant reported a male patient presenting for pre-operative cabg support using the impella cp optical device for hemodynamic support. during support, the patient developed hemolysis. as treatment, the patient was administered dialysis. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36374 | 1220648-2022-01419 | 14317929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14317929&pc=OZD | 04/12/2022 | 05/06/2022 |  | 05/07/2022 |  |  | 04/12/2022 | 05/07/2022 |  |  | 11/30/2023 |  | 05/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022157931 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/12/2022 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with cardiomyopathy for impella support. during support, the device's pigtail perforated the apex of the ventricle. surgical repair was required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36375 | 1220648-2022-01408 | 14309047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14309047&pc=OZD | 06/10/2021 | 05/05/2022 |  | 05/06/2022 |  |  | 04/05/2022 | 05/06/2022 |  |  | 12/31/2022 |  | 05/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022029825 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/05/2022 |  |  |  |  | The complainant reported an (b)(6) female patient presenting for hemodynamic support via the impella cp optical device. during support, the patient endured an ischemic stroke that which the impella could not be ruled out as a causal/contributory factor. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36376 | 1220648-2022-01407 | 14309258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14309258&pc=OZD | 07/26/2020 | 05/05/2022 |  | 05/06/2022 |  |  | 04/05/2022 | 05/06/2022 |  |  | 01/31/2022 |  | 05/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1465638 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 04/05/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp device. during support, the patient developed an impella related vascular injury. unspecified intervention was required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36377 | 1220648-2022-01411 | 14309368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14309368&pc=OZD | 08/27/2020 | 05/05/2022 |  | 05/06/2022 |  |  | 04/05/2022 | 05/06/2022 |  |  | 06/30/2022 |  | 05/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1484369 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/05/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp device. during support, the patient endured an ischemic stroke that which the impella could not be ruled out as a factor. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36378 | 1220648-2022-01410 | 14309390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14309390&pc=OZD | 03/20/2021 | 05/05/2022 |  | 05/06/2022 |  |  | 04/05/2022 | 05/06/2022 |  |  | 10/31/2022 |  | 05/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022028478 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/05/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for high risk percutaneous coronary intervention using the impella cp for hemodynamic support. during support, the patient endured an ischemic stroke. after 2 days, the patient's condition deteriorated and he was eventually withdrawn from care. the physician attributes the impella related stroke as a factor in the patient's death. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36379 | 1220648-2022-01409 | 14309393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14309393&pc=OZD | 06/22/2021 | 05/05/2022 |  | 05/06/2022 |  |  | 04/05/2022 | 05/06/2022 |  |  | 01/31/2023 |  | 05/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022038034 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/05/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient suffered an ischemic stroke that which the impella device could not be ruled out as a causal or contributory factor. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36380 | 1220648-2022-01412 | 14309794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14309794&pc=OZD | 04/05/2022 | 06/08/2022 |  | 05/06/2022 |  |  | 04/05/2022 | 05/06/2022 |  |  | 12/31/2023 |  | 05/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023172843 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2022 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support via impella cp optical device. an access site bleed and hematoma formed during use of the device that prompted 2 units of blood product delivered. | [5 empty] | The impella device was received and an investigation is underway. a supplemental mdr will be filed at the completion of our analysis. | The returned product was visually inspected and no abnormalities were observed. during testing, the guidewire repositioning unit was inspected and the silicone was found to have deformation. the cause of the access site bleeding was a damaged valve due to improper installation of the silicone valve.
36381 | 1220648-2022-01400 | 14285408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14285408&pc=OZD | 04/01/2022 | 05/22/2025 |  | 05/05/2022 |  |  | 04/05/2022 | 05/05/2022 |  |  | 12/31/2023 |  | 05/05/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023167847 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/05/2022 |  |  |  |  | The complainant in (b)(6) had an impella cp-ecmo support for just under 12 days. at explant the team observed thrombus burden and performed a fogarty embolectomy to the impella limb. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the thrombus was removed, the patient experienced limb ischemia in the contralateral lower limb. it was reported that the limb ischemia was resolved by troubleshooting, the type of medical intervention is unknown. no further information was provided. |  |  |  |  | The medical center in (b)(6) discarded all product at explant. further clinical details and imaging have not been provided. should more be shared and a definitive root cause be determined, a final report will be filed. the failure mode will be monitored and trended. | The investigation for the reported limb ischemia issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. however, the investigation was completed, and the following was concluded: the relationship between impella and ischemia is unknown as it occurred on the contralateral limb to the impella. the investigation for the thrombus is in process, when completed a supplemental will be filed. as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The relationship between impella and ischemia is unknown as it occurred on the contralateral limb to the impella. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation for thrombus is not determined. g1: corrected reporting contact first name, reporting contact last and reporting contact email.
36382 | 1220648-2022-01404 | 14285506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14285506&pc=OZD | 04/06/2022 | 05/02/2022 |  | 05/05/2022 |  |  | 04/06/2022 | 05/05/2022 |  |  | 04/30/2023 |  | 05/05/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022073153 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/06/2022 |  |  |  |  | The complainant in (b)(6) had a successful impella cp support in (b)(6) 2021. later in (b)(6) 2022 the (b)(6) registry noted a cva. the cva occurred the day of explant and was ischemic in nature. the medical team's entry to the (b)(6) registry could not deny nor proove impella use contributed to the cva. | [5 empty] | The medical center discarded all impella product in 2021. no analysis to root cause is possible. should more information be shared and root cause of the cva be determined, a final report will be filed. the failure mode will be monitored and trended.
36383 | 1035166-2022-00059 | 14279851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14279851&pc=DYB | 03/19/2022 | 05/03/2022 |  | 05/04/2022 |  |  | 04/13/2022 | 05/04/2022 |  |  | 04/01/2023 |  | 05/04/2022 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-19198 |  | 0052-3021 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that we received a complaint involving 2 consective oscor 23 fr introducers. a (b)(6) male underwent a post cardiotomy cardiogenic shock procedure as the patient was having chest pain, nausea and chronic heart failure. the physician alleged the introducer had split distal to the tip when tracking the rp through the sheath, attempting to enter the inferior vena cava. a new 23 fr introducer was used, although the physician endured the same failure. the physician noted stenosis in the left common iliac vein was causing friction between the introducer and catheter, thus, unable to advance the impella into proper placement. there was no harm to the patient and support was escalated to ecmo. | [5 empty] | The device was not returned for analysis, as it was discarded. the investigation was focused on a review of product documentation. the device used in treatment. the device was not returned for analysis, therefore, the exact cause of the product issue cannot be determined and the clinical observation could not be confirmed. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. however, the following controls are in place to mitigate the reported product issue. as per procedure (abiomed introducer sheath in-process and final inspection), the devices in this lot were inspected as follows: 9.2.2 visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. per section 12.2, visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. per ifu: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. abiomed provides the ifu for this product. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
36384 | 1220648-2022-01405 | 14281966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14281966&pc=OZD | 04/07/2022 | 05/03/2022 |  | 05/04/2022 |  |  | 04/07/2022 | 05/04/2022 |  |  | 01/31/2023 |  | 05/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022060017 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/07/2022 |  |  |  |  | The medical team in (b)(6) had an impella cp support that was successful in (b)(6) 2021. the team later in (b)(6) 2022 had a jpvad registry noted access site bleed that was treated by blood product infusion on day prior to explant. | [5 empty] | The medical center discarded all impella product at explant. no investigation or benchtop testing/analysis is possible. root cause determination is not possible. should more information be shared that leads to a definative root cause, a final report will be filed.
36385 | 1220648-2022-01406 | 14282057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14282057&pc=OZD | 04/19/2022 | 07/12/2022 |  | 05/04/2022 |  |  | 04/19/2022 | 05/04/2022 |  |  | 06/30/2023 |  | 05/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022098336 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia Stroke | [11 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2022 |  |  |  |  | The medical center in (b)(6) had a successful impella 5.0 support in (b)(6) 2022. later in (b)(6) 2022 the jpvad registry noted there had been an ischemic cva during the support. there was no noted treatment but, the allegation was made against possibly the impella contributing to the event. | [5 empty] | The medical center in (b)(6) discarded all impella product in (b)(6) 2022 at the pump explant. the support was one that had exceed the duration of use, as the pump was explanted at day 47. no pump analysis is possible. no data logs were returned, nor any imaging. the manufacturer is unable to determine root cause with known information to date. the failure mode will be monitored and trended. | Since the original report was filed more clinical details have been shared. the jpvad registry team updated there was suspected hemolysis during the impella support that was treated with blood products. as all product was discarded in february of 2022 no investigation to root cause is possible. the failure mode will be monitored and trended.
36386 | 1220648-2022-01403 | 14272142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14272142&pc=OZD | 10/14/2020 | 05/02/2022 |  | 05/03/2022 |  |  | 04/20/2022 | 05/03/2022 |  |  | 04/30/2020 |  | 05/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1488338 |  | 0048-0034-JP |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 04/20/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for hemodynamic support using the impella cp device. during support, the patient was suspected as enduring hemolysis due to urine color. haptoglobin was delivered as treatment. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36387 | 1220648-2022-01402 | 14272174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14272174&pc=OZD | 06/12/2020 | 05/02/2022 |  | 05/03/2022 |  |  | 04/07/2022 | 05/03/2022 |  |  | 05/31/2021 |  | 05/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1424953 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/07/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with cardiomyopathy for impella support. during support, hemolysis and renal failure were alleged. the patient was treated with dialysis and a blood transfusion, as a result. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36388 | 1220648-2022-01399 | 14250040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14250040&pc=OZD | 04/19/2022 | 04/29/2022 |  | 04/29/2022 |  |  | 04/05/2022 | 04/29/2022 |  |  | 10/31/2022 |  | 04/29/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED IN. |  |  | US |  | 2021019862 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/05/2022 |  |  |  |  | The medical team in (b)(6) had an impella cp support in (b)(6) 2022 that was successful until wean and explant on day 3. later in (b)(6) 2022 the (b)(6) registry noted an access site bleed treated by blood product infusion and a skin suture. | [5 empty] | The medical center in (b)(6) discarded all impella pump and introducer sheath product in 2021. no investigation to root cause is possible. should the medical center or registry team share more information, a final report with root cause of the bleed will be filed.
36389 | 1220648-2022-01401 | 14250139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14250139&pc=OZD | 11/13/2020 | 04/29/2022 |  | 04/29/2022 |  |  | 04/05/2022 | 04/29/2022 |  |  | 07/31/2022 |  | 04/29/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1493301 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [18 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/05/2022 |  |  |  |  | The medical team in (b)(6) had a successful impella cp support in (b)(6) 2020. later in (b)(6) 2022 the (b)(4) registry noted an ischemic cva occurring during the support. the medical team could not confirm the use of impella was related or not to the outcome/adverse event. | [5 empty] | The medical team in (b)(6) discarded all product in 2020. the (b)(6) registry team furnished no further details that could lead to root cause of the cva. the failure mode will be monitored and trended.
36390 | 1035166-2022-00058 | 14254682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14254682&pc=DYB | 03/13/2022 | 09/16/2022 |  | 04/29/2022 |  |  | 04/07/2022 | 04/29/2022 |  |  | 07/01/2024 |  | 04/29/2022 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | DP-13652 |  | 0052-3021 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that the customer was notified of a complaint involving an oscor 23 fr introducer. a (b)(6) male patient underwent a procedure for decompensated post aortic valve replacement and required hemodynamic support from impella rp. the physician reported the introducer splitting while preparing the patient for rp support. a new 23 fr oscor introducer was used without further issue. no patient harm was reported. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. | The device was not returned for analysis, as it was discarded by the customer; however,customer provided the picture of the sheath which showed that sheath was spilt at the distal end. probable cause of the sheath split could be because the physician utilized the left groin for access, which is not the best practice or sheath met with the resistance while advancing into the patient body. no further follow-up is required. no corrective or preventive action resulted after investigation of this event. the event will be reevaluated if additional information becomes available. oscor will continue to monitor this event type. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | The device was not returned for analysis, as it was discarded. the investigation was focused on a review of product documentation. device was used in treatment. the device was not returned for analysis, therefore, the exact cause of the product issue cannot be determined and the clinical observation could not be confirmed. however, following controls are in place to mitigate the reported product issue. the complaint was confirmed using a photo provided by the customer. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure (abiomed introducer sheath in-process and final inspections): measure dimensions: measure tip i.d. of sheath using appropriate pin gauges to be in range of 0.305" - 0.307" for 23f. visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. per procedure, visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. the instructions for use (ifu) informs the user: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. no further follow-up is required. no corrective or preventive action resulted after investigation of this event. the event will be reevaluated if additional information becomes available. oscor will continue to monitor this event type. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36391 | 1220648-2022-01393 | 14221412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14221412&pc=OZD | 10/14/2021 | 04/27/2022 |  | 04/27/2022 |  |  | 03/29/2022 | 04/27/2022 |  |  | 07/31/2023 |  | 04/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022103866 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Unspecified Vascular Problem | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/29/2022 |  |  |  |  | The complainant reported a patient enduring an access site infection and vessel damage during use of the device. the infection was treated with antibacterial drug administration and the vessel damage was treated via unspecified intervention. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36392 | 1220648-2022-01394 | 14221609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14221609&pc=OZD | 10/25/2021 | 04/27/2022 |  | 04/27/2022 |  |  | 04/01/2022 | 04/27/2022 |  |  | 06/30/2023 |  | 04/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022100674 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/01/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient endured ischemia and required use of lower limb blood feeding to resolve ischemia in the impella inserted leg. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36393 | 1220648-2022-01395 | 14221651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14221651&pc=OZD | 03/30/2022 | 04/27/2022 |  | 04/27/2022 |  |  | 03/30/2022 | 04/27/2022 |  |  |  |  | 04/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | [5 empty] | 20 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/30/2022 |  |  |  |  | While conducting a literature review, a us based publication of 6 5.0 cases was found. in case 4 there was noted hemolysis and acute kidney injury that prompted pump explant and replacement with ecmo. | [5 empty] | The medical center did not alert abiomed to the hemolysis when it occurred. there was no pump return or any benchtop hemolysis testing. the root cause of the hemolysis could not be determined. the failure mode will be monitored and trended.
36394 | 1220648-2022-01397 | 14225338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14225338&pc=OZD | 04/01/2022 | 06/16/2022 |  | 04/27/2022 |  |  | 04/01/2022 | 04/27/2022 |  |  | 11/30/2023 |  | 04/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023166575 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2022 |  |  |  |  | The medical center had an impella cp support placed via the femoral artery. the femoral was observed to be perforated. the team treated the perforation by blood products and covered stenting. the patient was stable after explant and the patient was sent to the icu on no hemodynamic support. | [5 empty] | The medical center returned the impella pump for analysis. the investigation is on-going at this time. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation has concluded. the product had no defect observed at inspection of the impella pump. when the clinical report was reviewed, the team had reported that excessive force was applied when advancing the pump. the patient injury of vascular damage was due to a user error. the failure mode will be monitored and trended.
36395 | 1220648-2022-01398 | 14225339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14225339&pc=OZD | 04/01/2022 | 04/27/2022 |  | 04/27/2022 |  |  | 04/01/2022 | 04/27/2022 |  |  | 04/30/2023 |  | 04/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2022075419 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/01/2022 |  |  |  |  | In (b)(6) 2021 the medical center in (b)(6) documented upon an impella cp support of over 3 days. later in (b)(6) 2022 the (b)(6) registry noted an access site bleed that was treated by blood product infusion in excess of 2 units. | [5 empty] | The medical center in (b)(6) discarded all product in 2021. only the minimal information provided in the (b)(6) registry could be evaluated. there was no determination of root cause possible. should more information be shared, and a root cause determined, a final report will be filed.
36396 | 1220648-2022-01396 | 14225355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14225355&pc=OZD | 04/01/2022 | 04/29/2022 |  | 04/27/2022 |  |  | 04/01/2022 | 04/27/2022 |  |  | 12/31/2023 |  | 04/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED INC. |  |  | US |  | 2023163145 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2022 |  |  |  |  | The medical center in the eu had an impella cp which was removed for hemolysis. the team explanted and there was no replacement of the pump. when the pump was explanted the pump was observed to have biomaterial adhered to the pump. | [5 empty] | The medical team in the eu returned the impella pump product. the analysis and benchtop testing is on-going at this time. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation has concluded. the medical team returned the impella pump and benchtop hemolysis testing was performed. the pump passed the hemolysis testing. the root cause was native patient condition. the failure mode will be monitored and trended.
36397 | 1220648-2022-01392 | 14214562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14214562&pc=OZD | 03/30/2022 | 04/26/2022 |  | 04/26/2022 |  |  | 03/30/2022 | 04/26/2022 |  |  |  |  | 04/26/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | [5 empty] | 38 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/30/2022 |  |  |  |  | While conducting a literature review, a us based publication of six impella 5.0 cases was found. in case #3 there was noted hemolysis that prompted pump explant and replacement with ecmo. this occurred on day 4 of the support for the patient with cardiomyopathy. | [5 empty] | The medical center did not alert abiomed to the hemolysis when it occurred. there was no pump return or any benchtop hemolysis testing. the root cause of the hemolysis could not be determined. the failure mode will be monitored and trended.
36398 | 1220648-2022-01391 | 14197150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14197150&pc=OZD | 04/28/2021 | 04/25/2022 |  | 04/25/2022 |  |  | 03/29/2022 | 04/25/2022 |  |  | 10/31/2022 |  | 04/25/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2021011333 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 53 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/29/2022 |  |  |  |  | In (b)(6) 2021 the medical center in (b)(6) had an impella cp that was supporting a patient successfully until wean and explant. later in (b)(6) 2022 the (b)(4) registry noted an impella access site bleed. the bleed was treated with blood product infusion. | [5 empty] | The medical center discarded all product at explant in 2021. no benchtop analysis and testing to root cause is possible. should the registry team or medical center share more details, which lead to the root cause, a final report will be filed. the failure mode will be monitored and trended.
36399 | 1220648-2022-01390 | 14168262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14168262&pc=OZD | 03/23/2022 | 04/21/2022 |  | 04/21/2022 |  |  | 03/24/2022 | 04/21/2022 |  |  | 10/31/2023 |  | 04/21/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022157374 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/24/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient developed signs of hemolysis and renal failure via red urine and hemolyzed labs. the pump was removed and hemolysis improved. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36400 | 1220648-2022-01384 | 14163529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14163529&pc=OZD | 03/28/2022 | 04/20/2022 |  | 04/20/2022 |  |  | 03/28/2022 | 04/20/2022 |  |  | 01/31/2024 |  | 04/20/2022 |  | TEMPORARY NON-ROLLER TYPED LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED INC. |  |  | US |  | 2023179486 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/28/2022 |  |  |  |  | The medical team in (b)(6) had an impella 5.5 support for 2 weeks when the pump was explanted due to hemolysis concerns. there was no further intervention needed. the pump when explanted was observed to have biomaterial adhered to the pump. | [5 empty] | The medical center in (b)(6) returned the impella pump product for benchtop analysis and hemolysis testing. the testing revealed the root cause to be biomaterial ingestion. when biomaterial was removed, the pump ran to specification. the failure mode will be monitored and trended.
36401 | 1220648-2022-01385 | 14164348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14164348&pc=OZD | 04/27/2021 | 04/20/2022 |  | 04/20/2022 |  |  | 03/23/2022 | 04/20/2022 |  |  | 12/31/2022 |  | 04/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022032158 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/23/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with myocarditis for hemodynamic support using the impella cp device. during support, the patient was suspected to have endured hemolysis and renal failure. dialysis was administered as a result. additionally, an access site bleed developed for which a blood transfusion was required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36402 | 1220648-2022-01386 | 14164351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14164351&pc=OZD | 02/28/2022 | 12/14/2024 |  | 04/20/2022 |  |  | 03/23/2022 | 04/20/2022 |  |  | 07/31/2023 |  | 04/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022105542 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/23/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with cardiomyopathy for hemodynamic support using the impella 5.0. during support, the patient developed an access site bleed. a blood transfusion was administered. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | Additional information indicated that the patient experienced a stroke during the use of impella. |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01386 was provided in sections b2, b5, h1, and h6. | Corrected information provided in b5 and h6.
36403 | 1220648-2022-01387 | 14164367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14164367&pc=OZD | 03/25/2022 | 10/14/2024 |  | 04/20/2022 |  |  | 03/30/2022 | 04/20/2022 |  |  | 11/30/2023 |  | 04/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023160848 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient developed ischemia in the impella inserted limb. as treatment, the device was prematurely removed. | It was also reported the patient experienced bleeding and hematoma while on support. a blood transfusion was given for treatment. no further information was provided. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation for the reported limb ischemia issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were sufficient to determine a root cause. the cause of the limb ischemia was most likely a combination of patient condition and impella since the blood flow improved after impella removal. the investigation for the access site bleeding is in process and a supplemental will be sent when completed. instructions for use for the related event are as follows: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. additional information for the initial mfr 1220648-2022-01387 was provided in sections b1,b2,b5, d6a, d6b, g1, h1, and h6.
36404 | 1220648-2022-01388 | 14164371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14164371&pc=OZD | 03/15/2020 | 10/13/2024 |  | 04/20/2022 |  |  | 03/29/2022 | 04/20/2022 |  |  | 09/30/2021 |  | 04/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1444716 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2022 |  |  |  |  | The complainant reported a (b)(6) female presenting with myocarditis for hemodynamic support using the impella cp device. during support, a vascular injury was endured in relationship to the device. intervention of unspecified nature was administered. the patient also developed an access site bleed requiring a blood transfusion. | It was also reported the patient experienced an infection at the insertion site. however, additional information could not be provided, due to the facility visiting restrictions. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation for the reported infection has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the infection could not be determined as it occurred 8 days after the removal of impella. the impella was sterilized under validated conditions. the failure modes will be monitored and trended. as noted in the impella ifu: impella cp with smart assist system section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. additional information for the initial mfr 1220648-2022-01388 was provided in sections b1, b2, b5, d6a, d6b, g1, h1, and h6.
36405 | 1220648-2022-01389 | 14164373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14164373&pc=OZD | 03/28/2022 | 06/22/2022 |  | 04/20/2022 |  |  | 03/29/2022 | 04/20/2022 |  |  | 11/30/2023 |  | 04/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023174939 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Fragmentation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/29/2022 |  |  |  |  | The complainant reported an (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical. while removing the device, the cannula detached. vascular surgery was required to remove the fragmented piece. | [5 empty] | The impella device has been received and an evaluation is underway. upon completion, a supplemental mdr will be filed. | This correction report is being filed to revise the event occurrence and aware date to march 28, 2022. originally, we had entered the occurrence and aware date in error as march 29, 2022. | The pump was returned and the inflow cage detachment was confirmed. evidence suggests the metal of the inflow had cut into the cannula during explant, thus, resulting in detachment.
36406 | 1220648-2022-01377 | 14144766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14144766&pc=OZD | 03/24/2022 | 05/27/2022 |  | 04/19/2022 |  |  | 03/24/2022 | 04/19/2022 |  |  | 09/30/2023 |  | 04/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET APAC |  |  | ABIOMED INC. |  |  | US |  | 2022142691 |  | 0048-0047 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | Distributor |  | 66 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/24/2022 |  |  |  |  | The medical team in singapore placed an impella cp to support a patient in cardiogenic shock. the team peeled away the introducer sheath and advanced the impella's own repositioning sheath. at this time the pump kinked and was snared out of the patient's native anatomy. a new second impella pump was successfully placed and pump support proceeded. | [5 empty] | The medical team in singapore returned the impella pump product for analysis. the investigation is on-going at this time. upon completion a final report will be filed. | Since the original report was filed, the investigation into the pump kinking and removal by snare has concluded. the pump was returned and kink was confirmed. the kink was due to improper positioning, and this necessitated the snared removal. the root cause is techique/user related. the failure mode will be monitored and trended.
36407 | 1220648-2022-01378 | 14151614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14151614&pc=OZD | 03/23/2022 | 04/19/2022 |  | 04/19/2022 |  |  | 03/23/2022 | 04/19/2022 |  |  | 11/30/2023 |  | 04/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023165838 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/23/2022 |  |  |  |  | The medical center in (b)(6) had an impella cp supporting a patient admitted in cardiogenic shock. the cp was placed via the femoral artery. the placement was via the femoral despite their knowledge of peripheral arterial disease at the site. the limb became ischemic with the placement of the cp. they placed a contralateral leg perfusion catheter. this did not fully perfuse the limb and the pump was explanted after 5 days of use. | [5 empty] | The medical team in japan discarded the impella product. no benchtop analysis is possible. it was known the patient had peripheral arterial disease, and as such native patient anatomy contributed to the outcome. should further clinical details be furnished that leads to root cause, beyond native condition, a final report will be filed.
36408 | 1220648-2022-01380 | 14151669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14151669&pc=OZD | 03/26/2022 | 04/19/2022 |  | 04/19/2022 |  |  | 03/26/2022 | 04/19/2022 |  |  | 01/31/2024 |  | 04/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023182760 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/26/2022 |  |  |  |  | The medical team had an impella cp supporting a patient when it was having positional issues and hemolysis observed. the lab levels were rising and urine color change had occurred. the team made choice to give 1 unit of blood products and explant the pump. the team then placed an iabp. | [5 empty] | The medical team discarded the impella pump product at the explant. no benchtop analysis is possible. should more clinical details be shared that lead to a definitive root cause, a final report will follow.
36409 | 1220648-2022-01379 | 14151675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14151675&pc=OZD | 03/23/2022 | 05/12/2022 |  | 04/19/2022 |  |  | 03/23/2022 | 04/19/2022 |  |  | 12/31/2023 |  | 04/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2023178248 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/23/2022 |  |  |  |  | The medical center had an impella cp placed to the left ventricle, for support of patient admitted with multiple cardiac comorbidities. when the cp was placed it was in poor position under the papillary. when the team attempted to reposition the pump, it kinked and fractured to 2 pieces and was fully snared out. the team then placed a 2nd cp pump for support. | [5 empty] | The medical center has not returned any impella product for analysis and benchtop testing. to date no root cause has been determined. should the investigation lead to root cause of the pump separation, a final report will be filed. | Since the original report was filed, the investigation into the pump damage has concluded. the cause of the damage and detachment was a user error and improper technique of advancement. the failure mode will be monitored and trended.
36410 | 1220648-2022-01381 | 14152247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14152247&pc=OZD | 03/21/2022 | 07/01/2022 |  | 04/19/2022 |  |  | 03/21/2022 | 04/19/2022 |  |  | 07/31/2023 |  | 04/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2022121676 |  | 005062 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella 5.0. while removing the device, the impella's pigtail became caught on the patient's axillary graft and separated. intervention was required to reopen the site and remove the fragmented device component. | [5 empty] | The impella device has been received and an evaluation is underway. upon completion, a supplemental mdr will be filed with the results. | The device was received and the pigtail was confirmed as fragmented. after review of the information provided, the cause of the removal issue was linked to use of a graft not indicated for use with impella.
36411 | 1220648-2022-01383 | 14152405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14152405&pc=OZD | 02/22/2021 | 04/19/2022 |  | 04/19/2022 |  |  | 03/23/2022 | 04/19/2022 |  |  | 09/30/2022 |  | 04/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1503408 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/23/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with myocarditis for hemodynamic support using the impella cp. during support, the patient developed an access site bleed. a blood transfusion was required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36412 | 1220648-2022-01382 | 14152407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14152407&pc=OZD | 03/22/2022 | 04/19/2022 |  | 04/19/2022 |  |  | 03/22/2022 | 04/19/2022 |  |  | 11/30/2023 |  | 04/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022157961 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/22/2022 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. after advancing the repositioning sheath, the patient developed an access site bleed. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36413 | 14134815 | 14134815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14134815&pc=OZD | 04/12/2022 | 04/13/2022 | 04/13/2022 | 04/18/2022 | 04/13/2022 | 04/18/2022 |  | 04/18/2022 |  |  |  |  | 04/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | Yes |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 25915 DA |  | Male | 93 KG |  | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Chest Pain | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Pulmonary Edema | Heart Block | [8 empty] | Death | [24 empty] | Unknown |  |  | Yes | Initial submission | Followup |  |  |  |  | Y | HOSPITAL |  |  | An elderly patient with a past medical history (pmh) of diabetes mellitus (dm), chronic obstructive pulmonary disease (copd), hypertension (htn), hyperlipidemia (hld), coronary artery disease (cad) with prior stent, who presented from rural town with 2 hours of chest pain. anterior stemi present on ekg. patient was given thrombolytics, aspirin and started on a heparin drip. after transfer to our emergency department (ed) continued to be hypotensive and in pulmonary edema. intubated and taken for left heart catheterization (lhc). stenting performed on an ostial left anterior descending artery (lad) and right coronary artery (rca) lesion. patient remained hypotensive during the procedure and had a 2.5 impella placed. brief episode of complete heart block with placement of a temporary pacer. he then resumed his underlying rhythm. remained hypotensive and with poor left ventricular (lv) function. taken to the operating room (or) for 5.5 impella placement. remained hemodynamically unstable throughout the case requiring ne and epinephrine drips. bleeding noted at the axilla insertion site with subsequent repair. 2 units of red blood cells (rbcs) and 1 of platelets given. estimated blood loss (ebl) approximately 500 ml. despite above remained hypotensive and started on dobutamine. several amps of bicarbonate and calcium given. essentially anuric throughout the case and had pulmonary edema with difficulty oxygenating requiring higher positive end-expiratory pressure (peep) settings. on arrival out of the or he is essentially totally dependent on the impella for cardiac output and has a pulse pressure only 3-4. patient with fluid resuscitation with crystalloid, blood products and colloids; some waking, able to decrease pressors. patient with decreasing oxygen requirements, clearing lungs. impella failed; patient with cardiopulmonary resuscitation (cpr) initiated, to cath lab with cpr for replacement initially of impella in the groin, then transfer to axillary location. patient with poor perfusion and metabolic panels. loss of contractility ensued; patient to icu and declared dead. impella leaking blood from both connection sites on red portion of tubing connected to reservoir. surgeon and representative made aware. | An elderly patient with a past medical history (pmh) of diabetes mellitus (dm), chronic obstructive pulmonary disease (copd), hypertension (htn), hyperlipidemia (hld), coronary artery disease (cad) with prior stent, who presented from rural town with 2 hours of chest pain. anterior stemi present on ekg. patient was given thrombolytics, aspirin and started on a heparin drip. after transfer to our emergency department (ed) continued to be hypotensive and in pulmonary edema. intubated and taken for left heart catheterization (lhc). stenting performed on an ostial left anterior descending artery (lad) and right coronary artery (rca) lesion. patient remained hypotensive during the procedure and had a 2.5 impella placed. brief episode of complete heart block with placement of a temporary pacer. he then resumed his underlying rhythm. remained hypotensive and with poor left ventricular (lv) function. taken to the operating room (or) for 5.5 impella placement. remained hemodynamically unstable throughout the case requiring ne and epinephrine drips. bleeding noted at the axilla insertion site with subsequent repair. 2 units of red blood cells (rbcs) and 1 of platelets given. estimated blood loss (ebl) approximately 500 ml. despite above remained hypotensive and started on dobutamine. several amps of bicarbonate and calcium given. essentially anuric throughout the case and had pulmonary edema with difficulty oxygenating requiring higher positive end-expiratory pressure (peep) settings. on arrival out of the or he is essentially totally dependent on the impella for cardiac output and has a pulse pressure only 3-4. patient with fluid resuscitation with crystalloid, blood products and colloids; some waking, able to decrease pressors. patient with decreasing oxygen requirements, clearing lungs. impella failed; patient with cardiopulmonary resuscitation (cpr) initiated, to cath lab with cpr for replacement initially of impella in the groin, then transfer to axillary location. patient with poor perfusion and metabolic panels. loss of contractility ensued; patient to icu and declared dead. impella leaking blood from both connection sites on red portion of tubing connected to reservoir. surgeon and representative made aware.
36414 | 1220648-2022-01374 | 14138055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14138055&pc=OZD | 03/21/2022 | 05/12/2022 |  | 04/18/2022 |  |  | 03/21/2022 | 04/18/2022 |  |  | 11/30/2023 |  | 04/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023164776 |  | 0550-0008 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2022 |  |  |  |  | The medical center had an impella 5.5 supporting an operative patient. she was having heart bypass (cabg) surgery and the impella pump was placed in surgical mode while in the operating suite. during the surgical procedure the pump stopped due to a user error, as they used a hard clamp on the pump motor, rather than the soft clamp that is recommended. the stopped pump was explanted and the patient remained stable on the ecmo alone. | [5 empty] | The medical center discarded the impella pump product at explant. further clinical details and data log retrieval are expected. upon completion of the investigation, a final report will be filed. the failure mode will be monitored and trended. | Since the original report was filed, the investigation into the pump stop has concluded. the medical center never did return the impella product as expected. only the clinical report was analyzed and it found the stop occurred due to user error of applying a hard clamp to the motor housing. the hard clamp was applied to the pump in surgery. the failure mode will be monitored and trended.
36415 | 1220648-2022-01375 | 14138067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14138067&pc=OZD | 12/15/2021 | 10/09/2025 |  | 04/18/2022 |  |  | 03/22/2022 | 04/18/2022 |  |  | 02/28/2023 |  | 04/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022064783 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 59 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2022 |  |  |  |  | The medical center in (b)(6) had an impella cp support in (b)(6) 2021 along with ecmo. the support was successful. later in (b)(6) 2022 the jpvad registry noted hemolysis occurring during support. the issue was not treated nor did it cause any lasting harm or injury. later in the month of (b)(6) 2022 the jpvad registry team further noted access site bleeding treated by 2+ units of blood. | [5 empty] | The medical center discarded all product in 2021, and as such no product investigation to root cause could be done. the failure mode will be monitored and trended.
36416 | 1220648-2022-01376 | 14138069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14138069&pc=OZD | 02/27/2021 | 04/18/2022 |  | 04/18/2022 |  |  | 03/22/2022 | 04/18/2022 |  |  | 10/31/2022 |  | 04/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2021014956 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/22/2022 |  |  |  |  | In 2021 the medical team in (b)(6) had a successful impella cp support. in (b)(6) 2022 the jpvad registry noted an access site bleed attributed to the use of impella treated with 2+ units of blood. | [5 empty] | The medical center discarded all impella pump and sheath product in 2021. no investigation to root cause is possible. the failure mode will be monitored and trended.
36417 | 1220648-2022-01364 | 14118191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14118191&pc=OZD | 03/18/2022 | 04/15/2022 |  | 04/15/2022 |  |  | 03/18/2022 | 04/15/2022 |  |  | 11/30/2023 |  | 04/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023166228 |  | 0550-0008 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/18/2022 |  |  |  |  | The complainant reported a (b)(6) year old female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella 5.5 device. an access site bleed/hematoma formed during support. as treatment, surgical repair was required. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36418 | 1220648-2022-01365 | 14120000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14120000&pc=OZD | 03/18/2022 | 06/03/2022 |  | 04/15/2022 |  |  | 03/18/2022 | 04/15/2022 |  |  | 01/31/2023 |  | 04/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022038010 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2022 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. after pump initiation, the cannula "kinked back on itself" and required a radial snare to straighten and remove the device. | [5 empty] | The impella device was received and an investigation is underway. a supplemental mdr will be filed at the conclusion of the investigation. | The device was received and a cannula kink was confirmed. after review of the clinical information provided, we feel the root cause of the kink was related to improper technique, as the physician advised the cannula may have been kinked during insertion through the introducer.
36419 | 1220648-2022-01366 | 14120033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14120033&pc=OZD | 03/11/2022 | 04/15/2022 |  | 04/15/2022 |  |  | 03/16/2022 | 04/15/2022 |  |  | 10/31/2023 |  | 04/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022148621 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/16/2022 |  |  |  |  | The complainant reported a (b)(6) old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical. during support, the patient was suspected as enduring hemolysis, as evident by dark urine color. the pump was removed and replaced with an intra-aortic balloon pump. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36420 | 1220648-2022-01367 | 14120079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14120079&pc=OZD | 01/01/2022 | 04/15/2022 |  | 04/15/2022 |  |  | 03/16/2022 | 04/15/2022 |  |  | 09/30/2023 |  | 04/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022132592 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/17/2022 |  |  |  |  | The complainant reported a (b)(6) patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient developed an access site bleed that prompted a blood transfusion. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36421 | 1220648-2022-01368 | 14120107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14120107&pc=OZD | 12/29/2021 | 09/08/2022 |  | 04/15/2022 |  |  | 03/17/2022 | 04/15/2022 |  |  | 08/31/2023 |  | 04/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022127537 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 16 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/17/2022 |  |  |  |  | The complainant reported a (b)(6) old female patient presenting with myocarditis for hemodynamic support using the impella cp optical device. during support, the patient developing an impella access site bleed. as treatment, a blood transfusion was administered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | We erroneously filed a supplemental mdr reporting thrombocytopenia. we have been made aware that this did not occur with this pump. we are filing this supplemental mdr as a correction to redact reference to thrombocytopenia. | New information obtained 2022-08-22 reports the patient endured thrombocytopenia resulting from impella use, prompting platelet delivery. the impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36422 | 1220648-2022-01369 | 14120167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14120167&pc=OZD | 03/17/2022 | 05/31/2022 |  | 04/15/2022 |  |  | 03/17/2022 | 04/15/2022 |  |  | 01/31/2024 |  | 04/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023177943 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/17/2022 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with cardiomyopathy for hemodynamic support using the impella 5.0 device. during support, the physician switched the device to surgical mode to complete required intervention related to procedure bleeding. after intervention was completed, the 5.0 failed to restart when attempting to switch out of surgical mode. the pump was implanted for approximately 3 hours. a new pump was used without further issue. | [5 empty] | The device was received and an evaluation is underway. a supplemental mdr will be submitted upon completion of our analysis. | The device was received and an investigation was completed. data logs confirmed the reported pump failing to restart. a visual inspection revealed a foreign piece of fractured blue tubing observed within the impeller blades, and thus, believed to be the root cause of the reported device issue.
36423 | 1220648-2022-01370 | 14120597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14120597&pc=OZD | 03/18/2022 | 05/27/2022 |  | 04/15/2022 |  |  | 03/18/2022 | 04/15/2022 |  |  | 10/31/2023 |  | 04/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED INC. |  |  | GM |  | 2022146154 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | Hospitalization |  |  |  |  | CARDIOHELP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2022 |  |  |  |  | The medical center in the eu had placed an impella cp for hemodynamic support. the pump's femoral artery access site had a bleed that was treated by blood product infusion. additionally, the team noted limb ischemia. this ischemia prompted a perfusion catheter to be placed, and pump explant. a new impella 5.0 was placed via the upper limb. | [5 empty] | The medical center in the eu has returned the impella pump product for analysis. this analysis is currently on-going. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation into the access site bleeding and limb ischemia has concluded. the medical team returned a portion of the impella pump. no defects were found. a clinical review and analysis lead to the root causes. the access site bleed was due to the use issue of an improper technique at insertion. the limb ischemia was due to patient condition, as the patient had known peripheral arterial disease. the failure modes will be monitored and trended.
36424 | 1220648-2022-01371 | 14120599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14120599&pc=OZD | 03/18/2022 | 10/25/2023 |  | 04/15/2022 |  |  | 03/18/2022 | 04/15/2022 |  |  | 11/30/2023 |  | 04/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED INC. |  |  | US |  | 2023161178 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Insufficient Flow or Under Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/18/2022 |  |  |  |  | The medical center in (b)(6) escalated care from the impella cp to the impella 5.5 pump. the pump was explanted due to concerns of hemolysis after just over 25 hours. after explant the team placed another cp pump and support proceeded. | In addition to the original report of hemolysis, the user facility reported the device exhibiting saturated flow. |  |  |  |  | The medical team in switzerland returned the impella pump and data logs for analysis. the investigation is on-going at this time. upon completion, a final report will be filed. | Since the original report was filed, the investigation into the hemolysis has concluded. the medical team in the eu returned the impella pump product and benchtop analysis was completed. the determination was made that ingested biomaterial was the cause of the hemolysis. the failure mode will be monitored and trended. | (h3) the investigation into the reported "hemolysis and has been completed since the original report was filed. product was returned for investigation. two pictures were received from the field and they confirmed the presence of a large mass of biomaterial ingested into the outflow cage upon explant. the data logs revealed strong evidence of biomaterial ingestion, including flat motor current and saturated flows. it is very likely that this also contributed to the false "impella position in aorta" alarms, as placement signal metrics were normal. the device was visually inspected and found significant amounts of biomaterial on and around the impeller and outflow cage. no other abnormalities were found. in conclusion, it was determined that the cause of both the hemolysis and the saturated flow was ingested biomaterial. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
36425 | 1220648-2022-01372 | 14120636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14120636&pc=OZD | 03/18/2022 | 04/15/2022 |  | 04/15/2022 |  |  | 03/18/2022 | 04/15/2022 |  |  | 09/30/2023 |  | 04/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022140698 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/18/2022 |  |  |  |  | The medical team placed an impella cp for hemodynamic support via the femoral artery. after just over 2 hours the pump was explanted due to an access site bleed. the bleed was treated by 2 units of blood cell products and the placement of a covered stent. the patient was then supported by an iabp via the contralateral leg. | [5 empty] | The medical center discarded all impella pump product at explant. a full analysis is not possible. should the medical team share more clinical details that lead to a root cause, a final report will be filed. the failure mode will be monitored and trended.
36426 | 1220648-2022-01373 | 14120654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14120654&pc=OZD | 01/27/2022 | 10/14/2024 |  | 04/15/2022 |  |  | 03/29/2022 | 04/15/2022 |  |  | 10/31/2023 |  | 04/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022143678 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Hospitalization | Required Intervention | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2022 |  |  |  |  | In (b)(6) 2022 the medical team in (b)(6) had a successful impella cp support. months later in (b)(6) 2022 the jpvad registry noted that there was limb ischemia attributed to the use of impella. the registry documented that the ischemia was treated by contralateral leg infusion catheter placement. | Additional information was reported indicating the patient experienced hemolysis and access site bleeding. it was reported the hemolysis was confirmed through blood sampling and was resolved by lowering the pressure level of the pump. gastrointestinal bleeding was also noted during support, resulting in the patient requiring an unknown amount of blood transfusion. no further information is available. the patient remained on impella support and was successfully weaned. |  |  |  |  | The medical center in (b)(6) discarded the impella pump product. the clinical details alone were analyzed for root cause determination. the cause of the limb ischemia was determined to be native patient condition. the failure mode will be monitored and trended. | Additional information for the initial mfr 1220648-2022-01373 was provided in sections b1, b2, b5, d6a, d6b, g1, h1, and h6. the investigation for the reported hemolysis and gastrointestinal bleed has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the hemolysis was unable to be determined as no product or logs were returned for analysis. the root cause of the gastrointestinal bleed was determined to be unrelated to the use of impella. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: limb ischemia. impella cp with smart assist system. section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ access site bleeding. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis: impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system. section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
36427 | 1220648-2022-01363 | 14114170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14114170&pc=OZD | 03/22/2022 | 04/14/2022 |  | 04/14/2022 |  |  | 03/22/2022 | 04/14/2022 |  |  | 01/31/2023 |  | 04/14/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2022039695 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/22/2022 |  |  |  |  | The medical center in (b)(6) had an impella cp support in (b)(6)2021. the support was successful, and the support was weaned and pump explanted after 9 hours. later in (b)(6)2022 the jpvad registry documented and access site bleed attributed to the use of impella. the team had treated the bleed with 2+ units of blood products. | [5 empty] | The impella product was not returned for analysis, as it was discarded in 2021. the clinical report alone was analyzed. the bleeding was determined to be use related. the user's management of the access site was the cause of the bleed. the failure mode will be monitored and trended.
36428 | 1220648-2022-01362 | 14094106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14094106&pc=OZD | 03/16/2022 | 04/12/2022 |  | 04/13/2022 |  |  | 03/16/2022 | 04/13/2022 |  |  | 07/31/2022 |  | 04/13/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1488085 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/16/2022 |  |  |  |  | The medical center in (b)(6) had an impella cp support in (b)(6) 2020 which was successful until explant. in (b)(6) 2022 the (b)(6) registry review noted an access site bleed attributed to impella that was treated by blood product infusion. | [5 empty] | The medical center discarded all impella pump and sheath product in 2020. no product investigation to root cause is possible. the (b)(6) registry did not furnish further clinical details. the access site bleeding failure will be monitored and trended.
36429 | 1220648-2022-01361 | 14094188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14094188&pc=OZD | 03/16/2022 | 11/29/2024 |  | 04/13/2022 |  |  | 03/16/2022 | 04/13/2022 |  |  | 01/31/2024 |  | 04/13/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED INC. |  |  | US |  | 2023174269 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/16/2022 |  |  |  |  | The medical center in the (b)(6) had an impella 5.0 supporting a patient for 6 days along with ecmo for hemodynamic support. during the final day of support there were purge issues and the pump itself stopped and was replaced. | It was reported that there were recurrent "purge pressure high" alarms. a few hours later, the pump stopped. it was able to be restarted, however the pump speed was unable to exceed p-1. |  |  |  |  | The investigation into the impella 5.0 pump stopping is on-going at this time. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation into this pump stop in the eu has concluded. the data logs and pump were returned for analysis and benchtop testing. the cause of the pump stop was determined to be blood ingress to the pump. the failure mode will be monitored and trended. | Additional information for b5, h6, and h10 for the purge pressure high issue. and d6a, d6b for implant and explant dates. . this report is being filed as part of a retrospective review of historical records.
36430 | 1220648-2022-01356 | 14084779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14084779&pc=OZD | 02/19/2022 | 06/03/2022 |  | 04/12/2022 |  |  | 03/14/2022 | 04/12/2022 |  |  | 11/30/2023 |  | 04/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022157890 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient endured limb ischemia and an access site bleed. lower limb blood feeding was performed for ischemia and a blood transfusion was completed for access site bleeding. | New information was obtained from the physician on 2022-05-11 indicating the patient had endured an access site infection that prompted antibiotic administration. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | New information received 2022-05-11 indicating the patient had endured an infection that required antibiotics. filing this supplemental to capture this new allegation. the impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36431 | 1220648-2022-01358 | 14084898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14084898&pc=OZD | 03/13/2022 | 05/24/2022 |  | 04/12/2022 |  |  | 03/13/2022 | 04/12/2022 |  |  | 12/31/2023 |  | 04/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023177740 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2022 |  |  |  |  | The complainant reported a male patient presenting for high risk percutaneous coronary intervention for hemodynamic support using the impella cp optical device. during support, the patient endured red urine. due to the patient's hemolyzed labs, hemolysis was suspected. as a result, the pump was prematurely removed and the patient was ultimately escalated to an impella 5.0 device. | [5 empty] | The device was been received and an investigation is underway. a supplemental mdr will be filed upon completion of our analysis. | The pump was found to have biomaterial entrained within the impeller. evidence suggests the root cause of the reported hemolysis was ingested biomaterial.
36432 | 1220648-2022-01359 | 14084904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14084904&pc=OZD | 07/31/2020 | 04/12/2022 |  | 04/12/2022 |  |  | 03/14/2022 | 04/12/2022 |  |  | 09/30/2021 |  | 04/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1446706 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/14/2022 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with myocarditis for hemodynamic support using the impella cp device. during support, the patient developed an impella related access site bleed that prompted blood product delivery. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36433 | 1220648-2022-01357 | 14084922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14084922&pc=OZD | 02/13/2022 | 04/12/2022 |  | 04/12/2022 |  |  | 03/14/2022 | 04/12/2022 |  |  | 09/30/2023 |  | 04/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022132445 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/14/2022 |  |  |  |  | The complainant reported a male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient developed an access site bleed that prompted a blood transfusion. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36434 | 1220648-2022-01360 | 14084972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14084972&pc=OZD | 03/16/2022 | 04/12/2022 |  | 04/12/2022 |  |  | 03/16/2022 | 04/12/2022 |  |  | 09/30/2022 |  | 04/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1503408 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | User facility | Company representation |  |  | 53 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/16/2022 |  |  |  |  | The medical center in (b)(6) on (b)(6) 2021, had an impella cp support. the patient case was reviewed on (b)(6) 2022 for the jpvad registry and noted an access site bleed that prompted blood product transfusion. | [5 empty] | All products was discarded in 2021 at the time of the impella support. no investigation of the product is possible that could lead to root cause. the jpvad registry did not furnish further clinical details. the access site bleeding cause is not determined. the failure will be monitored and trended.
36435 | 1220648-2022-01355 | 14058060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14058060&pc=OZD | 03/02/2022 | 10/24/2024 |  | 04/08/2022 |  |  | 03/11/2022 | 04/08/2022 |  |  | 11/30/2023 |  | 04/08/2022 |  | (01)813502011609(10)2022151471(21)359082(17)231130 |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022151471 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Death | Other | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2022 |  |  |  |  | The complainant reported a (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. as the result of support, the patient developed ischemia. a forgarty catheter was used after device removal to clear a thrombus. | Additional information reported that continuous renal replacement therapy (crrt) red blood cells administration as appropriate due to systemic bleeding tendency, no bleeding at the impella insertion site. also there that the patient had hemolysis. after removal of the impella, the patient's condition gradually deteriorated, switched to palliative treatment and the patient expired on the second day after the removal of impella. the cause of death was the primary disease of myocardial infarction. although it is not caused by lower limb ischemia, it cannot be said that it is not related at all because the patient's original condition was bad. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation for the reported hemolysis, purge cassette leak, limb ischemia and vascular damage - peripheral vessel issues has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the hemolysis was most likely due to patient condition as the patient's left ventricle was unable to sufficiently fill. the root cause of the purge cassette leak was unable to be determined as no product was returned for this investigation so the location of the leak could not be identified. the root cause of the leg ischemia was determined to be thrombus due to patient condition. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2022-01355 was provided in sections b2, b5, d6, h1, and h6.
36436 | 1220648-2022-01354 | 14058124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14058124&pc=OZD | 03/10/2022 | 04/08/2022 |  | 04/08/2022 |  |  | 03/11/2022 | 04/08/2022 |  |  | 02/28/2023 |  | 04/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022062638 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/15/2022 |  |  |  |  | The complainant reported a (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella optical. during support, the patient developed an access site hematoma that prompted 10 units of blood product delivery. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36437 | 1220648-2022-01352 | 14044369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14044369&pc=OZD | 01/08/2022 | 09/22/2024 |  | 04/07/2022 |  |  | 03/08/2022 | 04/07/2022 |  |  | 05/31/2023 |  | 04/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022083720 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2022 |  |  |  |  | The complainant suspected a (b)(6) male patient enduring hemolysis during use of the impella. dialysis was required as intervention. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
36438 | 1220648-2022-01353 | 14044523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14044523&pc=OZD | 05/10/2021 | 04/07/2022 |  | 04/07/2022 |  |  | 03/08/2022 | 04/07/2022 |  |  | 01/31/2023 |  | 04/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022048977 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/08/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient developed an ischemic stroke. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36439 | 1220648-2022-01350 | 14024232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14024232&pc=OZD | 02/23/2022 | 04/06/2022 |  | 04/06/2022 |  |  | 03/07/2022 | 04/06/2022 |  |  | 07/31/2022 |  | 04/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022041421 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/07/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock. the physician inserted the impella device without evaluating the properties of the patient's blood vessel anatomy and support was initiated. after 5 days of support the patient developed a cerebral infarction which resulted in patient death. the relationship between this event and the impella could not be determined. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36440 | 1220648-2022-01351 | 14024392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14024392&pc=OZD | 03/04/2022 | 10/15/2024 |  | 04/06/2022 |  |  | 03/08/2022 | 04/06/2022 |  |  | 12/31/2023 |  | 04/06/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023163709 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2022 |  |  |  |  | The complainant suspected a (b)(6) male patient developing hemolysis during support with an impella cp optical. as treatment, the device was removed and replaced with an intra-aortic balloon pump. | Additionally it was reported the patient experienced thrombocytopenia and access site bleeding which required a blood transfusion. weaning was successful, the device was removed, and the procedural outcome was the patient survived explant. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the cause of the hemolysis was not determined since product was not returned. the cause of the access site bleeding was not determined since product was not returned. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
36441 | 1220648-2022-01347 | 14013751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14013751&pc=OZD | 03/08/2022 | 04/05/2022 |  | 04/05/2022 |  |  | 03/08/2022 | 04/05/2022 |  |  | 12/31/2023 |  | 04/05/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2023163731 |  | 0048-0024-JP |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | User facility | Company representation |  |  | 47 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/08/2022 |  |  |  |  | The medical center in (b)(6) had an impella cp support in which there was an access site bleed. the bleed prompted the team to infuse greater than 2 units of rbcs. this was documented in the (b)(6) registry the month after explant. | [5 empty] | The manufacturer was unable to conduct any investigation to root cause. after explant all impella product was discarded. further details are not forthcoming from the (b)(6)registry. the failure mode will be monitored and trended.
36442 | 1220648-2022-01349 | 14018294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14018294&pc=OZD | 03/29/2022 | 04/05/2022 |  | 04/05/2022 |  |  | 03/29/2022 | 04/05/2022 |  |  | 10/31/2021 |  | 04/05/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1453055 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | Company representation |  |  | 84 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/29/2022 |  |  |  |  | The medical center in (b)(6) had a successful impella cp support in 2020. later in (b)(6) 2022 the (b)(6) registry document review further noted a cva. the cva was stated to be ischemic and occurred during the impella support. no treatment was given. | [5 empty] | The cause of the ischemic cva was not determined since product was not returned. the medical center long discarded the product in 2021. no further clinical details or imaging were shared. the failure mode will be monitored and trended.
36443 | 1220648-2022-01348 | 14018295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=14018295&pc=OZD | 03/08/2022 | 04/05/2022 |  | 04/05/2022 |  |  | 03/08/2022 | 04/05/2022 |  |  | 10/31/2023 |  | 04/05/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022148488 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 83 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/08/2022 |  |  |  |  | The medical center in (b)(6) had an impella cp support along with ecmo in 2021. later in (b)(6) 2022 the (b)(6) registry noted hemolysis, attributed to the use of impella, treated by the infusion of haptoglobin. | [5 empty] | The medical center discarded the impella pump product in 2021. no investigation of it or hemolysis benchtop testing is possible. no root cause can be determined. the failure mode will be monitored and trended.
36444 | 1220648-2022-01346 | 13994749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13994749&pc=OZD | 03/06/2022 | 04/15/2022 |  | 04/04/2022 |  |  | 03/06/2022 | 04/04/2022 |  |  | 11/30/2023 |  | 04/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023167091 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2022 |  |  |  |  | The medical center had an impella 5.5 supporting a patient admitted with multiple cardiac comorbidities. after just 20 hours of support, the impella pump stopped. the pump was explanted and replaced by another 5.5 pump which supported the patient successfully for over 7 days. | [5 empty] | The medical facility did return the impella 5.5 pump for analysis, which is on-going at this time. upon completion a final report will be submitted. | Since the original report was filed the investigation into the pump stop has completed. the pump and data logs were analyzed. the root cause was determined to be ingestion of biomaterial. the failure mode will be monitored and trended.
36445 | 1220648-2022-01345 | 13994750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13994750&pc=OZD | 03/04/2022 | 04/03/2022 |  | 04/04/2022 |  |  | 03/04/2022 | 04/04/2022 |  |  | 05/31/2023 |  | 04/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022089418 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/04/2022 |  |  |  |  | The medical team in (b)(6) had an impella cp support on (b)(6) 2021 for a patient suffering from heart failure. the patient had previously been supported by an intra-aortic balloon pump and had the care escalated. the pump support was successful for just over a day when it was explanted. the team later in 2022 reported in the (b)(6) registry that there was an access site bleed. the bleed was treated by infusion of red blood cells. | [5 empty] | The medical center in (b)(6) discarded all impella product in 2021. no analysis is possible. the failure mode will be monitored and trended.
36446 | 1220648-2022-01344 | 13994754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13994754&pc=OZD | 03/29/2022 | 04/03/2022 |  | 04/04/2022 |  |  | 03/29/2022 | 04/04/2022 |  |  | 09/30/2022 |  | 04/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC |  |  | US |  | 1503580 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/29/2022 |  |  |  |  | The medical center in (b)(6) had a successful cp support in (b)(6) 2020 for a patient with cardiogenic shock, and the pump was explanted after 72 hours. the jpvad registry review occurred in (b)(6) 2022 and noted dialysis was needed for hemolysis. | [5 empty] | The medical center discarded all product last year. no testing or analysis is possible for root cause determination. the failure mode will be monitored and trended.
36447 | 1220648-2022-01343 | 13994756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13994756&pc=OZD | 03/02/2022 | 04/03/2022 |  | 04/04/2022 |  |  | 03/04/2022 | 04/04/2022 |  |  | 09/30/2023 |  | 04/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022135444 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/04/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with post cardiotomy cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient formed an access site bleed that prompted a blood transfusion. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36448 | 1220648-2022-01341 | 13971211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13971211&pc=OZD | 09/04/2021 | 03/31/2022 |  | 03/31/2022 |  |  | 03/04/2022 | 03/31/2022 |  |  | 02/28/2023 |  | 03/31/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022067869 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/04/2022 |  |  |  |  | The complainant reported an (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical. during support, the patient developed an access site bleed that prompted a blood transfusion. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36449 | 1220648-2022-01342 | 13971276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13971276&pc=OZD | 01/15/2022 | 03/31/2022 |  | 03/31/2022 |  |  | 03/04/2022 | 03/31/2022 |  |  | 07/31/2023 |  | 03/31/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022114250 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/04/2022 |  |  |  |  | The complainant reported a (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient developed and infection at the impella access site. intervention in the form of unspecified anti-bacterial agents were required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36450 | 1035166-2022-00044 | 13953031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13953031&pc=DYB | 02/20/2022 | 08/25/2022 |  | 03/30/2022 |  |  | 03/07/2022 | 03/30/2022 |  |  |  |  | 03/30/2022 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | UNKNOWN |  |  | OSCOR INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A (b)(6) male patient was presented with nstemi and hypotensive. the procedure was for acute myocardial infarction/cardiogenic shock. the customer reported they were notified of a complaint involving an oscor 14fr x 25cm introducer (material/model # unknown, and lot # unknown). the physician "cracked" the introducer while it was still in the patient's groin. patient loss 50cc of blood, as bleeding could not be controlled, therefore, the supplied short introducer was used. there was no unplanned surgical intervention or blood products required. it was reported possible issue was caused by poor training, physician forgot the steps. the procedure was completed successfully with a new introducer. the device will not be returned for evaluation as it was discarded by the customer. patient outcome was reported as no harm. no additional information available. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete. | The device used in treatment. the device was not returned for analysis (discarded), therefore, the exact cause of the product issue cannot be determined and the clinical observation could not be confirmed. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. the following controls are in place to mitigate the reported product issue. per abiomed introducer sheath in-process and final inspections procedure: 9.2.1:b measure dimensions: measure tip i.d. of sheath using appropriate pin gauges to be in range of 0.305" - 0.307" for 23f. 9.2.2 visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. per section 12.2, visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. abiomed, inc. provides the instructions for use (ifu) for this product. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36451 | 1220648-2022-01338 | 13953432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13953432&pc=OZD | 12/09/2021 | 09/13/2024 |  | 03/30/2022 |  |  | 03/01/2022 | 03/30/2022 |  |  | 08/31/2023 |  | 03/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022127411 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/01/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient exhibited ischemia in the impella inserted limb. as treatment, lower limb blood feeding was configured. | It was reported while on impella cp support, the patient experienced an access site infection on (b)(6) 2021. it was reported no intervention was performed as a result of the infection. the patient remained on impella cp support and was successfully weaned. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
36452 | 1220648-2022-01340 | 13953582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13953582&pc=OZD | 08/04/2021 | 10/14/2024 |  | 03/30/2022 |  |  | 03/01/2022 | 03/30/2022 |  |  | 04/30/2023 |  | 03/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022073197 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/01/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for hemodynamic support using the impella cp optical device. during support, the patient developed an access site bleed that prompted multiple units of blood product to be delivered. | It was also reported the patient experienced hemolysis after insertion of the device. no additional information was provided and no further measures were taken. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information for the initial mfr 1220648-2022-01340 was provided in sections b5, d6a, d6b, g1, h1, and h6. the investigation for the reported hemolysis has been completed. the impella device was not received from the customer. however, due to insufficient clinical information the root cause of the hemolysis was not determined. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
36453 | 1220648-2022-01339 | 13953591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13953591&pc=OZD | 02/08/2022 | 03/30/2022 |  | 03/30/2022 |  |  | 03/01/2022 | 03/30/2022 |  |  | 12/31/2022 |  | 03/30/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022034969 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/01/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support. the patient developed an access site bleed that prompted a blood transfusion. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36454 | 1035166-2022-00038 | 13941095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13941095&pc=DYB | 03/03/2022 | 08/26/2022 |  | 03/29/2022 |  |  | 03/04/2022 | 03/29/2022 |  |  | 04/01/2023 |  | 03/29/2022 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-19198 |  | 0052-3021 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [7 empty] | Unknown |  |  |  | Injury | Yes | Yes | Break | Crack | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Patient was presented with right ventricular failure. the procedure was for femoral vein insertion for placement of the impella rp (rp). first attempted impella rp insertion from the left femoral vein and the 23 fr sheath kinked. the kink occurred because of the tortuosity of the iliac artery. then went to the right femoral vein with a new sheath. x-ray showed sheath was kinked as well, but were able to pass the rp through it. it is uncertain as to why it split, it was a straight shot and no issues were observed. rather than rewire and attempt re-seat the rp, the physician elected to pull the rp, put large dilator in, and pulled the split sheath without complication. when the sheath split and the rp was not placed, they went to the or for placement of other support devices. it is believed the patient was then sent to surgery for centrimag/protek duo. patient outcome was reported as no harm. no additional information available. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete. | The device was used for treatment. the device was not returned for evaluation, as a result, the exact cause of the product issue cannot be determined and the clinical observation could not be confirmed. however, the following controls are in place to mitigate the reported product issue. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections before shipping to the customer. per abiomed introducer sheath in-process and final inspections procedure: 9.2.1:b measure dimensions: measure tip i.d. of sheath using appropriate pin gauges to be in range of 0.305" - 0.307" for 23f. 9.2.2 visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. per section 12.2, visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. the instructions for use (ifu) informs the user: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. abiomed provides the ifu for this product. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Correction h1: serious injury. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36455 | 1220648-2022-01336 | 13943906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13943906&pc=OZD | 02/28/2022 | 03/26/2024 |  | 03/29/2022 |  |  | 02/28/2022 | 03/29/2022 |  |  | 04/30/2023 |  | 03/29/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022071881 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 79 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | Aortic Valve Insufficiency/ Regurgitation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/28/2022 |  |  |  |  | The medical center placed an impella 5.5 placed in a direct approach via the aorta, and not via a graft, while admitted to the operating suite for bypass (cabg) surgery. the pump placement caused/contributed to the perforation of the anterior aspect of the left ventricle and aortic valve damage. the team surgically repaired the ventricle, placed an edwards bovine valve, and finally placed a new, second impella 5.5 pump for support. | The narrative has been updated with additional information. the complainant reported that during direct insertion of an impella 5.5 pump for patient support, the device perforated the patient¿s aortic valve and left ventricle. the pump was removed due to the complication, surgical repairs were completed, and an edwards bovine valve was placed. a second impella pump was then implanted, but bleeding from the surgical repair remained profuse. continuous blood products had to be given (21 units) and the chest was left open as a result of the condition. the patient passed away the following day. |  |  |  |  | The medical center has not yet returned the impella pump product for any analysis. further clinical details and imaging have been requested, but not yet furnished. upon completion of the investigation a final report will be filed, should a root cause be determined. the failure mode will be monitored and trended. | Since the original report was filed, the investigation has completed. the pump was returned and a cannula kink was observed. the cause of the perforation was improper pump positioning, a user error, since the cannula was kinked. the failure modes will be monitored and trended. | Additional information added to b2, b5, h1, and h6 to coincide with the report of the patient's passing. should additional relevant information become available, a supplemental report will be submitted.
36456 | 1220648-2022-01337 | 13944018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13944018&pc=OZD | 03/03/2022 | 03/29/2022 |  | 03/29/2022 |  |  | 03/03/2022 | 03/29/2022 |  |  | 09/30/2023 |  | 03/29/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022135521 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/03/2022 |  |  |  |  | The medical center placed an impella cp to support a patient admitted in cardiomyopathy. when the team observed signs of hemolysis the team made the choice to explant the cp and place ecmo support instead. the explant was premature. | [5 empty] | The medical center discarded the impella pump product at explant. no benchtop testing is possible. the root cause for the hemolysis has not been determined. should more information be shared that leads to a root cause, a final report will be filed. the failure mode will be monitored and trended.
36457 | 1220648-2022-01334 | 13925400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13925400&pc=OZD | 02/24/2022 | 09/17/2024 |  | 03/27/2022 |  |  | 02/25/2022 | 03/27/2022 |  |  | 10/31/2023 |  | 03/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022144028 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/25/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient developed an access site bleed that prompted multiple units of blood product to be delivered. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information was received on 2022-10-11: the patient developed signs of hemolysis that prompted a blood transfusion. hemolysis subsequently resolved. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01334 was provided in sections b2, b5, h1, and h6.
36458 | 1220648-2022-01335 | 13925401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13925401&pc=OZD | 02/25/2022 | 03/27/2022 |  | 03/27/2022 |  |  | 02/25/2022 | 03/27/2022 |  |  | 10/31/2022 |  | 03/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021011243 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 02/25/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient developed a hematoma that required surgical repair. | [5 empty] | The impella device was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36459 | 1220648-2022-01332 | 13906169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13906169&pc=OZD | 02/25/2022 | 03/25/2022 |  | 03/25/2022 |  |  | 02/25/2022 | 03/25/2022 |  |  | 06/30/2023 |  | 03/25/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED INC. |  |  | US |  | 2022118283 |  | 005040 |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Distributor |  |  | 20 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/25/2022 |  |  |  |  | The medical center in the eu had an impella 2.5 which was supporting a patient admitted with myocarditis, explanted on day 7 due to hemolysis concerns. there was no other intervention noted. upon explant the medical team placed a new impella which supported the patient for another 13 days. | [5 empty] | The medical center in (b)(6) discarded the impella pump product at explant. no benchtop hemolysis testing and analysis is possible. should the team relay more information that leads to a root cause, a final report will be filed. the failure mode will be monitored and trended.
36460 | 1220648-2022-01330 | 13906188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13906188&pc=OZD | 02/24/2022 | 03/25/2022 |  | 03/25/2022 |  |  | 02/24/2022 | 03/25/2022 |  |  | 06/30/2023 |  | 03/25/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022093157 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission |  |  |  |  | 02/24/2022 |  |  |  |  | The complainant reported an (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella lower than expected flows during impella cp optical device. the pump was returned to investigate an allegation of low flow. there was no harm to the patient and no intervention required beyond removing and replacing the device. during our evaluation our team noted damage to the impeller blades. | [5 empty] | The device was returned for evaluation and broken blades were discovered. because the fluoroscopic image of the pump was not returned, a definitive root cause of the reported issue could not be determined.
36461 | 1220648-2022-01331 | 13906200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13906200&pc=OZD | 02/24/2022 | 03/25/2022 |  | 03/25/2022 |  |  | 02/24/2022 | 03/25/2022 |  |  | 10/31/2023 |  | 03/25/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022147641 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/24/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, hemolysis was suspected. as treatment, blood product delivery was required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36462 | 1220648-2022-01333 | 13906655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13906655&pc=PYX | 02/24/2022 | 06/30/2022 |  | 03/25/2022 |  |  | 02/24/2022 | 03/25/2022 |  |  | 10/31/2023 |  | 03/25/2022 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 2022142602 |  | 0046-0011 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2022 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella rp device. during an attempted insertion of the rp device, through the introducer, the patient's vein was perforated. the introducer was left in place and the patient was transferred to a new hospital for escalation of care. vascular surgery was performed, thereafter. | [5 empty] | The device has been received and an investigation is underway. upon completion of our analysis, a supplemental mdr will be filed. | The 23 fr introducer was returned for analysis and the tip of the device was found split at the peel-away score line and tip. no other damage or abnormalities were noted. after review of the information provided on this event, we feel the root cause of the reported vessel damage was related to an interaction between the sheath and the patient's vessels.
36463 | 1220648-2022-01328 | 13872003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13872003&pc=OZD | 07/31/2021 | 03/23/2022 |  | 03/24/2022 |  |  | 02/22/2022 | 03/24/2022 |  |  | 10/31/2022 |  | 03/24/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022029713 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/09/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with myocarditis for hemodynamic support using the impella cp. during support, a vascular injury formed in relationship to the device and unspecified intervention was required. additionally, the patient developed an access site bleed that prompted a blood transfusion. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36464 | 1220648-2022-01329 | 13896539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13896539&pc=OZD | 12/14/2021 | 03/24/2022 |  | 03/24/2022 |  |  | 02/22/2022 | 03/24/2022 |  |  | 10/31/2023 |  | 03/24/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022137424 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/22/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient developed a vascular injury that prompted unspecified intervention. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36465 | 1220648-2022-01325 | 13862380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13862380&pc=OZD | 02/22/2022 | 03/22/2022 |  | 03/23/2022 |  |  | 02/22/2022 | 03/23/2022 |  |  | 07/31/2023 |  | 03/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022114725 |  | 0048-0024-JP |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | User facility | Company representation |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/22/2022 |  |  |  |  | In (b)(6) in 2021 the medical center placed and impella cp for support for a patient admitted with cardiogenic shock, along with ecmo. the support was successful and then in 2022 there was a documentation of renal failure in the (b)(6) registry during the impella support. the team treated the patient with dialysis. | [5 empty] | The medical center in (b)(6) discarded the impella pump product last year. no analysis and testing is possible. should the medical team and/or registry team relay more information that leads to a definitive root cause of the renal injury, a final report will be filed.
36466 | 1220648-2022-01324 | 13862456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13862456&pc=OZD | 02/20/2022 | 05/02/2022 |  | 03/23/2022 |  |  | 02/20/2022 | 03/23/2022 |  |  | 10/31/2023 |  | 03/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2023161362 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 76 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2022 |  |  |  |  | The medical center had an impella 5.5 pump stop after 8 days of support. the 5.5 had successfully supported the patient postoperatively. when the pump stopped the pump was removed and not replaced as the patient was stable. | [5 empty] | The medical center has not yet returned the impella pump or data logs for investigation. the product is expected. when the investigation is completed, a final report will be filed. | Since the initial report was filed, the investigation into the pump stop has concluded. the medical center returned the impella pump product. upon investigation, the pump was found to have stopped due to ingested biomaterial. the failure mode will be monitored and trended.
36467 | 1220648-2022-01326 | 13862477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13862477&pc=OZD | 10/02/2021 | 10/16/2024 |  | 03/23/2022 |  |  | 02/22/2022 | 03/23/2022 |  |  | 05/31/2023 |  | 03/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 2022094914 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Positioning Failure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Thrombocytopenia | [8 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2022 |  |  |  |  | In 2021 a medical team in (b)(6) placed an impella cp for support of a patient suffering with cardiogenic shock. there was escalation to the impella 5.0 for additional flow. later in 2022 the (b)(6) registry noted that there had been an adverse event, of an access site bleed, while on impella support that was treated by unknown amount of infusion of blood products. | Additional information was reported that the patient developed hemolysis and access site bleeding (blood products given), episodes of ventricular tachycardia and thrombocytopenia. the pig tail may have been in contact due to hypocardiac function and positioning was used for troubleshooting. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | In 2021 the medical center in (b)(6) discarded the impella pump and introducer product. no product analysis is possible. should the medical team and/or the registry team reveal more details that can lead to a definitive root cause, a final report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2022-01326 was provided in sections b2, b3, b5, d4, d6, g2 and h6.
36468 | 1220648-2022-01327 | 13862576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13862576&pc=OZD | 03/04/2022 | 03/22/2022 |  | 03/23/2022 |  |  | 03/04/2022 | 03/23/2022 |  |  | 12/31/2022 |  | 03/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2022034973 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | User facility | Company representation |  |  | 45 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/04/2022 |  |  |  |  | The medical team in (b)(6) placed an impella cp for hemodynamic support in (b)(6) 2021. the support was successful. in (b)(6) 2022 the (b)(6) registry noted that the patient had hemolysis during the impella support treated by 4000 units of haptoglobin. | [5 empty] | The root cause of the presumed hemolysis could not be determined. the product was not returned from the medical center in (b)(6). no analysis or hemolysis testing could be done. the failure mode will be monitored and trended.
36469 | 1035166-2022-00039 | 13846572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13846572&pc=DYB | 02/24/2022 | 07/21/2022 |  | 03/22/2022 |  |  | 03/08/2022 | 03/22/2022 |  |  | 07/01/2024 |  | 03/22/2022 |  | INTRODUCER, CATHETER |  |  | ABIOMED 23F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | DP-13404 |  | 0052-3021 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Distributor |  |  |  |  | 63 YR |  | Female |  |  |  | Injury | Yes | Yes | Peeled/Delaminated | [6 empty] | Perforation of Vessels | Heart Failure/Congestive Heart Failure | [11 empty] | Hospitalization | Required Intervention | Death | Life Threatening | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during a procedure for an acute myocardial infarction/cardiogenic shock after insertion of the 23 fr oscor introducer, an opaque structure was seen under imagery. when the introducer was advanced, the structure moved with the device. the procedure was halted and an angiography was taken, revealing a peel in the lock of the introducer - one portion was found perforating the left femoral vein. the patient a (b)(6) female was then transferred for escalation of care, with the introducer still in the patient's body. once patient was transferred, vascular surgery was performed. the patient expired after therapy was discontinued. | Corrected event description submitted to bfarm from the physician: the following problem occurred during the implantation of the sheath required for the impella rp: after a simple puncture of the right femoral vein, first a guide wire and then a 6f sheath was inserted, over which the implantation wire of the impella rp sheath was inserted. sequential dilatation of the vessel via this wire. visual and haptic inspection of the 23f sheath. then insertion of the 23f impella rp sheath. this was done with moderate pressure under fluoroscopic control and wire guided. then a semi radiopaque structure is noticed. the company support, who was present during the implantation, could not comment on this. when trying to insert the sheath further, this structure moves with it. the procedure is stopped immediately and an angiography is performed. the angiography shows that the tip of the sheath is peeled and that one part of the sheath is in the vein and the other part is in the tissue after presumed perforation of the vein wall. the procedure is discontinued and the sheath is left in situ. presentation of the patient at the cardiac center and transfer for further therapy. the patient suffers from bilateral pump failure and should be supported biventricularly for this reason. biventricular support. prior to this, an impella cp left ventricular assist pump. with this pump the patient could be stabilized for transport." information received on (b)(6) 2022: patient was already in very poor condition on arrival at the hospital ((b)(6) 2022). patient had bilateral pump failure. patient probably would have died even with successful surgery with no incident. patient was very unstable on the operating table before surgery started. it is very likely that the patient did not die because of the misplaced/split introducer. it seems unlikely that the impella rp would have really helped the patient in the long term. diagnosis when patient arrived in hospital was so severe that this would have led to complete heart failure in the long term. one year ago patient had a good ventricular function and good ejection of the heart. on admission to the hospital, the patient had an ejection fraction of 10-15%. this indicates severe heart failure. additional information received on (b)(6) 2022: it was reported that the patient was unstable when presented to the physician. the patient was transferred to a second hospital with the introducer left in place and once received she underwent vascular surgery. on (b)(6) 2022, the patient expired after therapy was discontinued. |  |  |  |  | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. | One 23f abiomed introducer sheath was returned from the customer without the dilator. two blue non-vented luer lock caps were attached to the stopcock. there were no other accessories. blood was found on and inside the sheath. upon evaluation of the returned product, it was found that the sheath had two splits in the shaft along the scorelines. one split measured approximately 11cm from the distal tip and the other split measured approximately 7.5cm from the distal tip. both splits were found along the scoreline of the sheath shaft. the distal tip of the sheath was slightly damaged/irregular in shape around the edge most likely due to the splits in the sheath. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. returned device analysis revealed the sheath was split in two places along the scorelines. one split was located approximately 11cm from the distal tip and the other was approximately 7.5cm from the distal tip. there was also slight damage on the distal tip. the sheaths are inspected by qa for cracks and damage and are also leak tested in-process 100% per procedure. the break force data was reviewed and there were no rejects noted on the record. no manufacturing rejects or anomalies of this type were recorded in the dhr. the introducer sheath passed all in-process and qa final inspections before shipping to the customer, including visual, dimensional and leak testing. the root cause of the split occurrence could not be determined. no manufacturing defects were found. per procedure (adelante s2s introducer sheath in-process and final inspection) final assembly inspect 100% for steps 12.2 and 12.3 visual inspection: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. foreign matter (fm)/fiber is limited to 2 particles < 0.10 mm2 or 1 particle < 0.2 mm2 with naked eye at a distance of 12" to 18", verify the sheath hub and split caps are free of excess silicone oil applied between the seals during assembly. verify split caps are properly placed and secured onto sheath hub and free of cracks/damages or excessive adhesive. vacuum leak test perform pressure and vacuum leak test per procedure. the ifu was provided to the customer as part of the product. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36470 | 1220648-2022-01320 | 13806103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13806103&pc=OZD | 02/19/2022 | 03/17/2022 |  | 03/18/2022 |  |  | 02/19/2022 | 03/18/2022 |  |  | 06/30/2023 |  | 03/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022106267 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/19/2022 |  |  |  |  | The complainant reported a patient enduring a stroke (m2 occlusion) during use of the impella device. the medical team alleged the device was a causal or contributory factor in the patient's injury. intervention was required to repair the occlusion. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36471 | 1220648-2022-01321 | 13806149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13806149&pc=OZD | 01/19/2022 | 10/13/2023 |  | 03/18/2022 |  |  | 01/19/2022 | 03/18/2022 |  |  | 09/30/2023 |  | 03/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022147154 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup |  |  |  | 01/19/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for high risk percutaneous coronary intervention. the pump's side arm developed a leak, and after approximately 5 days of support, the pump stopped and required removal/replacement with a new impella. a user submitted medwatch was received 2022-02-28 which also alleged the impella stoppage may have been a factor in the patient's outcome (death), as the patient ultimately suffered an embolitic stroke and did not recover. | [5 empty] | The device was returned and an investigation was completed. the purge lumen was inspected and found damaged, with blood ingress noted. we believe the root cause of the reported pump stop was related to a damaged purge lumen which allowed for blood ingress, thus, leading to the eventual pump stop. due to lack of clinical information surrounding the embolitic stroke and death, we are unable to determine a definitive relationship to the impella or root cause. | The investigation into the reported "purge leak - pump, pump stop and stroke/cva" has been completed since the original report was filed. product was returned for investigation. the device was visually inspected and found that the purge sidearm was leaking and the pump ultimately stopped. during the visual inspection of the returned pump revealed dried blood within the outflow. inspection of the purge sidearm revealed blood ingress up to the distal end of the purge line kink protector where the blood appears to have clotted. the kink protector was removed from around the purge lumen and damage to the lumen was confirmed in this location. the root cause of the pump stop was determined to be blood ingress due to purge lumen damage within the purge line kink protector. furthermore, the lack of provided event information regarding the reported embolic stroke indicates an unknown relation to impella. section d4: unique identifier (udi).
36472 | 1220648-2022-01322 | 13806434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13806434&pc=OZD | 09/03/2021 | 03/17/2022 |  | 03/18/2022 |  |  | 02/16/2022 | 03/18/2022 |  |  | 09/30/2022 |  | 03/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1503408 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/16/2022 |  |  |  |  | The complainant suspected a (b)(6) male patient developing hemolysis and renal failure during impella support, as evident by red urine. haptoglobin was delivered as treatment. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36473 | 1220648-2022-01323 | 13806646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13806646&pc=OZD | 02/19/2022 | 05/24/2022 |  | 03/18/2022 |  |  | 02/19/2022 | 03/18/2022 |  |  | 11/30/2023 |  | 03/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2023161192 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/19/2022 |  |  |  |  | The complainant reported a patient of unknown age and gender presenting for hemodynamic support. the impella 5.5 device was selected for support and inserted without any reported issues. after approximately 48 hours of use, the device's purge pressure suddenly increased and the pump stopped. the device was removed after unsuccessful attempts to restart. | [5 empty] | The device was received and an investigation is underway. a supplemental mdr will be filed upon completion of our analysis. | The device was received and an investigation was conducted. simulated use testing successfully reproduced the reported pump stop. blood was found within the purge lumen and sidearm, thus, we believe the root cause was related to ingested biomaterial.
36474 | 1220648-2022-01319 | 13795994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13795994&pc=OZD | 11/24/2021 | 10/06/2025 |  | 03/17/2022 |  |  | 02/14/2022 | 03/17/2022 |  |  | 08/31/2023 |  | 03/17/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022123267 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/14/2022 |  |  |  |  | The complainant reported an (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient developed ischemia in the impella inserted limb. after the pump was removed, a thrombectomy was required to restore blood flow. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36475 | 13782158 | 13782158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13782158&pc=PBL | 01/22/2022 | 03/01/2022 | 03/01/2022 | 03/16/2022 | 03/01/2022 | 03/16/2022 |  | 03/16/2022 |  |  |  |  | 03/16/2022 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No answer provided |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 24090 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Power Up | [6 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y |  |  |  | Abiomed impella controller would not turn on.
36476 | 1220648-2022-01318 | 13765538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13765538&pc=OZD | 02/14/2022 | 03/14/2022 |  | 03/15/2022 |  |  | 02/14/2022 | 03/15/2022 |  |  | 08/31/2023 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022127418 |  | 0048-0024-JP |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | [12 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [44 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/14/2022 |  |  |  |  | The medical center in (b)(6) had an impella cp support in (b)(6) 2021. in (b)(6) 2022 the (b)(6) registry stated an access site bleed had occurred which prompted intervention. the vessel was sutured and blood transfusion was required. | [5 empty] | The medical center in (b)(6) discarded all impella product in (b)(6) 2021. an investigation and analysis of the product is not possible. should the team share more clinical details that lead to a definitive root cause, a final report will be filed. the failure mode will be monitored and trended.
36477 | 1220648-2022-01317 | 13727088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13727088&pc=OZD | 02/11/2022 | 10/25/2023 |  | 03/10/2022 |  |  | 02/08/2022 | 03/10/2022 |  |  | 04/30/2023 |  | 03/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022079864 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/08/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with cardiomyopathy. after approximately 60 days of support, the patient endured an ischemic stroke. | In addition to our original filing for stroke, the pump had stopped during support on day 63 of support. attempts to restart the pump failed, so the device was removed/replaced with an iabp. |  |  |  |  | The impella device was received and an investigation is underway. a supplemental mdr will be filed upon completion of our analysis. | The pump was returned and an investigation was completed. the cause of the ischemic stroke was determined to have an unknown relation to impella as there was no evidence nor or clinical imaging was provided. | The pump was returned and an investigation was completed. the product had been cut but the user, however, no other visual abnormalities were observed. data logs confirmed the pump had stopped, correlating with the reported pump stop date. because the pump was deformed by the user, we could not simulate testing to reproduce the pump stop. we believe the root cause was attributed to an open line due to duration of use well beyond the designed use life of the product.
36478 | MW5108022 | 13733733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13733733&pc=OZD | 03/02/2022 | 03/08/2022 |  | 03/10/2022 |  |  |  | 03/10/2022 |  |  |  |  | 03/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | CATH CP IMPELLA 14F |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  |  |  | 2023169057 |  |  |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Nurse | [6 empty] | 77 YR |  | Female |  |  |  | Death | Yes | Yes | Unexpected Shutdown | [6 empty] | Cardiac Arrest | [12 empty] | Life Threatening | Death | [27 empty] | Initial submission | [9 empty] | The patient coded and expired after impella device suddenly stopped and alarmed showing "impella stopped retrograded flow".
36479 | 1220648-2022-01312 | 13712995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13712995&pc=OZD | 02/08/2022 | 03/17/2022 |  | 03/09/2022 |  |  | 02/08/2022 | 03/09/2022 |  |  | 09/30/2023 |  | 03/09/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022140697 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2022 |  |  |  |  | The medical center had an impella 5.5 supporting a patient with many cardiac comorbidities while being worked up for transplant. the 5.5 stopped after 11 days of support. there was no lasting harm or injury from the pump stop as the pump was replaced. | [5 empty] | The investigation into the impella 5.5 stopping is on-going at this time. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation has concluded. the pump was not returned for analysis. the clinical details were reviewed. the root cause of the pump stop was most likely blood ingress due to sidearm damage. the failure mode will be monitored and trended.
36480 | 1220648-2022-01313 | 13713025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13713025&pc=OZD | 01/18/2022 | 11/29/2024 |  | 03/09/2022 |  |  | 02/08/2022 | 03/09/2022 |  |  | 12/31/2023 |  | 03/09/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022157611 |  | 005066-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 32 YR |  | Male |  |  | Asian | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Life Threatening | Death | Other | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2022 |  |  |  |  | The complainant reported a (b)(6) male patient was put on hemodynamic support using the impella 5.0. during support, the patient developed an ischemic stroke. no treatment was provided. | The user facility reported a patient was implanted with an impella 5.0 for mechanical circulatory support (mcs) and experience low pump flow, cerebral infarct and expired. approximately four days after start of impella the purge flow rate increased from 15 ml / h to 20 ml / h, and it was said that pp300-400 units had passed. it was not noted in the management so far that the purge flow rate fluctuated so much, but there was an inquiry as to what was the cause. it was noted, the value of the purge flow rate is in the range where there is no problem, but usually it does not fluctuate so much, therefore the purge set was replaced. following three days later, now day seven of support, the patient developed a cerebral infarction, and it was noted to be unclear whether there was a link between the cerebral infraction and impella. the patient was also on extracorporeal membrane oxygenation (ecmo). the physician comments around this event were noted as the following: when computed tomography (ct) scan was taken, there was sporadic infarction rather than ct scan findings such as post-resuscitation encephalopathy. it was further noted there was a strong possibility of an ischemic infarction. since the patient's cardiac function was not above an ejection fraction of 10% before the insertion of impella 5.0, and the heart was not moving, there was thrombus formation in the left chamber, and the possibility that the impella pump ejected it is not zero. further, since it was a cardiogenic shock, and there was abnormal coagulation, various factors can be considered. three days later, now 10 days after start of support the patient expired. it was noted the patient was supported extracorporeal membrane oxygenation (ecmo), but it was stopped due to blood clot in ecmo. patient died without maintaining hemodynamics. the impella 5.0 device performed with a high purge pressure and a purge flow rate of 2 or less. the cause of death was noted to be multiple organ failure and with a ¿not relevant¿ relationship to the impella device. the physician further commented due to cardiogenic shock, the situation was severe after the resuscitation, and it was difficult to save the patient's life; no relationship to the impella device. with the information at hand, including the physician's comment of unrelatedness, the impella 5.0 cannot be excluded as a possible contributing factor. there are unclear timeline details and some conflicting information as it was noted it was unclear whether there was a link between cerebral infarction and impella device. |  |  |  |  | The impella device was discarded by the customer and therefore, a physical examination of the device was not possible. as such, we are unable to definitively determine a root cause for the reported complaint. | The product and its data logs were not returned for analysis. investigation concluded the following: the site reported that the patient suffered an ischemic cerebral infarction and thrombus formation was reported in the left chamber. furthermore, the patient ultimately expired from multiple organ failure after ecmo (extracorporeal membrane oxygenation) support was disrupted due to a blood clot. therefore, the root cause of the reported cerebral infarction and patient death were determined to be related to patient condition and not related to impella. the reported fluctuating purge flows and values were noted to be within normal ranges. however, no product or data logs were provided. this report is being filed as part of a retrospective review of historical records. additional review of the information was completed, and correction/update was made to the respective populated fields in this report.
36481 | 1220648-2022-01311 | 13713032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13713032&pc=OZD | 03/02/2022 | 03/09/2022 |  | 03/09/2022 |  |  | 03/02/2022 | 03/09/2022 |  |  | 08/31/2023 |  | 03/09/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022135044 |  | 0550-0008 |  | Yes |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | Yes | Initial submission |  |  |  |  | 03/02/2022 |  |  |  |  | The medical center had an impella 5.5 support for a patient with cardiomyopathy. the medical team documented an observation of an impella access site hematoma, that prompted intervention, in a user filed a medwatch which was received in 3/2022. | [5 empty] | The user medwatch prompted investigation of the access site hematoma. the medical center did not return the impella pump or introducer product for analysis. without product or further clinical details, a root cause could not be determined. the failure mode will be monitored and trended.
36482 | 1220648-2022-01314 | 13714783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13714783&pc=OZD | 02/07/2022 | 03/09/2022 |  | 03/09/2022 |  |  | 02/10/2022 | 03/09/2022 |  |  | 11/30/2023 |  | 03/09/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022151522 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/10/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with myocarditis. the patient was found to have endured an ischemic stroke after explant of the device, suspected as occurring during support. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36483 | 1220648-2022-01316 | 13715031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13715031&pc=OZD | 02/11/2022 | 03/09/2022 |  | 03/09/2022 |  |  | 02/11/2022 | 03/09/2022 |  |  | 03/31/2023 |  | 03/09/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022063782 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | No | Difficult to Remove | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/11/2022 |  |  |  |  | The medical center had an impella 5.5 supporting a surgical patient for over 8 days when the team attempted to explant it via the 10mm graft. the pump had supported the patient during and after the coronary artery bypass surgery. when the team was unable to explant by normal practice, they surgically opened the chest for removal. during the surgical intervention they observed a tear at the aorta. a repair and blood product infusion was done. | [5 empty] | The investigation is on-going at this time. upon completion a final report will be filed.
36484 | 1220648-2022-01315 | 13715058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13715058&pc=OZD | 02/10/2022 | 03/09/2022 |  | 03/09/2022 |  |  | 02/10/2022 | 03/09/2022 |  |  | 04/30/2023 |  | 03/09/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022080984 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  |  |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/10/2022 |  |  |  |  | The medical center in japan had an impella cp supporting a patient along with ecmo when the team observed an lv perforation. the perforation may have been from guidewire or cp. the patient was treated by surgical repair and the impella cp was withdrawn. the patient survived the event. | [5 empty] | The medical center in japan has not returned any further clinical details, imaging, or product. the investigation is open at this time. should details be forthcoming, and a root cause be determined, a final report will be filed.
36485 | 1220648-2022-01310 | 13689431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13689431&pc=OZD | 08/02/2021 | 11/29/2024 |  | 03/08/2022 |  |  | 02/03/2022 | 03/08/2022 |  |  | 05/31/2023 |  | 03/08/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022092429 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 59 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Fluid/Blood Leak | Decrease in Pressure | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/03/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with cardiomyopathy. the impella 5.0 was selected for hemodynamic support. during support, the patient developed an access site infection that prompted antibiotic treatment. an access site bleed also formed requiring a blood transfusion. | Additional information, the pump exhibited low purge pressure and a purge leak that prompted device replacement. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation for the reported purge leak, low purge pressure or purge system leak, infection and access site bleeding issues has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2022-01310 was provided in sections b1, b5, d6, g2, h1, and h6.
36486 | 1220648-2022-01303 | 13675821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13675821&pc=OZD | 02/06/2022 | 09/18/2024 |  | 03/04/2022 |  |  | 02/06/2022 | 03/04/2022 |  |  | 08/31/2022 |  | 03/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022137458 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 43 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/06/2022 |  |  |  |  | The medical center had an impella cp support a patient with underlying cancer diagnosis admitted with angina. the patient had an impella access site bleed treated by blood product replacement and placement of a femstop. the team believed the patient to be suffering from a retroperitoneal bleed. | Procedural outcome noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The medical center discarded the impella pump product. further clinical details have been requested. upon completion of the investigation, a final report will be filed. | The investigation for the bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the bleeding and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01303 was provided in sections b2, b5, h1, and h6.
36487 | 1220648-2022-01306 | 13675847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13675847&pc=OZD | 01/19/2022 | 03/04/2022 |  | 03/04/2022 |  |  | 02/02/2022 | 03/04/2022 |  |  |  |  | 03/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022147705 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/14/2022 |  |  |  |  | The medical center had an impella cp optical placed via femoral access to give hemodynamic support for a patient undergoing pci. the team observed an access site bleed which prompted 4 units of blood product infused to the patient. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36488 | 1220648-2022-01305 | 13675881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13675881&pc=OZD | 02/02/2022 | 03/04/2022 |  | 03/04/2022 |  |  | 02/02/2022 | 03/04/2022 |  |  | 09/30/2023 |  | 03/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022128464 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 93 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/02/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for high risk percutaneous coronary intervention using the impella cp optical device for hemodynamic support. during support, the patient developed an access site bleed that prompted multiple units of blood product delivered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36489 | 1220648-2022-01307 | 13675933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13675933&pc=OZD | 02/02/2022 | 03/04/2022 |  | 03/04/2022 |  |  | 02/04/2022 | 03/04/2022 |  |  | 10/31/2023 |  | 03/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2022145463 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemothorax | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/04/2022 |  |  |  |  | The complainant reported an (b)(6) year old male patient presenting for high risk percutaneous coronary intervention using the impella cp optical device for hemodynamic support. resistance was reported during pump advancement and patient reporting feeling pain. during the procedure, the patient's hemoglobin dropped and hemodynamics became unstable. the patient expired due to hemorrhagic cardiogenic shock, thereafter. the physician was unsure whether the rotablation procedure or impella had caused the event. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36490 | 1220648-2022-01304 | 13676057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13676057&pc=OZD | 02/02/2022 | 03/04/2022 |  | 03/04/2022 |  |  | 02/02/2022 | 03/04/2022 |  |  | 10/31/2023 |  | 03/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023158891 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/02/2022 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with an acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, an access site bleed developed around the repositioning sheath that prompted 2 units of blood product delivered. | [5 empty] | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36491 | 1220648-2022-01308 | 13676138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13676138&pc=OZD | 02/08/2022 | 03/22/2022 |  | 03/04/2022 |  |  | 02/08/2022 | 03/04/2022 |  |  | 04/30/2023 |  | 03/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2022072566 |  | 0048-0034-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2022 |  |  |  |  | The medical center in (b)(6) had an impella cp supporting a patient hemodynamically. the patient had signs of hemolysis which prompted the team to explant after just over 2 days of support. | [5 empty] | The medical center in (b)(6) returned the impella pump product for analysis. the investigation and hemolysis testing of the pump is on-going at this time. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation into the hemolysis has concluded. the data logs and pump product were analyzed. the cause was determined to be the ingestion of biomaterial. the failure mode will be monitored and trended.
36492 | 1220648-2022-01309 | 13676176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13676176&pc=OZD | 02/04/2022 | 10/24/2023 |  | 03/04/2022 |  |  | 02/04/2022 | 03/04/2022 |  |  | 08/31/2023 |  | 03/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022135019 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/04/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with cardiomyopathy for hemodynamic support using the impella 5.5 device. during support, an access site bleed developed. intervention via wound-vac was performed. | It was also reported that during support, the 5.5 was accidently pulled into the aorta. the pump was removed as a result. there was no patient harm reported. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation for the positioning issues has been completed. the clinical report provided sufficient information to determine the root cause. therefore, the root cause of the positioning issues was user related as the pump was accidently pulled back across the valve.
36493 | 1220648-2022-01302 | 13662299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13662299&pc=OZD | 02/08/2022 | 03/03/2022 |  | 03/03/2022 |  |  | 02/08/2022 | 03/03/2022 |  |  | 10/31/2023 |  | 03/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022155598 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 77 YR |  | Male |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/08/2022 |  |  |  |  | The medical center placed an impella cp to provide hemodyamic support for a patient admitted with an acute myocardial infarction. after under 1/2 of an hour of support the pump was replaced due to low flows. the team placed a new second pump which successfully supported the patient. upon investigation of the first pump the impeller blade was observed to be broken. | [5 empty] | The investigation into the impella pump with low flows and broken impeller blade has concluded. because the fluoroscopic image of the pump was not returned, a definitive root cause of the fracture could not be determined. the failure mode will be monitored and trended.
36494 | 13653184 | 13653184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13653184&pc=PBL | 02/03/2022 | 02/17/2022 | 02/17/2022 | 03/02/2022 | 02/17/2022 | 03/02/2022 |  | 03/02/2022 |  |  |  |  | 03/02/2022 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP FILTER |  |  | PBL |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  | 343859 |  |  |  | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 28835 DA |  | Female |  |  |  | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella cp filter leaking. this is repaired by creating a "bypass" as advised by the impella representative. this filter is part of the pump and can not be removed or replaced without going to the operating room. the bypass is created externally to the patient and the patient did not require pump replacement in the operating room. no harm to patient.
36495 | 1220648-2022-01298 | 13653247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13653247&pc=OZD | 01/31/2022 | 03/02/2022 |  | 03/02/2022 |  |  | 01/31/2022 | 03/02/2022 |  |  | 10/31/2023 |  | 03/02/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022139863 |  | 0048-0024-JP |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/31/2022 |  |  |  |  | In (b)(6) the medical center placed an impella cp for support after the patient had undergone cpr and had ecmo placed. the cp supported the patient for 8 days when it was explanted. later the medical center informed abiomed of the patient having hemolysis presumed to be due to the use of the impella cp. the team had treated the hemolysis with infusion of multiple units of haptoglobin. | [5 empty] | The root cause of the presumed hemolysis could not be confirmed. the hemolysis was originally thought to be the cause of the observed hematuria. no product was returned for analysis. the data logs do no suggest any issues or malfunctions at the time that could cause or contribute to the presumed hemolysis. the relationship between hemolysis and impella use could not be determined. the failure mode will be monitored and trended.
36496 | 1220648-2022-01300 | 13653290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13653290&pc=OZD | 02/01/2022 | 03/02/2022 |  | 03/02/2022 |  |  | 02/01/2022 | 03/02/2022 |  |  | 12/31/2022 |  | 03/02/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2022055448 |  | 005048-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/01/2022 |  |  |  |  | The medical team in (b)(6) had an impella 2.5 supporting a patient admitted in cardiogenic shock, in combination with ecmo . the medical team observed signs of hemolysis and ischemic lower limbs bilaterally. the team placed an antegrade sheath for limb perfusion to the impella limb. the team also infused haptoglobin for the hemolysis. after over 4 days the impella pump was explanted after the ecmo had also been explanted. | [5 empty] | The investigation into the presumed hemolysis and limb ischemia is on-going at this time. upon completion of the investigation, a final report will be filed.
36497 | 1220648-2022-01299 | 13653292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13653292&pc=OZD | 02/02/2022 | 03/22/2022 |  | 03/02/2022 |  |  | 02/02/2022 | 03/02/2022 |  |  | 06/30/2023 |  | 03/02/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED INC. |  |  | US |  | 2022129980 |  | 005040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/02/2022 |  |  |  |  | The medical team in the eu had a failure of the delivery of the impella 2.5 pump. the team had to use surgical methods to remove it from the patient's native vasculature. the pump had become kinked. when the 2.5 pump was successfully removed, they placed a new impella and hemodynamic support successfully began. | [5 empty] | The investigation and product review is on-going at this time. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation has closed. the medical center in the eu returned the impella pump for analysis. the pump kink was observed as the medical team had reported. additionally they had reported the wire was difficult to remove at the explant. the root cause of the need to snare the pump at explant was most likely a guidewire kinking. the guidewire was not returned to investigate why the wire kinked during the impella case.
36498 | 1220648-2022-01301 | 13653545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13653545&pc=OZD | 02/01/2022 | 03/09/2022 |  | 03/02/2022 |  |  | 02/01/2022 | 03/02/2022 |  |  | 10/31/2023 |  | 03/02/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022156704 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 28 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2022 |  |  |  |  | The medical center placed an impella 5.5 to support, via the axillary graft, a younger postpartum patient suffering from cardiomyopathy. the team had the 5.5 support for over 13 days and then had difficulty at the explant. the patient had to undergo surgical procedure and placement of 2 viabahn stents to remove the pump. the team noted the arteriotomy had clamped down on the 5.5 motor. | [5 empty] | The investigation into the failure to explant per usual practice is on-going. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation has concluded. the impella pump was returned with biomaterial adhered to it. the clinical report cited the patient's arteriotomy had clamped down around the pump. the root cause of the difficulty with explant was patient condition. the failure mode will be monitored and trended.
36499 | 1220648-2022-01295 | 13633043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13633043&pc=OZD | 01/31/2022 | 03/01/2022 |  | 03/01/2022 |  |  | 01/31/2022 | 03/01/2022 |  |  | 10/31/2023 |  | 03/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022143709 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 50 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/01/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for impella support. during support, an access site bleed formed. a blood transfusion was required. | The complainant reported a (b)(6) male patient presenting for impella support. during support, an access site bleed formed. a blood transfusion was required. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36500 | 1220648-2022-01296 | 13633757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13633757&pc=OZD | 02/01/2022 | 03/01/2022 |  | 03/01/2022 |  |  | 02/01/2022 | 03/01/2022 |  |  | 06/30/2023 |  | 03/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022100442 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/01/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for hemodynamic support using the impella cp optical device. during support, the patient developed signs of ischemia. the device was removed and a thrombus was noted in the access site. an endovascular thrombectomy (evt) and application of vessel sutures was required as intervention. a vascular injury was also endured 4 days post device removal that required re-application of vessel sutures. | The complainant reported a (b)(6) male patient presenting for hemodynamic support using the impella cp optical device. during support, the patient developed signs of ischemia. the device was removed and a thrombus was noted in the access site. an endovascular thrombectomy (evt) and application of vessel sutures was required as intervention. a vascular injury was also endured 4 days post device removal that required re-application of vessel sutures. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36501 | 1220648-2022-01297 | 13633922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13633922&pc=OZD | 01/31/2022 | 09/18/2024 |  | 03/01/2022 |  |  | 01/31/2022 | 03/01/2022 |  |  | 10/31/2023 |  | 03/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022151404 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 84 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/31/2022 |  |  |  |  | An elderly patient was admitted to the medical center in cardiogenic shock and had an impella cp placed for hemodynamic support. after placement the patient was transferred to the icu for continued monitoring. upon admission to the icu an echo was performed which revealed an left ventricle perforation. the pigtail of the impella cp was seen to be in the pericardial space. after just under 5 hours of support, care was withdrawn and the patient expired. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the perforation of the left ventricle is on-going at this time. impella pump product was discarded, however video imaging and data logs have been retrieved. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation into the lv perforation has completed. the product was not returned for analysis. the clinical report and data logs were reviewed and lead to the root cause of the perforation being due to improper pump positioning. the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01297 was provided in sections b2, b5, h1, and h6.
36502 | 13590649 | 13590649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13590649&pc=OZD | 01/20/2022 | 02/17/2022 | 02/17/2022 | 02/24/2022 | 02/17/2022 | 02/24/2022 |  | 02/24/2022 |  |  |  |  | 02/24/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | Yes |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 21900 DA |  | Male |  |  |  | Malfunction | No | Yes | Crack | Appropriate Term/Code Not Available | [5 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The patient noted seeing a visible wire on the impella catheter near the red solid portion of device. it does appear that the sheath is cracked. the patient returned to the or to have the impella pump exchanged. the malfunctioning impella pump has been sequestered in pathology.
36503 | 1220648-2022-01292 | 13595053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13595053&pc=OZD | 12/16/2021 | 10/12/2024 |  | 02/24/2022 |  |  | 01/25/2022 | 02/24/2022 |  |  | 07/31/2022 |  | 02/24/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1490580 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/25/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for impella support, pre-operative to coronary artery bypass graft surgery. during support, the patient developed an access site hematoma. intervention was completed in the form of a blood transfusion. | It was reported the patient experienced hemolysis while on support. due to the patient's urine being red, volume was given for treatment. according to reports the treatment was able to resolve the issue. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information for the initial mfr 1220648-2022-01292 was provided in sections b5, d6a, d6b, g1, h1, and h6. the investigation for the reported delivery issues has been completed. no impella device nor data logs were received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the hemolysis was most likely patient condition related, since volume loading resolved the issue. instructions for use for the related event are as follows: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
36504 | 1220648-2022-01293 | 13595222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13595222&pc=OZD | 01/27/2022 | 02/24/2022 |  | 02/24/2022 |  |  | 01/27/2022 | 02/24/2022 |  |  | 06/30/2023 |  | 02/24/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022092864 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Valvular Insufficiency/ Regurgitation | Aortic Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/27/2022 |  |  |  |  | The complainant reported a patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical. during support, the patient developed aortic regurgitation. the patient's bioprosthetic valve was damaged and required replacement. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36505 | 1220648-2022-01294 | 13595249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13595249&pc=OZD | 08/22/2020 | 02/24/2022 |  | 02/24/2022 |  |  | 01/28/2022 | 02/24/2022 |  |  | 01/31/2022 |  | 02/24/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1465559 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/28/2022 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp. during support, the patient developed an access site bleed that prompted a blood transfusion. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36506 | 1220648-2022-01291 | 13586831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13586831&pc=OZD | 01/24/2022 | 02/23/2022 |  | 02/23/2022 |  |  | 01/24/2022 | 02/23/2022 |  |  | 10/31/2023 |  | 02/23/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022157075 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/24/2022 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic using the impella cp optical. during support, a loss of distal perfusion in the impella inserted leg was discovered. in response, the pump was removed. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36507 | 1220648-2022-01288 | 13565395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13565395&pc=OZD | 01/20/2022 | 02/19/2022 |  | 02/19/2022 |  |  | 01/20/2022 | 02/19/2022 |  |  | 07/31/2023 |  | 02/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022108938 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/20/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient endured ischemia in the impella inserted leg. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36508 | 1220648-2022-01290 | 13565396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13565396&pc=OZD | 01/19/2022 | 09/17/2024 |  | 02/19/2022 |  |  | 01/21/2022 | 02/19/2022 |  |  | 10/31/2023 |  | 02/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022139867 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/24/2022 |  |  |  |  | The complainant suspected a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient was suspected as enduring hemolysis via increased ldf lab values. dialysis was required as intervention. | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01290 was provided.
36509 | 1220648-2022-01289 | 13565397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13565397&pc=OZD | 01/20/2022 | 02/19/2022 |  | 02/19/2022 |  |  | 01/20/2022 | 02/19/2022 |  |  | 01/31/2023 |  | 02/19/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022048793 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/20/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, a loss of pedal pulses was discovered in the impella inserted leg. as a result, the pump was removed and pulses ultimately returned. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36510 | 13534117 | 13534117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13534117&pc=OZD | 01/14/2022 | 02/11/2022 | 02/11/2022 | 02/15/2022 | 02/11/2022 | 02/15/2022 |  | 02/15/2022 |  |  |  |  | 02/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 350674 |  | 0550-0008 |  | No |  |  | Not Available | [6 empty] | User Facility report | Yes | Other Health Care Professional | [6 empty] | 27010 DA |  | Male | 84 KG |  |  | Death | Yes | Yes | Fluid/Blood Leak | Protective Measures Problem | [5 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y |  |  |  | Patient is on impella 5.5. rn noticed a leak of purge fluid by the red impella plug and alarming "purge flow decreasing". there was a leak in the tubing at the connector and the cardiology team at the time decided not to change out the tubing knowing there was a pump failure risk. they decided they would change it out 6 days later, but on day 5 it failed anyway. patient received a new impella in the operating room that night of day 5. the old impella had a clot and a substance like a d5 that could harden. instead of sequestering the pump, the pump was immediately given to the rep in the middle of the night so the solution would not harden, and they could identify the problem. in the post-op period, the patient was discovered to have had an embolic stroke. he did not have appreciable recovery from the stroke and after a few days the impella was turned off and a day later the patient passed away. when the stroke was discovered, the team managing the patient was considering it a perioperative complication in the setting of his heparin induced thrombocytopenia. we were not classifying it as a device related stroke/death at that time. but if you take into consideration that the pump replacement was an end result of the initial leak in the purge tubing, then you can say there is a possibility that the device failure may have contributed to the ultimate outcome. it was discussed amongst care team and mutually agreed that the safest thing to do would be to report the incident.
36511 | 1220648-2022-01286 | 13539345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13539345&pc=OZD | 01/16/2022 | 02/15/2022 |  | 02/15/2022 |  |  | 01/16/2022 | 02/15/2022 |  |  | 09/30/2023 |  | 02/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022146011 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/16/2022 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient developed an access site hematoma that prompted 2 units of blood products and a surgical cutdown. | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient developed an access site hematoma that prompted 2 units of blood products and a surgical cutdown. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36512 | 1220648-2022-01287 | 13539379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13539379&pc=OZD | 01/17/2022 | 02/15/2022 |  | 02/15/2022 |  |  | 01/17/2022 | 02/15/2022 |  |  | 01/31/2023 |  | 02/15/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022048767 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/17/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical. during support, the patient experienced hemolysis. as a result, the pump was removed. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36513 | 1220648-2022-01283 | 13522790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13522790&pc=OZD | 01/25/2022 | 02/14/2022 |  | 02/14/2022 |  |  | 01/25/2022 | 02/14/2022 |  |  | 10/31/2023 |  | 02/14/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022138511 |  | 0048-0024-JP |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/25/2022 |  |  |  |  | The medical center in (b)(6) inserted an impella cp via the femoral artery for hemodynamic support of patient admitted in shock. the patient was known to have peripheral arterial disease and claudication but, the impella was inserted via the femoral and when the limb became ischemic the team placed an antegrade sheath for perfusion. | [5 empty] | The medical center in (b)(6) has not returned for analysis any impella product. the team did clarify that the patient had underlying peripheral disease. patient condition contributed to the outcome. should the team in japan furnish more details or product for testing, another report will be filed. the failure mode will be monitored and trended.
36514 | 1220648-2022-01284 | 13523244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13523244&pc=OZD | 01/25/2022 | 10/14/2024 |  | 02/14/2022 |  |  | 02/01/2022 | 02/14/2022 |  |  | 08/31/2023 |  | 02/14/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022122179 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2022 |  |  |  |  | The medical center in (b)(6) had an impella cp hemodynamically supporting a post surgical patient in cardiogenic shock. during the support hemolysis was noted. the connection between the use of impella and the hemolysis was unknown. the hemolysis was not confirmed by labs but, was treated by haptoglobin infusion. the infusion was determined to be in excess of 2 units. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the device exhibited a purge cassette leak and the medical staff replaced the purge cassette. during support, the patient experienced thrombocytopenia, the resolution for this issue is unknown. no further information is available. |  |  |  |  | The medical center did not return the impella pump product for analysis and benchtop hemolysis testing. the root cause of the hemolysis could not be determined. the failure mode will be monitored and trended. | The device was not returned for evaluation; however, the purge cassette was received on 2/10/2022. per the results of the clinical review and investigation: leaking from the cartridge cap was observed during leak reproduction testing. therefore, the cause of the purge leak was a cartridge cap leak. the cause of the hemolysis was not determined since neither the pump, nor the data logs were returned. as noted in the impella ifu: thrombocytopenia: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
36515 | 1220648-2022-01285 | 13523663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13523663&pc=OZD | 01/27/2022 | 03/09/2022 |  | 02/14/2022 |  |  | 01/27/2022 | 02/14/2022 |  |  | 08/31/2023 |  | 02/14/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022118846 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 38 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Valvular Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/27/2022 |  |  |  |  | The medical center in (b)(6) placed an impella cp for hemodynamic support of a patient with myocarditis. the care needed to be escalated to the impella 5.0 after just under 5 days of support. at the time of the care escalation the team observed new mitral valve regurgitation. the team suspected the chordae and/or papillary muscle had been damaged at the pump repositioning. the team had a successful surgery. | [5 empty] | The medical center in (b)(6) discarded the pump product at pump explant. the team was asked for imaging and further clinical details, which to date have not been shared. the root cause was not determined. should more details be shared that lead to root cause, a final report will be filed. | Since the original report was filed, the investigation has closed. the product was not returned for analysis. the clinical report alone was analyzed and the cause was determined to be improper technique. the heart valve damage was due most likely to improper positioning. the failure mode will be monitored and trended.
36516 | 1220648-2022-01279 | 13518938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13518938&pc=OZD | 10/26/2021 | 02/11/2022 |  | 02/11/2022 |  |  | 01/14/2022 | 02/11/2022 |  |  | 07/31/2023 |  | 02/11/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022108714 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/14/2022 |  |  |  |  | The complainant reported an (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. at the conclusion of support, a thrombus was identified on the device's pigtail and the patient's leg was ischemic. intervention via endovascular thrombectomy using a fogarty catheter was required to resolve the issue. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36517 | 1220648-2022-01278 | 13518939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13518939&pc=OZD | 03/30/2021 | 02/11/2022 |  | 02/11/2022 |  |  | 01/14/2022 | 02/11/2022 |  |  | 10/31/2022 |  | 02/11/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021020006 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/14/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with postcardiotomy cardiogenic shock. during support, an access site bleed developed and required delivery of blood products. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36518 | 1220648-2022-01280 | 13518943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13518943&pc=OZD | 01/13/2022 | 02/11/2022 |  | 02/11/2022 |  |  | 01/13/2022 | 02/11/2022 |  |  | 10/31/2023 |  | 02/11/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022149142 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/13/2022 |  |  |  |  | The complainant reported a (b)(6) patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient's impella inserted arm became ischemic. no alternate access points were available, therefore, support was discontinued and the pump was removed. | The complainant reported a (b)(6) patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, the patient's impella inserted arm became ischemic. no alternate access points were available, therefore, support was discontinued and the pump was removed. |  |  |  |  | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was not received from the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36519 | 1220648-2022-01281 | 13518945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13518945&pc=OZD | 01/13/2022 | 03/25/2022 |  | 02/11/2022 |  |  | 01/13/2022 | 02/11/2022 |  |  | 10/31/2023 |  | 02/11/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022154690 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 01/13/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with postcardiotomy cardiogenic shock. the impella 5.5 pump was inserted for support. there was a pump stop observed. | [5 empty] | The impella device was received and an investigation is underway. a supplemental mdr will be filed upon completion of our analysis. | The pump was returned and an evaluation revealed blood ingress found in the purge lumen. the cause of the pump stop was found to be related to blood ingress resulting from high purge pressure, however, we cannot definitively determine the cause of the high purge pressure as no catheter kinks were found, nor any biomaterial buildup in the purge gap.
36520 | 1220648-2022-01282 | 13518947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13518947&pc=OZD | 01/28/2022 | 02/11/2022 |  | 02/11/2022 |  |  | 01/28/2022 | 02/11/2022 |  |  | 09/30/2021 |  | 02/11/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1450243 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/28/2022 |  |  |  |  | The medical center in (b)(6) had an impella cp pump support in (b)(6) 2020. the support was successful for over 10 days in combination with ecmo. later in (b)(6) 2022 the (b)(6) registry noted limb ischemia during the support. an update was shared in which the limb had intervention in the form of external bypass. the limbs were ischemic with both impella and the ecmo placement in the other leg. | [5 empty] | The complainant in (b)(6) did not retain the product after explant for analysis. the (b)(6) registry team did not furnish other clinical details that lead to a definitive root cause, beyond the fact that both limbs were ischemic. the cause of the ischemia could not be determined. the failure mode will be monitored and trended.
36521 | 1220648-2022-01273 | 13508589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13508589&pc=OZD | 07/16/2021 | 10/13/2024 |  | 02/10/2022 |  |  | 01/12/2022 | 02/10/2022 |  |  | 02/28/2023 |  | 02/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022062653 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 77 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Aortic Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with myocarditis. during support, the patient endured aortic regurgitation that prompted premature removal of the device and replacement with extracorporeal membrane oxygenation combined with an intra-aortic balloon pump. | The complainant reported a (b)(6) male patient presenting with myocarditis. during support, the patient endured aortic regurgitation that prompted premature removal of the device and replacement with extracorporeal membrane oxygenation combined with an intra-aortic balloon pump. | During support, the patient experienced hemolysis. no intervention for this condition was specified, however, it was reported that the patient recovered. no further information is available. |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. as such, we are unable to definitively determine a root cause for the reported issue. | The impella device was discarded by the customer and therefore, investigation of device was not possible. as such, we are unable to definitively determine a root cause for the reported issue. | The investigation for the reported hemolysis has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the hemolysis could not be determined. as noted in the impella ifu: hemolysis: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. corrections: b5 updated to reflect additional reportable condition. g1 updated to correct contact.
36522 | 1220648-2022-01274 | 13508675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13508675&pc=OZD | 01/12/2022 | 02/10/2022 |  | 02/10/2022 |  |  | 01/12/2022 | 02/10/2022 |  |  | 03/31/2023 |  | 02/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022060018 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/12/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for impella support preoperative to a cabg procedure. after pump initiation, the patient's blood pressure dropped. imagery found the left ventricle perforated by impella pig tail. the pump was left running during repair, at which time bubbles were seen and the patient suffered a stroke. | The complainant reported a (b)(6) male patient presenting for impella support preoperative to a cabg procedure. after pump initiation, the patient's blood pressure dropped. imagery found the left ventricle perforated by impella pig tail. the pump was left running during repair, at which time bubbles were seen and the patient suffered a stroke. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36523 | 1220648-2022-01276 | 13508688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13508688&pc=OZD | 01/21/2022 | 11/29/2024 |  | 02/10/2022 |  |  | 01/21/2022 | 02/10/2022 |  |  | 10/31/2023 |  | 02/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED INC. |  |  | US |  | 2022145301 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/21/2022 |  |  |  |  | The medical center in the (b)(6) had an impella 5.0 supporting a patient admitted with cardiomyopathy. the patient was also supported by ecmo. the pump stopped on day 10 of the support and was restarted. the team was able to continue to support the patient hemodynamically for 4 more days before the pump was explanted. | It was additionally reported that just prior the pump stops, the pump alarmed for low purge flow and high purge pressure. it was reported troubleshooting was performed to resolve the pump stop and purge pressure high, however details on what troubleshooting was performed was not provided. |  |  |  |  | The investigation into the impella pump stop is on-going at this time. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation has concluded. the engineering team concluded that the pump pause was due to ingested biomaterial. the returned pump was tested and run upon the benchtop and the failure was reproduced. the failure mode will be monitored and trended. | Additional information added to b5 and h6 codes for purge pressure high issues and d6a/d6b for implant/explant dates. this report is being filed as part of a retrospective review of historical records.
36524 | 1220648-2022-01275 | 13508693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13508693&pc=OZD | 01/18/2022 | 02/10/2022 |  | 02/10/2022 |  |  | 01/18/2022 | 02/10/2022 |  |  | 10/31/2023 |  | 02/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022148525 |  | 0048-0024-JP |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/18/2022 |  |  |  |  | A patient was admitted in (b)(6) suffering from an acute myocardial infarction and was taken for angioplasty with hemodynamic support provided by ecmo and an impella cp. the patient was supported with two impella pumps, first a cp and then care was escalated to a 5.0 for over 3 days. the team observed a cerebral infarction at the time of pump explant. | [5 empty] | The investigation into the cerebral infarction is on-going at this time. no product, imaging, or further clinical details have been shared. should the medical center in japan share product and such, and the investigation lead to a root cause, a final report will be filed. the failure mode will be monitored and trended.
36525 | 1220648-2022-01277 | 13508698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13508698&pc=OZD | 01/24/2022 | 03/09/2022 |  | 02/10/2022 |  |  | 01/24/2022 | 02/10/2022 |  |  | 09/30/2023 |  | 02/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022135693 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/24/2022 |  |  |  |  | A patient was admitted with multiple cardiac symptoms and a diagnosis of mitral regurgitation was made. the patient was turned down for surgical repair but, was placed on impella cp support while care plan was made. during his support the patient lower limbs became ischemic. the impella pump was inserted into one of them and that limb was treated by intervention in the form of a distal reperfusion catheter being placed. | [5 empty] | The medical center has not yet returned any impella product for analysis. the root cause of the limb ischemia is not known at this time. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation has concluded. the root cause of the limb ischemia was most likely patient's native condition, as the ischemia was bilateral and the patient had a history of peripheral arterial disease. the failure mode will be monitored and trended.
36526 | 1220648-2022-01271 | 13471109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13471109&pc=OZD | 01/13/2022 | 10/19/2024 |  | 02/07/2022 |  |  | 01/13/2022 | 02/07/2022 |  |  | 09/30/2023 |  | 02/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022135402 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/13/2022 |  |  |  |  | The complainant reported a (b)(6) white female had impella cp pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had access site bleeding and as a result a total of four (4) units of packed red blood cells (prbcs) was administered. | Additional information from the complainant reported there was low pump flow. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | Since the original report was filed, the investigation into the access site bleeding has concluded. the impella pump product was returned for analysis, as were the data logs. the cause of the bleeding and rbc infusion was determined to be use/technique related. the bleed was due to a nonactive retroperitoneal bleed. the ct imaging confirmed the retroperitoneal bleed to not be active. the failure mode will be monitored and trended. | Additional information for the initial mfr 1220648-2022-01271 was provided. this report is being filed as part of a retrospective review of historical records. f6 and f8 dates were provided on manufacturer device report number 1220648-2022-01271 in error.
36527 | 1220648-2022-01272 | 13471167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13471167&pc=OZD | 01/13/2022 | 03/02/2022 |  | 02/07/2022 |  |  | 01/13/2022 | 02/07/2022 |  |  | 12/31/2021 |  | 02/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED INC. |  |  | US |  | 1467435 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  |  |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/13/2022 |  |  |  |  | In (b)(6) of 2021 a medical center in (b)(6) had an impella cp inserted to hemodynamically support a patient. in (b)(6) of 2022 a publication was discovered that was inclusive of the impella support. the publication noted aortic thrombus that was observed after explant of the impella and was surgically removed. at the time of the event, abiomed sales team in (b)(6) was not informed of the event. | [5 empty] | The investigation into the thrombosis and necessary intervention is on-going at this time. upon completion, a final report will be filed. | Since the original report was filed the investigation into the biomaterial found in the impella limb, that was explanted surgically, has concluded. the literature alone was analyzed as no product was returned for analysis, though the field representative was able to retrieve the data logs. the data logs revealed no abnormalities during the support. the root cause of the biomaterial was not determined. the failure mode will be monitored and trended.
36528 | 1220648-2022-01269 | 13451031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13451031&pc=OZD | 01/11/2022 | 02/15/2022 |  | 02/04/2022 |  |  | 01/11/2022 | 02/04/2022 |  |  | 07/31/2023 |  | 02/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022120178 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Unexpected Therapeutic Results | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/11/2022 |  |  |  |  | A patient with a cardiac history was admitted and placed an impella cp for hemodynamic support. the team had the pump supporting the patient for over 14 hours when the team explanted the pump due to concerns of hemolysis. the cp was replaced by an intra aortic balloon pump. | A patient with a cardiac history was admitted and placed an impella cp for hemodynamic support. the team had the pump supporting the patient for over 14 hours when the team explanted the pump due to concerns of hemolysis. the cp was replaced by an intra aortic balloon pump. |  |  |  |  | The investigation into the hemolysis is on-going at this time. upon completion of the investigation, a final report will be filed. | The investigation into the hemolysis is on-going at this time. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation into the hemolysis has concluded. the pump and logs were returned for analysis. the pump was hemolysis tested on the benchtop and passed. when the logs were reviewed, no malpositioning that could cause hemolysis was observed. the root cause of the hemolysis is patient condition.
36529 | 1220648-2022-01270 | 13451709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13451709&pc=OZD | 01/06/2022 | 02/04/2022 |  | 02/04/2022 |  |  | 01/07/2022 | 02/04/2022 |  |  | 03/31/2023 |  | 02/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022059104 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Valvular Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/07/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with cardiomyopathy, using the impella 5.0 for hemodynamic support. the patient endured a ruptured chordae tendineae during repositioning of the device. repair is planned to occur during lvad placement. | The complainant reported a (b)(6) male patient presenting with cardiomyopathy, using the impella 5.0 for hemodynamic support. the patient endured a ruptured chordae tendineae during repositioning of the device. repair is planned to occur during lvad placement. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36530 | 1220648-2022-01267 | 13442002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13442002&pc=OZD | 01/05/2022 | 02/03/2022 |  | 02/03/2022 |  |  | 01/05/2022 | 02/03/2022 |  |  | 09/30/2023 |  | 02/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022141344 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/05/2022 |  |  |  |  | A covid + patient was admitted with cardiogenic shock. the team placed an impella 5.5 pump for hemodynamic support due to poor condition and possible cardiac surgery. the 5.5 pump supported the patient for 8 days when the patient stepped upon the cable and separated the pump/cable, which lead to a pump stop. the pump was replaced. | [5 empty] | The medical center did not return the pump product for investigation. the team did relay that the patient had stepped upon the pump cable, and this caused the pump to stop. should more details be shared regarding the pump stop and root cause, another report will be filed.
36531 | 1220648-2022-01268 | 13442045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13442045&pc=OZD | 04/21/2021 | 11/29/2024 |  | 02/03/2022 |  |  | 01/06/2022 | 02/03/2022 |  |  | 07/31/2022 |  | 02/03/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2022041417 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | Thrombosis/Thrombus | [9 empty] | Required Intervention | Death |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2022 |  |  |  |  | The medical center in (b)(6) had an impella 5.0 placed for support escalated from an impella cp. the support was successful in (b)(6) 2021. in (b)(6) of 2022 the (b)(6) registry noted multiple injuries during the time of impella support. there was access site bleeding and limb ischemia which was noted to be treated and attribute to the use of impella. the bleeding was treated by red blood cell infusion and the limb thrombosis was treated by removal of thrombus. | Remediation review of the information supplied indicated that the patient had an episode of ventricular tachycardia upon returning to the intensive care unit (icu) after the impella 5.0 had been placed. it was reported that the position of the impella was adjusted and confirmed through echocardiogram. the information also noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The medical center in (b)(6) discarded the impella product a year prior. no investigation is possible. the data shared by the (b)(6) registry did not furnish all clinical details requested. should more information be shared that leads to a root cause, a final report will be filed. | This report is being filed as part of a retrospective review of historical records. additional information reviewed as part of remediation indicated the patient had an episode of tachycardia, thrombus and the patient expired. sections b2, b3, b5, h1 and h6 have been updated. g1 reporting contact email updated to current information. d6a, d6b implant and explant dates added. d7a single-use updated. e1 facility phone updated following correction have been made: f6 (date user facility or importer became aware of event) and f8 (date of this report) have been corrected to blank. g4 (pma/510(k)number) incorrectly reported as p1400003 and have been updated to p140003.
36532 | 1220648-2022-01266 | 13430233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13430233&pc=OZD | 01/20/2022 | 10/24/2023 |  | 02/02/2022 |  |  | 01/20/2022 | 02/02/2022 |  |  | 07/31/2023 |  | 02/02/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022113607 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/20/2022 |  |  |  |  | The medical center placed an impella 5.5 via a direct approach during open heart bypass surgery. the pump had placement signal and low flows alarming during the support of just over 12 hours, and so the pump was replaced. a 2nd 5.5 pump was placed and support proceeded without harm or injury. weeks after the event the abiomed field rep was notified of heart valve damage. the damage was observed by the surgical team when placing the heartmate after the 2nd pump was explanted. | In addition to the original report filed for heart valve injury, this complaint has since been deemed reportable for pump flow that was lower than expected during support. |  |  |  |  | No product, imaging, or further clinical details were provided. the root cause of the aortic valve damage could not be determined. the failure mode will be monitored and trended. | The investigation for the reported low flow and heart valve damage has been completed. the pump was returned for evaluation. no abnormalities were found upon product evaluation. the pump was able to run at p9, 5.3 l/min in the lab. prior to adjustment, the ao ps read -30 mmhg. after adjustment, the ao ps read 30 mmhg. the root cause of the low flow was not determined as it could not be reproduced and logs were not returned. the root cause of the valve damage was not determined as no clinical details are known. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
36533 | 1220648-2022-01263 | 13418411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13418411&pc=OZD | 12/16/2021 | 09/15/2024 |  | 02/01/2022 |  |  | 01/12/2022 | 02/01/2022 |  |  | 06/30/2022 |  | 02/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1484369 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 68 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Required Intervention | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2022 |  |  |  |  | The complainant reported a (b)(6) asian male had impella cp pump inserted in the left femoral for post cardiac-cardiogenic shock (pccs). the patient had ischemic cerebrovascular accident (cva). | The complainant reported a (b)(6) asian male had impella cp pump inserted in the left femoral for post cardiac-cardiogenic shock (pccs). the patient had ischemic cerebrovascular accident (cva). | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  | The impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed. | The impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01263 was provided in sections b2, b5, h1, and h6.
36534 | 1220648-2022-01264 | 13418473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13418473&pc=OZD | 01/20/2022 | 02/01/2022 |  | 02/01/2022 |  |  | 01/20/2022 | 02/01/2022 |  |  | 09/30/2022 |  | 02/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2021006027 |  | 0048-0024-JP |  | No |  |  | No |  |  |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | User facility | Company representation |  |  | 89 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/20/2022 |  |  |  |  | The medical center in (b)(6) had an impella cp supporting a patient along with ecmo on (b)(6) 2021 admitted in cardiogenic shock. months later on (b)(6) 2021 the jpvad registry noted hemolysis and renal failure during the support. there was no known treatment for the issues nor definitive causal relationship known and shared by any labs or testing and as such was not reported upon. later on (b)(6) 2022 an update was made of blood product replacement being given for the hemolysis. | [5 empty] | No product or imaging were shared for analysis. hemolysis testing on the benchtop could not be performed. a root cause could not be determined. the failure mode will be monitored and trended.
36535 | 1220648-2022-01265 | 13418516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13418516&pc=OZD | 02/03/2020 | 02/01/2022 |  | 02/01/2022 |  |  | 01/05/2022 | 02/01/2022 |  |  | 08/31/2021 |  | 02/01/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1430528 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/05/2022 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for support using the impella cp device. during support, the patient endured an ischemic stroke. | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for support using the impella cp device. during support, the patient endured an ischemic stroke. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36536 | 1220648-2022-01262 | 13401341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13401341&pc=OZD | 01/03/2022 | 01/31/2022 |  | 01/31/2022 |  |  | 01/03/2022 | 01/31/2022 |  |  | 10/31/2023 |  | 01/31/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022147949 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 01/03/2022 |  |  |  |  | The complainant reported a (b)(6) female presenting for high risk percutaneous coronary intervention using the impella cp optical device for hemodynamic support. during support, the patient lost distal pulses in the impella inserted leg that prompted premature removal of the device. a dissected vessel was discovered and surgical repair was completed. | The complainant reported a (b)(6) female presenting for high risk percutaneous coronary intervention using the (b)(6) cp optical device for hemodynamic support. during support, the patient lost distal pulses in the (b)(6) inserted leg that prompted premature removal of the device. a dissected vessel was discovered and surgical repair was completed. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The (b)(6) device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36537 | 1220648-2022-01261 | 13401401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13401401&pc=OZD | 01/02/2022 | 01/31/2022 |  | 01/31/2022 |  |  | 01/02/2022 | 01/31/2022 |  |  | 08/31/2023 |  | 01/31/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2022131606 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 01/02/2022 |  |  |  |  | The complainant suspected a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for support using the impella 5.0 device. approximately 6-12 hours into support, the patient was exhibiting signs of hemolysis as evident by dark red urine and elevated ldh values. there was no report of pfhg testing results to confirm hemolysis, however. as treatment, the pump was prematurely removed and replaced with a left side tandem heart and begun dialysis. hemolysis improved. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36538 | 1220648-2022-01260 | 13401404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13401404&pc=PYX | 01/02/2022 | 05/02/2022 |  | 01/31/2022 |  |  | 01/02/2022 | 01/31/2022 |  |  | 06/30/2023 |  | 01/31/2022 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022094393 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/02/2022 |  |  |  |  | The complainant suspected a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for support using the impella rp device. approximately 6-12 hours into support, the patient was exhibiting signs of hemolysis as evident by dark red urine and elevated ldh values. there was no report of pfhg testing results to confirm hemolysis, however. as treatment, the pump was prematurely removed and replaced with a left side tandem heart and begun dialysis. hemolysis improved. | The complainant suspected a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for support using the impella rp device. approximately 6-12 hours into support, the patient was exhibiting signs of hemolysis as evident by dark red urine and elevated ldh values. there was no report of pfhg testing results to confirm hemolysis, however. as treatment, the pump was prematurely removed and replaced with a left side tandem heart and begun dialysis. hemolysis improved. |  |  |  |  | The device was returned for analysis and 2 indents were found on the blades of the impeller. the markings were reproduced with a 0.027" guidewire on a sample impella rp device. we feel the root cause of the reported hemolysis was related to damaged impeller blades from guidewire interaction during removal of the wire. | The impella device has been received and an investigation is underway. upon completion, a supplemental mdr will be filed with the results of our analysis. | The impella device has been received and an investigation is underway. upon completion, a supplemental mdr will be filed with the results of our analysis.
36539 | 1220648-2022-01255 | 13374039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13374039&pc=OZD | 02/24/2021 | 01/27/2022 |  | 01/27/2022 |  |  | 01/07/2022 | 01/27/2022 |  |  | 12/31/2022 |  | 01/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022028941 |  | 0048-0034-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 73 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/07/2022 |  |  |  |  | The complainant reported a (b)(6) asian male had impella cp pump inserted in the right femoral. the patient also had extracorporeal membrane oxygenation (ecmo) inserted, had hemolysis and as a result haptoglobin was administered. | [5 empty] | Product evaluation revealed a slight kink in the cannula and no other pump defects. hemolysis was also reported but no evidence is provided showing that it is related to impella, and the relationship to impella could not be confirmed by the physician or testing. thus, the relation is unknown.
36540 | 1220648-2022-01256 | 13374046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13374046&pc=OZD | 12/25/2021 | 09/19/2024 |  | 01/27/2022 |  |  | 01/05/2022 | 01/27/2022 |  |  | 01/31/2023 |  | 01/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022043000 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 77 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2022 |  |  |  |  | The complainant reported a (b)(6) asian male had impella cp pump inserted in the right femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had access site bleeding and as a result ten (10) units of red blood cells (rbcs) and forty (40) units of fresh frozen plasma (ffp) was administered, and the patient expired due to do not attempt resuscitation (dnar). | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01256 was provided in sections b2, b5, d6a, d6b, h1, and h6.
36541 | 1220648-2022-01254 | 13374074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13374074&pc=OZD | 07/07/2020 | 03/14/2022 |  | 01/27/2022 |  |  | 01/05/2022 | 01/27/2022 |  |  | 01/31/2022 |  | 01/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1465638 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 84 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/05/2022 |  |  |  |  | The complainant reported an (b)(6) asian female had impella cp pump inserted in the left femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had ventricular rupture from the vsp area or ventricular septal defect (vsd). | [5 empty] | The impella cp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed. | The device was not received for evaluation, however, further review was conducted on the information provided. the physician could not determine a relationship between the impella device and the cardiac tamponade. furthermore, the patient presented with an existing ventricular septal defect.
36542 | 1220648-2022-01257 | 13374118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13374118&pc=OZD | 01/05/2022 | 01/27/2022 |  | 01/27/2022 |  |  | 01/05/2022 | 01/27/2022 |  |  | 09/30/2022 |  | 01/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1503409 |  | 0048-0034-JP |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | VA ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/05/2022 |  |  |  |  | The medical center in (b)(6) had an impella cp placed via the femoral artery for hemodynamic support. the patient was also supported by ecmo. the team observed an impella access site bleed that prompted the team to infuse blood products to the patient. the bleed was not controlled despite the team attempting to fix the pump position and angle at the access site. at the start of the support the ifu recommendations for the angulation of the pump at insertion was not maintained appropriately. the impella pump remained on for support, despite the bleed, for over 123 hours. the pump was then weaned and explanted. | [5 empty] | The medical center in (b)(6) did not furnish all case details, such as amount of blood products infused, nor did they return the impella pump product for analysis. only the clinical case details were reviewed to determine root cause. the cause was improper technique. the team had not maintained the angulation of the pump at the access site. this allowed for access site bleeding. the failure mode will be monitored and trended.
36543 | 1220648-2022-01258 | 13374134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13374134&pc=OZD | 01/06/2022 | 02/17/2022 |  | 01/27/2022 |  |  | 01/07/2022 | 01/27/2022 |  |  | 08/31/2023 |  | 01/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST PUMP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2022128132 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 20 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/07/2022 |  |  |  |  | The complainant reported a (b)(6) white female had impella cp smart assist pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). the pigtail perforated the ventricle. | The complainant reported a (b)(6) white female had impella cp smart assist pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). the pigtail perforated the ventricle. |  |  |  |  | The impella cp smart assist pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The impella cp smart assist pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | Since the original report was filed, the investigation into the ventricular perforation has concluded. imaging, pump, and data logs were returned from the eu for consideration and analysis. the perforation was determined to be due to improper pump positioning. the failure mode will be monitored and trended.
36544 | 1220648-2022-01259 | 13374178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13374178&pc=OZD | 01/07/2022 | 01/27/2022 |  | 01/27/2022 |  |  | 01/07/2022 | 01/27/2022 |  |  | 10/31/2024 |  | 01/27/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022154156 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/07/2022 |  |  |  |  | The complainant reported a (b)(6) white male had impella cp pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient was bleeding from the groin and hematoma was growing, the pump was removed, and four (4) units of packed red bloods cells (prbcs) was administered. | [5 empty] | The impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36545 | 1220648-2022-01252 | 13343103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13343103&pc=OZD | 12/27/2021 | 01/25/2022 |  | 01/25/2022 |  |  | 12/27/2021 | 01/25/2022 |  |  | 09/30/2023 |  | 01/25/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022132409 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/27/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for hemodynamic support using the impella cp optical. during insertion, the physician alleged an artery was damaged that resulted in a hematoma of the access site. blood products were delivered as intervention. | The complainant reported a (b)(6) male patient presenting for hemodynamic support using the impella cp optical. during insertion, the physician alleged an artery was damaged that resulted in a hematoma of the access site. blood products were delivered as intervention. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36546 | 1220648-2022-01253 | 13343111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13343111&pc=OZD | 12/27/2021 | 12/13/2024 |  | 01/25/2022 |  |  | 12/27/2021 | 01/25/2022 |  |  | 09/30/2021 |  | 01/25/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1444438 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  | Asian | Injury | Yes | No | Positioning Failure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2021 |  |  |  |  | The complainant reported an (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for support using the impella 5.0 device. during support, the patient developed an access site hematoma. as treatment, a blood transfusion was required. | The clinical team reported that hemolysis, thrombocytopenia, and positioning issues occurred during impella support. further information about any of these adverse events was not provided due to hospital restrictions. it was later reported in a publication that an intractable hematoma developed at the impella access site. the bleeding was eventually treated with blood transfusions and the placement of a heart band compression bandage. no further information was provided. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation into the reported issues has been completed. no pump was received for evaluation. the causes of the reported issues could not be determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records. corrected information provided in b5, h6 and h.10.
36547 | 1220648-2022-01251 | 13324686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13324686&pc=OZD | 12/24/2021 | 01/22/2022 |  | 01/22/2022 |  |  | 12/24/2021 | 01/22/2022 |  |  | 09/30/2023 |  | 01/22/2022 |  | IMPELLA 5.5 |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022144189 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/24/2021 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with post cardiotomy cardiogenic shock. during an attempted delivery of the device, the physician used "forceful pushing" and tore the axillary artery. the pump was removed, pledgeted stitches were applied, and a new impella 5.5 was placed via direct approach. | The complainant reported a (b)(6) female patient presenting with post cardiotomy cardiogenic shock. during an attempted delivery of the device, the physician used "forceful pushing" and tore the axillary artery. the pump was removed, pledgeted stitches were applied, and a new impella 5.5 was placed via direct approach. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36548 | 13314020 | 13314020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13314020&pc=OZD | 01/13/2022 | 01/18/2022 | 01/18/2022 | 01/21/2022 | 01/18/2022 | 01/21/2022 |  | 01/21/2022 |  |  |  |  | 01/21/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0042-0000-US |  |  | ABIOMED, INC |  |  | US |  |  |  | 40905 |  | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 27010 DA |  | Male | 79 KG |  |  | Malfunction | No | Yes | Pumping Stopped | Protective Measures Problem | [5 empty] | Cardiac Arrest | Low Blood Pressure/ Hypotension | Asystole | Respiratory Arrest | [34 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y |  |  |  | Staff checked impella device at 1100. noted no alarms going off. flows, co2, purge pressure and motor current numbers essentially unchanged from previous hourly checks. at around 1130, impella device alarming, so staff went to investigate. noted impella pump to be stopped. p level was p0. no motor waveform present. device alarming "restart impella." patient quickly became hypotensive. epinephrine given x 2. cardiologist at bedside. impella representative notified who advised to pull back impella catheter. spoke with impella representative, who said he was driving to hospital and would proceed to cath lab when he arrived. catheter adjusted by cardiologist. vasopressors adjusted. 3rd vasopressor added. patient taken to cath lab, expired at 1214.
36549 | 1220648-2022-01247 | 13305873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13305873&pc=OZD | 01/10/2022 | 09/21/2024 |  | 01/20/2022 |  |  | 01/10/2022 | 01/20/2022 |  |  | 08/31/2023 |  | 01/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022124524 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Fracture | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2022 |  |  |  |  | A patient was admitted suffering from an acute myocardial infarction and was placed on an impella cp for hemodynamic support. the impella cp stopped and restarted during support. the pump restarted with suction. the team made choice with the family not to escalate care or transfer the patient despite this level of pump performance/suction. the patient expired on support. in the pump analysis, the engineer discovered damage to the impeller blade. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The pump product was returned and analysis completed. the impeller blade damage root cause could not be determined. because fluoroscopic imaging of the pump/additional evidence was not returned, a definitive cause of the fracture could not be definitively determined. the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6.
36550 | 1220648-2022-01248 | 13305874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13305874&pc=OZD | 12/21/2021 | 01/20/2022 |  | 01/20/2022 |  |  | 12/21/2021 | 01/20/2022 |  |  | 04/30/2022 |  | 01/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1488338 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/21/2021 |  |  |  |  | The complainant reported a patient presenting with acute myocardial infarction and cardiogenic shock for impella support. the patient ultimately sustained an ischemic stroke (post explant) that allegedly was caused by the impella device. | The complainant reported a patient presenting with acute myocardial infarction and cardiogenic shock for impella support. the patient ultimately sustained an ischemic stroke (post explant) that allegedly was caused by the impella device. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36551 | 1220648-2022-01249 | 13305901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13305901&pc=OZD | 12/23/2021 | 10/27/2025 |  | 01/20/2022 |  |  | 12/27/2021 | 01/20/2022 |  |  | 08/31/2023 |  | 01/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022119917 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 59 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/27/2021 |  |  |  |  | The complainant reported a (b)(6) asian female had impella cp smart assist pump inserted. the patient had hemolysis after insertion of impella and as a result an unknown amount of haptoglobin was administered. | [5 empty] | Device discarded by customer. the impella cp smart assist pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36552 | 1220648-2022-01250 | 13305937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13305937&pc=OZD | 12/17/2021 | 01/20/2022 |  | 01/20/2022 |  |  | 12/28/2021 | 01/20/2022 |  |  | 09/30/2023 |  | 01/20/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP SMART ASSIST PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022129769 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 60 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/28/2021 |  |  |  |  | The complainant reported a (b)(6) asian had impella cp smart assist pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). after insertion of impella, lower limb vascular dissection was discovered and the pump was removed. | [5 empty] | Device discarded by customer. the impella cp smart assist pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36553 | 1035166-2022-00012 | 13306762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13306762&pc=DYB | 12/19/2021 | 05/20/2022 |  | 01/20/2022 |  |  | 12/29/2021 | 01/20/2022 |  |  | 07/01/2024 |  | 01/20/2022 |  | INTRODUCER, CATHETER |  |  | ABIOMED 23F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | DP-13404 |  | 0052-3021 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 76 YR |  | Female |  |  |  | Death | Yes | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that on (b)(6) 2021 during an acute myocardial infarction/cardiogenic shock procedure an impella rp pump's failure to start prompted the removal and attempted delivery of a 2nd impella rp. delivery of the 2nd rp failed after multiple attempts resulted in the introducer splitting at the distal tip and folding onto itself. the physician was unable to straighten the device. the patient's family withdrew support and made the patient cmo (comfort measures only) making the patient as comfortable as possible while the natural dying process takes place. patient passed on (b)(6) 2021 due to cardiogenic shock secondary to stemi. it was suspected that excessive manipulation of both the 1st and 2nd impella rp caused the damage to the introducer. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. | The following sections were updated in follow-up 1. b1, b2, b4, g3, g6, h1, h2, h6, h10, h11. b1: selected adverse event; b2: death; h1: other serious or important medical events; h6: selected impact code for death as patient outcome. device was used in treatment. the device was not returned for analysis as it was discarded, therefore, the exact cause of the product issue cannot be determined; however, it was suspected that excessive manipulation of both the 1st and 2nd impella rp caused the damage to the introducer. the patient's family withdrew support and made the patient comfort measures only status before the issue could be addressed any further and therefore; the clinical observation could not be confirmed. however, following controls are in place to mitigate the reported product issue. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure (abiomed introducer sheath in-process and final inspections): measure dimensions: measure tip i.d. of sheath using appropriate pin gauges. visual inspection: using 10x magnification, verify the tip is round, no flash protruding and no flash with width (around circumference), no splitting, cracks or other damages. per visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. per ifu never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. abiomed provides the ifu for this product. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36554 | 1035166-2022-00013 | 13290470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13290470&pc=DYB | 12/26/2021 | 05/20/2022 |  | 01/19/2022 |  |  | 01/10/2022 | 01/19/2022 |  |  | 07/01/2024 |  | 01/19/2022 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | DP-13641 |  | 0052-3006 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 50 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the physician attempted to deliver the impella 5.5 device through the introducer, at which point, the guide wire came out of position. the decision was made to remove the pump and start over. during 2nd attempt to deliver the device, the pump endured resistance when passing the distal tip of the introducer. the pump and introducer were removed to find a small tear at the end of the introducer. a new introducer was used without further issue. there was no harm to the patient. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. | The following sections were updated in follow-up. b4, g3, g6, h2, h6, h10. device was used in treatment. the device was not returned for analysis, therefore, the exact cause of the product issue cannot be determined and the clinical observation could not be confirmed. however, following controls are in place to mitigate the reported product issue. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections per procedure (abiomed introducer sheath in-process and final inspections): measure dimensions: measure tip i.d. of sheath using appropriate pin gauges. visual inspection: using 10x magnification, verify the tip is round, no flash protruding and no flash with width (around circumference), no splitting, cracks or other damages. per visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. per ifu never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. abiomed provides the ifu for this product. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36555 | 1220648-2022-01244 | 13279673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13279673&pc=OZD | 12/23/2021 | 02/04/2022 |  | 01/18/2022 |  |  | 12/23/2021 | 01/18/2022 |  |  | 08/31/2023 |  | 01/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022114394 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/23/2021 |  |  |  |  | In (b)(6) a medical center placed an impella cp for hemodynamic support. the patient had known coronary and peripheral arterial disease. during the time in the icu the medical team observed a hematoma. they explanted the impella cp and treated the access site with contralateral leg ballooning and blood product infusion. when this failed to achieve hemostasis the team thought there may be a retroperitoneal bleed and so they surgically closed the site. | [5 empty] | The medical center in (b)(6) did not return the product for analysis. the investigation is open pending more clinical details. should new information be shared that leads the root cause, a final report will be filed. | Since the original report was filed, the investigation has concluded. the medical center in japan did not return the impella pump or sheath product. the clinical details alone were shared. the root cause of the injury/bleed has been determined to be native patient condition. the failure mode will be monitored and trended.
36556 | 1220648-2022-01245 | 13279704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13279704&pc=OZD | 12/23/2021 | 02/07/2022 |  | 01/18/2022 |  |  | 12/23/2021 | 01/18/2022 |  |  | 06/30/2023 |  | 01/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022097234 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 21 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/23/2021 |  |  |  |  | The complainant reported a (b)(6) asian male had impella 5.0 pump inserted in the femoral for cardiomyopathy. the patient had blood clots and a fogarty catheter was used to remove the clots. | [5 empty] | The impella 5.0 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The device was returned for evaluation. visual inspection confirmed a clot on the device. the cause of the thrombus was attributed to the patient's condition.
36557 | 1220648-2022-01246 | 13279709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13279709&pc=OZD | 12/24/2021 | 02/07/2022 |  | 01/18/2022 |  |  | 12/24/2021 | 01/18/2022 |  |  | 10/31/2023 |  | 01/18/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022148253 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/24/2021 |  |  |  |  | The complainant reported a (b)(6) white male had impella cp pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). the pump prolapsed on itself as the repo sheath was inserted into the groin and a snare was used to straighten the pump out. | [5 empty] | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The device and data logs were returned for review. significant kinking was found on the cannula. the cause of the reported issue was attributed to a use-related issue, as the pump came out of position while attempting to insert the repositioning unit.
36558 | 1220648-2022-01242 | 13258750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13258750&pc=OZD | 12/15/2021 | 09/15/2024 |  | 01/14/2022 |  |  | 12/21/2021 | 01/14/2022 |  |  | 10/31/2023 |  | 01/14/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP SMART ASSIST PUMP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2022141984 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 80 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/21/2021 |  |  |  |  | The complainant reported an (b)(6) male had impella cp smart assist pump inserted in the right femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had bleeding complications around the repositioning sheath, surgical repair was successful, 3 (three) 5 (five) units of erythrocyte concentrates was administered but the patient expired. | The complainant reported an (b)(6) male had (b)(6) assist pump inserted in the right femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had bleeding complications around the repositioning sheath, surgical repair was successful, 3 (three) 5 (five) units of erythrocyte concentrates was administered but the patient expired. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  | The impella cp smart assist was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The (b)(6) assist pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The impella device was returned and an evaluation of the reported event was completed. the root cause was attributed to anticoagulation management not be followed by the hospital. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2022-01242 was provided.
36559 | 1220648-2022-01243 | 13259014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13259014&pc=OZD | 12/19/2021 | 01/14/2022 |  | 01/14/2022 |  |  | 12/19/2021 | 01/14/2022 |  |  | 09/30/2023 |  | 01/14/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022129793 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/20/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with cardiomyopathy for impella support. after the device was removed for escalation of care, a clot was found on the device. intervention to remove residual clotting within the access site was completed via a fogarty catheter. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36560 | 1220648-2022-01240 | 13254314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13254314&pc=OZD | 12/16/2021 | 01/13/2022 |  | 01/13/2022 |  |  | 12/16/2021 | 01/13/2022 |  |  | 09/30/2023 |  | 01/13/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022140429 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 12/16/2021 |  |  |  |  | The complainant reported an (b)(6) patient presenting for high risk percutaneous coronary intervention using the impella cp optical for hemodynamic support. the physician alleged the patient had endured a cerebrovascular accident as a result of using the device. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36561 | 1220648-2022-01241 | 13256271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13256271&pc=OZD | 12/17/2021 | 11/06/2023 |  | 01/13/2022 |  |  | 12/17/2021 | 01/13/2022 |  |  | 08/31/2023 |  | 01/13/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022140789 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | Pressure Problem |  |  |  |  | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/17/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with cargiomyopathy for impella support. during support, an access site hematoma developed. after decision was made to remove the pump, the hematoma was evacuated. a new pump was placed and support continued. | It was also reported that during support, there was a purge high pressure alarm and a low pump flow alarm. tissue plasminogen activator (tpa) was administered through the purge solution and the purge pressure high alarms resolved. the low flow alarms continued despite optimal pump placement. the pump was then removed. |  |  |  |  | The impella device was returned and an investigation is underway. a supplemental mdr will be filed upon conclusion of our analysis. | The pump was received and the event was analyzed. anticoagulation was found to be above the recommended range, therefore, evidence suggests the cause of the reported hematoma was related high levels of anticoagulation. | The investigation for the reported purge pressure issues and low flows has been completed. data logs showed the purge pressure rapidly rising after about 24 hours of support. the subsequent "purge pressure high" alarms were accompanied by simultaneous "placement signal low" and "suction" alarms. the high purge pressure eventually resolved, while the other two alarms persisted for the remainder of support. images sent by the clinical team confirmed the presence of a large clot ingested in the inflow cage upon explant. visual inspection of the returned pump revealed the presence of biomaterial on the optical sensor and blood in the purge gap. additionally, minor kinks were found. when the pump was run, biomaterial ejected from the outflow and all metrics were within normal limits and no alarms occurred. no other damage or abnormalities noted. the cause of the high purge pressure was biomaterial buildup occluding the purge gap. the cause of the suction/low flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records. h6-medical device codes, component code, investigation findings, investigation conclusions.
36562 | MW5106628 | 13251732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13251732&pc=OZD | 01/05/2022 | 01/10/2022 |  | 01/12/2022 |  |  |  | 01/12/2022 |  |  | 09/30/2023 |  | 01/12/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 21 FR |  |  | OZD |  |  | 05500008 |  |  | ABIOMED, INC. | [8 empty] | No answer provided | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 64 YR |  | Male |  |  |  | Injury | Yes | Yes | Pumping Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | Impella line damaged pt return to operating room. fda safety report id # (b)(4).
36563 | 13224112 | 13224112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13224112&pc=OZD | 01/01/2022 | 01/03/2022 | 01/03/2022 | 01/11/2022 | 01/03/2022 | 01/11/2022 |  | 01/11/2022 |  |  |  |  | 01/11/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | No answer provided |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 17155 DA |  | Male | 85 KG |  |  | Malfunction | No | Yes | Break | Obstruction of Flow | Therapeutic or Diagnostic Output Failure |  |  |  |  | Low Blood Pressure/ Hypotension | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The impella console alarmed ¿purge disconnect¿ circa 04:50. on inspection of the impella system, the purge system sidearm had broken off at the point where it enters the catheter, interrupting purge flow. the icu team contacted the manufacturer support line and were instructed to use a micropunture needle to temporarily replace the purge sidearm. with this temporizing measure, purge system flow was restored around 06:00. this exposed the patient to a risk of pump failure and hemodynamic compromise as the impella requires a constant purge fluid infusion and is at risk of pump stoppage if purge flow is compromised for as little as 2-3 minutes. the patient was urgently taken to the or (operating room) at 07:15 for impella replacement. during the time impella support was interrupted, blood pressure dropped to a mean of 45, requiring support with a second high-dose inotrope. because the impella insertion conduit contained intramural thrombus, the surgeon elective to place an iabp (intra-aortic balloon pump) and to not replace the impella. of note, the patient¿s purge solution was changed to sodium bicarbonate (for 3 days, approximately 2 months ago) after repeated urging by the manufacturer¿s representatives that this was the preferred therapy in the setting of heparin intolerance. five weeks ago, the manufacturer issued an advisory stating that this therapy had been associated with degradation of the purge system sidearm, resulting in 48 complaints and 5 cases of pump stoppage.
36564 | 1220648-2022-01238 | 13230990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13230990&pc=OZD | 12/14/2021 | 02/24/2022 |  | 01/11/2022 |  |  | 12/14/2021 | 01/11/2022 |  |  | 08/31/2023 |  | 01/11/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022125370 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2021 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with cardiomyopathy for hemodynamic support using the impella 5.5. during support, the patient endured hemolysis as confirmed via 3 elevated pfhg measurements over a 24 hour period. the pump was removed and replaced as a result. | [5 empty] | The impella device was received and an investigation is underway., a supplemental mdr will be filed at the completion of our analysis. | The device was returned and an investigation was completed. a visual inspection revealed a thrombus wrapped around the tip of the impeller. the root cause of the reported hemolysis was related to ingested biomaterial.
36565 | 1220648-2022-01239 | 13231023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13231023&pc=OZD | 12/15/2021 | 02/24/2022 |  | 01/11/2022 |  |  | 12/15/2021 | 01/11/2022 |  |  | 09/30/2023 |  | 01/11/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022138391 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 12/15/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella 5.0 device. approximately 226 hours into support, the pump stopped. troubleshooting was performed in an attempt to restart the device, to no avail. the pump was removed and a thrombus was noted on the outflow cage. the patient remained in stable condition. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Although the device was discarded, data logs were reviewed and a pump stop was confirmed followed by an elevated motor current spike. evidence suggests the root cause of the pump stop was due to an open line in all 3 motor phases, which happened dueto aggressive patient movement.
36566 | 1220648-2022-01232 | 13204194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13204194&pc=OZD | 12/11/2021 | 01/10/2022 |  | 01/10/2022 |  |  | 12/11/2021 | 01/10/2022 |  |  | 09/30/2023 |  | 01/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022142124 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/11/2021 |  |  |  |  | The complainant reported a (b)(6) white male had impella cp pump inserted in the right femoral artery (rfa) for high risk percutaneous coronary intervention(hrpci). the patient had hemolysis as urine was dark red/black color and the patient had groin bleeding/hematoma as a result the pump was removed and two (2) units of packed red blood cells (prbcs) was given. | [5 empty] | Device was not returned by customer. the impella cp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
36567 | 1220648-2022-01233 | 13204673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13204673&pc=OZD | 12/11/2021 | 01/10/2022 |  | 01/10/2022 |  |  | 12/11/2021 | 01/10/2022 |  |  | 07/31/2023 |  | 01/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022123906 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/11/2021 |  |  |  |  | The complainant reported a (b)(6) white male had impella cp pump inserted in the right femoral artery (rfa) for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient became extremely combative during removal of peel away sheath causing the catheter to advance aggressively and collapse completely into the left ventricle (lv) causing mitral valve damage, to resolve this the physician immediately began arranging surgical intervention. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36568 | 1220648-2022-01235 | 13204983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13204983&pc=OZD | 12/11/2021 | 02/24/2022 |  | 01/10/2022 |  |  | 12/11/2021 | 01/10/2022 |  |  | 09/30/2023 |  | 01/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022139816 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. after device was initiated, the pump fell out of the ventricle when removing the 14 fr introducer. upon attempting to re-advance the device past the aortic valve, the pump separated from the cannula and remained in the patient. intervention was required to snare and retrieve the device. a new impella was used, thereafter. | [5 empty] | The device has been received and an investigation is underway. upon completion of our analysis, a supplemental mdr will be filed. | The device was returned and an investigation for the reported failure was conducted. visual inspection confirmed a detached inflow cage from the remainder of the device. based on clinical review, we feel the root cause was related to the user attempting to wirelessly re-cross the aortic valve after the pump had fallen out of place, which resulted in severe cannula kinking and detachment of inflow cage from cannula.
36569 | 1220648-2022-01234 | 13205019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13205019&pc=OZD | 12/11/2021 | 01/10/2022 |  | 01/10/2022 |  |  | 12/11/2021 | 01/10/2022 |  |  | 02/28/2023 |  | 01/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022043394 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/13/2021 |  |  |  |  | The complainant reported an (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp device. during support, the patient endured an impella related access site hematoma. hemostasis was achieved after surgically opening the site to evaluate the injury, applying a tourniquet, and delivering multiple units of blood product. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36570 | 1220648-2022-01236 | 13205070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13205070&pc=OZD | 12/18/2021 | 02/07/2022 |  | 01/10/2022 |  |  | 12/18/2021 | 01/10/2022 |  |  | 08/31/2023 |  | 01/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022132097 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2021 |  |  |  |  | The complainant reported a (b)(6) white male had impella cp pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). after about eight (8) hours after implant, the patient had red urine and low kidney function with suspected hemolysis and as a result the pump was removed another one was inserted to try and resolve the hemolysis. | [5 empty] | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | Analysis of the provided field data logs confirmed numerous positioning alarms (impella position unknown, impella position in aorta, and impella position in ventricle). additionally, a sudden sustained rise in purge pressure was observed with a sustained drop in purge flow as well as multiple transient increases in motor current throughout the case; however, no alarms were triggered. the returned pump was visually inspected, and biomaterial was observed under the impeller, occluding approximately 50% of the purge gap. furthermore, minor epo-tek delamination on the sleeve bearing was noted. the pump was ran in a basin of water with elevated purge pressure and flow values and it was noted that motor current was above 900mmhg. the returned pump passed hemolysis testing per astm f841. a root cause for the reported positioning issues could not be determined; however, the reported hemolysis was likely due to the confirmed positioning issues.
36571 | 1220648-2022-01237 | 13208506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13208506&pc=OZD | 11/30/2021 | 05/22/2025 |  | 01/10/2022 |  |  | 01/07/2022 | 01/10/2022 |  |  | 02/28/2023 |  | 01/10/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022073207 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 57 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Unspecified Infection | Ischemia | Renal Failure | Ventricular Fibrillation | Vascular Dissection | Thrombocytopenia | [6 empty] | Life Threatening | Required Intervention | Hospitalization | Other |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 01/07/2022 |  |  |  |  | The medical center in (b)(6) had a successful cp support of over 11 days in (b)(6) 2021 in combination with ecmo for a patient admitted in cardiogenic shock. months later in (b)(6) 2021 the jpvad registry review team documented limb ischemia, hemolysis, and renal issues during the impella and ecmo support. there was no noted intervention for any of these events at the time of complaint creation. weeks later the medical team further detailed that after explant the limb was amputated to remedy the limb ischemia. the medical team noted that the patient was supported by both impella and ecmo and that there was a possibility that the ischemia was related to the use of impella, and that there was a possibility of peripheral arterial disease. | Additional information was reported indicating the patient experienced hemolysis, ventricular fibrillation, infection/sepsis, vascular damage (gastrointestinal bleed), thrombocytopenia and renal failure. the renal failure and hemolysis were treated with haptoglobin and dialysis. the ventricular tachycardia was reportedly due to the st elevated myocardial infarction. no further information was provided regarding further treatments or resolutions for the additional issues reported. the patient remained on impella support and was successfully weaned. |  |  |  |  | The medical center did not return the impella pump product for analysis. the cause of the limb ischemia was most likely the condition of the native anatomy. the failure mode will be monitored and trended. | Additional information for the initial mfr 1220648-2022-01237 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: hemolysis. impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system. section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ infection/sepsis. impella 5.5 with smart assist for use during cardiogenic shock & cp with smart assist section: table 3.2 impella catheter components. as noted in the impella ifu ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. as noted in the impella ifu ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ thrombocytopenia. impella cp with smart assist system. section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ vascular damage. impella cp with smart assist for use during cardiogenic shock and high risk pci & impella 5.5 with smart assist for use during. section: warnings & cautions: cautions. as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: use of impella with transcatheter aortic valves. as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ renal failure. impella cp with smart assist. section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
36572 | 1035166-2022-00001 | 13195122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13195122&pc=DYB | 11/30/2021 | 05/13/2022 |  | 01/07/2022 |  |  | 12/14/2021 | 01/07/2022 |  |  |  |  | 01/07/2022 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB | [5 empty] | OSCOR INC. |  |  | US | [5 empty] | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Distributor |  |  |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during insertion of 14f sheath, blood was reported leaking from the introducer after the dilator was removed. a new introducer was used and the impella was inserted without issue. there was no adverse impact to the patient. no additional information is available. | [5 empty] | Our investigation is still in progress, follow up report will be submitted if we find any further additional information. | One 14f abiomed introducer sheath was returned from the customer with the dilator. there were no other accessories. blood was found on and inside the sheath and dilator. according to the event description summary, blood leaked from the introducer after the dilator was removed. during the decontamination process, the sheath leaked immediately from the sheath shaft directly under the hub when water was flushed through the sideport tubing assembly with a syringe. when viewed under a 10x microscope, a kink and two 2-3mm cracks were found on the score lines of the sheath shaft directly under the hub. returned device analysis revealed the sheath was within manufacturing specifications. a kink and two 2-3mm cracks were found on the score lines of the sheath shaft directly under the hub which caused the sheath to leak when being used with an impella pump. although the exact cause is unknown, it is possible that resistance was encountered while in use which bent the sheath directly under the hub causing the sheath to crack and eventually leak. the sheaths are inspected by qa for cracks and damage and are also leak tested in-process 100%. since the lot number is unknown, the dhr cannot be reviewed to see if there's any related issues or anomalies. per procedure adelante s2s introducer sheath in-process and final inspection: final assembly. inspect 100% for following steps. visual inspection: - with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. foreign matter (fm)/fiber is limited to 2 particles < 0.10 mm2 or 1 particle < 0.2 mm2. - with naked eye at a distance of 12" to 18", verify the sheath hub and split caps are free of excess silicone oil applied between the seals during assembly. - verify split caps are properly placed and secured onto sheath hub and free of cracks/damages or excessive adhesive. vacuum leak test. perform pressure and vacuum leak test per procedure , latest revision. per ifu: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. no corrective or preventive action resulted after investigation of this event. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36573 | 1220648-2022-01228 | 13198855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13198855&pc=OZD | 12/09/2021 | 01/07/2022 |  | 01/07/2022 |  |  | 12/10/2021 | 01/07/2022 |  |  | 08/31/2023 |  | 01/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022129716 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/10/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for hemodynamic support prior to receiving cabg. during support, the patient was suspected as enduring hemolysis due to red color of urine. as treatment to prevent renal impact, haptoglobin was administered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36574 | 1220648-2022-01229 | 13198894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13198894&pc=OZD | 12/09/2021 | 05/19/2025 |  | 01/07/2022 |  |  | 12/15/2021 | 01/07/2022 |  |  | 02/28/2023 |  | 01/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 2022067866 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/15/2021 |  |  |  |  | The medical center in (b)(6) placed an impella cp via the femoral artery for hemodynamic support. the patient was critical with cardiovascular and surgical comorbidities and was also supported by ecmo. they observed blood loss at the access for the impella cp. to treat this the team infused blood products, both red blood cells and plasma (ffp). in addition to the access site bleed the patient was having considerable blood loss from the chest which was left open post-operatively from the coronary bypass. | The medical center in (b)(6) placed an (b)(6) via the femoral artery for hemodynamic support. the patient was critical with cardiovascular and surgical comorbidities and was also supported by ecmo. they observed blood loss at the access for the (b)(6). to treat this the team infused blood products, both red blood cells and plasma (ffp). in addition to the access site bleed the patient was having considerable blood loss from the chest which was left open post-operatively from the coronary bypass. | Additional information was reported indicating there was hemolysis that was treated by repositioning the pump. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  | The medical center in (b)(6) did not return the impella product nor any imaging for analysis. the root cause has not been determined. should more information be shared that leads to a root cause, a final report will follow. | The medical center in (b)(6) did not return the (b)(6) product nor any imaging for analysis. the root cause has not been determined. should more information be shared that leads to a root cause, a final report will follow. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2022-01229 was provided in sections b3, b5, d6, d7, h3 and h6. | The investigation for the reported access site bleeding ¿ major issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding was determined to be a result of anticoagulation management. this was determined because the bleeding was not only occurring at the impella access site but also at the surgical wound. this report is being filed as part of a retrospective review of historical records.
36575 | 1220648-2022-01230 | 13198904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13198904&pc=OZD | 12/15/2021 | 02/07/2022 |  | 01/07/2022 |  |  | 12/15/2021 | 01/07/2022 |  |  | 09/30/2023 |  | 01/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 SMART ASSIST PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022143612 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2021 |  |  |  |  | The complainant reported a (b)(6) white male had impella 5.5 pump inserted in the right femoral artery (rfa) for post cardiac-cardiogenic shock (pccs). the pump stopped and it was removed. | [5 empty] | The impella 5.5 smart assist pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The provided data logs showed 'impella stopped - motor current high' alarms triggered on (b)(6) 2021 beginning at 15:26. evaluation of the returned pump confirmed biomaterial around the impeller. the root cause of the pump stop was ingested biomaterial.
36576 | 1220648-2022-01231 | 13198928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13198928&pc=OZD | 12/11/2021 | 01/07/2022 |  | 01/07/2022 |  |  | 12/13/2021 | 01/07/2022 |  |  | 12/31/2021 |  | 01/07/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1461320 |  | 005066-JP |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 46 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/13/2021 |  |  |  |  | The complainant suspected a (b)(6) male patient presenting with cardiomyopathy for impella support using the impella 5.0 device. during support, the patient was suspected as enduring hemolysis, as evident by red colored urine. the device was removed prematurely and a permanent left ventricular assist device was implanted. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36577 | 1220648-2022-01226 | 13156156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13156156&pc=OZD | 12/09/2021 | 01/04/2022 |  | 01/04/2022 |  |  | 12/09/2021 | 01/04/2022 |  |  | 08/31/2023 |  | 01/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022132866 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/09/2021 |  |  |  |  | The complainant reported a (b)(6) black or african american female had impella cp pump inserted in the femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had limb ischemia as a result the pump was removed, and surgical repair was done. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36578 | 1220648-2022-01227 | 13156196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13156196&pc=OZD | 12/10/2021 | 01/04/2022 |  | 01/04/2022 |  |  | 12/10/2021 | 01/04/2022 |  |  | 07/31/2023 |  | 01/04/2022 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP SMART ASSIST PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022108789 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 50 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/10/2021 |  |  |  |  | The complainant reported a (b)(6) asian male had impella cp smart assist pump inserted in the right femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had hemolysis, p-level was changed and haptoglobin (2000u) was administered. | [5 empty] | Device discarded by customer. the impella cp smart assist pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36579 | MW5106425 | 13151764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13151764&pc=KFM | 12/28/2021 | 12/30/2021 |  | 01/03/2022 |  |  |  | 01/03/2022 |  |  |  |  | 01/03/2022 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA |  |  | KFM | [5 empty] | ABIOMED INC. |  |  | US |  | 024989 0000236419 1621 |  |  |  | No | [6 empty] | Health Professional |  |  | Voluntary report | Yes | Other Health Care Professional | [8 empty] | Unknown |  |  |  | Death | Yes | No | Therapeutic or Diagnostic Output Failure | [6 empty] | Failure of Implant | [12 empty] | Death | [28 empty] | Initial submission | [9 empty] | Impella pump failure. fda safety report id# (b)(4).
36580 | 1220648-2021-01219 | 13102725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13102725&pc=OZD | 11/30/2021 | 02/24/2022 |  | 12/29/2021 |  |  | 11/30/2021 | 12/29/2021 |  |  | 08/31/2023 |  | 12/29/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2022133896 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/30/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with cardiomyopathy. the impella 5.5 pump was selected for hemodynamic support. after approximately 185 hours (~8 days), a pump stop occurred. the pump was unable to be restarted, thus, removed. the patient was stable and did not require additional pump support. | [5 empty] | The device and data logs were returned for analysis. data logs confirmed the device had stopped and an alarm was triggered. testing of the device revealed open lines between all 3 motor phases. the root cause of the reported pump stop was related to an open line from excessive tensile force on the catheter. | The impella device has been received and an investigation is underway. upon completion of our analysis, a supplemental mdr will be filed.
36581 | 1220648-2021-01220 | 13102743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13102743&pc=OZD | 11/30/2021 | 01/27/2022 |  | 12/29/2021 |  |  | 11/30/2021 | 12/29/2021 |  |  | 01/31/2023 |  | 12/29/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022049054 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 12/01/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock. the impella cp optical was inserted and initiated without any reported issues. during support, the patient endured an ischemic stroke. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information was received from the physician on (b)(6) 2022 indicating the patient's cva was unrelated to the impella device. based on this information, this event was re-assessed and is no longer deemed a reportable event.
36582 | 1220648-2021-01221 | 13102746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13102746&pc=OZD | 12/21/2020 | 12/29/2021 |  | 12/29/2021 |  |  | 11/30/2021 | 12/29/2021 |  |  | 10/31/2022 |  | 12/29/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021011292 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/30/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp device. during support, the patient endured leg ischemia. as treatment, a bypass was performed and support continued. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36583 | 1220648-2021-01222 | 13109350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13109350&pc=OZD | 12/01/2021 | 12/29/2021 |  | 12/29/2021 |  |  | 12/01/2021 | 12/29/2021 |  |  | 08/31/2023 |  | 12/29/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022114762 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 81 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Eject | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/01/2021 |  |  |  |  | The complainant reported an (b)(6) male patient presenting for hemodynamic support. with the impella cp optical device. after insertion, the physician was unable to fully remove the guidewire. although the pump was able to be started, as the wire had cleared the outflow cage, vascular surgery was required to remove the wire and pump. the patient was inserted with pcps for further support. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36584 | 1220648-2021-01223 | 13109360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13109360&pc=OZD | 07/07/2021 | 12/29/2021 |  | 12/29/2021 |  |  | 12/07/2021 | 12/29/2021 |  |  | 02/28/2021 |  | 12/29/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022050281 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/07/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp device. during support, there were 2 instances of suspected hemolysis ((b)(6) 2021 and (b)(6) 2021) via red urine. as treatment, 2 units of red blood cells were delivered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36585 | 1220648-2021-01224 | 13109376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13109376&pc=OZD | 12/07/2021 | 12/29/2021 |  | 12/29/2021 |  |  | 12/07/2021 | 12/29/2021 |  |  | 09/30/2023 |  | 12/29/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022143974 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/07/2021 |  |  |  |  | The complainant reported an (b)(6) white male had impella cp pump inserted in the left common femoral for high risk percutaneous coronary intervention (hrpci). the patient had thrombus that was removed per aspiration catheter and a stent was placed post impella. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36586 | 1220648-2021-01225 | 13109413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13109413&pc=OZD | 11/03/2021 | 12/29/2021 |  | 12/29/2021 |  |  | 12/01/2021 | 12/29/2021 |  |  | 07/31/2023 |  | 12/29/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022108709 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 12/01/2021 |  |  |  |  | The complainant reported a (b)(6) patient presenting with cardiomyopathy. the impella cp optical device was selected for hemodynamic support. after less than 1 day of support, the patient was suspected of enduring hemolysis. as treatment, haptoglobin was delivered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36587 | 1035166-2021-00158 | 13095401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13095401&pc=DYB | 12/22/2021 | 12/28/2021 |  | 12/28/2021 |  |  | 12/22/2021 | 12/28/2021 |  |  |  |  | 12/28/2021 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | UNKNOWN |  | 0052-3021 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [7 empty] | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | No Apparent Adverse Event | [5 empty] | Heart Failure/Congestive Heart Failure | Mitral Valve Insufficiency/ Regurgitation | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | During an acute myocardial infarction (ami), cardiogenic shock (cgs) procedure, a 23fr introducer sheath kinked when the physician was trying to push the sheath into the body under fluoro. a second 23fr sheath was used and it kinked at the same spot where the iliac looked calcified. the 23fr introducers were exchanged for a gore sheath. the patient passed away in the intensive care unit due to overall heart and organ failure, death not related to device. the patient had decompesated heart failure due to severe mitral regurgitation. on (b)(6) 2021 the customer confirmed device is a 23f, but the lot number is unknown. | [5 empty] | The device was used in treatment. the device was not returned for analysis, therefore, the exact cause of the product issue cannot be determined and the clinical observation could not be confirmed. the lot number of the 23f abiomed introducer is unknown, however the following controls are in place to mitigate the reported product issue. per procedure (abiomed introducer sheath in-process and final inspections): measure dimensions: measure tip i.d. of sheath using appropriate pin gauges to be in range of 0.305" - 0.307" for 23f. visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. visual inspection: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. per ifu never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
36588 | 1220648-2021-01216 | 13091124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13091124&pc=OZD | 10/15/2019 | 12/27/2021 |  | 12/27/2021 |  |  | 12/07/2021 | 12/27/2021 |  |  | 07/31/2021 |  | 12/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 1434079 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Consumer | Health Professional | User facility | Company representation | 34 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/07/2021 |  |  |  |  | We received a publication from sage publication, 1-4, doi: 10.1177/0267659120941339 that a (b)(6) white female had impella cp pump inserted in the left femoral for cardiomyopathy. the patient had severe aortic regurgitation, the decision was made to proceed to median sternotomy and mechanical aortic valve replacement (avr). | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36589 | 1220648-2021-01218 | 13091131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13091131&pc=OZD | 08/03/2021 | 12/27/2021 |  | 12/27/2021 |  |  | 12/06/2021 | 12/27/2021 |  |  | 01/31/2023 |  | 12/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP SMART ASSIST PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022062677 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 62 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/06/2021 |  |  |  |  | The complainant reported a (b)(6) asian male had impella cp smart assist pump inserted in the left femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had access site bleeding and blood transfusion was administered. | [5 empty] | Device discarded by customer. the impella cp smart assist pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36590 | 1220648-2021-01210 | 13059722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13059722&pc=OZD | 11/24/2021 | 10/23/2024 |  | 12/22/2021 |  |  | 11/24/2021 | 12/22/2021 |  |  | 05/31/2023 |  | 12/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022083708 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | Asian | Death | Yes | No | Deformation Due to Compressive Stress | Adverse Event Without Identified Device or Use Problem | [5 empty] | Sepsis | Thrombocytopenia | [11 empty] | Other | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for impella support. during support, the patient developed thrombocytopenia and required platelet transfusion. the physician attributed the impella device as a causal/contributor factor. | Additional information was received indicating there was a kink found in the device's cannula upon removal. it was also reported that sepsis was noted and the patient expired. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation for the reported issues have been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the relationship between the thrombocytopenia and the use of impella was determined to be unknown as no evidence was provided to suggest a causal relationship. the relationship between sepsis and the use of impella was determined to be unknown as the impella device was confirmed to have been sterilized under validated conditions. the root cause of the cannula kink was unable to be determined as no product was returned for analysis. this report is being filed as part of a retrospective review of historical records. additional information for the initial mfr 1220648-2021-01210 was provided.
36591 | 1220648-2021-01211 | 13060126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13060126&pc=OZD | 11/24/2021 | 09/06/2022 |  | 12/22/2021 |  |  | 11/24/2021 | 12/22/2021 |  |  | 08/31/2023 |  | 12/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022119870 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia Stroke | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for support using the impella cp optical device. during support, the patient endured an access site bleed that prompted a blood transfusion of at least 2 units of blood products. | [5 empty] | New information received 2022-08-23 reports the patient endured an ischemic stroke during use of the device. the impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36592 | 1220648-2021-01212 | 13060194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13060194&pc=OZD | 12/07/2021 | 01/10/2022 |  | 12/22/2021 |  |  | 12/07/2021 | 12/22/2021 |  |  | 09/30/2023 |  | 12/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022132497 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2021 |  |  |  |  | The medical team in (b)(6) placed an impella cp for hemodynamic support of a patient admitted in cardiogenic shock. the support was ended and the pump explanted after just 19 hours due to concerns of limb ischemia. the team placed a balloon pump, iabp, for continued support. | [5 empty] | Since the original report was filed, the investigation into the limb ischemia has concluded. the sheath utilized was not indicated for use with impella. the sheath and pump were not returned for analysis from the medical center in japan. the root cause was the improper technique of utilizing the incorrect sheath. the failure mode will be monitored and trended. | The medical center in (b)(6) did not return the impella pump or introducer sheath for analysis. the further clinical details requested, as well as imaging, are not being furnished. should more details be shared that lead to root cause, a final report will be filed.
36593 | 1220648-2021-01213 | 13061104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13061104&pc=OZD | 10/27/2021 | 10/27/2025 |  | 12/22/2021 |  |  | 11/24/2021 | 12/22/2021 |  |  | 06/30/2023 |  | 12/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095403 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | Company representation |  |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2021 |  |  |  |  | Additionally, thrombocytopenia, access site bleeding and hemolysis, noted during support requiring blood transfusion were noted during support. | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. after successful use of the device, the pump was removed. at this time, a thrombus was found adhered to the device. an aspiration procedure was required the patient. |  |  |  |  | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output - use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis - ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction - ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2021-01213. b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2021-01213. d2 common device name revised on manufacturer device report 1220648-2021-01213 in accordance with updated procedures. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2021-01213. d6a and d6b revised from the initial manufacturer device report 1220648-2021-01213 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2021-01213 in accordance with updated procedures. g1 reporting contact first, last name, and email revised on manufacturer device report 1220648-2021-01213 in accordance with updated procedures. g2 report source; study and foreign was missing from manufacturer device report 1220648-2021-01213. h6 health effect - clinical code 1888, 1886, 4431 and 4643 was inadvertently omitted on the initial manufacturer device report number 1220648-2021-01213. h6 component code revised from the initial submission in accordance with updated procedures. h10 instructions for use inadvertently omitted on the initial manufacturer device report number 1220648-2021-01213. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36594 | 1220648-2021-01214 | 13061231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13061231&pc=OZD | 11/24/2021 | 01/27/2022 |  | 12/22/2021 |  |  | 11/24/2021 | 12/22/2021 |  |  | 08/31/2023 |  | 12/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022134525 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for high risk percutaneous coronary intervention. the impella cp optical was selected for support and inserted without any reported issues. during support, an access site bleed formed between the blue suture hub of the device and the tuohy borst lock. the patient also endured a retroperitoneal bleed. as treatment, the pump was removed, a covered stent was applied, and blood products were administered. | [5 empty] | Based on our investigation and review of pictures and videos provided from the customer, there was no problem with the device. the patient had a large body habitus. the device was pushed deeper into the skin in an attempt to stop the bleeding. the root cause of the reported complaint is related to the patient's size and lack of vessel recoil. | The impella device was received and an investigation is underway. a supplemental mdr will be filed upon completion of our analysis.
36595 | 1220648-2021-01207 | 13053426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13053426&pc=OZD | 12/03/2021 | 01/10/2022 |  | 12/21/2021 |  |  | 12/03/2021 | 12/21/2021 |  |  | 03/31/2023 |  | 12/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022069503 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/03/2021 |  |  |  |  | The team placed an impella cp for a patient with cardiogenic shock. the patient was actively coding and being resuscitated in the cardiac catheterization lab and needed the hemodynamic support. the cp was placed but under imaging the pump was seen to be kinked. due the patient condition, the pump was not replaced. the pump remained on with the kink and suboptimal support while awaiting transfer to a tertiary medical center. the limb became ischemic and the cp was explanted and replaced by a balloon pump (iabp). the team also closed the site with covered stents. | [5 empty] | Since the original report was filed, the investigation into the kinked pump and subsequent limb ischemia has concluded. the kink provided for lower than optimal flows. the cannula was observed to be kinked at investigation. the limb ischemia was due to the high dose of pressors and native patient condition. the failure modes will be monitored and trended. | The investigation into the limb ischemia is on-going at this time. upon completion of the investigation, the final report will be filed.
36596 | 1220648-2021-01208 | 13054093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13054093&pc=OZD | 12/05/2021 | 10/06/2025 |  | 12/21/2021 |  |  | 12/05/2021 | 12/21/2021 |  |  | 02/28/2023 |  | 12/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022055856 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Difficult to Fold, Unfold or Collapse | [6 empty] | Hematoma | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [10 empty] | Death | Other | Required Intervention | Hospitalization | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 12/05/2021 |  |  |  |  | The procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | A patient was admitted with known coronary artery disease suffering from an acute myocardial infarction. the team performed a cardiac catheterization and place an impella cp for hemodynamic support and planned for transfer to the tertiary medical center. upon removing the sheath and advancing the impella cp pump, the pump kinked upon itself. the team had to snare the product out via the previous radial artery access. after successfully snaring, the pump was explanted and a new impella pump was placed. |  |  |  |  | Since the original report was filed, the investigation has concluded. the fluoroscopic imaging shared by the medical center and the data log review revealed the pump was improperly positioned. the cause of the kinked pump and it's necessary retrieval, was the improper positioning of the pump while advancing the pump. the failure was use related. the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2021-01208 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The investigation into the kinked pump and necessary snaring is on-going at this time. upon completion of the investigation, a final report will be filed with root cause.
36597 | 1220648-2021-01199 | 13024762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13024762&pc=OZD | 11/29/2021 | 01/10/2022 |  | 12/17/2021 |  |  | 11/29/2021 | 12/17/2021 |  |  | 07/31/2023 |  | 12/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022114703 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  |  |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2021 |  |  |  |  | A patient was admitted to a medical center in (b)(6) suffering from an acute myocardial infarction. the team made the choice to give patient hemodynamic support by impella cp and ecmo. after just over 21 hours of support the team explanted the impella cp due to concerns of hemolysis. at explant they made the choice to place a balloon pump (iabp) instead of another impella. | [5 empty] | Since the original report was filed, the investigation into the reported hemolysis has concluded. the medical center in japan did not return the pump for benchtop hemolysis testing. the medical center did report that pump position was assessed and adjusted as well as volume/fluid delivered to treat the hemolysis, in this patient who had reported hypertrophy. the cause of the hemolysis was due to patient anatomy/condition. the failure mode will be monitored and trended. | The medical center in japan has not returned any impella pump product for analysis and benchtop hemolysis testing. should more details and clinical report be shared, that leads to a root cause, a final report will be filed.
36598 | 1220648-2021-01200 | 13024813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13024813&pc=OZD | 11/29/2021 | 12/17/2021 |  | 12/17/2021 |  |  | 11/29/2021 | 12/17/2021 |  |  | 05/31/2023 |  | 12/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED INC. |  |  | US |  | 2022088329 |  | 005062 |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/29/2021 |  |  |  |  | The medical team made the choice to escalate a patient's hemodynamic support from the impella cp to the impella 5.0. the patient had been admitted suffering from cardiogenic shock. after 345 hours/14 days of support the pump stopped. the pump stop occurred after previous alarms for rising motor current. they were beginning to wean and prepare the patient for explant due to these alarms, and so when the pump stopped, they made choice not to replace the pump with another impella. they explanted and the patient was stable. | [5 empty] | The medical center has not returned the impella pump. the data logs from the console were returned and analysis is on-going. upon completion of the investigation, a final report will be filed.
36599 | 1220648-2021-01201 | 13024959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13024959&pc=OZD | 11/29/2021 | 12/17/2021 |  | 12/17/2021 |  |  | 11/29/2021 | 12/17/2021 |  |  | 06/30/2023 |  | 12/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022100633 |  | 0048-004-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/29/2021 |  |  |  |  | The medical center in (b)(6) placed an impella cp to support a patient suffering from an acute myocardial infarction. the plan was to allow the impella to provide hemodynamic support during a coronary angioplasty procedure. the team observed a vascular damage to the ascending aorta after the insertion of the impella. the team performed a pericardiocentesis to remove the fluid and administered blood product replacement. the impella remained on for support for over 35 hours and was then removed. | [5 empty] | The medical team in japan discarded the impella pump after explant. no imaging or other clinical details have been shared. at this time the investigation is ongoing. should more information be shared that leads to a root cause of the vessel damage, a final report will be filed.
36600 | 1220648-2021-01202 | 13026083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13026083&pc=OZD | 12/01/2021 | 10/20/2024 |  | 12/17/2021 |  |  | 12/01/2021 | 12/17/2021 |  |  | 08/31/2023 |  | 12/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022127589 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemoptysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Thrombocytopenia | [9 empty] | Required Intervention | Death | Hospitalization |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 12/01/2021 |  |  |  |  | A smaller sized female patient was admitted to the medical center in (b)(6) suffering from an acute myocardial infarction. the team made the choice to hemodynamically support her by both the use of an impella cp and ecmo. the pump supported her for over 26 hours when she expired due to underlying condition. later, the jpvad registry documentation noted the patient to have had an access site bleed, thrombocytopenia, and ventricular tachycardia during her impella support. the impella access site bleed was noted to be treated by blood product infusion. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the access site bleeding is on-going at this time. no impella product has been returned from the medical center in (b)(6). should more clinical details be shared that lead to a root cause, a final report will be filed. | Since the original report was filed, the investigation into the jpvad registry, in japan, reported harms of access site bleeding and hemolysis (which prompted rbc infusion) has concluded. no product was returned for analysis or bench-top hemolysis testing. the cause of the access site bleeding was determined to be use related. the clinical team had not used any imaging during setup and insertion. the cause of the hemolysis was patient condition. the myocardial infarction and filling issues contributed to the hemolysis. the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided. | H.6 code 1887, 2095, and 4431 were inadvertently omitted from the initial manufacturer device report number 1220648-2021-01202 and was added.
36601 | 1220648-2021-01203 | 13026391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13026391&pc=OZD | 11/19/2021 | 04/15/2025 |  | 12/17/2021 |  |  | 11/19/2021 | 12/17/2021 |  |  | 08/31/2023 |  | 12/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022125201 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/19/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting for high risk percutaneous coronary intervention and peripheral vascular disease. the impella pump was selected for hemodynamic support. during support, limb ischemia developed in the impella inserted leg. as a result, the pump was removed. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported limb ischemia has been completed. the impella device has not been returned for evaluation. the patient reportedly had severe peripheral vascular disease, therefore, the root cause of the limb ischemia was most likely patient condition. this report is being filed as part of a retrospective review of historical records. g1 reporting contact and contact email revised on manufacturer device report 1220648-2021-01203 in accordance with updated procedures. | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36602 | 1220648-2021-01195 | 13017567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13017567&pc=OZD | 12/09/2021 | 01/20/2022 |  | 12/16/2021 |  |  | 12/09/2021 | 12/16/2021 |  |  | 10/31/2021 |  | 12/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED INC. |  |  | US |  | 1452333 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/09/2021 |  |  |  |  | The medical center in the(b)(6) placed an impella 5.0 via the axillary artery graft for hemodynamic support of a patient suffering from cardiomyopathy. the patient was transferred to the icu for continued monitoring. after 7 days of support the impella pump stopped and was unable to be restarted. the medical team explanted and replaced the pump. | [5 empty] | The investigation into the impella 5.0 stopping is ongoing at this time. no product has yet been returned from the medical center in the uk. upon completion of the investigation, a final report will be filed. | Since the filing of the original report for the pump stop, the investigation has concluded. the impella pump was returned and upon examination of it, there was ingested biomaterial which caused the pump to stop. the failure mode will be monitored and trended.
36603 | 1220648-2021-01196 | 13017586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13017586&pc=OZD | 11/18/2021 | 01/07/2022 |  | 12/16/2021 |  |  | 11/18/2021 | 12/16/2021 |  |  | 12/31/2021 |  | 12/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1466157 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/18/2021 |  |  |  |  | The medical team in (b)(6) placed an impella 5.0 via an axillary graft to supplement the support being provided by the ecmo. the access site for the impella had a hematoma develop. the team had to infuse blood products to treat the patient. the pump supported the patient despite the hematoma for over 52 hours. | [5 empty] | Since the original report was filed, the investigation has concluded. the root cause of the access site bleeding was use related. the medical team in japan had relayed that the blood vessel was likely injured during the skin incision. the failure mode will be monitored and trended. | The medical center in (b)(6) discarded the impella pump product. no analysis is possible. should more details be shared that leads to a root cause, a final report will be filed.
36604 | 1220648-2021-01197 | 13017633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13017633&pc=OZD | 11/17/2021 | 10/01/2025 |  | 12/16/2021 |  |  | 11/17/2021 | 12/16/2021 |  |  | 04/30/2023 |  | 12/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 2022079848 |  | 0048-0034-JP |  | Yes |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fungal Infection | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2021 |  |  |  |  | The medical team in (b)(6) placed an impella cp to further support the patient being supported by va-ecmo. the patient had been admitted in cardiogenic shock and had coronary angioplasty performed. during the support the team observed the impella cp to have a loosened ring in it's repositioning sheath. the ring had been fixed but was loose and allowed for pump movement. days later the team observed an infection that they attributed to the malfunctioning cp pump. the infection was treated by medical therapy (bactericide administration) and blood products. | [5 empty] | The investigation into the reported infection is on-going at this time. upon completion of the investigation, a final report with root cause will be filed.
36605 | 1220648-2021-01198 | 13017643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13017643&pc=OZD | 11/16/2021 | 12/16/2021 |  | 12/16/2021 |  |  | 11/16/2021 | 12/16/2021 |  |  | 05/31/2023 |  | 12/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022085961 |  | 0048-0024-JP |  | No |  |  | Not Available |  |  |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility |  |  |  |  | 0 YR |  | Unknown |  |  |  | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/16/2021 |  |  |  |  | A patient of unknown age was admitted to a medical center in (b)(6) in (b)(6) of 2021 suffering from cardiogenic shock. the medical team placed an impella cp for hemodynamic support during cardiac angioplasty. the support was successful, the pump was then weaned and explanted. months later the jpvad registry noted access site bleeding throughout the entire support of 69 hours. the bleed was treated by blood cell infusion of an unknown amount. | [5 empty] | The team in (b)(6) has not relayed any further clinical details as requested. the product was discarded. should more details be shared which lead to a root cause, a final report will follow.
36606 | 1220648-2021-01194 | 13011102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13011102&pc=OZD | 11/19/2021 | 01/10/2022 |  | 12/15/2021 |  |  | 11/19/2021 | 12/15/2021 |  |  | 08/31/2023 |  | 12/15/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  |  |  | 2022127659 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/19/2021 |  |  |  |  | A male patient of unknown age was admitted to the medical center in (b)(6) suffering from ventricular fibrillation. the team placed the patient on ecmo support and an impella cp. the team observed an access site bleed at the impella site. the bleed was treated by removal of the heparin from the pump purge, blood product infusion, and a skin suture. the patient was noted to have a hemoglobin drop as the patient was bleeding form multiple sites, not only the impella site. the impella pump remained on for hemodynamic support til the family made the choice to withdraw care. | [5 empty] | The medical center in (b)(6) has not yet returned any impella product for analysis. the root cause of the access site bleeding is not determined at this time. should more details be shared, which lead to root cause, a final report will be filed. the failure mode will be monitored and trended. | Since the original report was filed the investigation into the access site bleed has concluded. no product was returned from the medical center in japan. after reviewing the clinical information shared, it was determined that the cause of the access site bleeding was most likely use related. multiple sites were seen to be bleeding during the patient's support. the team observed a mediastinum and arm hematoma due to cardiac massage. the root cause was use related. the failure mode will be monitored and trended.
36607 | MW5106103 | 13019656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13019656&pc=DYB | 12/11/2021 | 12/14/2021 |  | 12/15/2021 |  |  |  | 12/15/2021 |  |  |  |  | 12/15/2021 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14FR IMPELLA CP |  |  | DYB | [5 empty] | ABIOMED, INC. | [8 empty] | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 50 YR |  | Male | 95 KG | Hispanic | White | Malfunction | No | Yes | Break | Material Deformation | [5 empty] | Insufficient Information | [41 empty] | Initial submission | [9 empty] | The impella device kinked during insertion and broke off when the physician attempted to pull back. fda safety report id # (b)(4).
36608 | 1220648-2021-01189 | 13001588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13001588&pc=OZD | 11/15/2021 | 01/04/2022 |  | 12/14/2021 |  |  | 11/15/2021 | 12/14/2021 |  |  | 07/31/2023 |  | 12/14/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMPPPPP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022118319 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2021 |  |  |  |  | An impella cp was placed to provide hemodynamic support for a patient admitted after collapsing outside the hospital and found to have multivessel coronary artery disease when taken to the cardiac catheterization lab. the team observed and access site bleed at the impella site. this prompted the team to replace the sheath and pump and infuse 2 units of blood products to the patient. | [5 empty] | The investigation has concluded since the original report was filed. no introducer product was returned for evaluation. the root cause for the access site bleeding was determined to be a use related one. the medical team had reported the introducer was punctured during the single access. | The investigation into the access site bleeding is on-going. no product has been returned yet for analysis. should product and clinical details be shared, which lead to a root cause for the failure, a final report will be filed.
36609 | 1220648-2021-01192 | 13002972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13002972&pc=OZD | 11/21/2021 | 12/14/2021 |  | 12/14/2021 |  |  | 11/22/2021 | 12/14/2021 |  |  | 08/31/2023 |  | 12/14/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP SMART ASSIST PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022119997 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 0 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/22/2021 |  |  |  |  | The complainant reported an asian patient had impella cp smart assist pump inserted in the femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had bleeding at the puncture site and blood transfusion was administered. | [5 empty] | Device discarded by customer. the impella cp smart assist pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36610 | 1220648-2021-01190 | 13003175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13003175&pc=OZD | 11/16/2021 | 12/21/2021 |  | 12/14/2021 |  |  | 11/16/2021 | 12/14/2021 |  |  | 08/31/2023 |  | 12/14/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022128071 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 76 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/16/2021 |  |  |  |  | The medical center admitted a patient with heart failure and other cardiac comorbidities. to support the patient hemodynamically they placed an impella cp, via the 14fr x 13cm sheath. the patient developed a hematoma at the site of the impella 14fr sheath and pump. the hematoma was treated by ballooning from the contralateral leg and blood cell infusion. | [5 empty] | The investigation is on-going at this time. no product has yet been returned for analysis. should more clinical details be shared that lead to a root cause, a final report will be filed. | Since the original report regarding the access site hematoma was made, the investigation has concluded. no product was ever returned for analysis. upon review of the clinical details, the root cause has been determined to be patient condition. the team followed best practices, but due to lack of vessel recoil the hematoma developed.
36611 | 1220648-2021-01193 | 13003274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=13003274&pc=OZD | 11/22/2021 | 09/18/2024 |  | 12/14/2021 |  |  | 11/22/2021 | 12/14/2021 |  |  | 08/31/2023 |  | 12/14/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022131964 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2021 |  |  |  |  | The complainant reported a (b)(6) female had impella cp pump for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient's left ventricle (lv) had laceration and perforated resulting in tamponade, surgical repair, and four (4) liters (l) of packed red blood cells (prbcs) was administered. | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
36612 | 1220648-2021-01183 | 12982849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12982849&pc=OZD | 12/06/2021 | 10/19/2024 |  | 12/13/2021 |  |  | 12/06/2021 | 12/13/2021 |  |  | 04/30/2023 |  | 12/13/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022075994 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Other | Foreign | User facility | Company representation |  |  | 42 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Priming Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Thrombocytopenia | [10 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2021 |  |  |  |  | The medical team in (b)(6) had a cp pump support in month of (b)(6) 2021. the team had reported purge cassette issues and an unrelated death that were previously investigated. the team reported in (b)(6) 2021,via the jpvad registry, of access site bleeding treated by blood product infusion | Additional information reported that the patient had a ventricular tachycardia (vt) storm and thrombocytopenia. it was also reported that there was an issue with the purge cassette. purge system failure alarm occurred during priming. the purge cassette was replaced, and priming was completed without any problems. the patient later expired due to acute heart failure. no allegations were made towards the impella for cause of death. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the access site bleeding that prompted blood product infusion is complete. there was a tendency to bleed as noted by the medical team. the root cause of the bleeding was patient condition related. the pump and introducer sheath were not returned for analysis. the failure mode will be monitored and trended. | Additional information for the initial mfr 1220648-2021-01183 was provided in sections b1, b5, d6a, d6b, d7a, d9, g3, h3, h6, and h10. the investigations for the reported issues; priming, tachycardia, bleeding and thrombocytopenia have been completed. the impella was received from the customer and the following were confirmed: the cause of the priming failure was determined to most likely be a use related damage purge tubing line resulting in a purge leak. any device contacting the heart wall in conjunction with a poor patient condition could result in vf/vt but since fluoroscopic imaging and/or sufficient clinical details were not provided, a relation between impella and the vf/vt is unable to be determined. the cause of the bleeding and thrombocytopenia were most likely patient condition related since the patient was reported to have a tendency to bleed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
36613 | 1220648-2021-01184 | 12987181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12987181&pc=OZD | 07/22/2020 | 09/21/2024 |  | 12/13/2021 |  |  | 12/03/2021 | 12/13/2021 |  |  | 04/30/2022 |  | 12/13/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 SMART ASSIST PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1477154 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/03/2021 |  |  |  |  | The complainant reported a (b)(6) male had impella 5.5 inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient was confused, got up, tripped and fell causing the 5.5 pump to get pulled out of the left ventricle (lv) and disconnected. the pump did not turn on. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | Neither the product nor the data logs were sent. after reviewing the clinical information, it was determined that the cause of the pump stop was likely an open line due to tensile stress on the impella stemming from the patient's fall, which generated enough force to disconnect the patient cable from the (b)(6). | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6.
36614 | 1220648-2021-01185 | 12987225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12987225&pc=OZD | 11/15/2021 | 01/10/2022 |  | 12/13/2021 |  |  | 11/15/2021 | 12/13/2021 |  |  | 08/31/2023 |  | 12/13/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022132886 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2021 |  |  |  |  | An impella cp was placed for hemodynamic support of a patient who had arrested at home. he was given cpr and treated for ventricular fibrillation. the cp pump came out of position after 6 hours of support. when the medical team attempted to re-position the pump, the pump became kinked and was removed by a snare. | [5 empty] | The investigation into the pump kinking and being removed by a snare is on-going. upon completion of the investigation, a final report will be filed. | Since the initial report was filed, the investigation has completed. the impella product was returned and damage to the catheter and the reported kink in the cannula could be seen. it is possible that force was used in the attempted repositioning of the cp. it is also possible that some of the damage occurred during the cpr itself. wirelessly recrossing the aortic valve is not permitted per ifu. the cause of the pump needing to be explanted via snare was the kink/damage at pump repositioning. this was a user error of improper technique. the failure mode will be monitored and trended.
36615 | 1220648-2021-01186 | 12987313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12987313&pc=OZD | 06/19/2021 | 12/13/2021 |  | 12/13/2021 |  |  | 11/16/2021 | 12/13/2021 |  |  | 04/30/2022 |  | 12/13/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1488338 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/18/2021 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with myocarditis. the impella cp was selected for hemodynamic support. during support, the patient developed an access site bleed that prompted a blood transfusion. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36616 | 1220648-2021-01187 | 12987324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12987324&pc=OZD | 03/12/2021 | 12/13/2021 |  | 12/13/2021 |  |  | 11/16/2021 | 12/13/2021 |  |  | 10/31/2022 |  | 12/13/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021017577 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/16/2021 |  |  |  |  | The complainant reported an (b)(6) female patient presenting for high risk percutaneous coronary intervention. during impella support, the patient's leg became discolored. after the procedure, the impella was removed as a result of the ischemia. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36617 | 1220648-2021-01188 | 12987561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12987561&pc=OZD | 12/07/2021 | 12/13/2021 |  | 12/13/2021 |  |  | 12/07/2021 | 12/13/2021 |  |  | 11/30/2021 |  | 12/13/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED INC. |  |  | US |  | 1449535 |  | 0048-0002 |  | No |  |  | Not Available |  |  |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/07/2021 |  |  |  |  | The eu complainant had a pump stop during cp support. the pump was explanted and not replaced as the patient was hemodynamically stable. the pump had support the patient for 5 days in 2020. in 2021 there was a publication discovered inclusive of this support that noted access site bleeding that prompted many units of blood replacement as well as vascular intervention. | [5 empty] | The impella was discarded after use in 2020. access site bleeding occurred, which was reported to be related to the high dose of heparin. the cause of the bleeding was high patient act values.
36618 | 1220648-2021-01179 | 12975602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12975602&pc=OZD | 04/10/2021 | 12/10/2021 |  | 12/10/2021 |  |  | 11/18/2021 | 12/10/2021 |  |  | 07/31/2022 |  | 12/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1493301 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 0 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/18/2021 |  |  |  |  | The complainant reported an asian male had impella cp pump inserted in the right coronary artery (rca) for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had hemolysis and the pump was removed earlier than expected. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36619 | 1220648-2021-01180 | 12976866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12976866&pc=OZD | 11/18/2021 | 12/10/2021 |  | 12/10/2021 |  |  | 11/18/2021 | 12/10/2021 |  |  | 08/31/2023 |  | 12/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP SMART ASSIST PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022119877 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 81 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/18/2021 |  |  |  |  | The complainant reported an (b)(6) male had impella cp smart assist pump inserted in the left femoral for high risk percutaneous coronary intervention (hrpci). the patient had bleeding from impella access site and two (2) or more units of blood transfusions was administered. | [5 empty] | Device discarded by customer. the impella cp smart assist pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36620 | 1220648-2021-01181 | 12977158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12977158&pc=OZD | 11/19/2021 | 10/30/2024 |  | 12/10/2021 |  |  | 11/19/2021 | 12/10/2021 |  |  | 06/30/2023 |  | 12/10/2021 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA RP PUMP |  |  | OZD |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2022109611 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/19/2021 |  |  |  |  | The complainant reported a white male had impella rp pump placed via the right (rt) femoral vein. the patient had suspected hemolysis about 3-4 days into the case, the pump was explanted as a result. | Additional information was received from the complainant that the impella rp pump was removed due to a suspicion of thrombus. in addition, it was noted that the patient was heparin-induced thrombocytopenia (hit) positive and bi-carb remained in the purge. |  |  |  |  | Device discarded by customer. the impella rp pump was discarded by the customer therefore a failure analysis investigation cannot be completed. | As noted in the impella ifu: impella rp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2021-01181 was provided in sections b5, d6a, d6b, h6 and h10.
36621 | 1220648-2021-01176 | 12961605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12961605&pc=OZD | 11/13/2021 | 09/18/2024 |  | 12/08/2021 |  |  | 11/13/2021 | 12/08/2021 |  |  | 08/31/2023 |  | 12/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022124849 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2021 |  |  |  |  | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock. the impella cp optical was inserted for hemodynamic support. during support, an access site hematoma developed. two units of blood product were delivered as treatment. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided. | The impella device was discarded by the customer and therefore, investigation of device was not possible. a supplemental mdr will be filed if new information is obtained.
36622 | 1220648-2021-01177 | 12961642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12961642&pc=OZD | 11/13/2021 | 01/27/2022 |  | 12/08/2021 |  |  | 11/13/2021 | 12/08/2021 |  |  | 08/31/2023 |  | 12/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022135040 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2021 |  |  |  |  | The complainant reported a (b)(6) male had impella 5.5 pump inserted in the left axillary. the impella 5.5 pump was inserted for post cardiac-cardiogenic shock (pccs), however, air bubbles was seen on the echo in the left ventricle (lv) and ascending aorta. no other adverse effects or additional intervention was reported. | [5 empty] | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | Clinical images and data logs were not sent for investigation. visual inspection of the returned pump revealed several small puncture holes in the catheter between the 30-cm and 31-cm marks. the pump plug was tightly bound in many layers of bloody gauze, surgical tape, and bone wax. no other damage or abnormalities were found. when the pump was run with the returned purge cassette, its performance was within specification and no relevant alarms occurred. in conclusion, it was determined that the cause of the pump handle leak were puncture holes in the catheter. the holes did not result in a pump failure.
36623 | 1220648-2021-01178 | 12961650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12961650&pc=OZD | 06/12/2021 | 12/08/2021 |  | 12/08/2021 |  |  | 11/16/2021 | 12/08/2021 |  |  | 08/31/2022 |  | 12/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1496671 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 49 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/16/2021 |  |  |  |  | The complainant reported a (b)(6) asian patient had impella cp pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had access site bleeding and blood transfusion was administered. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36624 | 1220648-2021-01175 | 12961661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12961661&pc=OZD | 11/13/2021 | 12/08/2021 |  | 12/08/2021 |  |  | 11/13/2021 | 12/08/2021 |  |  | 07/31/2023 |  | 12/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022123989 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/13/2021 |  |  |  |  | The complainant reported an (b)(6) male patient presenting for hemodynamic support pre-operative to a coronary artery bypass procedure. during support, an access site hematoma formed. the pump was removed and a balloon tamponade was required to control the groin site as a result. a new impella was placed in the opposing femoral artery and hemodynamic support continued. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. a supplemental mdr will be filed if new information is obtained.
36625 | 1220648-2021-01171 | 12952398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12952398&pc=OZD | 11/12/2021 | 09/18/2024 |  | 12/07/2021 |  |  | 11/12/2021 | 12/07/2021 |  |  | 07/31/2022 |  | 12/07/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1495607 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 68 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/12/2021 |  |  |  |  | Additional information the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | A patient had been admitted for coronary angioplasty secondary to the acute myocardial infarction. when his condition deteriorated, the team placed an impella 5.5 for hemodynamic support. after 8 days of successful support with the 5.5 the team diagnosed an embolic cva/stroke. the team performed a thrombectomy to remove the thrombotic debris. the patient had an underlying condition that predisposed him to thrombus development. the stroke and the relationship to the use of the impella pump was not confirmed or denied fully by the medical team. |  |  |  |  | The investigation into the stroke/patient injury has completed. the patient injury was determined to have an unknown and unrelated relationship to the impella 5.5 pump. the cause was underlying patient condition. the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided. | The investigation into the cerebral event during impella 5.5 support is ongoing at this time. upon completion of the analysis, a final report will be filed.
36626 | 1220648-2021-01172 | 12952448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12952448&pc=OZD | 11/12/2021 | 12/21/2021 |  | 12/07/2021 |  |  | 11/12/2021 | 12/07/2021 |  |  | 08/31/2023 |  | 12/07/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022124056 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 61 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2021 |  |  |  |  | An impella 5.5 was placed to support a patient who already had va-ecmo support after a mitral valve replacement was done in the operating suite. the 5.5 pump was placed via the axillary graft to the left ventricle. unfortunately, after just 17.5 hours of support the pump stopped. the pump had exhibited purge pressure and purge flow alarms prior to the stop. the team explanted the pump. | [5 empty] | The investigation into the pump stop is on-going at this time. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation has closed. the medical center returned the impella pump product and the data logs for analysis. the analysis determined the root cause of the pump stop to be due to blood ingress which was most likely a result of a kinked purge line during support. the failure mode will be monitored and trended.
36627 | 1220648-2021-01173 | 12952490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12952490&pc=OZD | 11/08/2021 | 12/07/2021 |  | 12/07/2021 |  |  | 11/08/2021 | 12/07/2021 |  |  |  |  | 12/07/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/08/2021 |  |  |  |  | The complainant published an article that reported a (b)(6) male patient presenting for high risk percutaneous coronary intervention. the impella cp device was selected for hemodynamic support. at some point either during or after the procedure, a perforation was noted in the left ventricle and attributed to the impella device. the pump was removed and intervention was performed to resolve the injury. follow up between a local company representative and the hospital noted the 0.018" guidewire was attributed to the perforation. the patient was noted as "asymptomatic" at his 12 month follow up. | [5 empty] | The device was not received, therefore, a physical analysis could not be performed. the event was discovered by our company via an online publication and we are actively attempting to gain further information. should we received the device, or any new information, we will file a supplemental report.
36628 | 1220648-2021-01174 | 12952835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12952835&pc=OZD | 11/11/2021 | 01/14/2022 |  | 12/07/2021 |  |  | 11/11/2021 | 12/07/2021 |  |  | 08/31/2023 |  | 12/07/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022125460 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2021 |  |  |  |  | The complainant reported a (b)(6) male patient enduring hemolysis during support. the pump was removed for escalation to an impella 5.5, at which time a thrombus was found in the same artery the pump had been inserted. a cut-down procedure was performed to remove the thrombus and 2 units of blood product were delivered. | [5 empty] | The device was received and an investigation was conducted. result of hemolysis testing identified blood clot within the cannula-to-outflow bond of the device and delaminated epotek was found. the reported thrombus found within the left femoral artery was deemed the result of the patient condition. | The impella device has been returned and an investigation is underway. upon completion, a supplemental mdr will be filed.
36629 | 1220648-2021-01166 | 12923458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12923458&pc=OZD | 11/11/2021 | 12/03/2021 |  | 12/03/2021 |  |  | 11/11/2021 | 12/03/2021 |  |  | 06/30/2023 |  | 12/03/2021 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA RP PUMP |  |  | OZD |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2022120048 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/11/2021 |  |  |  |  | The complainant reported a (b)(6) white male had impella rp pump inserted for post cardiac-cardiogenic shock (pccs). the patient had bleeding around the rp insertion site and a total of 5 (five) units of blood products was administered. | [5 empty] | Device was not returned by customer. the impella rp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
36630 | 1220648-2021-01165 | 12923491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12923491&pc=OZD | 10/29/2021 | 12/03/2021 |  | 12/03/2021 |  |  | 11/11/2021 | 12/03/2021 |  |  | 06/30/2023 |  | 12/03/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, OZD |  |  | IMPELLA CP SMART ASSIST PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022100472 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 70 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/11/2021 |  |  |  |  | The complainant reported a (b)(6) asian male patient had impella cp optical pump inserted in the right femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had bleeding from the puncture site after insertion and blood was transfused; albumin administration and at least two (2) units of blood. | [5 empty] | Device discarded by customer. the impella cp smart assist pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36631 | 1220648-2021-01167 | 12923527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12923527&pc=OZD | 11/08/2021 | 12/03/2021 |  | 12/03/2021 |  |  | 11/08/2021 | 12/03/2021 |  |  | 07/31/2023 |  | 12/03/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022105828 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 73 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/08/2021 |  |  |  |  | The complainant reported a (b)(6) male patient suspected of enduring hemolysis during impella support. as treatment, haptoglobin was delivered. | [5 empty] | The impella deivce was not received by the customer and therefore, investigation of device was not possible. a supplemental mdr will be filed if new information is obtained.
36632 | 1220648-2021-01168 | 12924435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12924435&pc=OZD | 11/10/2021 | 10/23/2024 |  | 12/03/2021 |  |  | 11/10/2021 | 12/03/2021 |  |  | 02/28/2023 |  | 12/03/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022046102 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Hospitalization | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/10/2021 |  |  |  |  | Additional information, when the pump was explanted, and the patient returned to the lab for penumbra aspiration of the thrombus. during the intervention the patient coded. the family made the choice to withdraw care, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | A patient was admitted for a protected angioplasty of the coronary arteries and intervention also for the carotid artery. an impella cp was placed for hemodynamic support via the femoral artery. the interventions were successful. when in the icu, for continued monitoring, the patient sat upright which allowed for hemolysis, access site bleeding, and subsequent limb ischemia due to clot development at the impella access site. the pump was explanted and the patient returned to the interventional lab for penumbra aspiration of the thrombus. |  |  |  |  | Since the original report was filed, the investigation into the reported access site bleeding, limb ischemia, and hemolysis has concluded. product and data logs were not returned for analysis, only clinical case details were shared for investigation. the access site bleed was due to patient movement. the limb ischemia was due to patient condition. the patient was noted to have smaller peripheral arteries and had sat up while on support. the hemolysis was due to the use error of the patient sitting up and the pump coming out of optimal position. | Impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2021-01168 was provided in sections b2, b5, d6, h1, and h6. | The investigation into the patient harm is on-going at this time. the medical center discarded the impella product, but more clinical details and imaging have been requested. upon completion of the investigation, a final report will be filed.
36633 | 1220648-2021-01169 | 12924527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12924527&pc=OZD | 11/10/2021 | 10/24/2023 |  | 12/03/2021 |  |  | 11/10/2021 | 12/03/2021 |  |  | 05/31/2023 |  | 12/03/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022086819 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Crack | Leak/Splash | Pumping Stopped |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/10/2021 |  |  |  |  | A patient with a complex medical history was admitted for bypass surgery. the medical team had complications in the operating suite which prompted insertion of the impella 5.5 pump for hemodynamic support. the patient was then admitted to the icu with an open and packed chest. after over 130 hours of support the impella 5.5 stopped. the pump was explanted and replaced with another pump and support continued. | We are filing this supplemental report retrospectively to indicate a purge side arm crack leak that occurred leading up to the pump stop. both pump stop and purge side arm damage/leakage are reportable malfunctions. |  |  |  |  | The investigation into the impella pump stop is on-going at this time. no product has been returned yet for analysis. upon completion of the investigation should a root cause be determined, a final report will be filed. | Since the original report was filed, the investigation into the pump stop of the impella 5.5 has concluded. the team did not return the pump product for analysis. the root cause could not be definitively determined, however the team did note the purge sidearm was broken, which could have allowed for blood ingress which would cause the pump to stop. the failure mode will be monitored and trended. | (h3) the investigation into the reported 'pump stop and purge leak: pump" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the exact cause of the purge reservoir leak could not be determined. the cause of the pump stop is most likely due to blood ingress as a result of the low purge pressure. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
36634 | 1220648-2021-01164 | 12917267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12917267&pc=OZD | 11/22/2021 | 09/18/2024 |  | 12/02/2021 |  |  | 11/22/2021 | 12/02/2021 |  |  | 05/31/2023 |  | 12/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED INC. |  |  | US |  | 2022079421 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Break | Deformation Due to Compressive Stress | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2021 |  |  |  |  | A patient was admitted in cardiogenic shock and had a coronary angioplasty performed in (b)(6). the team observed the need for hemodynamic support and placed an impella cp. the complainant had the cp exhibit low flows, kink, and stop. when the team attempted to un-kink and snare the pump the pump was damaged and in 2 pieces. one piece remained in the patient and one piece was successfully explanted. the patient later expired with multiorgan failure. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The medical center in (b)(6) has not yet released and returned any portion of the impella pump, imaging, or data logs. the investigation is awaiting information and data from the field hospital at this time. should information or product be shared that leads to a root cause, a final report will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
36635 | 1220648-2021-01159 | 12895665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12895665&pc=OZD | 08/21/2020 | 11/30/2021 |  | 11/30/2021 |  |  | 11/05/2021 | 11/30/2021 |  |  | 01/31/2022 |  | 11/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1465638 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | Company representation |  |  | 87 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/05/2021 |  |  |  |  | The complainant reported an (b)(6) year old asian male patient had impella cp pump inserted in the left femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had access site bleeding that required blood transfusion. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36636 | 1220648-2021-01160 | 12895694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12895694&pc=OZD | 11/03/2021 | 11/30/2021 |  | 11/30/2021 |  |  | 11/03/2021 | 11/30/2021 |  |  | 05/31/2023 |  | 11/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 SMART ASSIST PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022085227 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/03/2021 |  |  |  |  | The complainant reported a (b)(6) white male patient had impella 5.5 pump inserted for post cardiac-cardiogenic shock (pccs). the physician was manipulating the device, position changed, physician noticed that there was a perforation on the lateral wall of the left ventricle, pump was removed, and repaired was done to the perforation. | The complainant reported a (b)(6) male patient had (b)(6) 5.5 pump inserted for post cardiac-cardiogenic shock (pccs). the physician was manipulating the device, position changed, physician noticed that there was a perforation on the lateral wall of the left ventricle, pump was removed, and repaired was done to the perforation. |  |  |  |  | Device not returned by customer. the impella 5.5 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed. | Device not returned by customer. the (b)(6) 5.5 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
36637 | 1220648-2021-01161 | 12895714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12895714&pc=OZD | 04/19/2021 | 11/30/2021 |  | 11/30/2021 |  |  | 11/05/2021 | 11/30/2021 |  |  | 12/31/2022 |  | 11/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022034967 |  | 0048-0034-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 79 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/05/2021 |  |  |  |  | The complainant reported a (b)(6) year old asian male patient had impella cp pump inserted in the left femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had access site bleed and blood transfusion was required. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36638 | 1220648-2021-01162 | 12895813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12895813&pc=OZD | 07/19/2021 | 11/30/2021 |  | 11/30/2021 |  |  | 11/09/2021 | 11/30/2021 |  |  | 10/31/2022 |  | 11/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022028383 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 77 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/09/2021 |  |  |  |  | The complainant reported a (b)(6) year old asian male patient had impella cp pump inserted in the right femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had increased hematuria and bilirubin and hemolysis was observed, as a result the pump was removed the next day and infusion was added. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36639 | 1220648-2021-01163 | 12895933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12895933&pc=OZD | 11/08/2021 | 11/30/2021 |  | 11/30/2021 |  |  | 11/08/2021 | 11/30/2021 |  |  | 06/30/2023 |  | 11/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2022103908 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/08/2021 |  |  |  |  | The complainant reported a (b)(6) year old white male patient had impella 5.0 pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had hemolysis and five (5) units of packed red bloods cells (prbcs) was administered. | [5 empty] | Device discarded by customer. the impella 5.0 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36640 | 1220648-2021-01158 | 12887847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12887847&pc=OZD | 11/02/2021 | 12/10/2021 |  | 11/29/2021 |  |  | 11/02/2021 | 11/29/2021 |  |  | 03/31/2023 |  | 11/29/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2022057519 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2021 |  |  |  |  | In (b)(6) an impella cp was placed for hemodynamic support of a patient admitted with cardiogenic shock. the pump supported the patient successfully for 74 hours when it was then explanted. after 5 months the case was reviewed for the jpvad registry and the team documented an access site bleed attributed to the use of impella. the team had infused blood products to treat the bleed. | [5 empty] | Since the original report was filed, the investigation into the reported access site bleeding has concluded. the root cause for the bleeding was determined to be due to the movement of the patient. the patient was reported to be moving throughout the support. the failure mode will be monitored and trended. | The medical team in (b)(6) discarded the impella product. no analysis of it is possible. additional details and clarification on the amount of blood product infused the patient has not been shared. due to covid, lack of access has hampered the investigation and analysis. should more information be shared which leads to a root cause, a supplemental report will be filed.
36641 | 1220648-2021-01155 | 12869385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12869385&pc=OZD | 10/28/2021 | 11/24/2021 |  | 11/24/2021 |  |  | 10/29/2021 | 11/24/2021 |  |  | 11/30/2022 |  | 11/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022023900 |  | 0048-0024 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Use of Device Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/29/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with ami/cgs. the impella cp optical was selected for support. the pump was successfully delivered via the supplied a 14fr introducer. when escalating the patient's care to high support, a perforation in the shared femoral artery was noted. the perforation was alleged as a result of the 14fr sheath being inserted without fluoroscopic guidance and handled from the back (against recommendation). the vessel was repaired with 2 stitches and the injury resolved. | [5 empty] | The device was discarded by the customer, therefore, a physical evaluation of the product cannot be completed. a supplemental mdr will be filed if new information is received.
36642 | 1220648-2021-01154 | 12869386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12869386&pc=OZD | 10/28/2021 | 11/24/2021 |  | 11/24/2021 |  |  | 10/28/2021 | 11/24/2021 |  |  | 07/31/2023 |  | 11/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022117761 |  | 0550-0006 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/28/2021 |  |  |  |  | The complainant reported a (b)(6) male patient develop aortic insufficiency during impella support. the device was removed and a centrimag device was placed. | [5 empty] | The device and data logs were returned for review. data logs found the pump had operated to specification. the pump was visually inspected and no abnormalities were observed. the pump could not be run, however, because it was returned with a cut catheter. the cause of the aortic insufficiency and mitral regurgitation was determined to be use related.
36643 | 1220648-2021-01156 | 12869388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12869388&pc=OZD | 08/24/2021 | 11/24/2021 |  | 11/24/2021 |  |  | 11/05/2021 | 11/24/2021 |  |  | 04/30/2023 |  | 11/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022081077 |  | 0048-0024-JP |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 11/05/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock. the impella cp optical was selected for hemodynamic support. during support, the patient developed an access site bleed that required a blood transfusion as treatment. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36644 | 1220648-2021-01157 | 12870635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12870635&pc=OZD | 11/01/2021 | 12/31/2021 |  | 11/24/2021 |  |  | 11/01/2021 | 11/24/2021 |  |  | 06/01/2023 |  | 11/24/2021 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA RP PUMP |  |  | OZD |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022105451 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2021 |  |  |  |  | The complainant reported a (b)(6) black or african american male had impella rp pump inserted in the right femoral vein (rfv). the patient had hemolysis and the pump was removed. | [5 empty] | The rt log review shows a spike in mc coinciding with an increase in dp and subsequently pap. the dp/pap remain elevated following the spike. this is evidence of biomaterial ingestion, the likely root cause of the suspected hemolysis. the pump was hemolysis tested and passed per fda astm testing standards. | The impella rp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36645 | 1220648-2021-01153 | 12860148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12860148&pc=OZD | 10/27/2021 | 11/22/2021 |  | 11/23/2021 |  |  | 10/27/2021 | 11/23/2021 |  |  | 06/30/2023 |  | 11/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022095359 |  | 0048-0024-JP |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/27/2021 |  |  |  |  | In (b)(6) 2021 the medical team in (b)(6) placed an impella cp for support of a patient suffering from an acute myocardial infarction. the impella supported the patient successfully for 120 hours when it was explanted. months later the jpvad registry documentation of the patient was reviewed and hemolysis was alleged. the hemolysis was diagnosed by the sign of urine color change. the medical team treated this presume hemolysis with fluid delivery and dialysis. | [5 empty] | The registry data was reviewed but no further clinical details have been shared regarding the presumed hemolysis. should more details be shared by the medical center, a final report with root cause shall be made.
36646 | 1220648-2021-01151 | 12848771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12848771&pc=OZD | 10/26/2021 | 11/19/2021 |  | 11/19/2021 |  |  | 10/26/2021 | 11/19/2021 |  |  | 10/31/2022 |  | 11/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021017062 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/26/2021 |  |  |  |  | The complainant reported a (b)(6) male patient enduring a vascular injury occurring at the impella access site during support. as treatment, vascular sutures were required. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36647 | 1220648-2021-01152 | 12848777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12848777&pc=OZD | 10/26/2021 | 11/19/2021 |  | 11/19/2021 |  |  | 10/26/2021 | 11/19/2021 |  |  | 10/31/2022 |  | 11/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021011309 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/26/2021 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock. approximately 2 days after explant, of the device, the patient endured an ischemic stroke. the physician could not rule out the impella as a possible casual/contributory factor. additionally, the patient suffered an access site bleed in relationship to the impella which prompted blood product delivery. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36648 | 1220648-2021-01149 | 12836080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12836080&pc=OZD | 10/25/2021 | 12/19/2024 |  | 11/18/2021 |  |  | 10/25/2021 | 11/18/2021 |  |  | 03/31/2023 |  | 11/18/2021 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP, OZD |  |  | IMPELLA RP PUMP |  |  | OZD |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2022081662 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/25/2021 |  |  |  |  | Additional placement signal was lost, was reported. | The complainant reported a 44 year old black or african american male patient had impella rp pump inserted for right ventricle failure (rvf) post left ventricular assist device (lvad). the patient had hemolysis and pump was removed. |  |  |  |  | No physical abnormalities were seen on the pump that could contribute to hemolysis, and it is not known how/when the hemolysis resolved. thus the hemolysis was likely due to patient's condition, and relation to impella is unknown. | The investigation for the reported hemolysis and optical signal issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that: optical signal issue: log review during support shows a hard failure of the placement signal, with the raw dp sensor value measuring around 5000 mmhg. the product was returned and the sensor resistance values were measured within specification. the product was run overnight, but the intermittent short was not reproduced in the lab and the ps remained within specification. hemolysis: no physical abnormalities were seen on the pump that could contribute to hemolysis, and it is not known how/when the hemolysis resolved. the root cause of the placement signal issue was an intermittent sensor short. the root cause of the hemolysis was patient condition. this report is being filed as part of a retrospective review of historical records. | The impella rp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36649 | 1220648-2021-01150 | 12839729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12839729&pc=OZD | 10/28/2021 | 10/20/2024 |  | 11/18/2021 |  |  | 10/28/2021 | 11/18/2021 |  |  | 08/31/2023 |  | 11/18/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022127879 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | Hospitalization | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 10/28/2021 |  |  |  |  | An impella cp was placed to support a patient suffering from many cardiac and multiorgan comorbidities. the cp was placed via the right femoral artery in the cardiac catheterization lab. due to the multi-vessel coronary disease he was referred for bypass surgery. when in the icu awaiting bypass surgery he suffered from a cerebral vascular event. the patient had a previously observed a thrombus on the impella during a daily echo. this thrombus had prompted the team to adjust the anticoagulation regimen. the presumption was that the thrombus could have embolized and cause the cerebral event. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | Since the original report was filed, the investigation has concluded. there was no product or data logs returned. the clinical report of the thrombus and cva were reviewed. the thrombus development is attributed to patient condition. the cva was deemed to be unknown/unrelated to the use of impella as no evidence or clinical imaging was provided. the failure mode will be monitored and trended. | The investigation into the cerebral event is on-going at this time. the impella product and further clinical details have not yet been shared. upon completion of the investigation, a final report will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2021-01150 was provided in sections b2, b5, h1, and h6. | H.6 code 4440 was inadvertently omitted from the initial manufacturer device report number 1220648-2021-01150 and was added.
36650 | 1220648-2021-01142 | 12828352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12828352&pc=OZD | 10/25/2021 | 09/18/2024 |  | 11/17/2021 |  |  | 10/25/2021 | 11/17/2021 |  |  | 07/31/2023 |  | 11/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2022126732 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 55 YR |  | Female |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/25/2021 |  |  |  |  | The complainant reported a (b)(6) female patient had impella 5.5 pump inserted in the right axillary for acute myocardial infarction (ami)/cardiogenic shock (cgs). the pump suddenly stopped, pump was unable to be restarted, patient was heavily dependent on the pump, and patient expired shortly afterwards. a large thrombus was seen on the inlet when the pump was removed independently. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | Review of the provided data logs found that only data from console sn (b)(6) was provided. these data logs showed the pump running for approximately 55 seconds on (b)(6) 2021 from 21:29:45 to 21:30:40. the pump appeared to function as intended during this time. no pump stops were captured in the provided data log. evaluation of the returned pump found evidence of biomaterial around the impeller and in the purge gap. the pump was run in the lab and ran without issue. the root cause of the pump stop was likely ingested biomaterial. | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2021-01142 was provided.
36651 | 1220648-2021-01143 | 12828370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12828370&pc=OZD | 10/25/2021 | 01/10/2022 |  | 11/17/2021 |  |  | 10/25/2021 | 11/17/2021 |  |  | 07/31/2023 |  | 11/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022118208 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 56 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2021 |  |  |  |  | The complainant reported a (b)(6) white male patient had impella 5.0 pump inserted in the right axillary for post cardiac-cardiogenic shock (pccs). the pump stopped, another console was used but the pump did not start, and another pump was used. | [5 empty] | The data logs were not returned for investigation. visual inspection of the returned pump revealed significant damage to the pins of the pump handle; the pins were almost completely sheared off. the pump was unable to be detected when connected to the aic. no other damage or abnormalities were found. in conclusion, it was determined that the cause of the pump stop was broken pins. | The impella 5.0 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36652 | 1220648-2021-01144 | 12828384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12828384&pc=OZD | 10/13/2021 | 09/24/2024 |  | 11/17/2021 |  |  | 10/28/2021 | 11/17/2021 |  |  | 07/31/2023 |  | 11/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, OZD |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022108793 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 59 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/28/2021 |  |  |  |  | The complainant reported a (b)(6) asian male patient had impella cp optical pump inserted in the right femoral. the patient had rhabdomyolysis due to lower limb ischemia that occurred on the impella side. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2021-01144 was provided in sections b2, b5, d6, h1, and h6.
36653 | 1220648-2021-01145 | 12830513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12830513&pc=OZD | 10/24/2021 | 11/17/2021 |  | 11/17/2021 |  |  | 10/24/2021 | 11/17/2021 |  |  | 01/31/2023 |  | 11/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2022039748 |  | 0048-0034-JP |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/24/2021 |  |  |  |  | In (b)(6) the medical team escalated a patient's care from an intra-aortic balloon pump (iabp) to the impella cp. the patient had been on the iabp for a week following coronary angioplasty. the patient's care was poor as she was intubated and suffering from renal and liver issues in addition to her cardiac disease. the impella cp was placed via the right femoral artery. after a day of support the impella limb was noted to be ischemic and so the team placed external bypass to relieve the limb. this allowed for pulses to be found via doppler. after 5 days of impella cp support the pump was explanted and care further escalated to the impella 5.0. | [5 empty] | The medical team in (b)(6) discarded the impella product and additionally, no data logs or imaging were returned for analysis into the reported limb ischemia. the root cause was determined to be patient condition. the patient's vessel size as well as vasopressor dose contributed to the limb ischemia. the failure mode will be monitored and trended.
36654 | 1220648-2021-01147 | 12830518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12830518&pc=OZD | 10/26/2021 | 10/23/2024 |  | 11/17/2021 |  |  | 10/26/2021 | 11/17/2021 |  |  | 10/31/2022 |  | 11/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2022028394 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Sepsis | Thrombocytopenia | [9 empty] | Hospitalization | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2021 |  |  |  |  | In (b)(6) a patient was admitted in (b)(6) 2021 suffering from cardiogenic shock. the team placed an impella cp for hemodynamic support when the team performed cardiac angioplasty. the pump remained on for over 4 days of successful support. months later the jpvad registry noted adverse events of hemolysis and thrombocytopenia with bleeding. blood products were infused to the patient to treat these adverse events. further details were not shared by team who completed the registry. | Additional information that the patient also had a fever and sepsis and were doing catecholamine-free. the cause of the fever is unknown. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The medical center in japan discarded the impella product. there was no investigation of the product or benchtop hemolysis testing possible. should more information be shared that leads to root cause, a final report will be filed. | The investigation for the reported hemolysis, access site bleeding, thrombocytopenia and sepsis issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2021-01147 was provided in sections b5, d6, and h6.
36655 | 1220648-2021-01146 | 12830519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12830519&pc=OZD | 10/25/2021 | 12/02/2021 |  | 11/17/2021 |  |  | 10/25/2021 | 11/17/2021 |  |  | 08/31/2023 |  | 11/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022126477 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2021 |  |  |  |  | A female patient was being placed on impella hemodynamic support after being turned down for surgery. an impella cp was placed via single access. while attempting to place the interventional terumo sheath, though the impella's 14fr sheath, the team applied force and caused a vessel perforation. the products were all removed and the vessel was repaired by a the placement of a covered stent. the team then placed a new impella via the contralateral leg. | [5 empty] | Since the original report was filed, the investigation into the vascular damage has completed. product was returned for analysis. visual inspection was made and no abnormalities were observed. the root cause of the vessel damage was a use/technique issue. the failure mode will be monitored and trended. | The investigation into the vessel damage is on-going at this time. upon completion of the investigation, a final report will be filed.
36656 | 1220648-2021-01148 | 12830525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12830525&pc=OZD | 11/08/2021 | 01/04/2022 |  | 11/17/2021 |  |  | 11/08/2021 | 11/17/2021 |  |  |  |  | 11/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED INC. |  |  | US |  |  |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature |  |  |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2021 |  |  |  |  | A publication was located that was inclusive of the impella cp and utilized in the (b)(6), at an unknown date prior to (b)(6) 2021. the publication noted that an impella cp was placed for hemodynamic support of a male patient awaiting bypass surgery (cabg) whose condition had unfortunately deteriorated prior to cabg surgery. the pump migrated out of optimal position and became entrapped within the mitral subvalvular apparatus. to remove and remedy this issue the team snared the pump with the 7fr ensare. the pump was explanted and the groin access site was closed. | [5 empty] | The literature was reviewed and european union field representative contacted for further information. when more details are determined, a final report with root cause will be made. | The investigation into the cannula kink and retrieval by snare has completed since the original report was filed. the investigation has found that the cause of the kink was a user issue. the failure mode will be monitored and trended.
36657 | 1220648-2021-01137 | 12819246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12819246&pc=OZD | 10/19/2021 | 10/24/2024 |  | 11/16/2021 |  |  | 10/19/2021 | 11/16/2021 |  |  | 05/31/2023 |  | 11/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022089522 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia Stroke | [9 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/19/2021 |  |  |  |  | The medical center in (b)(6) had an impella cp placed when care needed to be escalated from the intra-aortic balloon pump for a patient admitted in cardiogenic shock. the pump supported the patient for 2 days till explant. prior to explant the team had to medically treat the patient due to an access site bleed at the impella insertion site. the pump had not been fixated appropriately which allowed for blood loss. the team infused blood products and sutured to close the artery. | Additional information reported the patient experienced hemolysis which had no noted treatment. there was a hematoma at the impella insertion site, that was surgically removed along with blood vessel sutures placed. in addition, information reported that there was an ischemic stroke that occurred after the explant of the device. information was also provided that noted the patient expired due to the patient's hypocardiac function and the original cardiac dysfunction. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | Since the original report was filed, the investigation into the access site bleeding in japan has concluded. no product was returned for analysis. the clinical report from the medical center was reviewed. the cause of the access site was determined to be improper technique. the center had not fixed the pump at appropriate angulation, which allowed for the harm/injury. the failure mode will be monitored and trended. | The investigation into the access site bleed is on-going at this time. no product or additional clinical information has been shared by the medical center in (b)(6). upon completion of the investigation, a final report will be filed. | The investigation for the additional reported issues have been completed. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the issues could not be determined. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. additional information for the initial mfr 1220648-2021-01137 was provided in sections b5, d6a, d6b, d7a, g2, h3, h6 and h10. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2021-01137 and should not have been.
36658 | 1220648-2021-01138 | 12819287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12819287&pc=OZD | 10/19/2021 | 11/16/2021 |  | 11/16/2021 |  |  | 10/19/2021 | 11/16/2021 |  |  | 01/31/2022 |  | 11/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1474083 |  | 005066-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/19/2021 |  |  |  |  | A patient was admitted suffering from cardiomyopathy in (b)(6) had an impella 5.0 inserted to supplement the support already being provided by ecmo. the 5.0 support allowed the team to explant the ecmo. after 66 hours of support the team weaned and explanted the 5.0. approximately 1 month later the abiomed team was notified via the (b)(6) registry documents that there had been an access site bleed attributed to the use of impella that was treated by infusion of blood products. | [5 empty] | The medical center in (b)(6) discarded the impella pump product. further clinical details have not been provided. should new details be shared that lead to root cause, another report will be filed.
36659 | 1220648-2021-01140 | 12820136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12820136&pc=OZD | 10/19/2021 | 11/16/2021 |  | 11/16/2021 |  |  | 10/19/2021 | 11/16/2021 |  |  | 06/30/2023 |  | 11/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022100693 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/20/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock. upon of weaning/removal, the impella inserted limb was found to have an iliac perforation that lead to ischemia. a stent was urgently inserted as treatment. the patient was also suspected as enduring hemolysis during support, therefore, haptoglobin was administered and the issue resolved. | [5 empty] | The impella device was discarded, therefore, a physical evaluation of this product cannot be completed. should new information be obtained, a supplemental mdr will be filed.
36660 | 1220648-2021-01141 | 12824473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12824473&pc=OZD | 10/21/2021 | 10/11/2024 |  | 11/16/2021 |  |  | 10/21/2021 | 11/16/2021 |  |  | 07/31/2023 |  | 11/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022110311 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/21/2021 |  |  |  |  | The medical center placed an impella cp in a younger male patient who had been admitted with respiratory syncytial virus (rsv) and had progressed to being in cardiogenic shock with angina. the access was made at the right femoral artery. after approximately 30 minutes, the accessed limb began to show signs of ischemia, which prompted the team to place external bypass to relieve the cool foot. the pump remained on for hemodynamic support for 5.5 hours and was then explanted. | During support, lows were not higher than 1l/minute. the impella cp device was explanted. the patient survived the procedure. |  |  |  |  | Since the original report was filed, the investigation into the reported limb ischemia has concluded. the analysis determined the root cause of the ischemia was patient condition. the failure mode will be monitored and trended. | The investigation into the limb ischemia is on-going at this time. upon completion of the investigation, a final report will be filed. | The investigation has been completed for the reported issue of low pump flow. the device was received for evaluation. the root cause of the low pump flow was most likely a cannula kink. additional information for the initial mfr 1220648-2021-01141 was provided in sections b1, b5, e1, g1, h1, and h6. this report is being filed as part of a retrospective review of historical records.
36661 | MW5105394 | 12829763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12829763&pc=OZD | 11/13/2021 | 11/15/2021 |  | 11/16/2021 |  |  |  | 11/16/2021 |  |  |  |  | 11/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | UNK |  |  | ABIOMED, INC. |  |  |  |  | 318049 |  | UNK |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 68 YR |  | Male | 123 KG | Non Hispanic | White | Injury | No | Yes | Backflow | Energy Output Problem | [5 empty] | Failure of Implant | Low Cardiac Output | [11 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | Impella cp device began to alarm "motor current to high", "retrograde flow", and "loss of power". he was returned to the cath lab for a successful exchange of impella cp via right femoral artery. clinical support and local rep were made aware. fda safety report id# (b)(4).
36662 | 1220648-2021-01132 | 12811384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12811384&pc=OZD | 10/16/2021 | 12/13/2024 |  | 11/15/2021 |  |  | 10/19/2021 | 11/15/2021 |  |  | 03/31/2023 |  | 11/15/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022059109 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 20 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/19/2021 |  |  |  |  | The complainant reported a (b)(6) year old asian male patient had impella 5.0 pump inserted in the right femoral for cardiomyopathy. the patient had thrombus that required thrombectomy procedure with fogarty at impella removal. | It was reported the patient experienced hemolysis on (b)(6) 2021. no further information was provided on treatments given to the patient for the hemolysis. it was additionally reported that the pump had low purge pressure alarms during the case on (b)(6) 2021. the purge cassette was changed, but this did not resolve the issue. due to the ongoing purge issues, the pump was replaced. |  |  |  |  | The pump was returned and biomaterial was observed. | Device evaluated by mfr. device was not returned by customer. the impella 5.0 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. | Additional information provided in b1, b2, b5, h1, and h6 for reported low purge pressure issues and hemolysis. this report is being filed as part of a retrospective review of historical records.
36663 | 1220648-2021-01133 | 12813765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12813765&pc=OZD | 09/28/2021 | 05/31/2022 |  | 11/15/2021 |  |  | 10/20/2021 | 11/15/2021 |  |  | 05/31/2023 |  | 11/15/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022083711 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2021 |  |  |  |  | New information was obtained on 2022-05-02 indicating the patient had developed hemolysis during impella support. haptoglobin was delivered as treatment. | The complainant reported a (b)(6) asian male patient had impella cp optical pump inserted. the patient had a thrombus, impella was removed, blood flow in the lower limbs was poor, so thrombectomy with fogarty catheter was performed. |  |  |  |  | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. b(5): new information received and added to narrative. | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36664 | 1220648-2021-01134 | 12813954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12813954&pc=OZD | 10/20/2021 | 11/15/2021 |  | 11/15/2021 |  |  | 10/20/2021 | 11/15/2021 |  |  | 06/30/2023 |  | 11/15/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022108071 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/20/2021 |  |  |  |  | The complainant reported a (b)(6) white male patient had impella cp pump inserted percutaneously in the right axillary artery for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had thrombus and a fogarty balloon was used. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36665 | 1220648-2021-01135 | 12814409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12814409&pc=OZD | 06/18/2020 | 11/15/2021 |  | 11/15/2021 |  |  | 10/21/2021 | 11/15/2021 |  |  | 09/30/2021 |  | 11/15/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1444438 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Consumer | Health Professional | User facility | Company representation | 33 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/21/2021 |  |  |  |  | We received a publication from journal of surgical case reports (jscr), 2021;10, 1-3 that a (b)(6) old asian male patient had impella 5.0 pump inserted in the right femoral artery for familial dilated cardiomyopathy. the aortic insufficiency occurred after impella insertion, requiring extra surgical intervention twice; heartmate iii implantation was performed two (2) weeks after and aortic valve (av) replacement was performed. | [5 empty] | Device discarded by customer. the impella 5.0 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36666 | 1220648-2021-01136 | 12814437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12814437&pc=PYX | 10/19/2021 | 11/24/2021 |  | 11/15/2021 |  |  | 10/19/2021 | 11/15/2021 |  |  | 10/31/2021 |  | 11/15/2021 |  | TEMPORARY NON-ROLLER TYPE RIGHT SIDED HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED INC. |  |  | US |  | 1457586 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/19/2021 |  |  |  |  | There was a patient admitted with a pulmonary embolism after a hip surgery. due to a need for hemodynamic support due to right sided shock, and impella rp was placed via the right femoral vein. the team observed an access site bleed at the rp insertion. the bleed was due to a dissection. the team made the choice to abort the rp and placed the patient on ecmo support and repaired the venous dissection and gave the patient 2 units of blood products. | [5 empty] | The medical center did not return any of the impella rp product or sheath. the root cause for the access site bleeding could not be determined. the failure mode will be monitored and trended. | The investigation into the access site injury is on-going at this time. upon completion of the investigation, a final report will be filed.
36667 | 1220648-2021-01130 | 12802838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12802838&pc=OZD | 10/14/2021 | 10/15/2024 |  | 11/12/2021 |  |  | 10/14/2021 | 11/12/2021 |  |  | 06/30/2023 |  | 11/12/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022100411 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Restricted Flow rate | Pumping Stopped | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/14/2021 |  |  |  |  | A patient was admitted in cardiogenic shock in (b)(6) and when his care needed escalation the intra-aortic balloon pump was replaced by an impella cp. during the hemodynamic support with the impella suction alarms were noted and after reviewing his medical history of atrial fibrillation, the team assumed there was thrombus within the impella pump causing the suction and hemolysis. the pump was explanted and the hemolysis (diagnosed by urine color change) was treated by infusion of haptoglobin. | Additional information was reported indicating there were low pump flows and a pump stop. it was reported upon removal of the pump that a clot was observed in the bowl on the sterile table. |  |  |  |  | Since the original report was filed, the investigation into the reported hemolysis has concluded. the pump product was returned with biomaterial wrapped around the impeller. the data logs also show a motor current spike indicative of biomaterial ingestion. the root cause of the hemolysis was biomaterial ingestion, which is native patient condition related. the failure mode will be monitored and trended. | The analysis of the presumed hemolysis caused by impella use is on-going at this time. upon completion of the investigation, a final report will be filed. | The cause of the pump stop was most likely ingested biomaterial. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2021-01130 was provided in sections b5, d6, and h6.
36668 | 1220648-2021-01131 | 12803795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12803795&pc=OZD | 10/11/2021 | 12/29/2021 |  | 11/12/2021 |  |  | 10/22/2021 | 11/12/2021 |  |  | 06/30/2023 |  | 11/12/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022100665 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 20 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2021 |  |  |  |  | The complainant reported a (b)(6) asian male patient had impella cp pump inserted in the right femoral for cardiomyopathy. when the pump was removed thrombus was seen and a fogati catheter was used and the patient had suspected hemolysis and the pump was removed. | [5 empty] | An image received from field confirmed there were several significant masses of thrombus removed from the patient upon explant. the returned data logs showed a brief ventricular drop in placement signal (ps) with a brief "impella position unknown" alarm that occurred around the reported time of hemolysis, indicating attempts at repositioning. visual inspection of the returned pump revealed a considerable amount of biomaterial wrapped around the impeller. when the pump was run, its performance was within specification. the pump passed hemolysis testing per astm f841. in conclusion, it was determined that the cause of the hemolysis was ingested biomaterial. | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36669 | 1035166-2021-00136 | 12787913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12787913&pc=DYB | 10/19/2021 | 04/15/2022 |  | 11/10/2021 |  |  | 10/28/2021 | 11/10/2021 |  |  | 10/01/2023 |  | 11/10/2021 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-19434 |  | 0052-3021 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 74 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a (b)(6) male underwent an acute myocardial infarction/cardiogenic shock procedure. the hospital staff reported that, during attempted placement of the introducer, the device tore "near the distal end of the dilator." a new 23 fr introducer from another kit was used without further issue. there was no harm to the patient reported. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. | Device was used in treatment. the device was discarded and not returned for analysis, therefore, the exact cause of the product issue cannot be determined and the clinical observation could not be confirmed. however, following controls are in place to mitigate the reported product issue. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure (abiomed introducer sheath in-process and final inspections): measure dimensions: measure tip i.d. of sheath using appropriate pin gauges. visual inspection: using 10x magnification, verify the tip is round, no flash protruding and no flash with width around circumference, no splitting, cracks or other damages. per visual inspection, with naked eye, verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. per ifu: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. abiomed provides the ifu for this product. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Additional information h2: added codes 3331, 4109, 4110. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36670 | 1035166-2021-00134 | 12790701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12790701&pc=DYB | 10/25/2021 | 03/25/2022 |  | 11/10/2021 |  |  | 10/26/2021 | 11/10/2021 |  |  | 10/01/2023 |  | 11/10/2021 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-19434 |  | 0052-3021 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a (b)(6) year old male underwent a procedure to access the right femoral vein for rp placement. after doing serial dilation with the rp introducer kit before loading 23 fr dilator and sheath onto 035 super stiff wire md and field rep. noticed a crack in 2 spots that made a sharp edge that could potentially tear the vein, decided to use another rp introducer kit. no patient harm reported. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. | The following sections were updated in follow-up 1. one 23f abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. blood was found on and inside the sheath. upon evaluation of the returned product, it was found that the sheath tube was split in two places. the sheath was split on the score line in two places; one was located approximately 3.0mm from the distal tip and the other was split approximately 2mm from the distal tip. the tip has damage around the outer circumference. blood residue was found on the sheath. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. returned device analysis revealed the sheath was within manufacturing specifications. potential contributing factors to the damage (splits in tip) could be if the sheath was advanced through an area of resistance or if the sheath was subjected to unusual stresses such as tortuous patient anatomy. another possibility is that the sheath was bumped or dropped prior to use. according to the device history record, the introducer sheath passed all in-process and qa final inspection steps before shipping to the customer including visual, dimensional and leak testing. no manufacturing defects were found. per procedure (abiomed introducer sheath in-process and final inspections): measure tip i.d. of sheath using appropriate pin gauges. visual inspection: using 10x magnification, verify the tip is round, no flash protruding, and no flash with width (around circumference), no splitting, cracks or other damages. per visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. per ifu never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. abiomed provides the ifu for this product. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36671 | 1035166-2021-00130 | 12782603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12782603&pc=DYB | 10/19/2021 | 04/11/2022 |  | 11/09/2021 |  |  | 10/19/2021 | 11/09/2021 |  |  | 04/01/2022 |  | 11/09/2021 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-16187 |  | 0052-3021 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 72 YR |  | Male |  |  |  | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported patient was presented to the physician, suffering from a pe post hip surgery. it was attempted to place rp for support. significant bleeding around sheath/ hematoma noted at site. contrast injected through r4 diagnostic catheter inserted in sheath. vessel tear/ dissection noted on fluoro. patient transferred to hybrid room. ecmo cannulation on left to support patient. surgical vessel repair with patch for venous tear. the patient was given 2 units of blood. angiographically the vessel did not appear calcified. there was some tortuosity on the right vein. hemostasis was achieved with surgical patch repair. no additional information is available. | [5 empty] | The following sections were updated in follow-up 1: the device was used for treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. the investigation was focused on a review of product documentation. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. however, following controls are in place to mitigate the reported product issue. per qa procedure adelante s2s introducer sheath in-process and final inspection: sample size: 100% inspection 12.3 vacuum leak test - perform pressure and vacuum leak test per procedure. per instructions for use (ifu): when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete. .
36672 | 1220648-2021-01129 | 12782907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12782907&pc=OZD | 10/12/2021 | 11/23/2021 |  | 11/09/2021 |  |  | 10/12/2021 | 11/09/2021 |  |  | 06/30/2023 |  | 11/09/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022096630 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/12/2021 |  |  |  |  | A patient with cardiomyopathy had his hemodyamic support escalated from an intra-aortic balloon pump (iabp) to the impella 5.5 as he awaited and bridged to an lvad device. after 3 days of impella support the pump stopped and was unable to be restarted. the patient had pulled back his impella pump while under-sedated and so the pump stopped and was replaced. | [5 empty] | Since the original report was filed, the investigation into the pump stop has concluded. the 5.5 pump's data logs were reviewed. they showed the motor current and placement signal changes. the clinical details revealed that the patient had pulled upon the pump. the root cause for the stop was an open line caused by the patient pulling on the pump. the failure mode will be monitored and trended. | The investigation into the pump stoppage is on-going at this time. upon completion of the investigation, a final report will be filed.
36673 | 1220648-2021-01126 | 12783521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12783521&pc=OZD | 10/20/2021 | 11/09/2021 |  | 11/09/2021 |  |  | 10/20/2021 | 11/09/2021 |  |  | 07/31/2023 |  | 11/09/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022121810 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/20/2021 |  |  |  |  | The complainant reported a (b)(6) white female patient had impella cp pump inserted in the right femoral artery (rfa). | [5 empty] | The returned product was visually inspected and both impeller blades were fractured. one of the fractured blades was wedged between the outflow cage and the impeller. because fluroscopic imaging of the pump/additional evidence was not returned, a definitive cause of the fracture could not be determined.
36674 | 1220648-2021-01127 | 12783523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12783523&pc=OZD | 10/12/2021 | 11/09/2021 |  | 11/09/2021 |  |  | 10/12/2021 | 11/09/2021 |  |  | 09/30/2022 |  | 11/09/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2021006157 |  | 0048-0034-JP |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/12/2021 |  |  |  |  | In (b)(6) an impella cp was placed for hemodynamic support of the patient in cardiogenic shock. the patient was critical suffering from ventricular tachycardia and fibrillation requiring cardioversion. the impella cp was placed and supported the patient for 2 days til the pump was weaned and explanted. months later in a registry document, the abiomed team was informed of an access site bleed from the impella insertion that was treated by infusion of blood products. | [5 empty] | The investigation into the access site bleeding is on-going at this time. upon completion of the investigation, if root cause can be determined, a final report will be filed. to date no analysis has occurred as the pump product was discarded, logs were not returned, and clinical details are lacking.
36675 | 1220648-2021-01128 | 12783527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12783527&pc=OZD | 10/13/2021 | 11/09/2021 |  | 11/09/2021 |  |  | 10/13/2021 | 11/09/2021 |  |  | 08/31/2022 |  | 11/09/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 1496671 |  | 0048-0034-JP |  | No |  |  | Not Available |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Hospitalization |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/13/2021 |  |  |  |  | The medical center in japan placed an impella cp to support a patient suffering from cardiogenic shock in the cardiac catheterization lab. the team also had ecmo placed for support. the team made the observation that the vascular path of the patient was complicated due to peripheral arterial disease and tortuosity. the impella placement was noted to be the cause of a vessel damage. the vessel damage was treated by infusion of blood products. | [5 empty] | The investigation into the access site bleeding and presumed vascular damage from the impella access is on-going at this time. the medical center in japan has not returned the product nor answered all clinical questions. should more information be shared that leads to a root cause, a final report will be filed.
36676 | 1220648-2021-01125 | 12783530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12783530&pc=OZD | 07/29/2021 | 11/09/2021 |  | 11/09/2021 |  |  | 10/12/2021 | 11/09/2021 |  |  | 10/31/2022 |  | 11/09/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021014938 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/19/2021 |  |  |  |  | The complainant reported a 46 year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp device. during support, the patient developed an access site bleed that prompted a blood transfusion. | [5 empty] | The impella device was discarded, therefore, an investigation could not be completed on the physical device. should we receive the device, or new information becomes available on this event, a supplemental mdr will be filed.
36677 | 1220648-2021-01117 | 12751939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12751939&pc=OZD | 10/08/2021 | 11/17/2021 |  | 11/04/2021 |  |  | 10/08/2021 | 11/04/2021 |  |  | 06/30/2023 |  | 11/04/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022095406 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/08/2021 |  |  |  |  | In (b)(6) a patient was admitted suffering from cardiogenic shock. the medical team placed an impella cp via the femoral artery for hemodynamic support for a planned repeat bypass surgery. the leg accessed by the impella became ischemic. the team removed thrombus via the use of fogarty balloon. the thrombus was noted to have embolized to the common femoral and superficial femoral arteries. | [5 empty] | The medical center in japan did not return the product for investigation. the root cause of the limb ischemia was determined to patient condition. vessel size and condition contributed to the ischemia. the failure mode will be monitored and trended. | The investigation into the reported limb ischemia is on-going at this time. no product was returned from the (b)(4) medical center for analysis. when the investigation is completed, a final report will be filed.
36678 | 1220648-2021-01118 | 12753340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12753340&pc=OZD | 10/08/2021 | 12/03/2021 |  | 11/04/2021 |  |  | 10/08/2021 | 11/04/2021 |  |  | 08/31/2023 |  | 11/04/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | UNKNOWN |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/08/2021 |  |  |  |  | A patient was admitted for a high risk coronary angioplasty. the team placed an impella cp via the femoral artery with a plan to use the site for single access for the angioplasty devices and impella pump. the impella sheath kinked at the attempted single access. when the kink occurred, the team placed the angioplasty sheath in the opposite leg/femoral. the anatomy of both iliacs and femorals was observed to be aneurysmal. later the team observed the impella sheath had caused a perforation that was treated by balloon angioplasty and placement of a covered stent. | [5 empty] | Since the original report was filed, the investigation has concluded. the introducer was returned for analysis. the sheath was kinked due to the native patient condition, of diseased iliac and femoral artery. the damage to the vessel was due to the removal of the kinked introducer. the failure mode will be monitored and trended. | The investigation into the vessel damage is on-going at this time. upon completion of the investigation, a final report will be filed.
36679 | 1220648-2021-01123 | 12753779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12753779&pc=OZD | 09/30/2021 | 10/14/2024 |  | 11/04/2021 |  |  | 10/10/2021 | 11/04/2021 |  |  | 12/31/2022 |  | 11/04/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2022042950 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/10/2021 |  |  |  |  | An impella cp was placed for hemodynamic support of a patient admitted in (b)(6) with cardiogenic shock. the cp was supporting the patient when there were attempts made at repositioning the pump, within the left ventricle, due to less than optimal placement along the papillary muscle. there were concerns of hemolysis and the pump was explanted prematurely after 12.5 hours. | The impella device was explanted due to hematoma. the patient survived the procedure. |  |  |  |  | The medical center in (b)(4) has not returned the impella product for analysis. the impella pump was discarded. no data logs have been returned for analysis either. the investigation of the clinical report shared with abiomed lead to a root cause of improper pump positioning by the medical team. the failure mode will be monitored and trended. | The investigation has been completed for the reported issue of access site bleeding - major. the device was not received for evaluation. the cause of the access site bleeding - major was patient condition and the relation to impella was determined to be unknown/unrelated. additional information for the initial mfr 1220648-2021-01123 was provided in sections a5, b1, b3, b5, d6a, d6b, g1, h1, h6.
36680 | 1220648-2021-01120 | 12753784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12753784&pc=OZD | 10/07/2021 | 12/15/2021 |  | 11/04/2021 |  |  | 10/07/2021 | 11/04/2021 |  |  | 02/28/2022 |  | 11/04/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1475720 |  | 005042-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 10/07/2021 |  |  |  |  | The complainant reported a (b)(6) female patient presenting with myocarditis for hemodynamic support using the impella 2.5 device. the pump was delivered successfully, however, bubbles were seen in the left ventricle and aorta during use. the pump was removed and replaced, however, the patient reportedly endured an ischemic stroke the following day. | [5 empty] | The pump was returned for analysis. visual inspection found no abnormalities or damage. simulated use testing found the pump operating as expected. we are unable to definitively determine the cause of the reported issue. | The device has been received and an investigation is underway. upon completion of our analysis, a supplemental mdr will be submitted.
36681 | 1220648-2021-01119 | 12753801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12753801&pc=OZD | 10/09/2021 | 11/04/2021 |  | 11/04/2021 |  |  | 10/09/2021 | 11/04/2021 |  |  | 07/31/2022 |  | 11/04/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN | [7 empty] | 2021019617 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 37 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male patient had impella 5.0 pump inserted in the right axillary for cardiomyopathy. the patient had cerebral infarction after replacement of impella 5.0 pump and there was no treatment. the patient is relatively prone to blood clots. | [5 empty] | The impella 5.0 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
36682 | 1220648-2021-01121 | 12753807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12753807&pc=OZD | 10/07/2021 | 11/04/2021 |  | 11/04/2021 |  |  | 10/07/2021 | 11/04/2021 |  |  | 07/31/2023 |  | 11/04/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP | [7 empty] | 2022108818 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/08/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock. the impella cp optical device was selected for hemodynamic support and inserted without issue. during support, patient movement was noted and the impella's access site developed a bleed. a suture was applied to stop the bleeding and blood products were delivered. | [5 empty] | The impella device was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36683 | 1220648-2021-01122 | 12753810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12753810&pc=OZD | 10/09/2021 | 12/09/2021 |  | 11/04/2021 |  |  | 10/09/2021 | 11/04/2021 |  |  | 07/31/2023 |  | 11/04/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022126626 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Movement Disorder | [11 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 10/09/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with cardiomyopathy for hemodynamic support using the impella 5.5 device. during support, the patient developed a pocket hematoma that required evacuation to resolve. additionally, the patient lost grip and feeling in their hand attached to the impella inserted limb that required premature removal of the device as intervention. after removing the device, the neurological impairment was resolved. | [5 empty] | The device and data logs were returned for analysis. review of the data logs found no notable alarms. a physical examination of the device found no damage that would have caused/contributed to the reported issues. as such, we are unable to determine a root cause for the reported hematoma and neurological event. | The impella device was received and an investigation is underway. upon completion of our evaluation, a supplemental mdr will be filed.
36684 | 1220648-2021-01124 | 12753820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12753820&pc=OZD | 10/15/2021 | 11/04/2021 |  | 11/04/2021 |  |  | 10/15/2021 | 11/04/2021 |  |  | 02/28/2023 |  | 11/04/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET | [7 empty] | 2022056182 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/15/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient had impella cp pump inserted for high risk percutaneous coronary intervention (hrpci). the patient had oozing at the impella access site, manual pressure was held but did not completely resolve, so three (3) units of packed red blood cells (prbcs) was administered. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36685 | 1220648-2021-01116 | 12739585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12739585&pc=OZD | 09/23/2021 | 10/13/2023 |  | 11/02/2021 |  |  | 10/12/2021 | 11/02/2021 |  |  | 04/30/2023 |  | 11/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022076909 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/12/2021 |  |  |  |  | The complainant reported a (b)(6) white male patient had impella 5.5 pump inserted in the right axillary for cardiomyopathy. the patient¿s access site was opened by a vascular surgeon to remove clots. | [5 empty] | The investigation into the reported purge leak and access site bleeding has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the root cause of the leak in the purge sidearm was most likely damage to the yellow luer. furthermore, the root cause of the access site hematoma was unable to be determined. section d4: unique identifier (udi) #. | The impella 5.5 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
36686 | 1220648-2021-01115 | 12725523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12725523&pc=OZD | 10/02/2021 | 10/19/2024 |  | 10/30/2021 |  |  | 10/02/2021 | 10/30/2021 |  |  | 07/31/2023 |  | 10/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022111881 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  |  |  | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/02/2021 |  |  |  |  | Additional information reported that the patient developed suction issues. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | An impella cp was placed for hemodynamic support for an obese patient (bmi of 40) suffering from cardiogenic shock. the cp was placed via the femoral artery. the patient developed a hematoma at the access site, and suction was alarming on the console. the team treated the bleed with infusion of red blood cells, approximately 8 units. the thought was to replace the pump and insert a longer sheath, that may have had a better ability to reach the arteriotomy and reduce the access site bleed, but the patient expired. the patient's obesity was thought to be the cause of the bleed as per the physician, as the tissue track was too long to achieve successful hemostasis. |  |  |  |  | Since the original report was filed, the investigation has closed. the team followed best practice, but due the large abdominal pannus the impella pump's repositioning sheath was unable to reach the vessel and prevent the blood loss from the arteriotomy. the cause was patient condition. the failure mode will be monitored and trended. | Additional information for the initial mfr 1220648-2021-01115 was provided. this report is being filed as part of a retrospective review of historical records. information provided in f6 and f8 were previously reported in error. | The investigation into the access site bleeding is on-going at this time. no impella product has been returned and clinical details are being clarified. upon completion of the investigation, a final report will be filed.
36687 | 1220648-2021-01110 | 12711332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12711332&pc=OZD | 10/04/2021 | 11/01/2021 |  | 10/28/2021 |  |  | 10/04/2021 | 10/28/2021 |  |  | 07/31/2023 |  | 10/28/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP | [7 empty] | 2022110914 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2021 | * | Invalid/Incomplete |  |  | A patient in (b)(6) was admitted in cardiogenic shock and had an impella cp placed for hemodynamic support. the support was successful for over 40 hours when the pump was explanted due to concerns of blood loss at the impella access site. the bleed was treated by infusion of blood products. at the pump explant the patient was stable and required no other support. | [5 empty] | Since the original report was filed, the investigation has concluded. no product or logs were returned from japan for analysis into the access site bleeding. the clinical details were reviewed. the act was only 200 seconds at the time of access. additionally the team did not do proper serial dilation or gain access with echo imaging. the root cause of the bleeding was improper user technique. the failure mode will be monitored and trended. | The medical center in (b)(6) discarded the impella product. no investigation of the product is possible. should more details be shared that lead to root cause, a final report will be filed.
36688 | 1220648-2021-01111 | 12711334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12711334&pc=OZD | 10/01/2021 | 11/08/2021 |  | 10/28/2021 |  |  | 10/01/2021 | 10/28/2021 |  |  | 06/30/2023 |  | 10/28/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022092835 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/01/2021 |  |  |  |  | An impella 5.5 was placed to support a surgical patient undergoing coronary bypass and maze procedures. after 4 days of support the team was attempting to explant the pump and the pump detached at the cannula. they were able to clamp it off at the cannula and remove the parts of the pump manually. the team believed that perhaps the graft sutures may have not been entirely removed and with excessive force, at the area of suture, the detachment occurred. after all pieces of the impella were removed the team was able to surgically close the graft. | [5 empty] | Since the original report was filed, the investigation into the pump damage and detachment has concluded. the pump was returned for analysis and clinical details were reviewed. it was stated in the clinical notes that the physician had sutures stuck on the pump. during explant there was excessive force applied on the pump, since the pump had sutures on it, this lead to pump being damage. the failure mode will be monitored and trended. | The investigation into the product malfunction is on-going at this time. upon completion of the investigation, a final report will be filed.
36689 | 1220648-2021-01112 | 12711337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12711337&pc=OZD | 09/29/2021 | 10/28/2021 |  | 10/28/2021 |  |  | 10/06/2021 | 10/28/2021 |  |  | 02/28/2023 |  | 10/28/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP | [7 empty] | 2022067916 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/06/2021 | * | Invalid/Incomplete |  |  | The complainant reported an asian patient had impella cp optical pump inserted in the right femoral for acute myocardial infarction (ami)-cardiogenic shock (cgs). when the peel-away introducer 14fr x 25cm was peeled off, it was not completely removed from the outside of the body, the blood vessel was damaged and bleeding from the puncture site occurred and blood transfusion of 4 (four) units of red blood cells (rbc) was administered. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36690 | 1220648-2021-01113 | 12711344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12711344&pc=OZD | 10/06/2021 | 12/17/2021 |  | 10/28/2021 |  |  | 10/06/2021 | 10/28/2021 |  |  | 07/31/2023 |  | 10/28/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC |  |  | US |  | 2022116529 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/06/2021 |  |  |  |  | The complainant reported a (b)(6) white female patient had impella cp pump inserted for high risk percutaneous coronary intervention (hrpci). the pump was pulled out of the ventricle and folded on itself in the aorta and a snare was used to grab the pigtail of the pump to pull it straight to remove the pump. | [5 empty] | Evaluation of the returned pump found two kinks in the cannula. the provided data logs showed that the pump was improperly positioned on (b)(6) 2021, first in the ventricle, as indicated by the ventricular placement signal waveform, flattened motor current waveform, and 'impella position in ventricle' alarm, and then in the aorta, as indicated by the aortic placement signal waveform and flattened motor current waveform. the cause of the cannula kink was improper technique. per ifu 0048-9007, if the impella catheter is completely out of the ventricle, repositioning the device across the aortic valve should not be attempted without a guidewire. | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36691 | 1220648-2021-01114 | 12711345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12711345&pc=OZD | 10/08/2021 | 10/28/2021 |  | 10/28/2021 |  |  | 10/08/2021 | 10/28/2021 |  |  | 02/28/2023 |  | 10/28/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN | [7 empty] | 2022053040 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/08/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) female patient had impella 5.0 pump inserted in the right axillary for myocarditis. the patient had access site bleeding and thrombocytopenia that required an unknown amount of blood transfusion. | [5 empty] | Device evaluated by mfr. device discarded by customer. the impella 5.0 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36692 | 1220648-2021-01109 | 12702677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12702677&pc=OZD | 10/19/2021 | 11/02/2021 |  | 10/27/2021 |  |  | 10/19/2021 | 10/27/2021 |  |  | 12/19/2020 |  | 10/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN | [7 empty] | 2021017106 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/19/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that during impella support vascular injury was noted. as a result, a thrombectomy was performed. no further information was provided. | [5 empty] | Since the original report was filed, the investigation has concluded. the root cause of the patient injury of limb ischemia could not be determined. the clinical report was lacking information. the medical center in japan did not furnish abiomed with clinical details. the failure mode will be monitored and trended. | The impella cp was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
36693 | 1220648-2021-01107 | 12696286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12696286&pc=OZD | 07/31/2021 | 10/17/2025 |  | 10/26/2021 |  |  | 10/08/2021 | 10/26/2021 |  |  | 06/30/2023 |  | 10/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095365 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 70 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Renal Failure | Thrombocytopenia | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | Followup |  | 10/08/2021 |  |  |  |  | The complainant reported a (b)(6) year-old male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported the patient developed limb ischemia during impella support. as a result, an fa bypass was performed to treat limb ischemia. subsequently, access site bleed was noted. as a result, an unknown amount of blood products was administered. no further information was provided. | Additional information reported the patient suffered from renal failure and thrombocytopenia. it is unknown how hemostasis was achieved. survival outcome at explant noted the patient expired and no causal relationship with the impella. the heart did not stand up well, leading to multiple organ failure. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | Since the original report was filed, the investigation into the limb ischemia, that required femoral bypass, and access site bleed has concluded. there was no product returned for analysis from japan. only clinical details were shared for investigation. the cause of the limb ischemia was patient condition. without product return, the cause of the access site bleed could not be determined. the failure modes will be monitored and trended. | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information for the initial mfr 1220648-2021-01107 was provided. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
36694 | 1220648-2021-01108 | 12696348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12696348&pc=OZD | 10/08/2021 | 10/09/2024 |  | 10/26/2021 |  |  | 10/08/2021 | 10/26/2021 |  |  | 06/30/2023 |  | 10/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022100407 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Skin Infection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 10/08/2021 |  |  |  |  | The complainant reported a patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). patient¿s age and gender were not provided. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported the patient developed limb ischemia during impella support. as a result, an fa bypass was performed to treat limb ischemia. no further information was provided. | The complainant reported a patient, ethnicity unknown, was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis, and an infection. the hemolysis was suspected due to the patient¿s urine color and was resolved with repositioning the impella. the infection was a skin ulcer at the insertion site and occurred after the planned removal of the impella on (b)(6) 2021, this alleged event occurred on 8/27/2021. the ulcer was reported to have occurred because the insertion site was stopped with a band. additionally, the ulcer reportedly occurred in connection with the management of the facility. |  |  |  |  | The impella cp has not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the hemolysis, limb ischemia and infection have been completed. hemolysis was reported on the second day of impella cp optical use, the hemolysis was suspected due to urine color. the pump was repositioned, and the suspected hemolysis resolved. leg ischemia was also reported in the leg that the cp was inserted. an ulcer at the insertion site was reported 16 days after the removal of the pump, the cause of the patient injury/infection was not determined the abiomed device is sterilized under validated conditions and the clinical details provided were insufficient to determine a root cause. the cause of the hemolysis was improper positioning resulting for improper technique by the end user. the cause of the limb ischemia was most likely the patient¿s condition of diseased or small patient vessels. the cause of the infection was not determined, as clinical details provided were insufficient to determine a root cause. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: hemolysis impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ infection/sepsis impella 5.5 with smart assist for use during cardiogenic shock & cp with smart assist section: table 3.2 impella catheter components as noted in the impella ifu ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ limb ischemia impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation. this report is being filed as part of a retrospective review of historical records.
36695 | 1220648-2021-01100 | 12689493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12689493&pc=OZD | 04/02/2021 | 10/25/2021 |  | 10/25/2021 |  |  | 10/04/2021 | 10/25/2021 |  |  | 10/31/2022 |  | 10/25/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN | [7 empty] | 2021020000 |  | 0048-0034-JP |  | No |  |  | Not Available |  | Yes | Invalid/Incomplete | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/04/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed access site bleed during impella support. as a result, large amounts of blood products were administered, amount unknown and the device was explanted. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36696 | 1220648-2021-01102 | 12689503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12689503&pc=OZD | 10/05/2021 | 10/25/2021 |  | 10/25/2021 |  |  | 10/05/2021 | 10/25/2021 |  |  | 02/28/2023 |  | 10/25/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP | [7 empty] | 2022070760 |  | 0048-0024-JP |  | No |  |  | Not Available |  | Yes | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 51 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/05/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old male patient had impella cp pump inserted in the right femoral for an unknown indication. it was reported that when removing the impella, the fixing plug of the indwelling sheath and the fixing wing (blue) were strongly pulled and damaged. there was bleeding from the indwelling sheath. surgical hemostasis was performed, and blood products were administered, amount unknown. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36697 | 1220648-2021-01101 | 12689516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12689516&pc=OZD | 05/10/2021 | 10/25/2021 |  | 10/25/2021 |  |  | 10/04/2021 | 10/25/2021 |  |  | 12/31/2022 |  | 10/25/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN | [7 empty] | 2022032121 |  | 0048-0034-JP |  | No |  |  | Not Available |  | Yes | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/04/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed access site bleed during impella support. as a result, large amounts of blood products were administered, amount unknown. additionally, compression hemostasis was performed, and the device was removed. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36698 | 1220648-2021-01103 | 12689517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12689517&pc=OZD | 10/06/2021 | 10/20/2024 |  | 10/25/2021 |  |  | 10/06/2021 | 10/25/2021 |  |  | 05/31/2023 |  | 10/25/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022083694 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 55 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | Hemolysis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup |  |  | 10/06/2021 |  |  |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for and unknown indication. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient developed and infection and became septic during impella support. as a result, antibiotic (vancomycin) was administered. the patient ultimately expired due to sepsis. the source of the infection is unknown. no further information was provided. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | Additional information noted hemolysis occurring during the impella cp support. there was no lasting harm or injury from the hemolysis that prompted any treatment. |  |  |  | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b5 and h1. | The investigation for the reported hemolysis has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the hemolysis and its relation to the impella were not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm wheth er a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ additional information for the initial mfr 1220648-2021-01103 was provided in sections b5, d6a, d6b, d7a, g3, h6, and h10.
36699 | 1220648-2021-01104 | 12689543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12689543&pc=OZD | 07/29/2021 | 10/14/2024 |  | 10/25/2021 |  |  | 10/08/2021 | 10/25/2021 |  |  | 05/31/2023 |  | 10/25/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022089398 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 10/08/2021 |  |  |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in an unknown location for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed access site bleed during impella support. as a result, blood products were administered, amount unknown. no further information was provided, | It was also reported the patient experienced hemolysis while on support. the hemolysis was confirmed within the patient's urine; however, the device was repositioned which rectified the issue. |  |  |  |  | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr | Additional information for the initial mfr 1220648-2021-01104 was provided in sections b5, d6a, d6b, g1, g2, h1, and h6. the investigation for the reported hemolysis has been completed. no impella device nor data logs were received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of hemolysis was most likely improper positioning. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
36700 | 1220648-2021-01105 | 12689557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12689557&pc=OZD | 09/28/2021 | 11/08/2021 |  | 10/25/2021 |  |  | 09/28/2021 | 10/25/2021 |  |  | 07/31/2023 |  | 10/25/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022118238 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | [5 empty] | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2021 |  |  |  |  | An impella 5.5 was placed for a postoperative patient experiencing postcardiotomy cardiogenic shock. after 64 hours of support the pump stopped and was unable to be restarted. the pump was replaced and support proceeded. | [5 empty] | Since the original report was filed, the investigation into the pump stop has concluded. the medical center returned the product and data logs for analysis. the cause of the pump stop was excessive force by the operator which caused breakage of the motor cables which lead to an open line. the failure mode will be monitored and trended. | The investigation into this impella 5.5 pump stop, that occurred within the indicated use, is on-going at this time. upon the completion of the investigation, a final report will be filed.
36701 | 1220648-2021-01106 | 12689579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12689579&pc=OZD | 09/28/2021 | 10/25/2021 |  | 10/25/2021 |  |  | 09/28/2021 | 10/25/2021 |  |  | 04/30/2023 |  | 10/25/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP | [7 empty] | 2022073194 |  | 0048-0024-JP |  | No |  |  | Not Available |  | Yes | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 88 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/28/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year-old female patient had impella cp pump inserted in a unknown location for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed hemolysis during impella support. as a result, the patient was administered haptoglobin and the pump position was adjusted. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36702 | 1220648-2021-01093 | 12684741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12684741&pc=OZD | 09/30/2021 | 10/22/2021 |  | 10/22/2021 |  |  | 10/01/2021 | 10/22/2021 |  |  | 01/31/2022 |  | 10/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN | [7 empty] | 1483647 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/01/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male patient had impella 5.0 pump inserted in the right axillary. the patient had hematuria before insertion of impella, the patient had worsening hemolysis, pump level was decreased and haptoglobin was administered. | [5 empty] | Device discarded by customer. the impella 5.0 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36703 | 1220648-2021-01094 | 12684867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12684867&pc=OZD | 09/22/2021 | 10/25/2024 |  | 10/22/2021 |  |  | 09/27/2021 | 10/22/2021 |  |  | 04/30/2023 |  | 10/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022079835 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 79 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 09/27/2021 |  |  |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that hemolysis was noted during impella support. as a result, the pumps position was adjusted and haptoglobin was administered twice. hemolysis did not improve; however, there was cardiac function improvement. no further information was provided. | Additional report of thrombocytopenia and ventricular tachycardia and occurring during support. there was no bleed attributed to the use of impella that could have caused the exsanguination and necessary treatment. the survival outcome at explant noted the patient expired. use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | This report is being filed as part of a retrospective review of historical records. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ corrected information for the initial mfr 1220648-2021-01094 was provided in sections b3, b5, d6, and h6.
36704 | 1220648-2021-01096 | 12684980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12684980&pc=OZD | 07/08/2020 | 10/22/2021 |  | 10/22/2021 |  |  | 09/29/2021 | 10/22/2021 |  |  | 10/31/2021 |  | 10/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN | [7 empty] | 1453164 |  | 0048-0034-JP |  | No |  |  | Not Available |  | Yes | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 49 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/29/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient developed limb ischemia during impella support. as a result, endovascular therapy (evt) enforcement was performed. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36705 | 1220648-2021-01098 | 12685045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12685045&pc=OZD | 09/26/2021 | 10/22/2021 |  | 10/22/2021 |  |  | 09/30/2021 | 10/22/2021 |  |  | 12/31/2021 |  | 10/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1449404 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 60 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 09/30/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) had impella 5.0 pump inserted in an unknown location for acute myocardial infarction and cardiogenic shock (ami/cgs). the gender of the patient was not provided. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient developed cerebral infarction which became large. ultimately, the patient expired due to cerebral herniation. it was noted that ecmo was also being used with impella. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36706 | 1220648-2021-01095 | 12685049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12685049&pc=OZD | 09/28/2021 | 10/15/2024 |  | 10/22/2021 |  |  | 09/28/2021 | 10/22/2021 |  |  | 07/31/2023 |  | 10/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022103881 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 76 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 09/28/2021 |  |  |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for pre-op cabg coronary artery bypass grafting (cabg). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that a ct scan revealed the onset of hemorrhagic cerebral infarction. treatment details were not provided. no further information was provided. | Additional information reported that the patient was bleeding and there was woozing from the extracorporeal membrane oxygenation and impella insertion sites. as a result, the device was replaced with impela 5.0. the impella insertion site was purse-string sutured, and the bleeding subsided. information reported there is no causal relationship with the impella. the patient couldn't keep the circulatory dynamics and died. the cause of death was deterioration of general condition. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation for the reported access site bleeding and cva has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. no product or supplemental clinical details were shared. the cause of the bleeding and cva were not determined as the relation to the use of the impella could not be determined. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. additional information for the initial mfr 1220648-2021-01095 was provided in sections b5, d6a, d6b, d7a, g3, h6, and h10.
36707 | 1220648-2021-01097 | 12685072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12685072&pc=OZD | 06/04/2021 | 10/22/2021 |  | 10/22/2021 |  |  | 09/29/2021 | 10/22/2021 |  |  | 03/31/2023 |  | 10/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN | [7 empty] | 2022069695 |  | 0048-0034-JP |  | No |  |  | Not Available |  | Yes | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/29/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for cardiomyopathy. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient experienced bleeding at impella access site. as a result, blood products were administered, amount unknown. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36708 | 1220648-2021-01090 | 12674483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12674483&pc=OZD | 09/24/2021 | 11/08/2021 |  | 10/21/2021 |  |  | 09/24/2021 | 10/21/2021 |  |  | 04/30/2023 |  | 10/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022078119 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 39 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/24/2021 |  |  |  |  | A patient with a complex cardiac history and heart failure had an intra-aortic balloon pump support escalated to the support of the impella cp. the cp was hemodynamically supporting the patient for 2 days when the team removed all anticoagulation prior to the patient receiving dialysis. the heparin was not returned to the impella purge and the patient's limb became ischemic. the impella leg was cool to touch, pulseless, and painful. the pump was explanted and the patient needed no further intervention. | [5 empty] | The investigation into the reported limb ischemia is on-going at this time. upon the completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation into the limb ischemia has concluded. no product or logs were returned for analysis. the clinical details were reviewed and the root cause was determined to be patient condition, as the team did not properly monitor the anticoagulation. the patient's ptt levels were low. the failure mode will be monitored and trended.
36709 | 1035166-2021-00122 | 12674623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12674623&pc=DYB | 09/14/2021 | 03/31/2022 |  | 10/21/2021 |  |  | 10/01/2021 | 10/21/2021 |  |  | 07/01/2024 |  | 10/21/2021 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | DP-13602 |  | 0052-3006 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 54 YR |  | Female |  |  | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | On (b)(6) 2021 a (b)(6) female patient underwent a procedure for axillary artery insertion of the impella 5.5. the axillary insertion sheath was bleeding after insertion into the graft despite graft locks in place. sheath removed and replaced with another without any issue. no patient harm reported. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. | The following sections were updated in followup 1. b4, d9, g3, g6, h2, h3, h6 & h10 one abiomed 23f sheath was returned for analysis. traces of blood were found inside and outside the sheath. the sheath was observed using a vertex at higher magnification and no cracks, holes or gaps were found. the sheath score lines were observed to be prominent on the outer surface. the sheath was leak tested per iso 11070 standard and procedure for leak test as follows: sheath was occluded at the distal tip using a pin gauge and pressurized to 38kpa and 300kpa. the pressure was held for 30 seconds at 38kpa and 300kpa. the valve was occluded using a seal cap and dilator hub to maintain hemostasis. further investigation into the extrusion tool showed that the extrusion die change was implemented and excuted. this modification allowed for the score line to break easier during peel and reduce the thickness between score line and outer diameter. this change caused the wall thickness of score line to be lower for easier peel. review of the drawing showed that lower specification for wall thickness of score line was not specified. the lack of lower specification for wall thickness of score line caused the score line to be prominent. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. upon investigation, there were no cracks or holes observed on the surface of the sheath. the sheath was leak tested as per iso 11070 and no leakage was observed. the sheath was found to be within specification. investigation of extrusion process revealed that the extrusion die change was implemented to improve the peel strength at the tip. this resulted in thinner score line. review of drawing showed that there is no lower specification for the score line wall thickness. procedure for adelante, adelante-s and adelante-s2 sheath extrusion appropriate extrusion tool is selected for the part to be extruded and is inspected prior to the extrusion run. peel strength and critical dimensions are verified on every 100th extruded sheath as per procedure. procedure for abiomed introducer sheath in-process and final sample size: 100% inspection the sheath is inspected with a naked eye at a distance of 12" to 18" to ensure the sheath is free of kinks, cracks, splits, sinks and any other damages. leak test is performed as per procedures. ifu is provided with the product. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36710 | 1220648-2021-01091 | 12676052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12676052&pc=OZD | 09/23/2021 | 11/09/2021 |  | 10/21/2021 |  |  | 09/23/2021 | 10/21/2021 |  |  | 07/31/2023 |  | 10/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022110416 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/23/2021 |  |  |  |  | An impella 5.5 was placed to support a younger patient for support due to his cardiomyopathy while awaiting heart transplant. the pump successfully supported the patient for 21 days when it was explanted in preparation for the heart transplant. when explanted via the axillary insertion the team observed clot burden on the pump. at the time of explant the team reported the patient to have suffered a stroke. two arteries were occluded. the team in interventional radiology was unable to successfully remove all clot via embolectomy, and so a craniotomy was performed. the heart transplant was successful and the patient survived. | [5 empty] | Since the original report was filed, the investigation has closed. the root cause of the stroke was not determined. the product was not returned for analysis. the logs were analyzed and no abnormalities were observed. the root cause could not be determined. the failure mode will monitored and trended. | The medical center discarded the impella pump product. the data logs have been retrieved and the investigation into the cerebral vascular event is on-going at this time. upon the completion of the investigation, a final report will be filed.
36711 | 1220648-2021-01092 | 12676083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12676083&pc=OZD | 04/26/2021 | 11/09/2021 |  | 10/21/2021 |  |  | 09/27/2021 | 10/21/2021 |  |  | 09/30/2022 |  | 10/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021006154 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 66 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | A |  |  | No | Initial submission | Followup |  |  |  | 09/27/2021 |  |  |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient had developed sepsis during impella support. the patient was diagnosed with carbapenem-resistant bacteria. as a result, antibiotics were administered. no further information was provided. | [5 empty] | Since the original report was filed the investigation into the reported sepsis has concluded. no product or data logs were returned for analysis from japan. the clinical reports noted that the patient had sepsis from the time of admission. the sepsis is unrelated to the use of impella. the failure mode will be monitored and trended. | The impella cp ha snot been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36712 | 1220648-2021-01085 | 12656528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12656528&pc=OZD | 09/21/2021 | 11/12/2021 |  | 10/19/2021 |  |  | 09/21/2021 | 10/19/2021 |  |  | 06/30/2023 |  | 10/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022102866 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/21/2021 |  |  |  |  | The complainant reported a (b)(6) white male patient had impella cp pump inserted. the patient had hemolysis, pump was removed, and another pump was used. | [5 empty] | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The returned pump was visually inspected and a bent outflow cage was observed. the outflow cage bend was likely a result of the improper positioning. no other abnormalities were observed. the cause of the hemolysis was likely improper positioning since the pump position was reported to be in the mitral valve.
36713 | 1220648-2021-01086 | 12656556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12656556&pc=OZD | 09/13/2021 | 10/19/2021 |  | 10/19/2021 |  |  | 09/22/2021 | 10/19/2021 |  |  | 06/30/2023 |  | 10/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP | [7 empty] | 2022092745 |  | 0048-0024-JP |  | No |  |  | Not Available |  | Yes | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 17 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/22/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the right femoral for myocarditis. the patient was implanted on (b)(6) 2021 and explanted (b)(6) 2021. it was reported that during pcps and impella insertion hemolysis was noted. the device was adjusted in order to improve hemolysis; however, this did not help. as a result, blood products and haptoglobin were administered. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36714 | 1220648-2021-01087 | 12656567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12656567&pc=OZD | 09/22/2021 | 12/14/2024 |  | 10/19/2021 |  |  | 09/22/2021 | 10/19/2021 |  |  | 12/31/2021 |  | 10/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1463041 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 46 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup |  |  | 09/22/2021 |  |  |  |  | The complainant reported a (b)(6) male patient had impella 5.0 pump inserted in the right axillary for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that an imaging of the rsca after the device was removed showed and unclear area. thrombus was noted and a fogarty catheter was used to remove the thrombus. no further information was provided. | It was reported that imaging of the rsca, after removal of the impella 5.0, revealed an unclear area, suggesting dissociation or thrombus. when fogarty catheterization was performed, the thrombus was removed, after which there was no difference between the abp of the left and right sides. it was reported that there was no hazard to the patient; however, the event was reported to be related to the impella device. additional information stated that the patient's condition was stable at the time of the pump's removal. the blood clot was removed by fogarty catheterization. it was later reported that the patient experienced an ischemic cerebrovascular event on (b)(6) 2021. it was unknown if the stroke was related to the impella device. it was reported that there was a possibility that the stroke occurred at the time of the impella device's removal. it was previously reported that the impella device was explanted on (b)(6) 2021. |  |  |  |  | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | No product was returned for evaluation and no data logs were provided for analysis. it was reported that the clot was noted in the patient's rsca after the impella device was removed. the root cause of the clot and its relation to the impella device was determined to be unknown. it was later reported that the patient experienced an ischemic cerebrovascular event on (b)(6) 2021 and that there was a possibility that the stroke occurred at the time of the impella device's removal. it was previously reported that the impella device was explanted on (b)(6) 2021. the root cause of the cva was determined to be unrelated to the impella device. | This report is being filed as part of a retrospective review of historical records. corrected information provided in b5 and h6.
36715 | 1220648-2021-01088 | 12656576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12656576&pc=OZD | 09/24/2021 | 11/03/2021 |  | 10/19/2021 |  |  | 09/24/2021 | 10/19/2021 |  |  | 06/30/2023 |  | 10/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET | [7 empty] | 2022092249 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 87 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/24/2021 | * | Invalid/Incomplete |  |  | The medical center admitted an elderly patient with a complex cardiac history suffering from an acute myocardial infarction. while in the cardiac catheterization lab the patient was requiring cpr and so the team placed an impella cp for hemodynamic support during an angioplasty procedure to the coronary arteries. the patient was later sent to the icu on the impella cp support but, he continued to decompensate. the team observed a growing impella access site hematoma. the estimate was 300-400cc of blood lost and the hemoglobin levels dropped from 12.5 to 9. the team manually held the hematoma to control its growth. the patient expired while awaiting family decision and dnr status. | [5 empty] | Since the original report was filed, the investigation into the bleeding has concluded. the product was not returned for analysis. the clinical review was done, and revealed that all appropriate access procedures were followed. best practices were done. the root cause of the bleeding was patient condition. the failure mode will be monitored and trended. | The investigation into the access site bleeding with hemoglobin drop and hematoma is on-going at this time. upon completion of the investigation, a final report will be filed.
36716 | 1220648-2021-01089 | 12656589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12656589&pc=OZD | 09/18/2021 | 10/19/2021 |  | 10/19/2021 |  |  | 09/21/2021 | 10/19/2021 |  |  | 06/30/2023 |  | 10/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP | [7 empty] | 2022100409 |  | 0048-0024-JP |  | No |  |  | Not Available |  | Yes | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/21/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021, explant date was not provided. it was reported that the patient developed hemolysis during impella support. as a result, haptoglobin was administered and the pumps position was adjusted. the patient had renal failure prior to impella; however, it is unknown if the deterioration of renal function was due to ami. dialysis was performed. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36717 | 1220648-2021-01084 | 12648653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12648653&pc=OZD | 08/23/2021 | 10/18/2021 |  | 10/18/2021 |  |  | 09/22/2021 | 10/18/2021 |  |  | 06/30/2023 |  | 10/18/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET | [7 empty] | 2022103697 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 74 YR | [5 empty] | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/22/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) hispanic/latino white female patient had impella cp pump inserted in the right femoral artery (rfa). the pump¿s flows dropped to <1 l/m, pump was explanted, another pump was inserted and the pump was returned for investigation. | [5 empty] | The provided data logs showed that the pump flow decreased on (b)(6) 2021 at 11:11 and never recovered. suction alarms were triggered shortly after. the flow remained decreased for the remainder of the run. the product was returned for evaluation. it was observed that both impeller blades were broken. the pump was run in the lab and low flows were reproduced. the cause of the low pump flow was fractured impeller blades. because fluroscopic imaging of the pump/additional evidence was not returned, a definitive cause of the fracture could not be determined. based on the characteristic of the fracture, it was likely caused by an interaction with the pci catheter.
36718 | 12638366 | 12638366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12638366&pc=OZD | 10/07/2021 | 10/13/2021 | 10/13/2021 | 10/15/2021 | 10/13/2021 | 10/15/2021 |  | 10/15/2021 |  |  |  |  | 10/15/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST |  |  | OZD |  |  | 0550-0008 | [7 empty] | 321231 |  | LOG1509062780 |  | Yes |  |  | Not Available |  | Not Available | No | Health Professional |  |  | User Facility report | [8 empty] | 14965 DA |  |  | 120 |  |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | Insufficient Information | [37 empty] | Unknown |  |  |  | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella temporary vad (ventricular assist device) purge cassette was noted to be cracked on the distal edge and leaking. after consulting with abiomed, the manufacturer, recommendation was to cut cassette off and splice line with 19-gauge needle. after performing this intervention, pump working correctly.
36719 | 12638367 | 12638367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12638367&pc=OZD | 10/09/2021 | 10/13/2021 | 10/13/2021 | 10/15/2021 | 10/13/2021 | 10/15/2021 |  | 10/15/2021 |  |  |  |  | 10/15/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST |  |  | OZD |  |  | 0550-0008 | [7 empty] | S320266 |  | LOG1509114004 |  | No answer provided |  |  | Not Available |  | Not Available | No | Health Professional |  |  | User Facility report | [8 empty] | 22265 DA |  |  | 129 |  |  | Malfunction | No | Yes | Crack | [6 empty] | Insufficient Information | [37 empty] | Unknown |  |  |  | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella temporary vad (ventricular assist device) purge cassette noted to be cracked while changing the purge tubing. consulted with abiomed manufacturer and 19g needle inserted per recommendations. purge pressures and flow normalized.
36720 | 1220648-2021-01083 | 12618942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12618942&pc=OZD | 02/24/2018 | 10/12/2021 |  | 10/12/2021 |  |  | 02/24/2018 | 10/12/2021 |  |  | 01/31/2020 |  | 10/12/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, US | [7 empty] | 1323464 |  | 0048-0003 |  | No |  |  | No |  | Yes | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | STUDY | [5 empty] | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/24/2018 | * | Invalid/Incomplete |  |  | In 2018 a patient enrolled in (b)(6) registry had an impella cp placed for hemodynamic support for a coronary angioplasty. the support was proceeding when the impella limb was noted to be ischemic. the team performed a femoral to femoral external bypass. years later, in 2021, the registry team reviewed this patient and documented the adverse event to be possibly due to the use of impella. | [5 empty] | An investigation of the limb ischemia was conducted. the product and data logs were not possible to be retrieved, as the event occurred in 2018. when the clinical updates were reviewed the cause of the limb ischemia was determined to be due to native patient condition. the patient had underlying disease and was on vasopressors, along with suffering from right heart failure and kidney dysfunction. the failure mode will be monitored and trended.
36721 | 1220648-2021-01081 | 12604786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12604786&pc=OZD | 09/16/2021 | 10/11/2021 |  | 10/11/2021 |  |  | 09/16/2021 | 10/11/2021 |  |  | 03/31/2023 |  | 10/11/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET | [7 empty] | 2022068451 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Impairment | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/16/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock. after approximately 6 days of support, the patient was noted to have increased creatinine and an unspecified "renal injury." the device was noted to have been malpositioned for an extended period of time leading up to this event and likely related to this injury. the pump was repositioned and ultimately removed the same day. | [5 empty] | The impella cp optical was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36722 | 1220648-2021-01082 | 12609102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12609102&pc=OZD | 09/13/2021 | 10/22/2021 |  | 10/11/2021 |  |  | 09/13/2021 | 10/11/2021 |  |  | 07/31/2023 |  | 10/11/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US | [7 empty] | 2022112647 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | Valvular Insufficiency/ Regurgitation | [11 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2021 | * | Invalid/Incomplete |  |  | A surgical patient was undergoing bypass surgery on pump (cabg) and had an impella 5.5 placed for hemodynamic support for the surgical intervention. the team successfully placed the bypass grafts to the coronary arteries. while positioning the 5.5 pump the team observed that the impella 5.5 had perforated the ventricle wall. the pump positioning also may have contributed to valve damage. the team performed an aortic valve replacement and ventricle wall repair. | [5 empty] | The investigation is on-going at this time. upon completion, a final report will be file. | Since the original report was filed, the investigation has concluded. no product, data logs, or imaging were returned for analysis. without these the root cause of the lv perforation and heart valve damage could not be determined. the failure modes will be monitored and trended.
36723 | 1220648-2021-01076 | 12598611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12598611&pc=OZD | 09/12/2021 | 10/08/2021 |  | 10/08/2021 |  |  | 09/13/2021 | 10/08/2021 |  |  | 05/31/2023 |  | 10/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022089526 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 90 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/13/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted six hours later. it was reported that when after the pci was completed, there was a blood loss at the impella insertion site. as a result, endovascular therapy (evt) was performed on the right superficial femoral artery (sfa). subsequently, there was discoloration of the lower limb after impella was inserted. a puncture from the opposite side of impella was performed to treat limb ischemia. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36724 | 1220648-2021-01078 | 12599479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12599479&pc=OZD | 09/15/2021 | 12/13/2024 |  | 10/08/2021 |  |  | 09/16/2021 | 10/08/2021 |  |  | 12/31/2021 |  | 10/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1466157 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional |  |  |  |  | 27 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup |  |  | 09/16/2021 |  |  |  |  | The complainant reported a (b)(6) male patient had impella 5.0 pump inserted in the right axillary for post-cardiotomy cardiogenic shock (pccs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that during an echo, the patients chordae was noted to be ruptured. the cause of the ruptured is unknown. no further information was provided. | It was additionally reported that the pump's purge flow gradually increased to 30 ml/hr and purge pressure decreased to 100mmhg on (b)(6) 2021. a purge cassette exchange was performed which did not resolve the issue. however, it was reported that the purge flow suddenly settled to 10ml/hr. no further information was provided. |  |  |  |  | The returned pump was visually inspected, and it was noted that the pigtail had been removed and there was a catheter kink at the 42 cm mark. as no evidence of the chordae rupture was provided, the relationship to impella is unknown. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event. | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information added to b5 and h6 for purge pressure issues. the field long term and real time data logs were reviewed, and high purge flow and low purge pressure were confirmed on (b)(6) 2021. the purge issues suddenly resolved the night of (b)(6) 2021, and the purge system appeared to function normally for the remainder of the case. the returned pump was visually inspected, and it was noted that the pigtail had been removed and there was a catheter kink at the 42 cm mark. microscopic inspection of the differential pressure sensor revealed damage on and immediately surrounding the sensor face, including what appeared to be etching of the silicone. the pump was plugged into a console and ran at p-9 with normal purge pressure and flow values. the reported purge issues were unable to be reproduced, a specific cause for these issues could not be determined. this report is being filed as part of a retrospective review of historical records.
36725 | 1220648-2021-01077 | 12599508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12599508&pc=OZD | 09/15/2021 | 10/26/2021 |  | 10/08/2021 |  |  | 09/15/2021 | 10/08/2021 |  |  | 07/31/2023 |  | 10/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP | [7 empty] | 2022105830 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/15/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old male patient had impella cp pump inserted in the right femoral for pre-op coronary artery bypass graft (cabg). the patient was implanted on (b)(6) 2021 and explanted on the same day. it was reported that due to patient movement, a hematoma formed at the impella access site. as a result, the angle was changed; however, bleeding at the access site was uncontrollable. the patient required 30 units of red blood cells, 20 units of fresh frozen plasma and 20 units of plt. surgical vascular repair was also performed. no further information was provided. | [5 empty] | The product was returned for analysis. the returned pump was visually inspected and holes in the repo unit sheath and a tear in the blue butterfly were observed. the cause of the access site bleeding was most likely use related. best practices were not followed resulting in suture holes in the sheath. the product was released with a completed and reviewed device history record under the abiomed quality system. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event. | The impella cp was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
36726 | 1220648-2021-01079 | 12599562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12599562&pc=OZD | 09/17/2021 | 11/12/2021 |  | 10/08/2021 |  |  | 09/17/2021 | 10/08/2021 |  |  | 06/30/2023 |  | 10/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022098094 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/17/2021 |  |  |  |  | The complainant reported a (b)(6) old asian female patient had impella 5.5 pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). the physician tried multiple times to reinsert the pump, the patient lost some blood, 2 (two) units of red blood cells (rbcs) was given and a second pump was used. | [5 empty] | Data logs from the case were not sent. visual inspection of the returned pump revealed the presence of biomaterial on and around the impeller. there was also some biomaterial in the purge gap. no other abnormalities were found. the pump was run and the low pump flow reproduced. additionally, there was no "impella position in aorta" alarm. in conclusion, it was determined that the cause of the low pump flow was ingested biomaterial. the patient had some blood loss and 2 (two) units of red blood cells (rbcs) was given but the source of the bleed was not from the impella access site. | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36727 | 1220648-2021-01080 | 12599568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12599568&pc=OZD | 09/13/2021 | 10/14/2024 |  | 10/08/2021 |  |  | 09/13/2021 | 10/08/2021 |  |  | 01/31/2023 |  | 10/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022057168 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | Other |  |  |  |  | PCPS ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/13/2021 |  |  |  |  | The patient in (b)(6) was supported by an intra-aortic balloon pump (iabp) and pcps (emco) when care was escalated. an impella cp was placed to support the patient and the iabp was removed. after 6 days of support the pump was explanted due to the hemolysis concerns. they had attempted to troubleshoot first by p-level adjustment, which did not resolve hemolysis signs. the team did not attempt to reposition the pump despite alarms for positioning being inadequate. at pump explant the team observed thrombotic material to be adhered to the pump. | It was reported also the patient experience a thrombus ingestion while on support. upon removal of the device biomaterial was observed on the pump. no further measures were taken. |  |  |  |  | Since the initial report was filed, the investigation into the hemolysis has completed. neither pump product nor data logs were returned for analysis. no hemolysis benchtop testing was possible. the clinical report was reviewed and cause of the hemolysis was determined. the cause for the hemolysis was ingested biomaterial due to the patient's condition of hypercoagulation. there was reported biomaterial observed at pump explantation. | The medical center in (b)(6) discarded the impella pump product. the data logs have not been returned for analysis. the root cause has not been determined. should more information be shared that leads to a root cause of the presumed hemolysis, a final report will be filed. | Additional information for the initial 1220648-2021-01080 was provided in sections b5, d6a, d6b, g1, g2, h1, and h6. the investigation for the reported thrombus has been completed. no impella device nor data logs were received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the thrombus issue was due to biomaterial was observed on the pump at the time of pump removal. the failure modes will be monitored and trended. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
36728 | MW5104432 | 12589612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12589612&pc=DWA | 08/29/2021 | 10/14/2021 |  | 10/07/2021 |  |  |  | 10/07/2021 | 10/07/2021 |  |  |  | 10/07/2021 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  | IMPELLA | IMPELLA |  | DWA | DWA |  |  |  |  | ABIOMED, INC. | ABIOMED, INC. | [7 empty] | No answer provided | No |  |  |  |  | Invalid/Incomplete | Health Professional | No Information |  | Voluntary report | [8 empty] | 61 YR |  |  | 75 |  |  | Death | Yes | Yes | Improper or Incorrect Procedure or Method | Patient Device Interaction Problem | [5 empty] | Exsanguination | [12 empty] | Death | [28 empty] | Initial submission | Followup |  |  |  |  | I | NO INFORMATION |  |  | My husband had a heart attack and was taken to hospital by ambulance. upon arrival until 90 minutes plus he was kept in er until taken to cath lab where doctor chose to use not one but two impella devices. he did not make it. at no time did doctor confer or even tell me his spouse what was going on and only said after his demise he did all he could. two days later receive a notice from my husbands insurance doctor wanting approval for the devices each over (b)(6). when i was taken in to see my husband they had made no attempt to clean up and there was blood every where and none of the techs would look me in the eye. i got my husbands medical records and have discovered he literally bled out from impella causing damage. i am angry upset and want this device taken off the market. if doctor would have chosen to use balloon pump my husband might still be here. he was only (b)(6) and in good health and deserved better care not based on greed. fda safety report id# (b)(4). | Additional information received from a reporter on 10/14/2021 for a report number mw5104432. doctor place not one but two impellas in my husband after he suffered a major heart attack and my husband died bled out. fda safety report id # (b)(4).
36729 | 1220648-2021-01070 | 12592167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12592167&pc=OZD | 03/06/2021 | 10/07/2021 |  | 10/07/2021 |  |  | 09/13/2021 | 10/07/2021 |  |  | 10/31/2022 |  | 10/07/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022029739 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 84 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/13/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the right femoral for an unknown indication. it was reported that the patient experienced bleeding at the impella access site. there were no issues with the procedure and the act/ptt were managed with no problem. blood products were administered, amount unknown. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36730 | 1220648-2021-01071 | 12592250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12592250&pc=OZD | 09/12/2021 | 11/08/2021 |  | 10/07/2021 |  |  | 09/12/2021 | 10/07/2021 |  |  | 07/31/2023 |  | 10/07/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022112641 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2021 |  |  |  |  | The complainant reported a (b)(6) year old white male patient had impella 5.5 pump inserted through the left supraclavicular space. the device went directly towards the large papillary muscle, after removal, bleeding was noticed and a small hole was found in the portion of the left ventricle (lv) and as a result a patch was surgically placed over the hole. | [5 empty] | The data logs showed fluctuating placement signal waveforms between ventricular and aortic, even triggering numerous "impella position in aorta" alarms. there were also persistent "suction" and "placement signal low" alarms at this time. no damage or abnormalities were found on the returned pump or guidewire. in conclusion, after reviewing the clinical information, it was determined that the cause of the ventricle perforation was likely improper positioning. | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36731 | 1220648-2021-01072 | 12592417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12592417&pc=OZD | 09/09/2021 | 11/12/2021 |  | 10/07/2021 |  |  | 09/13/2021 | 10/07/2021 |  |  | 05/31/2023 |  | 10/07/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, OZD |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022082675 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 32 YR |  | Female |  |  | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2021 |  |  |  |  | The complainant reported a (b)(6) female patient had impella 5.0 pump inserted for myocarditis. the patient was brought to the operating room (or) due to bleeding on the graft axillary anastomosis but during the procedure, the pump stopped, and during pump exchange there was a tear on the great vessel. stitches were made to the artery to repair the tear. the patient expired due to unknown reasons. | [5 empty] | The returned pump was visually inspected and the red pump plug was opened, confirming blood ingress. further inspection of the pump catheter revealed three small pinholes between the 2cm and 4cm marks. therefore, the root cause of the pump stop was determined to be blood ingress into the pump plug caused by damage to the catheter. however, a specific cause for the reported access site bleeding could not be determined through this evaluation. | The impella 5.0 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36732 | 1220648-2021-01073 | 12592487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12592487&pc=OZD | 06/21/2021 | 10/16/2024 |  | 10/07/2021 |  |  | 09/13/2021 | 10/07/2021 |  |  | 10/31/2022 |  | 10/07/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021019858 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 09/13/2021 |  |  |  |  | Additional information was reported indicating that the patient developed hemolysis for which the position was adjusted. this seems to have resolved the issue. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The complainant reported a patient had impella cp pump inserted in an unknown location for acute myocardial infarction and cardiogenic shock (ami/cgs). patient's age and gender were not provided. it was reported that the patient developed limb ischemia during impella support. as a result, a reverse sheath stab was performed to improve limb ischemia. it was noted that the impella was inserted using a medi-kit 16fr sheath. no further information was provided. |  |  |  |  | The investigation for the reported limb ischemia and hemolysis issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on the provided information that the reported hemolysis improved after position adjustment, the root cause of the hemolysis was most likely related to improper positioning. no product or data logs were provided. additionally, the cause of the reported limb ischemia was most likely due to the patient¿s condition. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr was provided in sections b3, b5, d4, d6, d7a, g2, h3 and h6. | The impella cp has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36733 | 1220648-2021-01074 | 12592508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12592508&pc=OZD | 02/21/2021 | 10/24/2024 |  | 10/07/2021 |  |  | 09/13/2021 | 10/07/2021 |  |  | 08/31/2022 |  | 10/07/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1496670 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 09/13/2021 |  |  |  |  | Additional information received that patient developed limb ischemia during impella support. treatment details were not provided. ultimately the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the was reported there was bleeding at the impella insertion site. as a result, blood products were administered, amount unknown. it was noted that the introducer was peeled outside the body and there were no issues with the angle or sutures. no further information was provided. |  |  |  |  | Additional information for the initial mfr 1220648-2021-01074 was provided in sections b2, b5, d6a, d6b, h1, and h6. the investigation for the reported limb ischemia has been completed. the impella device has not been returned for evaluation. however, clinical details provided were sufficient to determine a root cause. the cause of the reported limb ischemia was most likely related to patient condition. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The impella cp was not return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36734 | 1220648-2021-01075 | 12592528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12592528&pc=OZD | 08/17/2021 | 10/07/2021 |  | 10/07/2021 |  |  | 09/22/2021 | 10/07/2021 |  |  | 01/31/2023 |  | 10/07/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2022049011 |  | 0048-0014 |  | Device was returned to manufacturer | [5 empty] | No | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | DISTRI |  |  |  |  | 75 YR | [5 empty] | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission |  |  |  |  | 09/22/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting for high risk percutaneous coronary intervention. the impella cp optical was used for hemodynamic support. during support, unresolvable suction alarms and low pump flow was exhibited. the pump was explanted. upon receipt of the device, our investigation team identified a broken impeller blade. | [5 empty] | The device and data logs were returned and analyzed. data logs confirmed suction alarms and reduced pump flow. visual inspection of the device revealed a fracture in one of the impeller blades. no other damage or deficiency was found. we believe the low pump flow and suction were caused by the fractured impeller blade. no fluoroscopic imaging of the pump/additional evidence was returned, therefore, a definitive cause of the fracture could not be determined.
36735 | 1220648-2021-01069 | 12582554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12582554&pc=OZD | 09/10/2021 | 10/12/2021 |  | 10/06/2021 |  |  | 09/10/2021 | 10/06/2021 |  |  | 12/31/2022 |  | 10/06/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN | [7 empty] | 2022039787 |  | 0048-0034-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/10/2021 | * | Invalid/Incomplete |  |  | A patient was admitted to a medical center in (b)(6) with a history of heart failure and prior multivessel coronary artery disease. the team placed an intra-aortic balloon pump (iabp) but when his care needed to be escalated, an impella cp was placed to support the patient and allow for coronary intervention. the impella cp supported the patient for 2 days when it was explanted due to concerns of hemolysis. the hemolysis was treated by dialysis. | [5 empty] | Since the initial report was filed, the investigation has concluded. the medical center in japan did not return the pump and full clinical details. without the pump for hemolysis testing and clinical details, the root cause was not able to be determined. | The investigation into the reported hemolysis is ongoing at this time. not all product have been returned, nor have all clinical questions have been answered. when the investigation is completed, a final report will be filed.
36736 | 12575878 | 12575878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12575878&pc=PBL | 04/27/2021 | 09/23/2021 | 09/23/2021 | 10/05/2021 | 09/23/2021 | 10/05/2021 |  | 10/05/2021 |  |  |  |  | 10/05/2021 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PBL |  |  | 0043-0002 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No answer provided | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 23725 DA | [5 empty] | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Abiomed impella left ventricular assist device was defective (broken).
36737 | 1220648-2021-01060 | 12555614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12555614&pc=OZD | 08/12/2021 | 09/30/2021 |  | 09/30/2021 |  |  | 09/01/2021 | 09/30/2021 |  |  | 05/31/2023 |  | 09/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022095052 |  | 0048-0045 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/01/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella cp pump inserted in the right femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had thrombosed iliac that was removed with thrombectomy and endarterectomy, patient also had leg ischemia in both legs but only the impella leg was surgically intervened. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed. the site reported that the patient¿s femoral and iliac arteries were thrombosed requiring surgical intervention. also, the patient had bilateral cold feed and absent distal pulses. the root cause for the reported thrombus is likely patient condition. the root cause for the reported limb ischemia was likely related to the reported thrombus.
36738 | 1220648-2021-01061 | 12555620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12555620&pc=OZD | 09/01/2021 | 10/28/2021 |  | 09/30/2021 |  |  | 09/01/2021 | 09/30/2021 |  |  | 11/30/2021 |  | 09/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US | [7 empty] | 1466927 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 18 YR | [5 empty] | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) white female patient had impella cp pump inserted in the right axillary. after pump implantation large volume of bubbles was visualized in the aorta and left ventricle (lv) concerns of a faulty catheter promoted pump to be removed. | [5 empty] | The field lt and rt logs were reviewed, and no alarms or abnormalities were noted. the returned pump was visually inspected and multiple indentation marks on either side of the pump catheter were observed at approximately the 30cm mark. further examination confirmed that the catheter had been punctured in this location. therefore, the root cause of the aortic and ventricular bubbles was determined to be damage to the catheter which likely occurred during use. | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36739 | 1220648-2021-01062 | 12556084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12556084&pc=OZD | 09/05/2021 | 10/21/2021 |  | 09/30/2021 |  |  | 09/05/2021 | 09/30/2021 |  |  | 05/31/2023 |  | 09/30/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET | [7 empty] | 2022081400 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2021 | * | Invalid/Incomplete |  |  | A patient with shortness of breath and cardiomyopathy had an impella cp placed for hemodynamic support. the cp was placed via the femoral artery through a very deep tissue track and much adipose tissue. after the team placed the 14fr sheath for impella placement, an access site a hematoma developed. vascular surgery was consulted, and the pump was explanted and replaced. additionally blood products were infused due to the bleed and hematoma. | [5 empty] | Since the original report was filed, the investigation into the access site bleed has completed. the bleed was due to patient condition. the medical center's report was that due to body habitus the re-positioning sheath was not long enough for the patient. the patient panus was an issue, as was the depth to the arteriotomy. the failure mode will be monitored and trended. | The investigation into the access site bleeding and hematoma is on-going at this time. upon completion of the investigation, a final report will be filed.
36740 | 1220648-2021-01063 | 12556252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12556252&pc=OZD | 09/03/2021 | 10/07/2023 |  | 09/30/2021 |  |  | 09/03/2021 | 09/30/2021 |  |  | 07/31/2023 |  | 09/30/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022112648 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 67 YR |  | Male |  |  |  | Malfunction | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/03/2021 |  |  |  |  | The patient was transferred to a tertiary center for escalation of care and surgery. he was to have bypass (cabg) and left atrial appendage repair done in the or with hemodynamic support provided by an impella 5.5. the team made choice to re-position the 5.5 in the ventricle (lv) by manually moving/manipulating the heart and caused the lv to be perforated. the team explanted the impella 5.5, surgically repaired the ventricle, and infused many units of blood and blood products. | [5 empty] | The investigation into the root cause of the left ventricle perforation was found to be user related. the operator was manually manipulating the heart while simultaneously re-positioning the impella. this allowed for the impella to perforate the ventricle. the failure mode will be monitored and trended. | (h3) the investigation into the reported ¿ventricle perforation and difficulty inserting/removing introducer¿ has been completed since the original report was filed. product was returned for investigation. the device was visually inspected and no abnormalities were observed. the catheter had been cut and biomaterial was observed in the outflow cage. the cause of the lv perforation was use related since the device and heart were reported to be manipulated by the user. the returned graft was visually inspected, and biomaterial was observed inside it. as per clinical investigation, the complainant was utilizing a 5.5 for support when taking the surgical patient off the pump. he had a cabg x3 and laa done in the operating room (or). the medical team had difficulty delivering the pump into appropriate position, via the tunneled graft, to the lv. the pump would hang up on the mitral. the team made the choice to manipulate the pump by physically moving the heart and caused the lv to be perforated by the 5.5 two times. the team surgically repaired and infused many units of blood and blood products. the patient did survive and was released from the or to the intensive care unit (icu). the most likely cause of the inability to explant through the introducer was most likely biomaterial.
36741 | 1220648-2021-01064 | 12556322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12556322&pc=OZD | 08/30/2021 | 11/29/2024 |  | 09/30/2021 |  |  | 09/01/2021 | 09/30/2021 |  |  | 05/31/2023 |  | 09/30/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED INC. |  |  | US |  | 2022094421 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/01/2021 |  |  |  |  | A patient admitted in cardiogenic shock in the eu had an impella 5.0 placed to support the patient hemodynamically in addition to the already placed ecmo circuit. the pump ran for over 7 days when the impella pump stopped. the team was unable to re-start the pump and it was explanted. the impella was not replaced, and the patient remained on the ecmo circuit alone. | It was also reported that on day 6 of support, there were alarms for high purge pressure and low purge flow. it was reported that the purge flow dropped to <2ml and the purge pressure was >1000mmhg. it was reported that changing the purge cassette resolved the issue. |  |  |  |  | The medical center did return the impella pump for analysis. the analysis is on-going at this time. upon completion of the investigation, a final report will be submitted. | Since the initial report was filed, the investigation into the pump stop has concluded. the pump was returned as were the data logs and analysis was performed. the pump stop was due to blood in the connection/lemo pins. the blood allowed for corrosion. the blood would be present at this connection due to user error and allowing the pump connection to sit within blood. the failure mode will be monitored and trended. | Additional information was provided in b5 and h6 for the high purge pressure alarm issue. the data logs were returned and confirmed the high purge pressure alarm. the root cause of the high purge pressure was determined to be most likely a kink the purge cassette line as replacing the cassette fixed the issue and no high purge pressure was reproduced with the returned cassette.
36742 | 1220648-2021-01065 | 12557669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12557669&pc=OZD | 09/01/2021 | 10/23/2024 |  | 09/30/2021 |  |  | 09/02/2021 | 09/30/2021 |  |  | 10/31/2022 |  | 09/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021017619 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 77 YR |  | Female |  |  | Asian | Death | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Hypovolemia | [10 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 09/02/2021 |  |  |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the left femoral for post-cardiotomy cardiogenic shock (pccs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the physician suspected lower limb ischemia during impella support. an anterograde sheath was inserted to improve lower limb blood flow. an escharotomy was performed on (b)(6) 2021 because there was increased in tension and creatine kinase (ck) in the lower limbs; however, there was no ck peak observed. it was noted that lower limb blood flow was observed after treatment. no further information was provided. | It was reported on (b)(6) 2021 the patient was taken to the operating room for a thoracotomy that was performed due to suspected cardiac tamponade. no information was provided on cause of the suspected cardiac tamponade. after hemostasis was achieved, there was bleeding and a hematoma which was removed from the anastomosis site of the bypass. it was noted that the patient had been heparinized for impella use, and heparin was reduced after bleeding was noted. due to circulatory failure following the bleeding and thoracotomy, venoarterial extracorporeal membrane oxygenation (va ecmo) was cannulated for additional hemodynamic support. after the patient returned to the icu, it was reported there were alarms for the position signal is unstable on the impella. the p level was lowered to 2 and the position was verified with an echocardiogram. it was suspected the pump was refluxing from the outlet, and it was reported the patient was too unstable for replacement. due do the positioning alarms and suspected flow issues the pump was removed. the patient's condition was critical, and it was decided to monitor the patient while they remained on va ecmo. it was reported the patient expired 3 days after explant of the impella. it was reported the patient's circulatory dynamics were poor prior to the insertion of impella, and there was no improvement even after impella was inserted and no change after removal of the impella. it was unknown what caused the patient's cause of death, however it was stated it was possibly related to the impella. |  |  |  |  | The impella cp has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information for the initial mfr 1220648-2021-01065 was provided in sections b5, d6a, d6b, and h6. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2021-01065 was provided in sections b1, b2, g1, and h1. the initial mfr was reported as a serious injury. after review of additional information, it was determined the impella could not be disassociated from the patient's outcome of death.
36743 | 1220648-2021-01067 | 12558085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12558085&pc=OZD | 09/07/2021 | 09/30/2021 |  | 09/30/2021 |  |  | 09/10/2021 | 09/30/2021 |  |  | 04/30/2023 |  | 09/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022073747 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/10/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for high-risk percutaneous coronary intervention (hrpci). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient developed hemolysis while on impella support. as a result, haptoglobin was administered. subsequently, when the impella was removed, a blood clot of about 1.5 x 1.5 cm was noted to be adhered to the impella pigtail. as a result, the blood vessel was surgically sutured. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36744 | 1220648-2021-01066 | 12558088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12558088&pc=OZD | 07/30/2021 | 11/29/2024 |  | 09/30/2021 |  |  | 09/07/2021 | 09/30/2021 |  |  | 04/30/2023 |  | 09/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022082486 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 64 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 09/07/2021 |  |  |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the left femoral for post-cardiotomy cardiogenic shock (pccs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported the patient developed am ischemic cerebrovascular accident (stroke). the ischemic stroke was intensive and therefore treatment was not rendered. no further information was provided. | Additional information was received that the patient ultimately expired. |  |  |  |  | The impella 5.0 has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation for the reported issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2021-01066 was provided in sections b2, b5, d6, d8, h1, h3, and h6.
36745 | 1220648-2021-01068 | 12558090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12558090&pc=OZD | 09/10/2021 | 10/30/2021 |  | 09/30/2021 |  |  | 09/11/2021 | 09/30/2021 |  |  | 05/31/2023 |  | 09/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN | [7 empty] | 2022083756 |  | 0048-0034-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/11/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the device was implanted on (b)(6) 2021 and explanted (b)(6) 2021. it was reported blood vessel damage with vascular contusion was noted during impella removal. there were issues with the angle during detention and bleeding was noted at the insertion site during indwelling. surgical anastomosis was performed. no further information was provided. | [5 empty] | Since the original report was filed, the investigation has concluded. the medical center in japan did return the impella pump but, not the introducer sheath. the root cause of the injury/contusion noted at the access site could not be determined. the failure mode will be monitored and trended. | The impella cp was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
36746 | 12532072 | 12532072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12532072&pc=OZD | 09/07/2021 | 09/10/2021 | 09/10/2021 | 09/27/2021 | 09/10/2021 | 09/27/2021 |  | 09/27/2021 |  |  |  |  | 09/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1411424 |  | 0042-0010-US |  | No answer provided | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 24090 DA | [5 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella 5.5 cassette leaking within cassette holder in the impella controller. new cassette obtained. when attempting to change cassette, controller was not reading new cassette insertion (stuck at step 4 of the process). impella representative on the phone with nurse. charge nurse and np at the bedside. new controller obtained. pt reconnected to new controller. patient stable after and throughout change of cassette.
36747 | 1220648-2021-01053 | 12533943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12533943&pc=OZD | 12/10/2017 | 09/18/2024 |  | 09/27/2021 |  |  | 12/18/2017 | 09/27/2021 |  |  | 09/30/2019 |  | 09/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | CP C-7 14FR PUMP SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1310936 |  | 0048-0020 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  | White | Injury | Yes | Yes | Fracture | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/18/2017 |  |  |  |  | The procedural outcome states patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | A (b)(6) female was admitted to the cardiac cath lab for placement of an impella cp for hemodynamic support. during the procedure, impella flows dropped. the team attempted to troubleshoot and eventually had to replace it with another impella cp. upon receipt of the device, broken impeller blades were identified. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2021-01053 was provided in sections b2, b5, h1, and h6. | The impella cp product and data logs were returned for analysis. the data logs confirm the clinical account of low pump flow. the pump was analyzed, and no signs of biomaterial were noted. simulated use testing was able to reproduce the low pump flow originally reported. further investigation found impeller blade damage. the root cause of the low pump flows was broken impeller blades. the root cause of the breakage was unable to be determined. the failure mode will be trended and monitored. note: this mdr is being reported late dueto retrospective analysis of prior broken impeller blade complaints.
36748 | 1220648-2021-01054 | 12533949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12533949&pc=OZD | 10/03/2016 | 09/27/2021 |  | 09/27/2021 |  |  | 10/14/2016 | 09/27/2021 |  |  | 06/30/2018 |  | 09/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1250794 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 62 YR | [5 empty] | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission |  |  |  |  | 10/07/2016 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting for high risk percutaneous coronary intervention. the impella cp was selected for hemodynamic support. shortly after support was initiated, the pump exhibited continuous suction alarms. the pump was removed and replaced without further issue. an investigation was completed on the returned device and broken impeller blades were identified. | [5 empty] | The device was returned and evaluated. the lower portion of the impeller blade was found to be fractured. simulated use testing resulted in lower than expected pump flow. the cause of the broken impeller blades could not be definitively determined, however, evidence suggests the blades came into contact with a foreign object inside the pump housing windows. note: this mdr is being reported late dueto retrospective analysis of prior broken impeller blade complaints.
36749 | 1220648-2021-01055 | 12533961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12533961&pc=OZD | 11/20/2018 | 09/27/2021 |  | 09/27/2021 |  |  | 11/20/2018 | 09/27/2021 |  |  | 07/31/2020 |  | 09/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1361612 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission |  |  |  |  | 11/20/2018 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting with cardiomyopathy for impella support. during the procedure, continuous suction alarms emitted and could not be resolved. the pump was removed and replaced without further issue. upon returned product investigation, damaged impeller blades were identified. | [5 empty] | The device and data logs were returned for analysis. the data logs confirmed suction alarms during support and simulated use testing successfully reproduced this failure. upon visual inspection of the device, the edges of the impeller blades appeared "abraded." because the fluoroscopic image of the pump was not returned, a definitive root cause of the fracture could not be determined. note: this mdr is being reported late dueto retrospective analysis of prior broken impeller blade complaints.
36750 | 1220648-2021-01056 | 12533976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12533976&pc=OZD | 09/02/2021 | 11/08/2021 |  | 09/27/2021 |  |  | 09/02/2021 | 09/27/2021 |  |  | 01/31/2023 |  | 09/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022036558 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/02/2021 |  |  |  |  | The complainant reported an 80 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. the impella 5.5 pump was inserted and initiated without any reported issues. during patient transfer to a new hospital (and subsequent transfer of pump to new hospital's console), the pump would not restart. attempts to restart the console were made (including transferring the pump back to the original console) but the pump could not be restarted. the decision was made remove the pump. no further impella support was required. | [5 empty] | The pump has been received for analysis. upon completion of our investigation, a supplemental mdr will be filed. | The device and data lots were returned for analysis. the data logs found the pump operating within specification until an "impella disconnected while running" alarm occurred. an inadvertent pump disconnect lead to a hole in the catheter which allowed blood to leak into the red pump handle and corrode the eeprom board.
36751 | 1220648-2021-01057 | 12533979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12533979&pc=OZD | 09/06/2021 | 09/27/2021 |  | 09/27/2021 |  |  | 09/06/2021 | 09/27/2021 |  |  | 06/30/2023 |  | 09/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022104381 |  | 0048-0045 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/06/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support. after explanting the device, the patient's access site was noted as "unstable" despite manual pressure being held. due to blood lost, 2 units of blood products were delivered. | [5 empty] | The device was not received from the customer. should we receive the device, or obtain any additional information on this event, a supplemental mdr will be filed.
36752 | 1220648-2021-01058 | 12533992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12533992&pc=OZD | 09/09/2021 | 12/15/2024 |  | 09/27/2021 |  |  | 09/09/2021 | 09/27/2021 |  |  | 04/30/2023 |  | 09/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET APAC |  |  | ABIOMED, INC. |  |  | US |  | 2022087344 |  | 0046-0037 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Other |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | Distributor |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Detachment of Device or Device Component | Loosening of Implant Not Related to Bone-Ingrowth | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/09/2021 |  |  |  |  | Additional information noted that the site found blue suture hub was separated from the tuohy/repositioning unit connection. the product damage contributed to the patient's blood loss. | The complainant reported a (b)(6) male presenting for hemodynamic support using the impella 5.0. during support, the patient endured blood loss of approximately 300ml. a nurse discovered the patient's access site bleed and was ultimately transfused with 2 units of blood products and escalated to venous-arterial ecmo. |  |  |  |  | The device was returned for analysis. a visual inspection found no damage or other abnormalities with the device. evidence suggests a user related issue was the cause of the bleed. | Additional information added to b5, h6 to capture device complaint for blue suture hub separation from the tuohy/repositioning unit connection. | The impella device has been received and an investigation to this event is underway. once the investigation has completed, a supplemental mdr will be filed with the results.
36753 | 1220648-2021-01059 | 12533997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12533997&pc=OZD | 09/09/2021 | 10/02/2025 |  | 09/27/2021 |  |  | 09/09/2021 | 09/27/2021 |  |  | 05/31/2023 |  | 09/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022091034 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2021 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical. during support, the patient developed a hematoma and purulent discharge at the impella site. antibiotics were given and a wound vacuum assisted closure was required. | [5 empty] | The device was discarded, therefore, an investigation could not be performed. should we received any additional information on this event, a supplemental mdr will be filed.
36754 | 12525469 | 12525469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12525469&pc=OZD | 09/01/2021 | 09/10/2021 | 09/10/2021 | 09/24/2021 | 09/10/2021 | 09/24/2021 |  | 09/24/2021 |  |  |  |  | 09/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No answer provided | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 14235 DA | [5 empty] | Malfunction | No | Yes | Defective Device | Protective Measures Problem | [5 empty] | Insufficient Information | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | In evening, impella 5.5 alarmed air in system. when trying to solve the issue nurse found that dials on the monitor was not working. impella monitor switched. charge nurse, nurse practitioner, and impella team aware. impella console exchanged with unit backup and removed from service. alarm review being conducted by impella representative.
36755 | 1220648-2021-01052 | 12529943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12529943&pc=OZD | 09/02/2021 | 09/24/2021 |  | 09/24/2021 |  |  | 09/02/2021 | 09/24/2021 |  |  | 05/31/2023 |  | 09/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022089510 |  | 0048-0024-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/02/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian female patient had impella cp optical pump inserted in the left femoral. the patient had access site bleeding and hematoma, as a result two (2) units of red blood cells (rbcs) was given, and surgical repair was done. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36756 | 1220648-2021-01046 | 12512384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12512384&pc=OZD | 08/25/2021 | 10/29/2021 |  | 09/22/2021 |  |  | 08/25/2021 | 09/22/2021 |  |  | 06/30/2023 |  | 09/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US | [7 empty] | 2022099128 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/25/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient suspected of enduring hemolysis during use of the device. as a result, the pump was removed. the patient was ultimately placed on an intra-aortic balloon pump for further support. | [5 empty] | The device was returned with the catheter severed. visual inspection found no abnormalities with the device. based on our analysis of the clinical information provided, we feel the cause of the reported hemolysis caused or contributed to by the patient's condition as the patient had previously undergone a bentall procedure. | The impella 5.5 was received and an investigation is underway. a supplemental mdr will be filed upon completion of our analysis.
36757 | 1220648-2021-01047 | 12512836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12512836&pc=OZD | 08/27/2021 | 11/17/2021 |  | 09/22/2021 |  |  | 08/27/2021 | 09/22/2021 |  |  | 06/30/2023 |  | 09/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022104958 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2021 |  |  |  |  | The complainant suspected a (b)(6) male patient presenting with cardiomyopathy enduring hemolysis during impella support. as a result, the pump was removed prematurely to be replaced with a new impella. | [5 empty] | The device and data logs were returned for analysis. a visual inspection found no damage or abnormalities with the device. data log analysis and simulated use testing found the pump to have operated within specification. we are unable to definitively determine the root cause of the reported hemolysis as we were unable to replicate the issue. | The impella 5.5 was received and an investigation is underway. a supplemental mdr will be filed upon completion of our analysis.
36758 | 1220648-2021-01048 | 12514384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12514384&pc=OZD | 08/30/2021 | 09/22/2021 |  | 09/22/2021 |  |  | 08/30/2021 | 09/22/2021 |  |  | 06/30/2023 |  | 09/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022100444 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/30/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in an unknown insertion site for myocarditis. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient developed bleeding from ecmo and impella insertion sites, for which was surgically opened to stop the bleeding. blood products were also administered. amount was not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36759 | 1220648-2021-01049 | 12516055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12516055&pc=OZD | 08/26/2021 | 11/27/2024 |  | 09/22/2021 |  |  | 08/27/2021 | 09/22/2021 |  |  | 12/31/2022 |  | 09/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022055482 |  | 005048-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 77 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombosis/Thrombus | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup |  |  | 08/27/2021 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | Additional information noted that the patient had also experienced thrombogenic complications (non-central nervous system). | The complainant reported a (b)(6) female patient had impella 2.5 pump inserted in the right axillary for high-risk percutaneous coronary intervention (hrpci). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that part of the 13fr peel away sheath had an issue, causing vascular damage (access site) and bleeding at the patients femoral access site, for which blood products were given along with catecholamine. insertion site was then switched from femoral to axillary. after surgical insertion of impella from the right subclavian artery, ischemia of the right upper limb was observed. the physician stated that limb¿s color wasn¿t very good when the patient arrived in the icu. no further information was provided. |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | Additional information for the initial mfr 1220648-2021-01049 was provided in sections b5, d6, and h6. this file is being submitted as part of a retrospective review of historical records. | The impella 2.5 was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
36760 | 1220648-2021-01050 | 12516153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12516153&pc=OZD | 08/31/2021 | 10/14/2024 |  | 09/22/2021 |  |  | 08/31/2021 | 09/22/2021 |  |  | 04/30/2023 |  | 09/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022073753 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  |  |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 69 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombocytopenia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 08/31/2021 |  |  |  |  | It was also reported that two days prior to pump removal, a pump pause occurred. the pump was successfully restarted at p-2. no abnormal events or alarms were noted prior to the pump pause. there was no reported patient injury from this issue. due to the clot that was observed in the patient, the pump was removed as a precautionary measure to avoid a potential pump stop. | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for cardiomyopathy. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that because of the patient developed hit, anticoagulant was switched from heparin to argatroban. act was managed in about 180 seconds. it was noted that a thrombus had formed at the site below the 3-branch and above the descending artery. surgical removal of the clot was performed while protecting the renal arteries with an occlusion balloon. after removal, any residual thrombus was removed using a fogati. no further information was provided. |  |  |  |  | The investigation for the reported thrombocytopenia, thrombus, and temporary pump stop has been completed. neither the product nor the data logs were returned for investigation. therefore, the cause of the pump pause was unable to be determined. after reviewing the clinical information, it was concluded that a clot was observed in the patient and the pump was removed as a precautionary measure to avoid a potential pump stop. the root cause of the thrombocytopenia was the observed formation of thrombus in the patient. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | The impella cp has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36761 | 1220648-2021-01044 | 12506779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12506779&pc=OZD | 08/24/2021 | 10/26/2021 |  | 09/21/2021 |  |  | 08/24/2021 | 09/21/2021 |  |  | 05/31/2023 |  | 09/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US | [7 empty] | 2022088097 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient enduring suspected hemolysis during impella support. the pump was removed earlier than anticipated, as a result. | [5 empty] | The device was returned and an investigation was conducted. visual inspection found no abnormalities. simulated use testing was conducted for hemolysis and the device operated without any issues. based on our findings, we feel the root cause was related to the patient's condition. | The impella 5.5 was received and an investigation is underway. a supplemental mdr will be filed at the completion of our analysis.
36762 | 1220648-2021-01045 | 12506786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12506786&pc=OZD | 08/24/2021 | 10/26/2021 |  | 09/21/2021 |  |  | 08/24/2021 | 09/21/2021 |  |  | 02/28/2023 |  | 09/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET | [7 empty] | 2022043710 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient suspected of enduring hemolysis during use of the device. multiple repositioning attempts were made, however, the device was ultimately removed and replaced. | [5 empty] | The device was returned and no abnormalities were observed during evaluation. we believe the root cause of the reported hemolysis was due to improper positioning. | The impella cp optical was received and an investigation is underway. a supplemental mdr will be filed at the completion of our analysis.
36763 | 12494791 | 12494791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12494791&pc=OZD | 08/15/2021 | 09/09/2021 | 09/09/2021 | 09/20/2021 | 09/09/2021 | 09/20/2021 |  | 09/20/2021 |  |  |  |  | 09/20/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP WITH SMARTASSIST |  |  | OZD |  |  | 0048-0024 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 23360 DA | [5 empty] | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Nurse assisted transfer from bed to chair using sara steady in early afternoon. no issues were noted during patient transfer with impella cp smart assist device. centimeter markings were checked after transfer to chair and the device was assessed to be in the same location as prior to the transfer. approximately 30 minutes later, impella device alarm sounded re: "no placement signal." staff immediately paged mechanical circulatory support (mcs) team, notified charge rn of the issue. assessed patient who was found to be hemodynamically stable and without symptoms. impella appeared to still be functioning with appropriate motor current, no other active alarms. centimeter markings were checked and found to be in the same position, no obvious kinks in catheter tubing were noted at that time. mcs arrived to bedside within minutes, primary service was called to bedside as well. writer and charge rn assisted patient back to bed. a bedside echo was done. rn checked again around this time with mcs and noticed a slight "bend" in the impella sheath, just distal to equipment stabilized on an armboard. impella continued to function and primary rn resumed care of patient while the issue continued to be assessed by mcs and chf services. patient remained hemodynamically stable throughout the event and was kept informed by mcs and chf of ongoing troubleshooting and plan. pump was replaced the following week. manufacturer response for heart pump, (brand not provided) (per site reporter). pending evaluation by manufacturer.
36764 | 1220648-2021-01043 | 12501369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12501369&pc=OZD | 07/01/2020 | 09/20/2021 |  | 09/20/2021 |  |  | 07/01/2020 | 09/20/2021 |  |  | 03/31/2022 |  | 09/20/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1477153 |  | 0550-0008 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/01/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 5.5 pump inserted for cardiomyopathy. there was difficulty placing the wire through introducer valve. the patient had moderate mitral regurgitation (mr), severe aortic insufficiency (ai), and small patent foramen ovale (pfo). as a result, the patient received two (2) units of packed red blood cells (prbcs) in the operating room (or). | [5 empty] | Device discarded by customer. the impella 5.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36765 | 1220648-2021-01037 | 12487912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12487912&pc=OZD | 08/14/2021 | 09/30/2021 |  | 09/17/2021 |  |  | 08/23/2021 | 09/17/2021 |  |  | 05/31/2023 |  | 09/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022085938 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Unspecified Infection | Ischemia | Renal Failure | [9 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 08/23/2021 | * | Invalid/Incomplete |  |  | Additional information was provided indicating renal failure occurred. as a result, chdf was performed. subsequently, infection at the access site was noted for which antibacterial agents were administered. | The complainant reported a male patient, unknown age had impella cp pump inserted in the right femoral for myocarditis. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient developed limb ischemia dueto thrombus, hemolysis (haptoglobin was administered four times) and diagnosed with heparin-induced thrombocytopenia (hit) after removal of impella. no further information was provided. |  |  |  |  | The product was returned for analysis. the pump was returned for review and did not show any signs of damage or biomaterial ingestion. the data logs show frequent suction throughout the case. the clinical states the leg ischemia was caused by the patient's hit + and unrelated to impella. the most likely root cause of the suspected hemolysis was patient condition of low lv volume. the renal failure was determined to have unknown/unrelated relation to impella. the infection was determined to be unknown/unrelated to impella. the impella device was confirmed to have been sterilized under validated conditions. the product was released with a completed and reviewed device history record under the abiomed quality system. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event. | The impella cp was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
36766 | 1220648-2021-01038 | 12488018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12488018&pc=OZD | 02/12/2021 | 09/17/2021 |  | 09/17/2021 |  |  | 08/23/2021 | 09/17/2021 |  |  | 10/31/2022 |  | 09/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2021011999 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 32 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/23/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the right femoral for myocarditis. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient developed an infection at the impella insertion site. the patient developed fever and redness was observed around the insertion site. the patient was administered antibiotics. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36767 | 1220648-2021-01039 | 12491040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12491040&pc=OZD | 04/09/2021 | 09/17/2021 |  | 09/17/2021 |  |  | 08/23/2021 | 09/17/2021 |  |  | 12/31/2022 |  | 09/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022023303 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/23/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient experienced bleeding at the impella access site. pressure hemostasis was applied, and the patient was administered blood products. amount is unknown. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36768 | 1220648-2021-01040 | 12491486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12491486&pc=OZD | 06/28/2021 | 09/17/2021 |  | 09/17/2021 |  |  | 08/24/2021 | 09/17/2021 |  |  | 10/31/2022 |  | 09/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021011900 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/24/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right axillary for post-cardiotomy cardiogenic shock (pccs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that after inserting of impella, numbness and cold sensation occurred on the peripheral side. when the device was removed, the numbness continued. the patient was diagnosed with vascular injury. as a result, the impella insertion site underwent dissection. a percutaneous transluminal angioplasty was used to suture the blood vessel. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36769 | 1220648-2021-01041 | 12491924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12491924&pc=OZD | 08/20/2021 | 09/17/2021 |  | 09/17/2021 |  |  | 08/24/2021 | 09/17/2021 |  |  | 06/30/2023 |  | 09/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022092747 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/24/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that soon after the impella device was implanted the patient showed signs of hemolysis. the patient¿s urine was the color of red wine. the impella was adjusted, and haptoglobin and blood products were administered. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36770 | 1220648-2021-01042 | 12491956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12491956&pc=OZD | 07/29/2021 | 09/17/2021 |  | 09/17/2021 |  |  | 08/25/2021 | 09/17/2021 |  |  | 10/31/2022 |  | 09/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021011981 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 49 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/25/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that an echo showed that the impella was in too deep which caused hemolysis. the device was adjusted, and the patient was administered blood products. the physician stated that the patient developed renal failure which may have been related to hemolysis. no treatment was rendered for renal failure. no further information was provided. | [5 empty] | The impella cp has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36771 | 1220648-2021-01036 | 12480661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12480661&pc=OZD | 02/26/2021 | 09/16/2021 |  | 09/16/2021 |  |  | 08/26/2021 | 09/16/2021 |  |  | 12/31/2022 |  | 09/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022028930 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for an unknown indication. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that on second day of impella support, hemolysis was confirmed via hemolytic urine. as a result, the impella's position was adjusted and haptoglobin was administered. hemolysis resolved after treatment. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36772 | 1220648-2021-01033 | 12484381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12484381&pc=OZD | 08/21/2021 | 10/22/2021 |  | 09/16/2021 |  |  | 08/21/2021 | 09/16/2021 |  |  | 09/30/2022 |  | 09/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 EU | [7 empty] | 1501106 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella 5.0 pump inserted in the right axillary artery. the patient had a suspected thrombus during the pump exchanged and patient suffered from a middle cerebral artery (m1) obstruction. | [5 empty] | Visual inspection of the returned product revealed biomaterial ingestion in the outflow cage and biomaterial on the outflow near the pressure sensor. there was also moderate deformation of the inflow cage. the entire pump was covered in mold. no other relevant damage or abnormalities were found. when the pump was run with a purge cassette, a mass of mold/biomaterial was ejected and the pump purged normally with no purge pressure alarms reproduced. the pump ran with no issues. in conclusion, after reviewing the clinical information, it was confirmed that a clot was suspected to be ingested during support and the pump was removed as a precautionary measure to avoid a potential pump stop. the cause of the cva/stroke was determined to be unrelated to this impella. given the clinical narrative, the stroke most likely resulted from using the same graft during exchange for the third pump. updated health effect - clinical code from 1770 to 4440. | The impella 5.0 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36773 | 1220648-2021-01035 | 12484463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12484463&pc=OZD | 02/04/2021 | 09/15/2021 |  | 09/16/2021 |  |  | 08/26/2021 | 09/16/2021 |  |  | 09/30/2022 |  | 09/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 1502424 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that on day one of impella support, hemolysis was confirmed via hemolytic urine. as a result, the impella's position was adjusted and haptoglobin was administered. hemolysis resolved after treatment. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36774 | 1035166-2021-00090 | 12474860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12474860&pc=DYB | 08/11/2021 | 09/29/2021 |  | 09/15/2021 |  |  | 08/27/2021 | 09/15/2021 |  |  | 10/01/2023 |  | 09/15/2021 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-19434 |  | 0052-3021 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The customer reported that second oscor sheath from new pump kit was used to complete the procedure (lot number unknown). the cause of patient death was acute respiratory rest followed by cardiogenic shock, and was not attributed to the oscor device. | It was reported that on (b)(6) 2021 a (b)(6) year old female patient under went an acute myocardial infarction (ami)/ cardiogenic shock (cgs) procedure. the sheath was found to be kinked at the groin where the access into the vein occurred. there was a 5 min delay in the procedure. another impella rp set was used to complete the procedure. the left femoral vein was more tortuous than the right in most patients, as was the case with this patient. it took a dip toward the patients right side and there was a sharp turn up into the inferior vena cava (ivc). patient expired within an hour after transfer to cardiac care or critical care units (ccu). the cause of the death is unknown at this time but was not attributed to the device. the device is not returning because it was discarded by the user facility. |  |  |  |  | Oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | The device was used in treatment. the device was not returned for analysis, therefore, the exact cause of the product issue cannot be determined and the clinical observation could not be confirmed. however, the following controls are in place to mitigate the reported product issue. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure (abiomed introducer sheath in-process and final inspections): measure dimensions: measure tip i.d. of sheath using appropriate pin gauges to be in range of 0.305" - 0.307" for 23f. visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. visual inspection: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. per ifu never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
36775 | 1220648-2021-01034 | 12475779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12475779&pc=OZD | 07/16/2021 | 09/30/2021 |  | 09/15/2021 |  |  | 08/20/2021 | 09/15/2021 |  |  | 10/31/2022 |  | 09/15/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021014912 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 90 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 08/20/2021 | * | Invalid/Incomplete |  |  | Additional information was provided indication there was and access site infection. as a result, antibacterial agents were administered. | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient developed access site bleed due to fixation method. the issue was corrected by readjusting the fixation of the pump to the patient. the patient required administration of blood product when the patient was in shock from bleeding. amount of blood products that was administered was not provided. no further information was provided. |  |  |  |  | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36776 | 1035166-2021-00105 | 12466746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12466746&pc=DYB | 08/08/2021 | 01/31/2022 |  | 09/14/2021 |  |  | 08/17/2021 | 09/14/2021 |  |  | 12/01/2023 |  | 09/14/2021 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | C1-19947 |  | 0052-3006 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 59 YR |  | Female |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported the patient was presented to the physician with the condition of depressed left ventricle function and moderate aortic insufficiency (ai). the 23f introducer was used for a cardiomyopthy procedure, for a bridge to transplant. a bleed occurred between the 23fr introducer and graft once the graft locks were deployed, not a full hemostasis. it is not known how much blood was lost, however, no blood products were required and the patient was not harmed as a result of the bleed. no additional information is available. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete. | The following sections were updated in follow-up 1: b4, g3, g6, h2, h6, and h10 the device was used for treatment. the device was not returned for analysis as device was discarded, therefore, the clinical observation could not be confirmed. the investigation was focused on a review of product documentation. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. however, following controls are in place to mitigate the reported product issue. per qa procedure adelante s2s introducer sheath in-process and final inspection: sample size: 100% inspection. 12.3 vacuum leak test - perform pressure and vacuum leak test per procedure. per instructions for use (ifu): when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report isbased upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36777 | 12459442 | 12459442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12459442&pc=OZD | 08/28/2021 | 09/07/2021 | 09/07/2021 | 09/13/2021 | 09/07/2021 | 09/13/2021 |  | 09/13/2021 |  |  |  |  | 09/13/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST SET |  |  | OZD |  |  | LOG1508821329 |  |  | ABIOMED INC. |  |  | US |  | 315419 |  |  |  | Yes | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 20075 DA | [5 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Insufficient Information | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | After changing purge cassette on impella vad, a leak was noted on connector of purge line and purge cassette. after consulting with abiomed, the manufacturer, cassette was changed again with no resolution in problem. second recommendation from the manufacturer was to cut cassette off and splice line with 19-gauge needle. after performing this intervention, purge pressure and flow reestablished. pump working correctly.
36778 | 1220648-2021-01030 | 12461397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12461397&pc=OZD | 08/20/2021 | 10/12/2021 |  | 09/13/2021 |  |  | 08/20/2021 | 09/13/2021 |  |  | 05/31/2023 |  | 09/13/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU | [7 empty] | 2022091682 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/20/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient had impella cp pump inserted in the femoral for cardiogenic shock (cgs). the patient had hemolysis and the pump was removed. | [5 empty] | Data logs from the case were not returned for investigation. visual inspection of the returned pump revealed multiple significant kinks and shearing on the catheter around the 85-cm mark. all motor phase resistances were within specification. the cause of the hemolysis could not be determined since the pump was unable to be tested. | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36779 | 1220648-2021-01031 | 12461402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12461402&pc=OZD | 08/16/2021 | 09/21/2024 |  | 09/13/2021 |  |  | 08/16/2021 | 09/13/2021 |  |  | 08/31/2021 |  | 09/13/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1444435 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 74 YR |  | Male |  |  | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Insufficient Information | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 08/16/2021 |  |  |  |  | It was additionally reported the patient experienced a severe gastrointestinal bleed. no information was provided if any interventions were performed for the bleeding. the decision was made to not provide further treatment to the patient, and the patient was made a do not resuscitate order. although the medical team was advised the pump was in danger of a pump stop as the pump was in use well beyond the design days of use, the medical team advised the pump would not be replaced. it was reported the pump stopped after 30 days of use and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a (b)(6) year-old (b)(6) male patient had impella 5.0 pump inserted in the right axillary for post-cardiotomy cardiogenic shock (pccs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. the complainant reported that the impella 5.0 stopped on (b)(6) 2021 after 30 days of use. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2021-01031 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6. | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36780 | 1220648-2021-01032 | 12461404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12461404&pc=OZD | 02/21/2021 | 11/30/2024 |  | 09/13/2021 |  |  | 08/17/2021 | 09/13/2021 |  |  | 05/31/2021 |  | 09/13/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1418897 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | Asian | Death | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 08/17/2021 |  |  |  |  | The complainant reported a (b)(6)-year-old male patient had impella 5.0 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient developed bleeding at the developed bleeding at the artificial blood vessel anastomosis. as a result, the patient required blood products. amount was not provided. no further information was provided. | It was additionally reported that the patient experienced thrombocytopenia on (b)(6) 2021. the patient received platelets; however, the number of units is unknown. on (b)(6) 2021 there was a report of low purge flow. it was noted there was no clear harm to the patient. there was no information provided on how the purge issue was resolved. it was also reported that the patient expired on (b)(6) 2021, approximately 2 months after the removal of the impella. it was noted that the patient's outcome was related to the impella. there were no further details received regarding the patient's condition, type of death and cause of death. given the limited information surrounding the details the patient's condition and cause of death the impella cannot be excluded as a possible contributing factor. |  |  |  |  | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation for the reported issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2021-01032 was provided in sections b5, d4, d6, e1, g1, h1, and h6.
36781 | 1220648-2021-01029 | 12456578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12456578&pc=OZD | 08/14/2021 | 09/10/2021 |  | 09/10/2021 |  |  | 08/14/2021 | 09/10/2021 |  |  | 06/30/2023 |  | 09/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022095906 |  | 0550-0008 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 40 YR | [5 empty] | Death | Yes | No | Use of Device Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 08/14/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with cardiomyopathy. the impella was inserted for hemodynamic support as the patient deteriorated. during support, echo imagery noted clotting. the patient suffered a cerebrovascular accident. the patient ultimately was withdrawn from care and expired. clot burden was found within the entire cannula upon removal. it was noted that anticoagulation had been removed from the purge system without the surgeon's knowledge. use of this device without anticoagulation is not recommended, per instructions for use. | [5 empty] | The impella 5.5 was discarded by the customer and therefore, investigation of device was not possible. a supplemental report will be filed should we receive any new information on this event.
36782 | 1220648-2021-01024 | 12438737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12438737&pc=OZD | 08/07/2021 | 09/22/2024 |  | 09/08/2021 |  |  | 08/11/2021 | 09/08/2021 |  |  | 04/30/2023 |  | 09/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022082485 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 59 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 08/11/2021 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a (b)(6) male patient had impella 5.0 pump inserted in the right axillary for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021. explant date was not provided. it was reported that a ct scan showed mediastinal bleeding. surgical treatment was difficult because the location of the bleeding could not be determined. as a result, the patient was administered red blood cells, fresh frozen plasma, and platelet count. it was noted that the patient was also on pcps, and therefore the physician was unable to determine if impella caused the bleeding. no further information was provided. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The impella 5.0 has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36783 | 1220648-2021-01025 | 12438743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12438743&pc=OZD | 02/09/2021 | 09/08/2021 |  | 09/08/2021 |  |  | 08/17/2021 | 09/08/2021 |  |  | 09/30/2022 |  | 09/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021006125 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/17/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that during impella support the patient developed access site bleed. it was noted that the angle did fall due to hanging and that the patient also developed bleeding from ecmo site. as a result, the patient received blood products, amount was not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36784 | 1220648-2021-01026 | 12438859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12438859&pc=OZD | 08/02/2021 | 09/17/2024 |  | 09/08/2021 |  |  | 08/09/2021 | 09/08/2021 |  |  | 06/30/2023 |  | 09/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022095904 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2021 |  |  |  |  | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a (b)(6) year old male patient presenting for hemodynamic support with the impella 5.5 device. the device perforated the patient's left ventricle when the physician flipped the heart to evaluate the coronary sinus. the perforation was repaired successfully. it was noted that the patient's atrial septal defect tissue was "frail." |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided. | The impella 5.5 was not received from the customer and therefore, investigation of device was not possible. should we receive the device, or if any new information be obtained, a supplemental mdr will be filed.
36785 | 1220648-2021-01027 | 12438866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12438866&pc=OZD | 06/05/2021 | 09/08/2021 |  | 09/08/2021 |  |  | 08/25/2021 | 09/08/2021 |  |  | 01/31/2023 |  | 09/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022033033 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 89 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/25/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that on day one of impella support, hemolysis was confirmed via hemolytic urine. as a result, the impella's position was adjusted and haptoglobin was administered. hemolysis resolved after treatment. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36786 | 1220648-2021-01028 | 12440626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12440626&pc=OZD | 12/10/2020 | 09/08/2021 |  | 09/08/2021 |  |  | 12/10/2020 | 09/08/2021 |  |  | 08/31/2022 |  | 09/08/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1496975 |  | 0550-0008 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | OTHER |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 12/10/2020 | * | Invalid/Incomplete |  |  | In (b)(6) of 2020 a patient was admitted and placed on impella cp hemodynamic support due to cardiomyopathy and septic and cardiogenic shock. when his care needed escalation an impella 5.5 replaced the cp. after 3 days of support on the 5.5 the medical team observed a hematoma at the access site which prompted blood replacement of 3 units of red blood cells. | [5 empty] | The cause of the access site bleeding was not determined as no product was returned for investigation. the team discarded the product in (b)(6) of 2020 and the complaint was opened months later after crf review. the failure mode will be monitored and trended.
36787 | MW5103796 | 12446641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12446641&pc=OZD | 09/04/2021 | 09/04/2021 |  | 09/08/2021 |  |  |  | 09/08/2021 |  |  | 03/31/2023 |  | 09/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 51 YR |  |  | 80 |  |  | Injury | No | Yes | Output Problem | [6 empty] | Insufficient Information | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Impella catheter did not work, it had to be replaced. the cause of the error is uncertain. fda safety report id # (b)(4).
36788 | 1035166-2021-00088 | 12432959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12432959&pc=DYB | 08/11/2021 | 01/27/2022 |  | 09/07/2021 |  |  | 08/11/2021 | 09/07/2021 |  |  | 03/01/2024 |  | 09/07/2021 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3025-JP |  |  | OSCOR INC. |  |  | US |  | DP-12994 |  | 0052-3025-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Distributor |  |  |  |  | 56 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Fluid/Blood Leak | Material Too Rigid or Stiff | Device Damaged by Another Device |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a (b)(6) male underwent a ami/cgs procedure and at the time of the impella insertion, a long 14fr introducer was inserted into the body, and when trying to insert impella, the hemostatic valve was damaged and there was a side leak of blood, so the sheath was to be reinserted. since it was not possible to pass through the bent part in the sheath, it was decided to issue a new 14fr introducer. there was no delay in the procedure and there was no patient harm reported. the issue was resolved using another introducer. | [5 empty] | This event was submitted in error. oscor does not own the 510k and does not have reporting responsibility. please void the initial mdr submitted by oscor on 09/07/2021 under mdr #1035166-2021-00088. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete.
36789 | 1035166-2021-00099 | 12421682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12421682&pc=DYB | 01/02/2020 | 09/03/2021 |  | 09/03/2021 |  |  | 09/01/2021 | 09/03/2021 |  |  | 09/01/2022 |  | 09/03/2021 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-17561 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 70 YR | [5 empty] | Malfunction | No | Yes | Material Too Rigid or Stiff | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that on (b)(6) 2020 a (b)(6) male underwent a procedure for insertion of an impella cp at the right femoral artery via the 14fr x 13cm introducer. the physician had a difficult time cracking the peel away sheath outside of the body. the physician used a scalpel to carefully cut the peel away sheath. the issue was resolved by placement of the impella and removal of the introducer. the patient was admitted with cancer, flu like symptoms, and a vf arrest. there was no patient harm reported, however, the patient expired when the family chose to withdraw care on (b)(6) 2020. | [5 empty] | The device was used in treatment. one 14f 13cm long abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. blood was found on and inside the sheath as well as the sideport tubing. the sheath was returned already peeled/cut apart upon receipt. upon evaluation of the returned device, it was found that the sheath failed to peel properly along the scoreline. the peel end tip of the sideport portion of the sheath started to look wavy and stopped at approximately 4cm from the distal tip. the sideport portion of the sheath tube was cut down the rest of the length of the tube. the beginning of the peel, directly under the hub looks jagged. both halves of the inside of the hub were inspected under a microscope at 0.63 power and peeled apart smoothly with no excess material found on either side. the hemostatic valve tore in half normally as well. the break force data of the returned lot was reviewed and it was found that all samples were within manufacturing specification. possible root cause for peel failure could be tooling wear and tear at the extrusion process. retraining will be conducted for procedures to prevent recurrence of this issue. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure for adelante s2s introducer sheath in-process and final inspections, ensures sheaths are inspected for visual, dimensional, and peel functional compliance. destructive testing sampling plan. ansi z 1.4, special level iv, and aql 1.0 reduced. manually break the sheath and hub and verify that the seals split easily without extreme elongation. also verify that the split cap and seals remain secure and do not break free or loosen from the sheath hub. a) manually peel the sheath and verify the sheath peels easily along the sheath body and tip. per procedure for adelante, adelante-s and adelante-s2 sheath extrusion the two score line model sheath tubes: perform spc charting on the peel strength and critical dimensions. this is to be performed on every 100th extruded sheath. measure the peel strength of the extruded sheath per procedure. measure dimensions a-f using the vertex or optical comparator and the laser micrometer for the ovality. peel test results and dimensional measurements are to be recorded in the extrusion spc data collection sheet. the spc data collection sheet is to be saved under each work order number in the appropriate folder. select the extrusion tooling (pin/die) for the part to be extruded. prior to installing the tooling, measure and inspect the tooling to ensure its suitability for use. document tooling measurement and condition per procedure. per ifu flush the sheath with 5 cc of saline immediately before peeling the sheath away in order to minimize backbleeding. first, completely withdraw the sheath from the vessel. then remove the sheath by sharply snapping the tabs of the valve housing in a plane perpendicular to the long axis of the sheath to split the valve and peel the sheath apart. no further follow-up is required. based on the investigation, a capa is not required. qa & manufacturing training was conducted to prevent further occurrence of this event. oscor will continue to monitor this event type and risk.
36790 | 1220648-2021-01022 | 12422995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12422995&pc=OZD | 10/06/2020 | 09/03/2021 |  | 09/03/2021 |  |  | 08/23/2021 | 09/03/2021 |  |  | 06/30/2022 |  | 09/03/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1486359 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 57 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/23/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male patient had impella cp pump inserted in the left femoral for acute myocardial infarction (ami)/ cardiogenic shock (cgs). the patient had renal failure that was due to hemolysis and hemodialysis was required. | [5 empty] | Neither the product nor the data logs were returned for investigation. after reviewing the clinical information, it was determined that the cause of the reported hemolysis was likely patient condition, specifically the patient's insufficient volume. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36791 | 1220648-2021-01023 | 12423013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12423013&pc=OZD | 01/05/2021 | 09/03/2021 |  | 09/03/2021 |  |  | 08/11/2021 | 09/03/2021 |  |  | 01/31/2022 |  | 09/03/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1474554 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/11/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) female patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient experienced a vascular injury. at the time of removal, there was tearing of the blood vessel at the insertion site. as a result, the blood vessel was sutured to repair damage. no further information was provided. | [5 empty] | The impella cp was has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36792 | 1035166-2021-00098 | 12423733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12423733&pc=DYB | 09/20/2019 | 09/03/2021 |  | 09/03/2021 |  |  | 09/01/2021 | 09/03/2021 |  |  | 10/01/2020 |  | 09/03/2021 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | CR-02070 |  | 0052-3021 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [11 empty] | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that a complaint involving 2 oscor 23fr introducers while attempting to pass the impella rp through the first introducer, the device kinked which prevented the impella from advancing. the introducer was removed and replaced with another oscor 23fr introducer, however, this device kinked in two locations and could not advance the rp through the end of the introducer, past the device's outflow cage. upon removal of the rp device, the teardrop detached from the impella and required a surgical snare for removal. the patient was in stable condition, therefore, the physician planned to rest the patient and potentially try again with a new device in the morning. no additional information available. affected lots reported: c1-14501 & cr-02070. reference mdr 1035166-2021-00096 for lot number c1-14501. | [5 empty] | The device was used in treatment. two abiomed 23f introducer sheaths were returned from the customer along with various sized dilators. there were no other accessories. blood was found on all components. the introducer sheaths were from two different lots, but they were received back from the customer in the same bag. since they are the same model number, it is not possible to differentiate one lot from the other, so they were referred as sheath 1 and sheath 2. upon evaluation of sheath #1, it was found that there was a kink in the sheath tube at approximately 13cm from the distal tip. also, the hemostatic valve was torn and the distal tip looked normal. upon evaluation of sheath #2, it was found that there were 2 kinks in the sheath tube at approximately 12cm from the distal tip and 6cm from the distal tip. the hemostatic valve was heavily damaged, had a hole in it and was starting to dislodge from the hub. also, the circumference around the distal tip edge is damaged. according to the event description summary, the physician could not pass two devices through both sheaths without them kinking. per drawing the id of the sheath should be .306" ± .001". the sheath was measured with a pin gauge and was within manufacturing specifications at .305". other devices used in the procedure with oscor sheath were not returned to oscor so a fit check could be performed. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure (abiomed introducer sheath in-process and final inspection). measure dimensions: measure tip i.d. of sheath using appropriate pin gauges to be in range of 0.305" - 0.307" for 23f. visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. this is performed by qa on the first 5 consecutive properly tipped parts. per visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. this is performed by qa on 100% of the sheaths. per ifu never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
36793 | 1035166-2021-00096 | 12423768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12423768&pc=DYB | 09/20/2019 | 09/03/2021 |  | 09/03/2021 |  |  | 09/01/2021 | 09/03/2021 |  |  | 07/01/2021 |  | 09/03/2021 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-14501 |  | 0052-3021 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [11 empty] | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that a complaint involving 2 oscor 23fr introducers while attempting to pass the impella rp through the first introducer, the device kinked which prevented the impella from advancing. the introducer was removed and replaced with another oscor 23fr introducer, however, this device kinked in two locations and could not advance the rp through the end of the introducer, past the device's outflow cage. upon removal of the rp device, the teardrop detached from the impella and required a surgical snare for removal. the patient was in stable condition, therefore, the physician planned to rest the patient and potentially try again with a new device in the morning. affected lots reported: c1-14501 & cr-02070. reference mdr 1035166-2021-00098 for lot number cr-02070. | [5 empty] | The device was used in treatment. two abiomed 23f introducer sheaths were returned from the customer along with various sized dilators. there were no other accessories. blood was found on all components. the introducer sheaths were from two different lots, but they were received back from the customer in the same bag. since they are the same model number, it is not possible to differentiate one lot from the other, so they were referred as sheath 1 and sheath 2. upon evaluation of sheath #1, it was found that there was a kink in the sheath tube at approximately 13cm from the distal tip. also, the hemostatic valve was torn and the distal tip looked normal. upon evaluation of sheath #2, it was found that there were 2 kinks in the sheath tube at approximately 12cm from the distal tip and 6cm from the distal tip. the hemostatic valve was heavily damaged, had a hole in it and was starting to dislodge from the hub. also, the circumference around the distal tip edge is damaged. according to the event description summary, the physician could not pass two devices through both sheaths without them kinking. per drawing the id of the sheath should be .306" ± .001". the sheath was measured with a pin gauge and was within manufacturing specifications at .305". other devices used in the procedure with oscor sheath were not returned to oscor so a fit check could be performed. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure (abiomed introducer sheath in-process and final inspection). measure dimensions: measure tip i.d. of sheath using appropriate pin gauges to be in range of 0.305" - 0.307" for 23f. visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. this is performed by qa on the first 5 consecutive properly tipped parts. per visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. this is performed by qa on 100% of the sheaths. per ifu never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
36794 | 1220648-2021-01020 | 12414316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12414316&pc=OZD | 07/05/2021 | 05/23/2025 |  | 09/02/2021 |  |  | 08/11/2021 | 09/02/2021 |  |  | 04/30/2023 |  | 09/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022081084 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 61 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup |  |  | 08/11/2021 |  |  |  |  | Additional information noted the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that during impella support the patient developed limb ischemia. to treat the issue, a reverse insertion from the ecmo circuit was performed. no further information was provided. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. | The investigation for the reported limb ischemia issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the limb ischemia was most likely the condition/size of the patient vessel. this report is being filed as part of a retrospective review of historical records. | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36795 | 1220648-2021-01021 | 12414320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12414320&pc=OZD | 08/12/2021 | 09/30/2021 |  | 09/02/2021 |  |  | 08/12/2021 | 09/02/2021 |  |  | 09/30/2021 |  | 09/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1461770 |  | 005042 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 08/12/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) white female patient had impella 2.5 pump inserted in the right femoral artery for high risk percutaneous coronary intervention (hrpci). the introducer 13fr was placed, pump inserted over it, but once in the aorta the aic alarmed, the physician attempted to remove the catheter and peel away sheath but the angiogram showed perforation of the iliac artery. a vascular surgeon performed cut down and closure at bedside in cardiac catheterization (cath) laboratory (lab). | [5 empty] | The returned 13fr introducer was peeled away, but large kinks were seen in the introducer, likely a result of the patient's highly calcified vessels. the root cause of the vascular damage was likely due the sheath kink as a result of the calcification, leading to difficulties inserting and removing the pump. no abnormalities were found on the returned pump. | The impella 2.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36796 | 1220648-2021-01016 | 12406336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12406336&pc=OZD | 08/05/2021 | 09/30/2021 |  | 09/01/2021 |  |  | 08/05/2021 | 09/01/2021 |  |  | 05/31/2023 |  | 09/01/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022091935 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 65 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 08/05/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) american indian or alaska native male patient had impella 5.5 pump inserted. the pump stopped and another pump was used. | [5 empty] | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The root cause of the pump stop was due to ingested biomaterial.
36797 | 1220648-2021-01017 | 12406683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12406683&pc=OZD | 02/24/2021 | 09/01/2021 |  | 09/01/2021 |  |  | 08/05/2021 | 09/01/2021 |  |  | 08/31/2022 |  | 09/01/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1502522 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/05/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient developed a hematoma at the impella access site. as a result, blood products were administered. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36798 | 1220648-2021-01018 | 12406849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12406849&pc=OZD | 08/04/2021 | 10/11/2021 |  | 09/01/2021 |  |  | 08/04/2021 | 09/01/2021 |  |  | 06/30/2023 |  | 09/01/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP EU | [7 empty] | 2022092371 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 62 YR | [5 empty] | Injury | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock. the impella cp optical was selected for hemodynamic support. resistance was reported when sliding the pump over the guidewire, resulting in a kink to the wire and cannula. both were removed, however, balloon insertion was required to relieve the kink in the pump's cannula so it could be safely removed. there was no injury sustained as a result. | [5 empty] | The device was returned for analysis. a visual inspection confirmed the reported of a kinked cannula. evidence suggests the cause of the cannula kink was due to cardiopulmonary resuscitation (cpr) during pump insertion. | The impella device has not yet been received, but is in en-route to our facility for analysis. a supplemental mdr will be filed upon a completed investigation.
36799 | 1220648-2021-01019 | 12406862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12406862&pc=OZD | 08/02/2018 | 09/18/2024 |  | 09/01/2021 |  |  | 08/02/2018 | 09/01/2021 |  |  | 04/30/2020 |  | 09/01/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | CP C-7 NO ACCESS PUMP SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1339701 |  | 0048-0026 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Material Twisted/Bent | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2018 |  |  |  |  | The complainant reported a (b)(6) male patient presenting with cardiomyopathy. during support, the pump fell out of the ventricle an attempted repositioning. during attempted re-advancement over the aortic valve, the pump "folded" in the aortic root. a snare was required to straighten and remove the device. the pump was then re-inserted without further issue. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella cp was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. per an analysis clinical details and results of the investigation, "the clinical states the physician tried to wirelessly cross the valve which is against the ifu 0048-9007". the impella device is not indicated for wireless re-access. the root cause is improper technique. this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr was provided in sections b2, b5, h1, and h6.
36800 | 1220648-2021-01015 | 12395875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12395875&pc=OZD | 02/05/2021 | 12/15/2024 |  | 08/31/2021 |  |  | 08/03/2021 | 08/31/2021 |  |  | 08/31/2021 |  | 08/31/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1436980 |  | 005048-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Hospitalization | Required Intervention | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/03/2021 |  |  |  |  | It was additionally reported that the patient experienced hemolysis during impella 2.5 support. the p level was decreased, which improved the hemolysis, however the hemolysis was present until the device was explanted. | In (b)(6) of 2021 an impella 2.5 was placed for hemodynamic support of a high risk coronary angioplasty patient in (b)(6). the 2.5 allowed for successful angioplasty. the pump remained on for continued support and was weaned and explanted after 3 days. six months later the case was reviewed in the jpvad registry and an access site bleed with hematoma was reported. the bleed was treated by infusion of red blood cells. |  |  |  |  | The medical center in japan did not return for analysis the impella pump product. as such, the root cause could not be determined. the access site bleeding failure mode will be monitored and trended. | Additional information was added for reported hemolysis. the cause of the hemolysis was not determined as no product or data was returned for review. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleed with hematoma and blood transfusion in on-going at this time. further clinical details have been requested. the impella pump product was discarded. when the investigation is completed, a final report will be submitted.
36801 | 1220648-2021-01010 | 12383973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12383973&pc=OZD | 08/02/2021 | 08/30/2021 |  | 08/30/2021 |  |  | 08/02/2021 | 08/30/2021 |  |  | 04/30/2023 |  | 08/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022078031 |  | 0048-0045 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | USER F |  |  |  |  | 56 YR | [5 empty] | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/02/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient developing an access site bleed during use of the device. a hematoma formed and required evacuation and surgical repair to close the site. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36802 | 1220648-2021-01011 | 12384499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12384499&pc=OZD | 02/01/2021 | 08/30/2021 |  | 08/30/2021 |  |  | 08/03/2021 | 08/30/2021 |  |  | 10/31/2022 |  | 08/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2021012002 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 08/03/2021 | * | Invalid/Incomplete |  |  | The complainant reported a patient whose age and gender were not provided had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that during impella support the patient developed access site bleed. as a result, hand pressure hemostasis was performed and blood products were administered. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36803 | 1220648-2021-01012 | 12384518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12384518&pc=OZD | 12/18/2020 | 08/30/2021 |  | 08/30/2021 |  |  | 08/03/2021 | 08/30/2021 |  |  | 03/31/2022 |  | 08/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1476729 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Cardiac Tamponade | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/03/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for post-cardiotomy cardiogenic shock (pccs). the patient was implanted on (b)(6)2021 and explanted on (b)(6) 2021. it was reported that during impella support cardiac tamponade was observed, treatment details were not provided. subsequently, the patient experienced a vascular injury for which treatment was required. details of the treatment were not provided. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed
36804 | 1220648-2021-01013 | 12384609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12384609&pc=OZD | 06/18/2021 | 08/30/2021 |  | 08/30/2021 |  |  | 08/04/2021 | 08/30/2021 |  |  | 07/31/2021 |  | 08/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1432562 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/04/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported acute lower extremity thrombosis was noted after removal of impella. arterial thrombectomy with fogarty catheter was performed. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36805 | 1220648-2021-01014 | 12384632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12384632&pc=OZD | 08/04/2021 | 08/30/2021 |  | 08/30/2021 |  |  | 08/04/2021 | 08/30/2021 |  |  | 08/31/2021 |  | 08/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1436980 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 42 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/04/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella 2.5 pump inserted in the right axillary for an unknown indication. it was reported that the patient developed limb ischemia during impella support. as a result, bypass was performed with the ecmo circuit, a reverse sheath was inserted in the brachial artery. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36806 | 1220648-2021-01006 | 12374773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12374773&pc=OZD | 07/09/2021 | 10/14/2024 |  | 08/26/2021 |  |  | 08/03/2021 | 08/26/2021 |  |  | 04/30/2023 |  | 08/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022079771 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 56 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Fracture | Defective Device | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 08/03/2021 |  |  |  |  | The complainant reported a (b)(6) year-old male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that during impella support the patient developed sepsis. as a result, antibiotics were administered. no further information was provided. | It was reported the physician had difficulty inserting the guidewire as it was bent and the 14 french introducer was damaged during insertion attempts. a short sheath was unable to be used due to the bend of the blood vessel so a new 14 french introducer was used. the 14 french introducer and guidewire were both replaced. |  |  |  |  | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information for the initial mfr 1220648-2021-01006 was provided in sections b1, b5, e1, g1, h1, and h6. the investigation for the reported guidewire issues, valve damage, and sepsis has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the guidewire issues was unable to be determined as no product or data was returned for this investigation. the cause of the valve damage was unable to be determined as no product was returned for this investigation. the root cause of the sepsis has an unknown relationship with impella. the impella device was confirmed to have been sterilized under validated conditions and there was no evidence to indicate that the sepsis was related to impella. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
36807 | 1220648-2021-01007 | 12374806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12374806&pc=OZD | 05/02/2021 | 08/26/2021 |  | 08/26/2021 |  |  | 08/03/2021 | 08/26/2021 |  |  | 11/30/2022 |  | 08/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022032137 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/03/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in an unknown location for an unknown indication. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that during impella support the patient developed hemolysis. as a result, haptoglobin was administered. subsequently, the patient developed access site bleed for which blood products were administered. vascular sutures were also added. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36808 | 1220648-2021-01008 | 12374885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12374885&pc=OZD | 01/09/2021 | 08/26/2021 |  | 08/26/2021 |  |  | 08/03/2021 | 08/26/2021 |  |  | 10/31/2022 |  | 08/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2021012005 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/03/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that hemolysis and kidney failure were noted during impella support. as a result, haptoglobin was administered. it was noted that the decrease in renal function was deemed related to hemolysis. the patient did require escalation due to kidney failure. access site bleed was also reported. as a result, hand pressure hemostasis was performed and blood products were administered, amount unknown. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36809 | 1220648-2021-01009 | 12374905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12374905&pc=OZD | 07/23/2021 | 08/26/2021 |  | 08/26/2021 |  |  | 08/03/2021 | 08/26/2021 |  |  | 07/31/2022 |  | 08/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1490580 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/03/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old male patient had impella cp pump inserted in the right femoral for high-risk percutaneous coronary intervention (hrpci). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that hemolysis was noted during impella support. as a result, haptoglobin was administered. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36810 | 12364846 | 12364846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12364846&pc=OZD | 07/12/2021 | 08/02/2021 | 08/02/2021 | 08/25/2021 | 08/02/2021 | 08/25/2021 |  | 08/25/2021 |  |  |  |  | 08/25/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA PURGE CASSETTE |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | 0043-0001 |  | No answer provided | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 23360 DA | [5 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella cassette was due to be changed, when new cassette primed and attached, noted leaking at yellow luer connection and appeared to have small crack at hub. impella rep called, and told to do cassette bypass. np and md above at bedside to complete procedure per instructions, pt tolerated well with resolution of leaking and alarms.
36811 | 12352765 | 12352765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12352765&pc=OZD | 08/11/2021 | 08/17/2021 | 08/17/2021 | 08/23/2021 | 08/17/2021 | 08/23/2021 |  | 08/23/2021 |  |  |  |  | 08/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | No answer provided | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 22265 DA | [5 empty] | Malfunction | No | Yes | Crack | [6 empty] | Insufficient Information | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella cassette cracked. bypass performed by icu team as directed by abiomed.
36812 | 1220648-2021-00999 | 12356525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12356525&pc=OZD | 07/02/2021 | 10/09/2024 |  | 08/23/2021 |  |  | 08/02/2021 | 08/23/2021 |  |  | 01/31/2023 |  | 08/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022057097 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Sepsis | Rupture | [9 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2021 |  |  |  |  | The complainant reported an asian patient had impella cp pump inserted. the patient's blood vessel was damage during removal, as a result ten (10) units of blood transfusion was given, and a beyer-byrne stent graft was used. | It was reported on that the patient experienced access site bleeding with hematoma at an unspecified location. additionally, reportedly the patient also experienced sepsis from an access site infection. the patient remained on impella support and was successfully weaned. |  |  |  |  | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. the product was released with a completed and reviewed device history record under the abiomed quality system. damage to the iliac artery occurred upon removal of the impella. the md noted the relationship to the impella to be unknown, and there is no evidence or imaging provided to suggest impella damaged the iliac artery. thus, this event is unrelated. | The impella device was not received from the customer, therefore, evaluation of the device was not possible. the investigation into the infection/sepsis has been completed. the infection/sepsis was determined to be unknown/unrelated to the impella. the impella device was confirmed to have been sterilized under validated conditions. the investigation into the bleeding/hematoma determined the cause of the patient injury was patient condition and the relationship to the impella was determined to be unrelated. the cause of the sepsis and access site bleeding were determined to be related to the patient¿s condition and unrelated to the device, which did not fail to meet specifications. the investigation into the sepsis and access site bleeding with hematoma determined the product did not fail to meet specifications and the issues were unrelated to the device. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
36813 | 1220648-2021-01001 | 12356552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12356552&pc=OZD | 08/02/2021 | 09/17/2024 |  | 08/23/2021 |  |  | 08/02/2021 | 08/23/2021 |  |  | 04/30/2023 |  | 08/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022076907 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/02/2021 |  |  |  |  | The complainant reported a (b)(6) female patient had impella 5.5 pump inserted in the right axillary artery for cardiomyopathy and bridge to left ventricular assist device (lvad). the pump stopped during use. | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The clinical states that the event occurred during patient movement. the data logs confirmed that the placement signal and the motor current suddenly spiked at the same time indicating a catastrophic failure of the system. the device was returned for review and confirmed to have a severed optical signal and the motor wires were open. upon opening the pump plug there was evidence that showed the wires had been stretched to failure and that the kink protector was also stretched. all glue bonds held together and there was no evidence of a manufacturing defect of this device. all evidence indicates that a sudden, excessive force was applied to the catheter during patient movement that damaged the motor lines and optical fiber. the root cause of the pump stop was excessive pulling force applied to the catheter during patient movement. | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
36814 | 1220648-2021-01002 | 12357167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12357167&pc=OZD | 06/18/2021 | 08/23/2021 |  | 08/23/2021 |  |  | 08/03/2021 | 08/23/2021 |  |  | 01/31/2023 |  | 08/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022057089 |  | 0048-0024-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/03/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) asian male patient had impella cp pump inserted for acute myocardial infarction (ami)/ cardiogenic shock (cgs). the patient had cardiac tamponade while on support. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36815 | 1220648-2021-01004 | 12357170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12357170&pc=OZD | 06/13/2021 | 10/09/2024 |  | 08/23/2021 |  |  | 08/02/2021 | 08/23/2021 |  |  | 02/28/2023 |  | 08/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022052451 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 50 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Sepsis | [9 empty] | Required Intervention | [24 empty] | A |  |  | No | Initial submission | Followup |  |  |  | 08/02/2021 |  |  |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for cardiomyopathy. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient developed limb ischemia while on impella support. as a result, a reverse stab from ecmo was performed. subsequently, the patient developed sepsis and as a result bactericide was administered. access site bleed was also noted. pressure hemostasis was performed, and blood products were administered, amount unknown. no further information was provided. | On (b)(6) 2021 it was reported the patient experienced hemolysis which was treated by adjusting the axillary level. the patient remained on impella support and was successfully weaned. |  |  |  |  | The impella cp ha snot been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the limb ischemia, hemolysis, infection and access stie bleeding have been completed. no data logs were provided for review and no product was returned for evaluation. the root cause of the hemolysis was unable to be determined as no product was returned for evaluation and no further information was provided. the cause of the limb ischemia was most likely the condition/size of the patient vessel as it was reported that the patient was on high doses of pressors, and no other devices were inserted on the patient's left leg (impella insertion site). the infection/sepsis was determined to be unknown/unrelated to impella. the impella device was confirmed to have been sterilized under validated conditions. the cause of the access site bleeding was not determined and the relation to impella was determined to be unknown as it was reported that the reason for the blood transfusion could not be attributed to a specific device. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: hemolysis. impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ infection/sepsis. impella 5.5 with smart assist for use during cardiogenic shock & cp with smart assist section: table 3.2 impella catheter components as noted in the impella ifu ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. as noted in the impella ifu ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ limb ischemia. impella cp with smart assist system. section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ access site bleeding. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
36816 | 1220648-2021-01003 | 12357172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12357172&pc=OZD | 08/03/2021 | 10/09/2025 |  | 08/23/2021 |  |  | 08/03/2021 | 08/23/2021 |  |  | 06/30/2023 |  | 08/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022097055 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2021 |  |  |  |  | The complainant reported a (b)(6) white male patient had impella cp pump inserted in the left femoral artery (lfa) for cardiogenic shock (cgs). the pump was unable to recross the aortic valve (av), the patient had hemolysis, pump was removed, and another pump was used. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36817 | MW5103455 | 12359685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12359685&pc=OJE | 08/18/2021 | 08/19/2021 |  | 08/23/2021 |  |  |  | 08/23/2021 |  |  |  |  | 08/23/2021 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA |  |  | OJE |  |  | 004334 |  |  | ABIOMED INC. |  |  |  |  | 307956 |  |  |  | Yes | [5 empty] | Invalid/Incomplete | No Information |  |  | Voluntary report | [8 empty] | 80 YR | [5 empty] | Malfunction | No | Yes | Inappropriate Waveform | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Impella device placed on (b)(6) 2021. on (b)(6), impella signal having negative spikes. on (b)(6)impella not functioning no waveform on r sided impella, provider spoke with manufacturer. possible thrombus/clots. device removed. fda safety report id# (b)(4).
36818 | 1220648-2021-00998 | 12349498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12349498&pc=OZD | 07/27/2021 | 08/20/2021 |  | 08/20/2021 |  |  | 07/27/2021 | 08/20/2021 |  |  | 06/30/2023 |  | 08/20/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022094896 |  | 0048-0045 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/27/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) patient had impella cp pump inserted for high risk percutaneous coronary intervention (hrpci) with intra-aortic balloon pump (iabp). the patient had limb ischemia and as a result the pump was removed. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36819 | 1220648-2021-00997 | 12349516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12349516&pc=OZD | 07/14/2021 | 10/19/2024 |  | 08/20/2021 |  |  | 07/28/2021 | 08/20/2021 |  |  | 02/28/2023 |  | 08/20/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022083563 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 79 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 07/28/2021 |  |  |  |  | The complainant reported a (b)(6) year-old female patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that during impella insertion, there a rupture and oozing, cardiac tamponade was suspected. as a result, pericardial drainage was performed. no further information was provided. | Additional information provided from the complainant reported that the patient had thrombocytopenia and indicated the patient expired due to hypo-cardiac function and multiple organ failure. per the physician, there is no causal relationship with the impella and death. medical safety officer review of the event noted the use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella cp has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information for the initial mfr 1220648-2021-00997 was provided. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2021-00997 and should not have been. e.4 should have been left blank on the initial manufacturer device report number 1220648-2021-00997 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.4 revised as pma/510(k) was inadvertently left blank and was inserted incorrectly in the ind number field on the initial manufacturer device report number 1220648-2021-00997 .
36820 | 12335757 | 12335757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12335757&pc=OZD | 07/31/2021 | 08/17/2021 | 08/17/2021 | 08/18/2021 | 08/17/2021 | 08/18/2021 |  | 08/18/2021 |  |  |  |  | 08/18/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | No answer provided | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 17885 DA | [5 empty] | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | It was noted that purge flow rates were twice as high upon returning from or. impella purge reservoir/filter found to be cracked and leaking.
36821 | 12335788 | 12335788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12335788&pc=PBL | 06/28/2021 | 08/17/2021 | 08/17/2021 | 08/18/2021 | 08/17/2021 | 08/18/2021 |  | 08/18/2021 |  |  |  |  | 08/18/2021 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PBL |  |  | IMPELLA CP PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2022079180 |  | 0048-0003.0814 |  | No answer provided | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 23360 DA | [5 empty] | Malfunction | No | Yes | Unexpected Shutdown | [6 empty] | Insufficient Information | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella pump failure. at 2010 pump shut off. product rep called at this time, informed that there was a device failure and that the pump would be unable to be restarted.
36822 | 1220648-2021-00994 | 12336653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12336653&pc=OZD | 03/31/2021 | 08/18/2021 |  | 08/18/2021 |  |  | 08/03/2021 | 08/18/2021 |  |  | 11/30/2022 |  | 08/18/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022032127 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/03/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old male patient had impella cp pump inserted in an unknown location for high-risk percutaneous coronary intervention (hrpci). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient developed bleeding from impella and ecmo access sites. as a result, blood transfusion was required, and pressure hemostasis was performed. amount of blood products was not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36823 | 1220648-2021-00995 | 12337870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12337870&pc=OZD | 07/29/2021 | 10/12/2024 |  | 08/18/2021 |  |  | 07/29/2021 | 08/18/2021 |  |  | 05/31/2023 |  | 08/18/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022083816 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Perforation | Vascular Dissection | [11 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/29/2021 |  |  |  |  | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, lows were not higher than 1l/minute. the impella cp device was explanted. the patient survived the procedure. | An impella cp was placed to support a patient undergoing a high risk coronary angioplasty. the patient had multivessel coronary artery disease and aortic valve stenosis. the patient was sent for a ct scan to evaluate her anatomy and was found to have a porcelain aorta and peripheral vascular disease. the patient's access at the femoral artery for the impella placement was complicated due to the peripheral disease. when access was made there was great difficulty advancing through the iliac artery. the vessel was imaged and the team questioned if there was an arterial dissection of the iliac. when the team successfully placed the introducer sheath for the impella, the team continue to have difficulty advancing up the native aorta and to the left ventricle (lv). when the team did advance the guidewire and impella pump to the lv the team observed the pump to have low flows and attributed this to a possible guidewire caused perforation of the lv. a pericardiocentesis was done to remove the fluid. later the team observed clot burden to be present in the pericardium, and so surgical window was performed. while in the operating room the team also repaired the iliac artery dissection and limb complication. |  |  |  |  | The medical center did return product for investigation and product analysis. the cause of the left ventricle perforation was due to the guidewire perforating, but without the analysis of fluoroscopic imaging, the root cause of the guidewire perforation was not able to be determined. the vascular damage to the iliofemoral anatomy was most likely due to patient condition, as it was noted the physician had difficulty advancing the guidewire through the native vasculature. the failure modes will be monitored and trended. | The investigation has been completed for the reported issue of low pump flow. the device was received for evaluation. the root cause of the low pump flow was patient condition. additional information for the initial mfr 1220648-2021-00995 was provided in sections b1, b5, e1, g1, h1, and h6. this report is being filed as part of a retrospective review of historical records. | The investigation into the left ventricular perforation and iliac artery dissection is on-going at this time. upon completion of the investigation, a final report will be filed.
36824 | 1220648-2021-00996 | 12337873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12337873&pc=OZD | 08/04/2021 | 09/21/2021 |  | 08/18/2021 |  |  | 08/04/2021 | 08/18/2021 |  |  | 05/31/2023 |  | 08/18/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022086806 |  | 0550-0008 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | Vascular Dissection | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 08/04/2021 | * | Invalid/Incomplete |  |  | The medical center placed an impella 5.5 for hemodynamic support for a bypass (cabg) surgery patient. the plan was to perform the cabg with impella cross-clamped. while the 5.5 was clamped and the physician was performing the cabg procedure, the physician observed the 5.5 to have perforated the left ventricle. the observation was made when the heart was lifted to inspect for surgical/graft sites. the team completed the cabg, repaired the ventricle and explanted the impella 5.5. post operatively the patient was instead supported by an impella cp and has survived the adverse event. | [5 empty] | Since the original report was filed, the investigation into the left ventricle perforation has concluded. the pump was not returned for analysis, though the data logs from the console were. the logs and clinical case review lead to the conclusion that the root cause of the perforation was patient condition. there was a free wall rupture due to the underlying myocardial infarction. the failure mode will be monitored and trended. | The investigation into the left ventricle perforation is on-going at this time. upon completion of the investigation, a final report will be filed.
36825 | 1220648-2021-00990 | 12328306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12328306&pc=OZD | 07/23/2021 | 08/16/2021 |  | 08/16/2021 |  |  | 07/23/2021 | 08/16/2021 |  |  | 06/30/2023 |  | 08/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022094957 |  | 0048-0045 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/23/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient had impella cp pump inserted for acute myocardial infarction (ami)/ cardiogenic shock (cgs). impella perforated the left ventricle (lv) and as a result 10 packs of blood products was given and fixing the perforation was required. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36826 | 1220648-2021-00991 | 12328323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12328323&pc=OZD | 07/23/2021 | 05/20/2025 |  | 08/16/2021 |  |  | 07/23/2021 | 08/16/2021 |  |  | 06/30/2023 |  | 08/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022093983 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/23/2021 |  |  |  |  | The complainant reported a (b)(6) white male patient had impella cp pump inserted from groin for acute myocardial infarction (ami)/ cardiogenic shock (cgs). impella perforated the aorta at the arch twice and the surgeon immediately repaired both times. | [5 empty] | No product was returned for evaluation. it was reported that the aorta was perforated twice, once during each insertion attempted. an ld was then placed across the new valve without issue. the root cause of the perforation was most likely related to the patient's condition. | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36827 | 1220648-2021-00993 | 12328337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12328337&pc=OZD | 07/24/2021 | 10/21/2021 |  | 08/16/2021 |  |  | 07/24/2021 | 08/16/2021 |  |  | 05/31/2023 |  | 08/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET | [7 empty] | 2022085950 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 51 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/24/2021 | * | Invalid/Incomplete |  |  | The complainant reported a 51 year old male patient enduring suspected hemolysis during impella support. the pump was removed and replaced as a result. | [5 empty] | The device was returned and an evaluation was performed. visual inspection found biomaterial wrapped around the impeller. no other abnormalities or issues were observed. simulated use testing for hemolysis passed using the returned pump and purge cassette. evidence suggests the reported hemolysis was caused by ingested biomaterial. | The impella cp optical was received and an investigation is underway. a supplemental report will be file upon completion of the investigation.
36828 | 1220648-2021-00992 | 12328348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12328348&pc=OZD | 07/23/2021 | 08/16/2021 |  | 08/16/2021 |  |  | 07/23/2021 | 08/16/2021 |  |  | 05/31/2023 |  | 08/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022089818 |  | 0048-0045 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 52 YR | [5 empty] | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [41 empty] | Initial submission |  |  |  |  | 07/23/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient developed an access site bleed. as treatment, surgical intervention was performed to remove the impella and multiple blood products were delivered. | [5 empty] | The impella cp optical pump was discarded by the customer. upon any new information received, a supplemental report will be filed.
36829 | 1220648-2021-00988 | 12313009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12313009&pc=OZD | 07/18/2021 | 08/12/2021 |  | 08/12/2021 |  |  | 07/19/2021 | 08/12/2021 |  |  | 04/30/2023 |  | 08/12/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022081043 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/19/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for cardiomyopathy. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that hemolysis was suspected due to blood sampling data (ldh increased) and urine color being reddish brown. hemolysis was improved by upgrading to 5.0. there was no problem with the position of impella and there is no intracardiac volume or right heart failure, it is presumed that hemolysis was caused by the rotation of the impella. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36830 | 1220648-2021-00989 | 12315612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12315612&pc=OZD | 11/30/2020 | 10/15/2024 |  | 08/12/2021 |  |  | 07/21/2021 | 08/12/2021 |  |  | 06/30/2022 |  | 08/12/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1494424 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 0 DA |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 07/21/2021 |  |  |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for cardiomyopathy. the patient was implanted on (b)(6) 2020 and explanted on (b)(6) 2020. it was reported that the patient developed hemolysis during impella support. as a result, the patient was administered haptoglobin, the dose is unknown. no further information was provided. | It was also reported the patient experience renal failure; however, no additional information was provided. |  |  |  |  | The impella cp has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Additional information for the initial mfr 1220648-2021-00989 was provided in sections b5, d6a, d6b, g1, g2, h1, and h6. the investigation for the reported renal failure has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of renal failure could not be determined. the failure modes will be monitored and trended. instructions for use for the related event are as follows: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
36831 | 1220648-2021-00985 | 12304985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12304985&pc=OZD | 07/13/2021 | 11/29/2024 |  | 08/10/2021 |  |  | 07/13/2021 | 08/10/2021 |  |  | 01/31/2022 |  | 08/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1483647 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 47 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | Mitral Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [24 empty] | A |  |  | No | Initial submission | Followup | Followup |  |  | 07/13/2021 |  |  |  |  | The complainant reported a (b)(6)-year-old male patient had impella 5.0 pump inserted in the left axillary for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that due to the patients pre-existing medical history it was difficult to move to vad. physician was aiming to weaned impella; however, cardiac function was quite poor, and weaning did not progress easily. impella was in for long-term support. on the 48th day after insertion, the pump suddenly stopped. troubleshooting was attempted; however, not successful. the device was explanted. subsequently, the patient developed mitral regurgitation (mr) due to the impella getting caught in the mitral valve. since the mr did not become worse, no treatment was rendered. due to the patient's medical history, no active treatment will be given. the patient's condition is stable with low hemodynamics. no further information was provided. | [5 empty] | The device was returned for analysis. the root cause of the pump stop was likely due to bearing wear at 48 days which is beyond the 28 day design life for the impella 5.0 per design. the cause of the heart valve damage was not determined. the product was released with a completed and reviewed device history record under the abiomed quality system. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event. | The impella 5.0 is expected to be returning and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion. | This report is being filed as part of a retrospective review of historical records. additional information added to d6a/d6b for implant/explant dates. corrected h6 codes from original mfr 1220648-2021-00985. health effect - clinical code added for mitral regurgitation and cardiovascular insufficiency reported. health effect - impact code correct to device explanted.
36832 | 1220648-2021-00984 | 12305026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12305026&pc=OZD | 02/15/2021 | 08/10/2021 |  | 08/10/2021 |  |  | 07/13/2021 | 08/10/2021 |  |  | 12/31/2022 |  | 08/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022028933 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/13/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient experienced a vascular injury which required treatment. treatment details were requested but not provided. no further information was provided. | [5 empty] | The impella cp has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36833 | 1220648-2021-00987 | 12305049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12305049&pc=OZD | 07/04/2021 | 08/10/2021 |  | 08/10/2021 |  |  | 07/16/2021 | 08/10/2021 |  |  | 09/30/2021 |  | 08/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1503580 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/16/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2020 and explanted on (b)(6) 2020. it was reported that the patient developed an infection at the impella access site. as a result, antibacterial agent were administered. no further information was provided. | [5 empty] | The impella cp has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36834 | 1220648-2021-00986 | 12305101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12305101&pc=OZD | 12/13/2020 | 08/10/2021 |  | 08/10/2021 |  |  | 07/15/2021 | 08/10/2021 |  |  | 09/30/2022 |  | 08/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2021006047 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | Sepsis | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/15/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2020 and explanted on (b)(6) 2020. it was reported the patient developed an access site infection; blood cultures tested positive for infection. as a result, the patient was administered antibiotics. subsequently, the impella device was replaced. blood cultures tested positive for infection. subsequently, on the day of impella removal, the patient was diagnosed with sepsis which required administration of antibiotics and pressor agents. no further information was provided. | [5 empty] | The impella cp has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36835 | 12287022 | 12287022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12287022&pc=OZD | 08/01/2021 | 08/05/2021 | 08/05/2021 | 08/06/2021 | 08/05/2021 | 08/06/2021 |  | 08/06/2021 |  |  |  |  | 08/06/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | Yes | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 18980 DA | [5 empty] | Death | Yes | Yes | Protective Measures Problem | Unexpected Shutdown | [5 empty] | Cardiac Arrest | Low Blood Pressure/ Hypotension | Ventricular Fibrillation | Diminished Pulse Pressure | [9 empty] | Death | [28 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | In the evening, the patient finished a bowel prep and was scheduled for a computed tomography (ct) colonoscopy. upon transfer from the patient's bed to the ct scan table the patient¿s impella device stopped working. alarm messages of placement signal not reliable, impella stopped, retrograde flow, motor current high displayed on console. a backup impella console was immediately retrieved and changed however there was no motor function on the backup. the abiomed representative was contacted to monitor the impella remotely. the impella representative recommended that urgent replacement of the impella motor be initiated. the patient was noted to still be conscious, but her blood pressure was compromised. the teams were at the bedside immediately. the impella insertion site was subclavian via graft. the cardio thoracic surgery team was consulted and bedside extracorporeal membrane oxygenation (ecmo) cannulation was arranged. epinephrine and levophed drips were started immediately however the patient's hemodynamics deteriorated rapidly without impella support. a code blue was called due to the loss of a pulse, vf/vt arrest with subsequent pulseless electrical activity (pea) arrest. the patient was cannulated onto ecmo however the ecmo circuit was unable to flow above 2lpm and efforts were unsuccessful. a transesophageal echocardiogram (tee) conducted demonstrated an akinetic heart. the decision was made to stop resuscitation after poor perfusion with no electrical or mechanical heart activity. in examining the pump catheter, there was notable thrombic material around the discharge ports of the pump, but it is unknown if this was present at the time of explant or formed later while sitting in a bag for over a day. the catheter will be given to the manufacturer for their evaluation. all pieces were removed, and an x-ray was taken. results were negative.
36836 | 1220648-2021-00983 | 12291168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12291168&pc=OZD | 04/16/2021 | 08/06/2021 |  | 08/06/2021 |  |  | 07/14/2021 | 08/06/2021 |  |  | 08/31/2022 |  | 08/06/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1502456 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/14/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian female patient had impella cp pump inserted for acute myocardial infarction (ami) with shock. the patient had bleeding at impella insertion site and as a result blood transfusion was given. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36837 | 1220648-2021-00982 | 12275297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12275297&pc=OZD | 07/07/2021 | 08/04/2021 |  | 08/04/2021 |  |  | 07/07/2021 | 08/04/2021 |  |  | 12/31/2022 |  | 08/04/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 2022027382 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/07/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old patient presenting with acute myocardial infarction and cardiogenic shock. the impella cp was selected for hemodynamic support during cardiac surgery. after placement of the impella, a hematoma had developed and was treated with a pressure bandage. after the cardiac procedure, the impella required removal due to bleeding around the impella's repositioning sheath. surgical intervention was required to remove the hematoma and close the site. an unspecified amount of blood products were then delivered. | [5 empty] | The device was discarded by the customer, therefore, it is unable to be physically analyzed. upon any new information received, a supplemental report will be filed.
36838 | 1220648-2021-00981 | 12259548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12259548&pc=OZD | 07/12/2021 | 09/10/2021 |  | 08/02/2021 |  |  | 07/12/2021 | 08/02/2021 |  |  | 01/31/2023 |  | 08/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022036539 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 25 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 07/12/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella 5.5 pump inserted. the pump stopped and the patient was rapidly brought to the operating room (or) for a 5.5 pump change out. | [5 empty] | The returned data logs indicated the purge pressure showed normal fluctuations until it exhibited a sudden, sharp increase on (b)(6) 2021. subsequently, the motor current began to rise and a pump stop occurred. the returned product showed blood in the purge filter, indicative of blood ingress. thus, the root cause of the pump stop is due to blood ingress as a result of a purge line kink which was not addressed. d device returned to manufacturer is actually 30-jul-2021 not 30-jun-2021 which was initially reported. | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36839 | 1220648-2021-00973 | 12252380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12252380&pc=OZD | 07/06/2021 | 07/30/2021 |  | 07/30/2021 |  |  | 07/06/2021 | 07/30/2021 |  |  | 09/30/2022 |  | 07/30/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2021007738 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Use of Device Problem | Patient Device Interaction Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 07/06/2021 | * | Invalid/Incomplete |  |  | A patient with cardiogenic shock and an acute myocardial infarction was admitted and had a bowel resection for a bowel perforation. the team placed an impella cp for hemodynamic support due to the shock. the team did not follow ifu recommendations to add heparin to the purge nor remove the 14fr peel away sheath. the patient was supported for approximately 4 days and was transferred to a tertiary hospital for continued care. when transferred there was a discovery of limb thrombosis in the impella accessed limb. the team explanted the impella pump, performed thrombectomy, and finally closed the access via a perclose. | [5 empty] | The medical center did return the impella product for analysis. the cause of the limb thrombosis and following thrombectomy, was the physician's choice to not remove the 14fr introducer and advance the impella cp. the ifu states the following for patient's like this one undergoing transport: "to prevent failure of the peel-away introducer, remove the peel-away introducer prior to transport when activated clotting time (act) is less than 150 seconds". in addition the patient was not properly anticoagulated per abiomed recommendations. the failure mode will be monitored and trended.
36840 | 1220648-2021-00975 | 12252821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12252821&pc=OZD | 06/16/2020 | 07/30/2021 |  | 07/30/2021 |  |  | 07/21/2021 | 07/30/2021 |  |  | 01/31/2022 |  | 07/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1465559 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/21/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2020 and explanted on (b)(6) 2020. it was reported that the patient developed hemolysis during impella support. as a result, haptoglobin was administered. amount was not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36841 | 1220648-2021-00976 | 12253272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12253272&pc=OZD | 06/11/2021 | 07/30/2021 |  | 07/30/2021 |  |  | 07/01/2021 | 07/30/2021 |  |  | 02/28/2023 |  | 07/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022050185 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/01/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian female patient had impella cp pump inserted in the right femoral. the patient had hemolysis and as a result some units of blood transfusion was given. | [5 empty] | Device not returned by customer. the impella cp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
36842 | 1035166-2021-00085 | 12256178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12256178&pc=DYB | 07/01/2021 | 12/28/2021 |  | 07/30/2021 |  |  | 07/06/2021 | 07/30/2021 |  |  | 10/01/2023 |  | 07/30/2021 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-19434 |  | 0052-3021 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 82 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | No Apparent Adverse Event | [5 empty] | Heart Failure/Congestive Heart Failure | Mitral Valve Insufficiency/ Regurgitation | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported on (b)(6) 2021 during an acute myocardial infarction (ami)/ cardiogenic shock (cgs) procedure, a 23 fr introducer sheath kinked when the physician was trying to push the sheath into the body under fluoro. a second 23 fr sheath was used and it kinked at the same spot. it was reported the spot in the iliac looked calcified. the 23 fr introducers were exchanged for a gore sheath. the patient passed away in the intensive care unit due to overall heart and organ failure, death not related to device. the patient had decompensated heart failure due to severe mitral regurgitation. | [5 empty] | Corrected brand name - abiomed introducer kit for impella® the device was used in treatment. the device was not returned for analysis, therefore, the exact cause of the product issue cannot be determined and the clinical observation could not be confirmed. however, the following controls are in place to mitigate the reported product issue. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure (abiomed introducer sheath in-process and final inspections): measure dimensions: measure tip i.d. of sheath using appropriate pin gauges to be in range of 0.305" - 0.307" for 23f. visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash. with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. visual inspection: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. per ifu never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | H11 corrected data - initial report passed, however, section g6 was inadvertently left blank, it should have stated "type of report" - initial. conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete.
36843 | 1220648-2021-00978 | 12256257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12256257&pc=OZD | 10/16/2020 | 07/30/2021 |  | 07/30/2021 |  |  | 07/06/2021 | 07/30/2021 |  |  | 07/31/2022 |  | 07/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1492166 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/06/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old male patient had impella 5.0 pump inserted in the right axillary for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed and access site bleed during impella support. as a result, the patient was administered blood products, amount was not provided. no further information was provided. | [5 empty] | The impella cp has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36844 | 1220648-2021-00979 | 12256259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12256259&pc=OZD | 07/30/2020 | 07/30/2021 |  | 07/30/2021 |  |  | 07/08/2021 | 07/30/2021 |  |  | 12/31/2020 |  | 07/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1400793 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 47 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/08/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old male patient had impella 5.0 pump inserted in the right axillary for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2020 and explanted on (b)(6) 2020. it was reported that the patient developed bleeding at the anastomotic site of the artificial blood vessel where impella was inserted. as result, the patient was administered red blood cells and surgical repair was performed. no further information was provided. | [5 empty] | The impella 5.0 is expected to be returning and an investigation will be underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
36845 | 1220648-2021-00977 | 12256301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12256301&pc=OZD | 06/21/2021 | 07/30/2021 |  | 07/30/2021 |  |  | 07/06/2021 | 07/30/2021 |  |  | 04/30/2023 |  | 07/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022073760 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/06/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that immediately after impella insertion the patient developed hemolysis. as a result, red blood cells were administered. amount was not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36846 | 1220648-2021-00980 | 12256381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12256381&pc=OZD | 07/15/2021 | 07/30/2021 |  | 07/30/2021 |  |  | 07/15/2021 | 07/30/2021 |  |  | 06/30/2022 |  | 07/30/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1486359 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 81 YR | [5 empty] | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/15/2021 | * | Invalid/Incomplete |  |  | A patient in (b)(6) (in (b)(6) 2020) was admitted suffering from an acute myocardial infarction and cardiogenic shock. when taken to the cardiac catheterization lab he was seen to have a totally occluded left main coronary artery. the team placed an impella cp for hemodynamic support. the pump supported the patient for over 3 days when it was weaned and explanted. months later, in (b)(6) 2021, the patient's registry data was documented and reviewed. in the registry hemolysis was noted, and attributed to the use of the impella. the team had treated the hemolysis by volume delivery, pump repositioning, dialysis, and pump explantation. later the patient expired when family choice of dnr was made. | [5 empty] | The medical center in (b)(6) discarded the impella product in 2020. no further clinical details or data logs are being able to be retrieved. the definitive root cause of the impella attributed hemolysis is not known. should any new clinical details be shared, a supplemental report will be filed.
36847 | 1220648-2021-00970 | 12241898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12241898&pc=OZD | 07/01/2021 | 09/02/2021 |  | 07/28/2021 |  |  | 07/01/2021 | 07/28/2021 |  |  | 10/31/2021 |  | 07/28/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 1451378 |  | 005060 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | DISTRI | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Mitral Valve Insufficiency/ Regurgitation | Renal Impairment | [10 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 07/01/2021 | * | Invalid/Incomplete |  |  | The complainant reported a patient of unknown age presenting with acute myocardial infarction and cardiogenic shock. after approximately 9 days of impella support, the patient exhibited signs of hemolysis via increased lactate dehydrogenase levels. the patient's hemoglobin dropped and kidney function was impacted. as treatment, the pump was removed and replaced with an intra-aortic balloon pump. additionally, multiple units of blood product were transfused and the patient was placed on continuous renal replacement therapy. mitral regurgitation from a ruptured chordae was noted and suspected as a result of positioning issues with the impella. as treatment, a mitraclip procedure was completed. the device was removed and the patient received multiple units of blood products as a result. | [5 empty] | The device was returned with the pump's catheter severed, therefore, we could not perform simulated use testing of the device. visual inspection found no abnormalities with the device related to the reported issues. based on our analysis of the clinical account, we feel the root cause of the reported issues is related to improper technique/positioning. | The device was received and an evaluation is underway. upon completion of our analysis, a supplemental mdr will be filed.
36848 | 1035166-2021-00082 | 12242377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12242377&pc=DYB | 06/22/2021 | 12/16/2021 |  | 07/28/2021 |  |  | 06/30/2021 | 07/28/2021 |  |  |  |  | 07/28/2021 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  |  |  | 0052-3006 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 51 YR |  | Female |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that patient was presented to physician for worsening chest pain, exertional angina/dyspnea. during procedure, sheath was leaking after the graft locks were placed. hemostasis could not be achieved because of this issue, therefore, the introducer was removed and a new introducer was placed from an impella 5.0 kit. no blood product delivery and no surgical intervention was required. patient outcome is stable. no additional information is available. | [5 empty] | The device used in treatment. one abiomed 23f sheath was returned for analysis. traces of blood were found inside and outside the sheath. upon returned product evaluation, the sheath was observed using a vertex at higher magnification and no cracks, holes or gaps were found. the sheath score lines were observed to be prominent on the outer surface. the sheath was leak tested per iso (b)(4) standard and procedure and sheath did not show any leakage through the graft or sheath score lines. the sheath was found to be within specification. further investigation revealed that wall thickness of the scoreline has been reduce to improve the peel strength. reduced wall thickness of the scoreline could have made the sheath susceptible to collapsing when used with graft locks, and that could have caused leakage. however, all devices that have been evaluated have met the current dimensional, visual, and functional requirements in the 23f sheath product specification. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per adelante, adelante-s and adelante-s2 sheath extrusion process, appropriate extrusion tool is selected for the part to be extruded and is inspected prior to the extrusion run. peel strength and critical dimensions are verified on every 100th extruded sheath as per procedure. perabiomed introducer sheath in-process and final inspection, (sample size: 100% inspection), the sheath is inspected with a naked eye at a distance of 12" to 18" to ensure the sheath is free of kinks, cracks, splits, sinks and any other damages. leak test is performed as per procedure based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress, follow up report will be submitted if we find any further additional information.
36849 | 1035166-2021-00084 | 12242380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12242380&pc=DYB | 06/21/2021 | 12/30/2021 |  | 07/28/2021 |  |  | 07/06/2021 | 07/28/2021 |  |  | 04/01/2024 |  | 07/28/2021 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | PQ0C11882 |  | 0052-3006 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was presented to the physician post-arrest while undergoing cpr. during procedure, 23fr introducer leaked after placement. after attaching to graft, physician noted that the introducer was cracked out of the package. a second 23 fr axillary introducer was opened and used without issue. there was no harm to patient reported. no adverse event was reported. no additional information is available. | [5 empty] | The device was used in treatment. the device was not returned for analysis. however, evaluation of the device from the different event for the same lot showed reduced wall thickness of the scoreline, which could have made the sheath susceptible to collapsing when used with graft locks, which could have caused leakage. the wall thickness of the scoreline has been reduced to improve the peel strength. however, all devices that have been evaluated have met the current dimensional, visual, and functional requirements in the 23f sheath product specification. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per abiomed introducer sheath in-process and final inspection :(sample size: 100% inspection) the sheath is inspected with a naked eye at a distance of 12" to 18" to ensure the sheath is free of kinks, cracks, splits, sinks and any other damages. leak test is performed as per procedure. ifu - is provided with the product. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing nonconformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress, follow up report will be submitted if we find any further additional information.
36850 | 1220648-2021-00963 | 12231314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12231314&pc=OZD | 07/04/2021 | 07/27/2021 |  | 07/27/2021 |  |  | 07/04/2021 | 07/27/2021 |  |  | 09/30/2021 |  | 07/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA LD PUMP |  |  | OZD |  |  | IMPELLA LD AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1441125 |  | 005082 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 83 YR | [5 empty] | Malfunction | No | Yes | Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/04/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b(6) year old white male patient had impella ld pump inserted. while attempting to remove the ld pump in the operating room (or), the ld came apart at the catheter and motor junction leaving the motor and cannula in the patient¿s descending aorta, the device was snared with assistance from an interventional radiology (ir) doctor. | [5 empty] | Device not returned by customer. the impella ld pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
36851 | 1035166-2021-00080 | 12233426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12233426&pc=DYB | 06/23/2021 | 12/21/2021 |  | 07/27/2021 |  |  | 06/30/2021 | 07/27/2021 |  |  | 04/01/2024 |  | 07/27/2021 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | PQOC11880 |  | 0052-3006 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 74 YR |  | Male |  |  |  | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Appropriate Term / Code Not Available | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported the patient was presented to the physician with multiple comorbidities (presented with stemi and pea arrest) . the physician reported being unable to obtain hemostasis after securing the graft locks. the sheath was noted as "more pliable/collapsible." intervention was taken to suture around the sheath, apply extra clamping, and delivery of 2 units of blood (30 - 45 minute delay). this case was completed via right axillary artery insertion via a 10/30 hemashield gold graft. the patient did have a history of pad of iliac arteries which is one of the reasons why the physician decided to place the pump axillary. | [5 empty] | Oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report isbased upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | The device was used in treatment. one half of 23f abiomed axillary was returned for analysis. traces of blood was found on the sheath. upon returned product evaluation, the sheath was observed to be peeled through the score line. the score lines were observed to be prominent. leak test could not be performed as the sheath was returned peeled. the sheath surface was observed on higher magnification using a vertex and no holes, cracks and gaps were observed. the sheath was observed to be within specifications. further investigation revealed that wall thickness of the scoreline has been reduce to improve the peel strength. reduced wall thickness of the scoreline could have made the sheath susceptible to collapsing when used with graft locks, and that could have caused leakage. however, all devices that have been evaluated have met the current dimensional, visual, and functional requirements in the 23f sheath product specification. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per operations procedure(s) adelante, adelante-s and adelante-s2 sheath extrusion · appropriate extrusion tool is selected for the part to be extruded and is inspected prior to the extrusion run. · peel strength and critical dimensions are verified on every 100th extruded sheath as per procedure. per abiomed introducer sheath in-process and final sample size: 100% inspection · the sheath is inspected with a naked eye at a distance of 12" to 18" to ensure the sheath is free of kinks, cracks, splits, sinks and any other damages. leak test is performed as per procedure. abiomed instructions for use (ifu) is provided with the product. based on the investigation, a capa was not required. there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report isbased upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete.
36852 | 1035166-2021-00081 | 12233565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12233565&pc=DYB | 06/23/2021 | 12/20/2021 |  | 07/27/2021 |  |  | 06/29/2021 | 07/27/2021 |  |  | 04/01/2024 |  | 07/27/2021 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | PQOC11880 |  | 0052-3006 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 45 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Appropriate Term / Code Not Available | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported there was a 23 fr introducer collapsed under graft locks. sheath was noted to have ridges on side (not smooth in appearance) prior to inserting into 10/30 vascutek gelweave. attempted to place sheath inside graft and graft locksx2 attached in standard fashion. hemostasis was not obtained, bleeding oozing around sheath between graft. the 23f sheath was exchanged for another axillary sheath kit was used without issue (approximately 5 minute delay). after impella was successfully inserted, coronary artery bypass (cab) grafting was completed & patient came off bypass to impella 5.5 support; chest closed. | [5 empty] | The device used in treatment. one abiomed 23f sheath was returned for analysis. traces of blood were found inside and outside the sheath. upon returned product evaluation, the sheath was observed using a vertex at higher magnification and no cracks, holes or gaps were found. the sheath score lines were observed to be prominent on the outer surface. the sheath was leak tested per iso (b)(4) standard and procedure, the sheath did not show any leakage through the graft or sheath score lines. the sheath was found to be within specification. further investigation revealed that wall thickness of the scoreline has been reduce to improve the peel strength. reduced wall thickness of the scoreline could have made the sheath susceptible to collapsing when used with graft locks, and that could have caused leakage. however, all devices that have been evaluated have met the current dimensional, visual, and functional requirements in the 23f sheath product specification. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per adelante, adelante-s and adelante-s2 sheath extrusion process, appropriate extrusion tool is selected for the part to be extruded and is inspected prior to the extrusion run. peel strength and critical dimensions are verified on every 100th extruded sheath as per procedure. perabiomed introducer sheath in-process and final inspection, (sample size: 100% inspection), the sheath is inspected with a naked eye at a distance of 12" to 18" to ensure the sheath is free of kinks, cracks, splits, sinks and any other damages. leak test is performed as per procedure. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete.
36853 | 1220648-2021-00965 | 12234300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12234300&pc=OZD | 06/29/2021 | 07/27/2021 |  | 07/27/2021 |  |  | 06/29/2021 | 07/27/2021 |  |  | 03/31/2023 |  | 07/27/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022068403 |  | 0048-0045 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/29/2021 | * | Invalid/Incomplete |  |  | A patient suffering from an acute myocardial infarction was admitted and placed on hemodynamic support with an impella cp via access at the femoral artery. after about 11 hours of support, the pump was explanted due to concerns of bleeding at the access site. the team explanted the pump from the femoral artery and placed a new, second cp, via the axillary artery. vascular surgery later placed a viabahn stent to the femoral artery to achieve successful hemostasis. | [5 empty] | The medical center discarded the impella product. no product analysis and testing is possible. the definitive root cause can not be determined. the failure mode will be monitored and trended.
36854 | 1220648-2021-00964 | 12234377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12234377&pc=OZD | 07/07/2021 | 07/27/2021 |  | 07/27/2021 |  |  | 07/07/2021 | 07/27/2021 |  |  | 04/30/2023 |  | 07/27/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022079567 |  | 0048-0045 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 82 YR | [5 empty] | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | [28 empty] | Initial submission |  |  |  |  | 07/07/2021 | * | Invalid/Incomplete |  |  | An elderly patient was admitted with multivessel coronary artery disease and was turned down by the cardiothoracic team for surgical intervention. as a cabg was not permitted the patient had a planned high risk coronary angioplasty. the team made the choice to place an impella cp for hemodynamic support. due to low pump flows during the support, the pump was explanted and replaced. the angioplasty proceeded, the pump was explanted, and the impella access site closed. when the cp pump was returned for analysis of the low flows, the impeller was observed to be damaged. | [5 empty] | The root cause of the low pump flow was determined to be a fractured impeller blade. however, as fluoroscopic imaging of the pump was not provided, a definitive root cause for the blade fracture could not be determined. the failure mode will be monitored and trended.
36855 | 1035166-2021-00070 | 12225480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12225480&pc=DYB | 06/24/2021 | 12/15/2021 |  | 07/26/2021 |  |  | 06/29/2021 | 07/26/2021 |  |  |  |  | 07/26/2021 |  | INTRODUCER, CATHETER, PRODUCT CODE: DYB |  |  | INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  |  |  | 0052-3021 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 49 YR |  | Female |  |  |  | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that rp dilators were used to dilate up to 23 fr, however unable to fully insert 23 fr catheter. the dilator was taken out of the sheath and successfully inserted to dilate track. removed and reinserted into 23 fr sheath and inserted into left groin of a (b)(6) year old female. once the dilator was removed and wire placed for catheter insertion, a steady flow of blood came out of the sheath around the wire. impella back loaded onto wire and inserted. however, once the impella was in the body, a mass amount of blood poured out of the hemostatic valve of impella sheath. shoved impella forward and into position and removed wire and sheath quickly to control blood volume. unfortunately, patient lost approximately 800-1000 ml. rp ramped up quickly. suction on both rp and cp. 3 l ns given in the cvl with 2 units prbc controlled bleeding at rp site with 5 mattress sutures and figure 8 suture. patient care withdrawn/patient expired 06/24/2021 due to multiple organ failure. | [5 empty] | The following sections were updated in follow-up 1. b4, d9, g3, g6, h2, h3, h6, h10. one 23f abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. blood was found on the sheath and inside the sideport tubing. the sheath was already peeled apart upon receipt of the product. according to the event description summary, once the dilator was removed and the wire placed for catheter insertion, a steady flow of blood came out of sheath around the wire. since the introducer sheath was received back from the field already peeled apart, a leak test could not be performed in order to confirm this. the hemostatic valve was examined under magnification and no anomalies were found. the hub looked normal with no cracks and the hub cap looked normal and was seated in the correct position. the lot number was not provided by the customer, therefore the device history records could not be reviewed, however, inspection procedures require any oscor product pass all in-process and qa final inspection before shipping to the customer. multiple attempts were made to obtain the product. returned device analysis revealed the sheath was within manufacturing specifications. according to the customer, the sheath was leaking. the sheath was returned to oscor already peeled apart, so a leak test could not be performed to confirm this, therefore, the exact cause of the leak could not be determined or confirmed. the lot number was unknown, so it was not possible to check the device history record for any related anomalies. no manufacturing defects were found. per manufacturing procedure (abiomed introducer final assembly) sample size: 100% the sheath assemblies are leak tested per procedures. per qa procedure (adelante s2s introducer sheath in-process and final inspection): sample size: 100% the sheaths are pressure and vacuum leak tested per procedures. per ifu: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. infusion through the sideport can be done only after all air is removed from the unit. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete.
36856 | 1220648-2021-00962 | 12216757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12216757&pc=OZD | 01/14/2021 | 07/23/2021 |  | 07/23/2021 |  |  | 07/13/2021 | 07/23/2021 |  |  | 09/30/2022 |  | 07/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 1502424 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 86 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | [41 empty] | Initial submission |  |  |  |  | 07/13/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old female patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported the patient developed cardiac tamponade during impella support. treatment details were requested but not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36857 | 1220648-2021-00955 | 12212137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12212137&pc=OZD | 07/21/2020 | 07/22/2021 |  | 07/22/2021 |  |  | 06/23/2021 | 07/22/2021 |  |  | 08/31/2020 |  | 07/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1367668 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/23/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old male patient had impella 5.0 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2020 explanted on (b)(6) 2020. it was reported the patient developed hemolysis during impella support. as a result, the patient received blood product, amount was not provided. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36858 | 1220648-2021-00956 | 12212139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12212139&pc=OZD | 06/29/2021 | 08/30/2021 |  | 07/22/2021 |  |  | 06/29/2021 | 07/22/2021 |  |  | 09/30/2022 |  | 07/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 1503999 |  | 0048-0014 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 06/29/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient had impella cp optical pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). during the night due to bleeding on the right groin femostop was installed, in the morning the patient had limb ischemia and the pump was removed. | [5 empty] | The field lt and rt logs were reviewed, and no abnormal alarms or events were noted. the root cause of the reported leg ischemia was likely related to patient condition. no product was returned for evaluation. the returned pump was visually inspected and no abnormalities were noted. | Device evaluated by mfr: device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36859 | 1220648-2021-00957 | 12212149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12212149&pc=OZD | 06/11/2021 | 07/22/2021 |  | 07/22/2021 |  |  | 06/26/2021 | 07/22/2021 |  |  | 12/31/2022 |  | 07/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022023838 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 explanted on (b)(6) 2021. it was reported that after transition to pci a hematoma formed at the impella insertion site. the hematoma was removed using a hemostatic device in the catheter room. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36860 | 1220648-2021-00958 | 12212152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12212152&pc=OZD | 06/28/2021 | 07/22/2021 |  | 07/22/2021 |  |  | 06/28/2021 | 07/22/2021 |  |  | 12/31/2022 |  | 07/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022039808 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/28/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for an unknown indication. the patient was implanted on (b)(6) 2021 explanted on (b)(6) 2021. it was reported that hemolysis was noted during impella support. as a result, haptoglobin was administered. the physician stated that it is unknown whether the event is related to the use of the device. no further information was provided. | [5 empty] | The impella cp has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36861 | 1220648-2021-00959 | 12212993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12212993&pc=OZD | 06/16/2021 | 07/22/2021 |  | 07/22/2021 |  |  | 06/29/2021 | 07/22/2021 |  |  | 02/28/2023 |  | 07/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022052336 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/29/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 explanted on (b)(6) 2021. it was reported the patient developed an ischemic cerebrovascular accident (cva) or stroke during impella support. treatment details were not provided. no further information was provided. | [5 empty] | The impella cp was not return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36862 | 1220648-2021-00961 | 12213059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12213059&pc=OZD | 06/25/2021 | 07/22/2021 |  | 07/22/2021 |  |  | 06/30/2021 | 07/22/2021 |  |  | 04/30/2023 |  | 07/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022075923 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/30/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year-old female patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs) the patient was implanted on (b)(6) 2021 explanted on (b)(6) 2021. it was reported that the patient experienced access site bleed. the issue was improved by applying purse string sutures (x6) to the insertion part and adjusting the angle of the sheath. the patient was also administered 20 units of red blood cells and fresh frozen plasma. subsequently, the patient experienced hemolysis during impella support. as a result, haptoglobin was administered on a daily due to persistent hemolysis. cardiac function recovered and the patient was successfully weaned and explanted. no further information was provided. | [5 empty] | The impella cp has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36863 | 1220648-2021-00960 | 12213069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12213069&pc=OZD | 01/07/2021 | 07/22/2021 |  | 07/22/2021 |  |  | 07/01/2021 | 07/22/2021 |  |  | 09/30/2022 |  | 07/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1503408 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/01/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that when impella was inserted from the left femoral artery, there was a change in color tone in the left lower limb. lower limb ischemia occurred. as a result, a 6" fr sheath was antegrade inserted into the left sfa, and the 4" fr sheath kept in the right femoral artery. ecmo was perfused into the sheath of the left sfa. after that, no progression of lower limb ischemia was observed. no further information was provided. | [5 empty] | The impella cp has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36864 | 1035166-2021-00071 | 12203355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12203355&pc=DYB | 06/23/2021 | 12/10/2021 |  | 07/21/2021 |  |  | 06/24/2021 | 07/21/2021 |  |  | 04/01/2024 |  | 07/21/2021 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | PQ0C11882 |  | 0052-3006 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 67 YR |  | Female |  |  |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was presented to a physician for cardiomyopathy. during the access for impella 5.5, axillary sheath starts bleeding prior to insertion of a catheter. the sheath was cracked on the side, and the patient lost 30cc of blood. bleeding was observed through the graft. there was no medical intervention required to control the blood loss. the sheath was replaced and the procedure was completed successfully. no adverse patient side effects were reported. no additional information is available. | [5 empty] | Our investigation is still in progress, follow up report will be submitted if we find any further additional information. | The device was used in treatment. the device was not returned for analysis. however, evaluation of the device from the different event for the same lot showed reduced wall thickness of the scoreline, which could have made the sheath susceptible to collapsing when used with graft locks, which could have caused leakage. the wall thickness of the scoreline has been reduced to improve the peel strength. however, all devices that have been evaluated have met the current dimensional, visual, and functional requirements in the 23f sheath product specification. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections per abiomed introducer sheath in-process and final inspection :(sample size: 100% inspection) the sheath is inspected with a naked eye at a distance of 12" to 18" to ensure the sheath is free of kinks, cracks, splits, sinks and any other damages. leak test is performed as per procedure ifu - ab-17-001z-x is provided with the product. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report
36865 | 1035166-2021-00073 | 12205049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12205049&pc=DYB | 06/22/2021 | 12/16/2021 |  | 07/21/2021 |  |  | 06/23/2021 | 07/21/2021 |  |  | 04/01/2024 |  | 07/21/2021 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | PQ0C11882 |  | 0052-3006 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The patient's heart was arrested and was on bypass, and they were transitioning off-pump to impella when bleeding was observed between graft and sheath. 5 cc of blood was lost. no blood products or medical intervention is required to treat blood loss. the physician removed the peel-away sheath and trimmed the graft. the patient was kept on support. no additional information is available. | [5 empty] | Our investigation is still in progress, follow up report will be submitted if we find any further additional information. | The following sections were updated in follow-up b4,g3,g6,h2,h3,h6,h10. the device used in treatment. one half of 23f abiomed axillary was returned for analysis. three other accessories were returned for analysis. there was blood found occluded in the side port tube. upon returned product evaluation, the sheath was observed to be peeled through the score line. the score lines were observed to be prominent. leak test could not be performed as the sheath was returned peeled. the sheath surface was observed on higher magnification using a vertex and no holes, cracks and gaps were observed. the sheath was observed to be within specifications. further investigation revealed that wall thickness of the scoreline has been reduce to improve the peel strength. reduced wall thickness of the scoreline could have made the sheath susceptible to collapsing when used with graft locks, and that could have caused leakage. however, all devices that have been evaluated have met the current dimensional, visual, and functional requirements in the 23f sheath product specification. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per adelante, adelante-s and adelante-s2 sheath extrusion process, appropriate extrusion tool is selected for the part to be extruded and is inspected prior to the extrusion run. peel strength and critical dimensions are verified on every 100th extruded sheath as per procedure. perabiomed introducer sheath in-process and final inspection, (sample size: 100% inspection), the sheath is inspected with a naked eye at a distance of 12" to 18" to ensure the sheath is free of kinks, cracks, splits, sinks and any other damages. leak test is performed as per procedure. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report | Correction : h1 : initial mdr should read as malfunction not serious injury. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36866 | 1035166-2021-00072 | 12205065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12205065&pc=DYB | 06/22/2021 | 12/16/2021 |  | 07/21/2021 |  |  | 06/23/2021 | 07/21/2021 |  |  | 04/01/2024 |  | 07/21/2021 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | PQOC11880 |  | 0052-3006 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The patient's heart was arrested and was on bypass, and they were transitioning off-pump to impella when bleeding was observed between graft and sheath. 5 cc of blood was lost. no blood products or medical intervention is required to treat blood loss. the physician removed the peel-away sheath and trimmed the graft. the patient was kept on support. no additional information is available. | [5 empty] | Correction: h1: initial mdr should read as malfunction not serious injury. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress, follow up report will be submitted if we find any further additional information. | The device used in treatment. one half of 23f abiomed axillary was returned for analysis. three other accessories were returned for analysis. there was blood found occluded in the side port tube. upon returned product evaluation, the sheath was observed to be peeled through the score line. the score lines were observed to be prominent. leak test could not be performed as the sheath was returned peeled. the sheath surface was observed on higher magnification using a vertex and no holes, cracks and gaps were observed. the sheath was observed to be within specifications. further investigation revealed that wall thickness of the scoreline has been reduce to improve the peel strength. reduced wall thickness of the scoreline could have made the sheath susceptible to collapsing when used with graft locks, and that could have caused leakage. however, all devices that have been evaluated have met the current dimensional, visual, and functional requirements in the 23f sheath product specification. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per adelante, adelante-s and adelante-s2 sheath extrusion process, appropriate extrusion tool is selected for the part to be extruded and is inspected prior to the extrusion run. peel strength and critical dimensions are verified on every 100th extruded sheath as per procedure. perabiomed introducer sheath in-process and final inspection, (sample size: 100% inspection), the sheath is inspected with a naked eye at a distance of 12" to 18" to ensure the sheath is free of kinks, cracks, splits, sinks and any other damages. leak test is performed as per procedure. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
36867 | 1220648-2021-00953 | 12200072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12200072&pc=OZD | 01/28/2020 | 07/20/2021 |  | 07/20/2021 |  |  | 07/08/2021 | 07/20/2021 |  |  | 07/31/2022 |  | 07/20/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1492166 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/08/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old male patient had impella cp pump inserted in the left femoral for high-risk percutaneous coronary intervention (hrpci). the patient was implanted on (b)(6) 2020 and explanted on (b)(6) 2020. it was reported that the patient had hemodynamic disruption during hrpci. the patient was also on ecmo with act's at 200per second. the physician stated that the pressure on the other side was not so good when the peel away was removed. the indwelling sheath was advanced as much as possible and a hematoma appeared, it later disappeared. the patient was administered blood products as a result, amount was not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36868 | 1220648-2021-00954 | 12200073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12200073&pc=OZD | 09/06/2020 | 07/20/2021 |  | 07/20/2021 |  |  | 03/09/2021 | 07/20/2021 |  |  | 06/30/2022 |  | 07/20/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1484369 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old female patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2020 and explanted on (b)(6) 2020. it was reported that the patient developed and access site bleed during impella support. as a result, the patient was administered red blood count and fresh frozen plasma. per the reported information the event was caused by impella. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36869 | 1220648-2021-00945 | 12189132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12189132&pc=OZD | 06/30/2021 | 07/19/2021 |  | 07/19/2021 |  |  | 06/30/2021 | 07/19/2021 |  |  | 04/30/2023 |  | 07/19/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022071871 |  | 0048-0045 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | [28 empty] | Initial submission |  |  |  |  | 06/30/2021 | * | Invalid/Incomplete |  |  | A patient was admitted with known coronary artery disease and an ef of just 10%. a plan was made for a high risk coronary angioplasty with the hemodynamic support of the impella cp. the pump was explanted after an hour of use for low flows. the team had attempted to troubleshoot by removal of the pump, flushing of it, and then re-inserting. when these measures had failed, the pump was removed and replaced. the second impella cp was placed and successfully supported the patient. | [5 empty] | The medical center returned the impella cp for analysis of the low flows. in the investigation the impeller was seen to be damaged. the cause of the low pump flow was fractured impeller blades. because fluoroscopic imaging of the pump/additional evidence was not returned, a definitive cause of the fracture could not be determined. based on the characteristic of the fracture, it was most likely caused by an interaction with another device.
36870 | 1220648-2021-00946 | 12189859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12189859&pc=OZD | 06/25/2021 | 07/19/2021 |  | 07/19/2021 |  |  | 06/25/2021 | 07/19/2021 |  |  | 04/30/2023 |  | 07/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022074949 |  | 0048-0045 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/25/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female patient presenting with acute myocardial infarction and cardiogenic shock for support with the impella cp optical device. during support, the patient developed suspected hemolysis as evident by urinalysis and hemolized labs. there was no indication plasma-free hemoglobin testing was performed to confirm suspected hemolysis. the family made the decision to withdraw care after hemolysis did not improve. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36871 | 1220648-2021-00947 | 12189947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12189947&pc=OZD | 07/06/2021 | 07/19/2021 |  | 07/19/2021 |  |  | 07/06/2021 | 07/19/2021 |  |  | 01/31/2023 |  | 07/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022038577 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 83 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/06/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock for support using the impella cp device. during support, a hematoma formed at the impella access site. as treatment, a blood transfusion was required. | [5 empty] | Neither the pump nor introducer was received for evaluation. we believe the access site bleed was related to high patient act values.
36872 | 1220648-2021-00949 | 12194163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12194163&pc=OZD | 10/19/2020 | 07/19/2021 |  | 07/19/2021 |  |  | 03/09/2021 | 07/19/2021 |  |  | 12/31/2021 |  | 07/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1461512 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6)-year-old female patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2020 and explanted on (b)(6) 2020. it was reported the patient experienced an access site bleed during impella support. as a result, blood products were administered, amount is unknown. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36873 | 1220648-2021-00948 | 12194183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12194183&pc=OZD | 06/23/2021 | 09/10/2021 |  | 07/19/2021 |  |  | 06/23/2021 | 07/19/2021 |  |  | 04/30/2023 |  | 07/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022081795 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Material Integrity Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 06/23/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting for impella support pre-operative percutaneous coronary intervention. the physician had difficulty obtaining hemostasis after inserting the 23 fr introducer as the introducer was noted as "more collapsible/pliable" than normal. immediately following the procedure, the patient required 2 units of blood products due to insertion site bleeding. | [5 empty] | The 23 fr introducer was returned and confirmed to have reduced the radial wall strength (hoop strength). the reduced radial strength resulted in a bleeding path between the graft and sheath. the reduced wall strength was attributed to supplier manufacturing lot. | The 23 fr introducer was returned and an investigation is underway. a supplemental mdr will be filed with the results of our analysis.
36874 | 1220648-2021-00951 | 12194185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12194185&pc=OZD | 02/10/2021 | 07/19/2021 |  | 07/19/2021 |  |  | 06/23/2021 | 07/19/2021 |  |  | 07/31/2022 |  | 07/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1489399 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/23/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 explanted on (b)(6) 2021. it was reported the patient experienced an ischemic cva two days after the device was explanted. furthermore, the patient developed hemolysis for which blood products were administered. amount of blood products were not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36875 | 1220648-2021-00950 | 12194189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12194189&pc=PYX | 06/23/2021 | 09/19/2024 |  | 07/19/2021 |  |  | 06/23/2021 | 07/19/2021 |  |  | 07/31/2021 |  | 07/19/2021 |  | TEMPORARY NONROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED INC. |  |  | US |  | 1449198 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hypervolemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/23/2021 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | A patient was admitted after arresting with ventricular tachycardia. after the patient was stabilized she was taken to the cardiac catheterization lab. the coronary arteries were stable, but her right heart was not. she was suffering from shock due to right ventricular failure, low potassium level, and supraventricular tachycardia. an impella rp and impella cp were placed, and not ecmo, due to her underlying neurological cancer and ongoing chemotherapy. when the rp was placed and the repositioning sheath advanced she had blood loss at the impella access site. the team treated the bleed of approximately 800-1,000ml with 2 units of blood product infused, and additional suturing. the patient remained on the impella rp and cp support till the family made the choice to withdraw care, and she expired. |  |  |  |  | Since the original report was filed, the investigation has come to a close. the access site bleed was due to a damaged valve at impella rp's introducer sheath. the cause of the damage was not determined. the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial 1220648-2021-00950 was provided. | The investigation is on-going at this time. upon completion, the final report will be submitted.
36876 | 1220648-2021-00952 | 12194193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12194193&pc=OZD | 01/30/2021 | 07/19/2021 |  | 07/19/2021 |  |  | 06/23/2021 | 07/19/2021 |  |  | 07/31/2021 |  | 07/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1430738 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/23/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old asian female patient had impella 5.0 pump inserted in the right axillary. the patient had hemolysis and as a result blood transfusion was administered. | [5 empty] | Device not returned by customer. the impella 5.0 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
36877 | 12180218 | 12180218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12180218&pc=OZD | 05/18/2021 | 07/01/2021 | 07/01/2021 | 07/16/2021 | 07/01/2021 | 07/16/2021 |  | 07/16/2021 |  |  |  |  | 07/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 6205 DA |  |  | 62 |  |  | Malfunction | No | Yes | Difficult to Advance | Material Twisted/Bent | Physical Resistance/Sticking |  |  |  |  | Insufficient Information | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Act/surgical teams chose impella 5.5 as preferred mode of mechanical support. cardiac planned impella 5.5 pump insertion. as this requires live fluoroscopy and insertion of the pump over a wire, had cut down on right axillary artery and anastomosed in end-to-side fashion with beveled 10 mm goretex graft. using fluoro-guidance (c-arm), the aortic valve was crossed retrograde with a 5 fr pigtail catheter and 0.035" straight wire. the pigtail was placed in the apex and then the wire was exchanged for the impella kit wire. an attempt to insert the impella but the motor portion kept getting stuck at the junction of the graft and the artery. a 12 fr flexor sheath to use as a pusher but this was not successful. removed the impella and advanced a 22fr dry-seal sheath into the proximal innominate artery. the impella was noted to have some kinking of the drive line so a second device was sourced and primed. advanced the impella but again were catching up before entering the left ventricle.
36878 | 1220648-2021-00944 | 12180866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12180866&pc=OZD | 06/17/2021 | 07/16/2021 |  | 07/16/2021 |  |  | 06/17/2021 | 07/16/2021 |  |  | 10/31/2022 |  | 07/16/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2021011066 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | USER F |  |  |  |  | 56 YR | [5 empty] | Injury | No | No | Use of Device Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/17/2021 | * | Invalid/Incomplete |  |  | A patient was admitted in cardiogenic shock after arresting and having cpr in the emergency department. he had walked into the hospital in poor condition. an impella cp was placed for hemodynamic support. the patient had multivessel coronary artery disease and was not able to have emergent angioplasty. the cp was supporting the patient for his transfer out to a tertiary medical center for continued monitoring and care. the medical team at the first center did not remove the 14fr introducer from the femoral access and advance the cp repositioning sheath for support in the icu. this is not recommended per the ifu. the limb became mottled and the tertiary center removed the impella after just under 5 hours of support. surgical repair for the femoral artery was done post explant. | [5 empty] | The medical center did return the product for evaluation, as well as the data logs. the product had no defects observed that could have caused the limb ischemia. the cause was the physician's choice to not remove the 14fr introducer and advance the impella cp. the ifu states the following for patient's like this one undergoing transport: "to prevent failure of the peel-away introducer, remove the peel-away introducer prior to transport when activated clotting time (act) is less than 150 seconds". the failure mode will be monitored and trended.
36879 | 1220648-2021-00940 | 12165795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12165795&pc=OZD | 06/16/2021 | 09/19/2024 |  | 07/14/2021 |  |  | 06/17/2021 | 07/14/2021 |  |  | 01/31/2023 |  | 07/14/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022057173 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 76 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 06/17/2021 |  |  |  |  | Additional information received that patient expired as the procedural outcome. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a (b)(6) female patient had impella cp pump inserted in the right femoral for an unknown indication. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the impella¿s pigtail hit the side wall, which caused blood oozing. a hole at the wall was also noted. the patient's physique is small, and the left ventricle was small. the length from the aortic valve to the left ventricular apex was only 4.5 cm. during surgery, the tip of the pigtail was surgically cut off. the impella did work with no further issues. the perforation was repaired as part of the treatment for vsp during left ventriculoplasty during vsp closure. no further information was provided. |  |  |  |  | The root cause of the ventricular perforation was the patient's anatomy, specifically, the patient has a small left ventricle. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. as noted in the impella ifu: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ corrected information for the initial 1220648-2021-00940 was provided. | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36880 | 1220648-2021-00941 | 12165810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12165810&pc=OZD | 04/07/2021 | 07/14/2021 |  | 07/14/2021 |  |  | 06/16/2021 | 07/14/2021 |  |  | 10/31/2022 |  | 07/14/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022028381 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [28 empty] | Initial submission |  |  |  |  | 06/16/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female patient presenting with acute myocardial infarction and cardiogenic shock for support with the impella cp device. on day 3 of support, the patient developed an ischemic stroke. the physician felt the impella may have had a causal or contributory relationship. | [5 empty] | The impella cp was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36881 | 1220648-2021-00942 | 12166065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12166065&pc=OZD | 07/07/2021 | 07/14/2021 |  | 07/14/2021 |  |  | 07/07/2021 | 07/14/2021 |  |  | 03/31/2023 |  | 07/14/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022057763 |  | 0048-0045 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR | [5 empty] | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/07/2021 | * | Invalid/Incomplete |  |  | A patient was admitted suffering from an acute myocardial infarction. after being stabilized by cpr and intervention she was admitted to the cardiac catherization lab. the patient was diagnosed with multivessel coronary artery disease and had an impella cp placed for hemodynamic support. when the pump exhibited irregular and reduced flows the pump was replaced. support proceeded on the second impella pump. the second pump successfully supported the patient for over 4 days, was weaned and explanted. the first pump was returned for analysis and found to have damage to the impeller. | [5 empty] | The impella pump was returned and analysis was conducted into the reported low flows. the cause of the low pump flow was a damaged impeller. because the fluoroscopic image of the pump/ additional evidence was not returned, a definitive root cause of the fracture could not be determined.
36882 | 1035166-2021-00067 | 12167863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12167863&pc=DYB | 06/16/2021 | 12/10/2021 |  | 07/14/2021 |  |  | 06/18/2021 | 07/14/2021 |  |  | 04/01/2024 |  | 07/14/2021 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | PQOC11882 |  | 0052-3006 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Cardiogenic Shock | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Patient is suffering from post cardiotomy cardiogenic shock. it was reported there was a 23f introducer that was leaking between the sheath and graft. sheath was flushed and prepped with dilator inserted into 10/30 vascutek graft and 2 graft locks attached. leaking noted out of graft. there was a five minute delay in procedure to find replacement sheath and exchange into graft. blood loss was reported, amount of blood loss is unknown. procedure was completed with another sheath. no adverse patient effects reported. | [5 empty] | The device was used in treatment. one 23f abiomed axillary sheath was returned. traces of blood was found inside and outside of the sheath. the returned sheath was observed using a vertex at higher magnification and no cracks, holes or gaps were found. the sheath score lines were observed to be prominent on the outer surface. the sheath was leak tested per iso 11070 standard and procedure for leak test. sheath leak test results did not show any leakage through the sheath score lines or the valve. further investigation revealed that wall thickness of the scoreline has been reduce to improve the peel strength. reduced wall thickness of the scoreline could have made the sheath susceptible to collapsing when used with graft locks, and that could have caused leakage. however, all devices that have been evaluated have met the current dimensional, visual, and functional requirements in the 23f sheath product specification.further investigation revealed that wall thickness of the scoreline has been reduce to improve the peel strength. reduced wall thickness of the scoreline could have made the sheath susceptible to collapsing when used with graft locks, and that could have caused leakage. however, all devices that have been evaluated have met the current dimensional, visual, and functional requirements in the 23f sheath product specification. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per operations procedure(s) adelante, adelante-s and adelante-s2 sheath extrusion appropriate extrusion tool is selected for the part to be extruded and is inspected prior to the extrusion run. peel strength and critical dimensions are verified on every 100th extruded sheath as per procedure. per abiomed introducer sheath in-process and final sample size: 100% inspection the sheath is inspected with a naked eye at a distance of 12" to 18" to ensure the sheath is free of kinks, cracks, splits, sinks and any other damages. leak test is performed as per procedure. abiomed instructions for use (ifu) is provided with the product. based on the investigation, a capa was not required. there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report isbased upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete.
36883 | 1220648-2021-00943 | 12169806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12169806&pc=OZD | 06/19/2021 | 07/14/2021 |  | 07/14/2021 |  |  | 06/21/2021 | 07/14/2021 |  |  | 01/31/2023 |  | 07/14/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022060060 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 44 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/21/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old male patient had impella cp pump inserted in an unknown location for myocarditis. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported the patient experienced bleeding at the impella access site. since bleeding could not be controlled, impella was surgically removed. the cardiac surgeon sutured the blood vessel and hemostasis was obtained. the patient did have physical movement; however, the relationship between physical movement and bleeding is unknown. furthermore, it was indicated that the patient experienced a vascular injury. the injury required treatment; however, details of the treatment were not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36884 | 1220648-2021-00939 | 12147501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12147501&pc=OZD | 06/15/2021 | 07/09/2021 |  | 07/09/2021 |  |  | 06/16/2021 | 07/09/2021 |  |  | 10/31/2022 |  | 07/09/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2021006132 |  | 0550-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 64 YR | [5 empty] | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Cardiac Output | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/16/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella 5.5 pump inserted in the right axillary. the pump failed due to flow dropping, cardiopulmonary resuscitation (cpr) was needed, and the patient was hemodynamically instable so the physician implanted extracorporeal life support (ecls). | [5 empty] | Device not returned by customer. the impella 5.5 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
36885 | 1220648-2021-00929 | 12134344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12134344&pc=OZD | 02/02/2021 | 07/08/2021 |  | 07/08/2021 |  |  | 06/10/2021 | 07/08/2021 |  |  | 10/31/2022 |  | 07/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021017022 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 50 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/10/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for an unknown indication. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported the patient developed sepsis during use of impella support. various samples were collected including blood cultures. the patient was administered antimicrobials and the catheter was replaced. supplement at3 (antithrombin iii) were given as needed. small amount of treprostinil was also administered. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36886 | 1220648-2021-00930 | 12134367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12134367&pc=OZD | 06/11/2021 | 07/08/2021 |  | 07/08/2021 |  |  | 06/11/2021 | 07/08/2021 |  |  | 10/31/2022 |  | 07/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021017581 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 86 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [41 empty] | Initial submission |  |  |  |  | 06/14/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male patient had impella cp pump inserted in the left femoral for high-risk percutaneous coronary intervention (hrpci). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that after imella was placed, it was replaced with a new sheath after pci was performed. after the replacement, there was subcutaneous hematoma, so manual compression was performed. hemostasis was possible, and the puncture angle was maintained and fixed. however, the patient was relieved and had wrinkled bleeding because it was difficult to maintain the angle. as a result, blood transfusion was performed, amount is unknown. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36887 | 1220648-2021-00931 | 12134384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12134384&pc=OZD | 03/24/2021 | 07/08/2021 |  | 07/08/2021 |  |  | 06/08/2021 | 07/08/2021 |  |  | 10/31/2022 |  | 07/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021017015 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/08/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male patient had impella cp pump inserted in the right femoral for post-cardiotomy cardiogenic shock (pccs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that five days after impella was placed there was an increase in pericardial fluid and shock were observed. the patient was diagnosed with cardiac tamponade. the physician stated that this event was likely because of the surgery, the effects of subsequent heparin use. as a result, pericardium was placed after the surgery, and the mediastinal was drained. per the physician the event is likely not related to the device; however, it is unclear. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36888 | 1220648-2021-00932 | 12134460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12134460&pc=OZD | 06/07/2021 | 07/08/2021 |  | 07/08/2021 |  |  | 06/09/2021 | 07/08/2021 |  |  | 07/31/2021 |  | 07/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1432562 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 45 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted in the right femoral myocarditis. the patient was implanted on (b)(6) 2021 explanted on (b)(6) 2021. it was reported impella was caught in the chordae tendineae or papillary muscles when the device was inserted. this made it difficult to adjust the position. as a result, bilirubin rose which resulted in impella being explanted prematurely. the physician stated there was no change to the patient¿s status due to also having pcps inserted. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36889 | 1220648-2021-00933 | 12134464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12134464&pc=OZD | 05/26/2021 | 07/08/2021 |  | 07/08/2021 |  |  | 06/09/2021 | 07/08/2021 |  |  | 09/30/2022 |  | 07/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1502523 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that after inserting impella, pci was performed with the peel-away sheath, there was no hematoma at that time. after the pci was completed, the repositioning sheath was replaced, and during other treatments, the right inguinal region increased rapidly. as a result, lower limb angiography was performed. no blood vessels were broken; however, bleeding was observed. this was judged to be due to the sheath gap. subsequent hematoma growth resulted in hemostasis. the patient received blood products due to insufficient circulating blood volume which was related to the bleed. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36890 | 1220648-2021-00934 | 12134469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12134469&pc=OZD | 04/25/2021 | 07/08/2021 |  | 07/08/2021 |  |  | 06/14/2021 | 07/08/2021 |  |  | 06/30/2022 |  | 07/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1484369 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/14/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for an unknown indication. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported the patient developed limb ischemia during impella support. as a result, the device was explanted early. no further information was provided. | [5 empty] | The impella cp was has not been returned for analysis by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36891 | 1220648-2021-00935 | 12134476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12134476&pc=OZD | 06/13/2021 | 07/08/2021 |  | 07/08/2021 |  |  | 06/14/2021 | 07/08/2021 |  |  | 01/31/2023 |  | 07/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022043002 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/14/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient developed hemolysis during impella management. haptoglobin was administered; however, there was no improvement. the position on the echo was problematic; however, there was no interference with the mitral valve. additionally, the left chamber was not collapse; however, there seemed to be no volume or right heart failure. due to this the physician believed that the number of revolutions of the impella may have an effect. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36892 | 1220648-2021-00936 | 12134501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12134501&pc=OZD | 06/03/2021 | 07/08/2021 |  | 07/08/2021 |  |  | 06/15/2021 | 07/08/2021 |  |  | 01/31/2023 |  | 07/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022062674 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/15/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the right femoral for myocarditis. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that during impella support, the patient experienced access site bleed. surgical hemostasis was performed; however, there was no improvement even after adjusting the fixation angle. as a result, vascular sutures were performed. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36893 | 1220648-2021-00937 | 12134520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12134520&pc=OZD | 02/08/2021 | 07/08/2021 |  | 07/08/2021 |  |  | 06/15/2021 | 07/08/2021 |  |  | 09/30/2022 |  | 07/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1502523 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/15/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported the patient developed access site bleed during impella support. as a result, the patient was administered red blood cells and fresh frozen plasma, amount was not reported. no further information was provided. | [5 empty] | The impella cp ha snot been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36894 | 1220648-2021-00938 | 12140067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12140067&pc=OZD | 06/14/2021 | 07/08/2021 |  | 07/08/2021 |  |  | 06/16/2021 | 07/08/2021 |  |  | 02/28/2023 |  | 07/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022050187 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/16/2021 | * | Invalid/Incomplete |  |  | The complainant reported a patient whose age and gender were not provided had impella cp pump inserted in an unknown location and for an unknown indication. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient experienced limb ischemia during impella support. reverse puncture was performed; however, blood flow was poorly improved, and the patient was removed the day after implant. limb ischemia did improve, and the patient's treatment was continued with an intra-aortic balloon pump (iabp). no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36895 | 1220648-2021-00923 | 12100830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12100830&pc=OZD | 06/05/2021 | 07/01/2021 |  | 07/01/2021 |  |  | 06/05/2021 | 07/01/2021 |  |  | 02/28/2023 |  | 07/01/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022049832 |  | 0048-0045 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | USER F |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [41 empty] | Initial submission |  |  |  |  | 06/05/2021 | * | Invalid/Incomplete |  |  | The patient was admitted in cardiogenic shock and angina. when the cardiac catherization revealed coronary artery disease, the team placed and impella cp for hemodynamic support. the pump remained in for over 8 hours and then was explanted. the explant was due to signs and concerns of hemolysis. the team had attempted to reposition the pump but, the hemolysis was not remedied. the pump was replaced. the second pump successfully supported the patient for another 2 days, and was then weaned and explanted. | [5 empty] | The medical center did return the impella pump for analysis. the pump was inspected and found no damage or biomaterial that could have caused hemolysis. as the team did relay that the pump had been in poor position during support, the improper pump position is the root cause of the hemolysis. the failure mode will be monitored and trended.
36896 | 1220648-2021-00924 | 12102335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12102335&pc=OZD | 06/07/2021 | 08/16/2021 |  | 07/01/2021 |  |  | 06/07/2021 | 07/01/2021 |  |  | 04/30/2023 |  | 07/01/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022071662 |  | 0048-0045 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 47 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 06/07/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted. repositioning sheath wired, pump removed and exchanged for a fresh 14fr sheath, upon removal of the pump, a hole was found in the left atrial appendage, and the physician had to repair it during surgery. | [5 empty] | Review of the provided data logs found that the purge system was deaired at 10:34 on (b)(6) 2021. the pump was then unplugged from the console at 11:30 on (b)(6) 2021. the pump was reconnected at 11:38 and unplugged again at 11:55. the data logs did not indicate any issues with the pump or purge that would have contributed to the reported air in the left ventricle. evaluation of the returned impella pump found a kink in the cannula proximal to the outflow cage along with several catheter kinks distal to the pump house. the pump was run in the lab and no alarms or leaks were noted. there was no damage to the device that would have contributed to the reported air in the left ventricle. the root cause of the reported air in the left ventricle was most likely due to the hole found in the patient's left atrial appendage; however, the root cause of the hole in the left atrial appendage was most likely related to guidewire use. | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36897 | 1220648-2021-00925 | 12102568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12102568&pc=OZD | 06/09/2021 | 09/17/2024 |  | 07/01/2021 |  |  | 06/09/2021 | 07/01/2021 |  |  | 03/31/2023 |  | 07/01/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMD, INC. |  |  | US |  | 2022061889 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/09/2021 |  |  |  |  | It was noted that the patient expired in the operating room. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a (b)(6) female patient had impella cp pump inserted for cardiogenic shock (cgs). due to the large habitus of the patient when the physician did the sheath exchange, the repositioning sheath was too short to reach the arteriotomy and the patient bled profusely. the physician force the pump down inside with the hub and suture ribs all inside the skin tract to reach the artery. |  |  |  |  | The cause of the bleeding was large skin tissue preventing the repo unit from reaching the artery. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2021-00925 was provided. | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36898 | 1220648-2021-00926 | 12102629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12102629&pc=OZD | 06/11/2021 | 07/01/2021 |  | 07/01/2021 |  |  | 06/11/2021 | 07/01/2021 |  |  | 03/31/2023 |  | 07/01/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022059592 |  | 0048-0045 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/11/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) white male patient had impella cp pump inserted in the right femoral artery (rfa) for acute myocardial infarction (ami)/cardiogenic shock (cgs). the physician did an angiography of the right groin which showed distal leg perfusion despite mottled foot and absent pulse, and the pump was removed. after removal color was improving in the foot. | [5 empty] | Device not returned by customer. the impella cp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
36899 | 1220648-2021-00927 | 12102703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12102703&pc=OZD | 06/22/2021 | 07/01/2021 |  | 07/01/2021 |  |  | 06/22/2021 | 07/01/2021 |  |  | 03/31/2023 |  | 07/01/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022056618 |  | 0048-0045 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 82 YR | [5 empty] | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission |  |  |  |  | 06/22/2021 | * | Invalid/Incomplete |  |  | The medical team placed an impella cp for hemodynamic support. earlier in the day the patient had a coronary angioplasty and had coded during the intervention. the cp pump ran for approximately 34 minutes and was replaced by a new cp due to low pump flows. the pump was returned for analysis and the impeller blade was found to be damaged. | [5 empty] | The root cause of the low flow was fractured impeller blades. because the fluoroscopic image of the pump/ additional evidence was not returned, a definitive root cause of the fracture could not be determined. based on our assessment it was most likely that the blade fracture was caused by pump - other device interaction. the failure mode will be monitored and trended.
36900 | 1220648-2021-00928 | 12102729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12102729&pc=OZD | 06/07/2021 | 07/01/2021 |  | 07/01/2021 |  |  | 06/07/2021 | 07/01/2021 |  |  | 09/03/2022 |  | 07/01/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2021006153 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [5 empty] | ECMO | [22 empty] | Initial submission |  |  |  |  | 06/07/2021 | * | Invalid/Incomplete |  |  | A patient was admitted in (b)(6) suffering from an acute myocardial infarction. the patient had an intra-aortic balloon pump (iabp) placed when the coronary arteries were evaluated and determination made that the patient needed bypass (cabg) done to perfuse the heart. when the patient's condition continued to deteriorate, the care was escalated to an impella cp. the cp supported the patient for 5 days. during the 5 days there were signs of hemolysis and limb ischemia. the limb ischemia was bilaterally observed, both in the impella limb and ecmo cannulated legs. to treat the hemolysis the team administered haptoglobin. to treat the limb ischemia external bypass was placed. when the patient went to the or for the cabg, the impella cp was explanted and care was escalated to the impella 5.0. | [5 empty] | The medical center in (b)(6) discarded the impella. no investigation and hemolysis testing was possible. further clinical details and data logs have not been returned due to covid hospital access. at this time no root cause determination is possible. should more details be made possible, that lead to the definitive root cause, attributed to the use of impella, a final report will be filed.
36901 | 1220648-2021-00919 | 12084049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12084049&pc=OZD | 04/26/2021 | 06/29/2021 |  | 06/29/2021 |  |  | 06/01/2021 | 06/29/2021 |  |  | 09/30/2022 |  | 06/29/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1503409 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Ischemia | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/01/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported the patient developed an infection at insertion site during impella support. as a result, antibacterial agents were administered. subsequently, the patient was at risk of ischemia with almost no contrast enhancement of sfa before surgery. the impella was left in place without bypass and was removed due to signs of lower limb ischemia. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36902 | 1220648-2021-00920 | 12084205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12084205&pc=OZD | 05/31/2021 | 06/29/2021 |  | 06/29/2021 |  |  | 05/31/2021 | 06/29/2021 |  |  | 03/31/2023 |  | 06/29/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022062050 |  | 0048-0045 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/31/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support. during support, the patient exhibited signs of hemolysis as evident by pink urine. despite efforts to resolve with repositioning, the issue persisted and the pump was prematurely removed. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36903 | 1220648-2021-00921 | 12084217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12084217&pc=OZD | 02/25/2021 | 06/29/2021 |  | 06/29/2021 |  |  | 06/07/2021 | 06/29/2021 |  |  | 09/30/2021 |  | 06/29/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1446706 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 50 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [41 empty] | Initial submission |  |  |  |  | 06/07/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that during impella support the patient showed signs and symptoms of hemolysis. hemolysis was later confirmed via visual findings seen on blood test. as a result, the patient was administered haptoglobin. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36904 | 1220648-2021-00922 | 12084232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12084232&pc=OZD | 06/04/2021 | 06/29/2021 |  | 06/29/2021 |  |  | 06/04/2021 | 06/29/2021 |  |  | 02/28/2023 |  | 06/29/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022070790 |  | 0048-0024-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [40 empty] | Initial submission |  |  |  |  | 06/04/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for support with the impella cp optical device. after advancing the repositioning sheath, during placement of the device, the access site developed a persistent bleed. the blood vessels around the access site were noted as torn. as a result, the pump was surgically removed and the injury was repaired. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36905 | 1220648-2021-00917 | 12078913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12078913&pc=OZD | 06/04/2021 | 06/28/2021 |  | 06/28/2021 |  |  | 06/04/2021 | 06/28/2021 |  |  | 02/28/2023 |  | 06/28/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED INC. |  |  | US |  | 2022064806 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 83 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hematuria | [10 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/04/2021 | * | Invalid/Incomplete |  |  | A patient was admitted suffering from an acute myocardial infarction and cardiogenic shock in (b)(6). the team placed ecmo and impella cp support for her hemodynamic and respiratory needs. the patient suffered from a blood loss a the impella access site that prompted surgical intervention in the form of vessel suture. the impella repositioning sheath was also observed to be torn, per the medical team in japan. the team believed hemolysis to be occurring during the impella support. the hemolysis was diagnosed by the color change in the patient's urine output and need for dialysis. after 11 days of support the impella was weaned and explanted. | [5 empty] | The medical center in (b)(6) discarded the impella product at explant. there will be no analysis or hemolysis testing done. should more information be shared that leads to a root cause, a final report will be filed. the failure mode will be monitored and trended.
36906 | 1220648-2021-00918 | 12078915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12078915&pc=OZD | 06/03/2021 | 02/23/2022 |  | 06/28/2021 |  |  | 06/04/2021 | 06/28/2021 |  |  | 07/31/2022 |  | 06/28/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1503732 |  | DANVERS |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional |  |  |  | 30 YR |  | Female |  |  | Asian | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Myocarditis | Unspecified Vascular Problem | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup |  |  | 06/04/2021 |  |  |  |  | New information was received on 2022-01-24, via online publication: during removal of the impella, the patient endured a vascular injury resulting in 2 units of blood products transfused. | The complainant reported a (b)(6)-year-old female patient had impella 5.0 pump inserted in the right axiallry for myocarditis. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that on the 23rd day of use there was a pump stop. the physician was unsuccessful at restarting the device. there were no other devices being used on the patient. no further information was provided. |  |  |  |  | The device was returned for analysis. no data logs were provided. the returned pump was plugged into a console and ran in a basin of water and would not start. the pump electrical resistances were all within specification. the cannula was removed, and the impeller was visualized during which it was noted that the impeller required considerable force to rotate initially and would only rotate approximately 135 degrees thereafter before stopping. therefore, the root cause of the pump stop was most likely bearing retainer wear. the product was released with a completed and reviewed device history record under the abiomed quality system. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event. | The device is no longer available for investigation as it was discarded after the initial investigation for pump stop. therefore, we are unable to provide a definitive root cause for the newly reported vascular injury. | The impella 5.0 was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
36907 | 1220648-2021-00916 | 12073168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12073168&pc=OZD | 01/15/2021 | 06/27/2021 |  | 06/27/2021 |  |  | 06/10/2021 | 06/27/2021 |  |  | 08/31/2022 |  | 06/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1496670 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/10/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported the patient developed sepsis during use of impella support. various samples were collected including blood cultures. the patient was administered antimicrobials and the catheter was replaced. supplement at3 (antithrombin iii) were given as needed. small amount of treprostinil was also administered. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36908 | MW5102127 | 12066348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12066348&pc=PBL | 04/21/2021 | 06/22/2021 |  | 06/24/2021 |  |  |  | 06/24/2021 |  |  |  |  | 06/24/2021 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA PURGE CASSETTE |  |  | PBL | [5 empty] | ABIOMED EUROPE GMBH | [8 empty] | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Voluntary report | [8 empty] | 72 YR | [5 empty] | Injury | No | Yes | Improper Flow or Infusion | No Flow | Unexpected Shutdown |  |  |  |  | Insufficient Information | [12 empty] | Life Threatening | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Impella purge cassette malfunctioned. multiple attempts to stabilize the purge flow were marginally successful. after several hours the purge flow stopped and the impella malfunctioned and shut off. pt placed on ecmo. fda safety reported id# (b)(4).
36909 | 1220648-2021-00908 | 12051979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12051979&pc=OZD | 05/01/2021 | 06/23/2021 |  | 06/23/2021 |  |  | 05/26/2021 | 06/23/2021 |  |  | 01/31/2023 |  | 06/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022038039 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/27/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that when pcps were inserted to transfer blood from the side branch of the pcps circuit to prevent lower limb ischemia. when pcps was removed signs and symptoms of lower limb ischemia appeared. as a result, a by-pass to the brachial-femoral artery was poor, and the impella device was explanted after several days of support. it was noted that the patient's blood vessels were initially thin; however, impella relation to leg ischemia was suspected. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36910 | 1220648-2021-00909 | 12051986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12051986&pc=OZD | 05/30/2021 | 07/12/2021 |  | 06/23/2021 |  |  | 05/30/2021 | 06/23/2021 |  |  | 03/31/2023 |  | 06/23/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022062043 |  | 0048-0045 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | USER F |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 05/30/2021 | * | Invalid/Incomplete |  |  | A patient was admitted via the emergency department and was actively arresting. he was taken to the cardiac catheterization lab where a diagnosis of coronary artery disease was made. the team placed an impella cp for hemodynamic support. the pump ran for 26 hours. during that time his labs were rising and the team determined the impella cp was causing hemolysis. the pump was explanted and cvvh dialysis was begun. | [5 empty] | The impella product was returned and benchtop analysis was conducted. the pump was found to have no abnormalities that could have contributed to the hemolysis. the pump passed hemolysis testing. the root cause of the hemolysis is patient condition. the failure mode will be monitored and trended. | The investigation into the hemolysis is on-going at this time. upon completion of the investigation, a final report will be filed.
36911 | 1220648-2021-00910 | 12052135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12052135&pc=OZD | 05/30/2021 | 06/23/2021 |  | 06/23/2021 |  |  | 05/30/2021 | 06/23/2021 |  |  | 02/28/2023 |  | 06/23/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022056155 |  | 0048-0045 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | USER F |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/30/2021 | * | Invalid/Incomplete |  |  | The medical center admitted a patient and found her to have multivessel coronary artery disease. an impella cp was placed for hemodynamic support. the team then made a choice to escalate her care to a tertiary medical center. she was transferred on impella support. at the second tertiary center they observed signs of hemolysis. the team explanted her impella as the urine color had changed and labs were hemolyzing. the team placed a new, second pump and support proceeded. | [5 empty] | The medical center did not return the pump for analysis and benchtop hemolysis testing. without this analysis a definitive root cause of the presumed hemolysis can not be determined. it is assumed to be due to native patient condition.
36912 | 1220648-2021-00907 | 12045780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12045780&pc=OZD | 05/28/2021 | 06/22/2021 |  | 06/22/2021 |  |  | 05/28/2021 | 06/22/2021 |  |  | 08/31/2022 |  | 06/22/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED INC. |  |  | US |  | 1497960 |  | 005060 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 56 YR | [5 empty] | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission |  |  |  |  | 05/28/2021 | * | Invalid/Incomplete |  |  | The complainant in the eu had an impella 5.0 supporting a patient due to instability after bypass (cabg) surgery. the 5.0 was place in addition to ecmo. after 9 days the pump stopped and was explanted. this was within the indicated use timeframe. | [5 empty] | The medical center discarded the impella pump, but did however return the data logs from the console. the logs revealed the pump had become disconnected during use. the team did note that the patient was sedated and during a time of confusion may have disconnected his own pump, as a "use" error. the failure mode will be monitored and trended.
36913 | 1220648-2021-00903 | 12022097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12022097&pc=OZD | 05/27/2021 | 07/21/2021 |  | 06/17/2021 |  |  | 05/27/2021 | 06/17/2021 |  |  | 03/31/2023 |  | 06/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022062561 |  | 0048-0045 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 05/27/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella cp pump inserted without issues. the patient had a left ventricle (lv) perforation and pericardial effusion as a result pericardiocentesis was done in cardiac catheterization laboratory (ccl) and patient was sent to emergent surgery. | [5 empty] | Analysis of the returned product revealed a large kink in the catheter and no other abnormalities. data log review did not show the pump in the ventricle at any time, but did show placement signal low alarms indicating the outlet may have been close to the valve. while the pump was not deep in the ventricle, the patient's anatomical variations and the fact the pump was close to the valve may have caused the kink in the catheter. as there is no imaging to show the impella pierced directly through the ventricle, the root cause of the ventricle perforation could not be definitively determined. however, the patient's anatomical variations along with the outlet on the valve may have been a contributing factor. | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36914 | 1220648-2021-00904 | 12022098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12022098&pc=OZD | 05/22/2021 | 09/17/2024 |  | 06/17/2021 |  |  | 05/25/2021 | 06/17/2021 |  |  | 02/28/2023 |  | 06/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 2022046944 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 82 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2021 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported an (b)(6) male patient had impella cp pump inserted in the femoral for cardiogenic shock (cgs). the patient had severe bleeding and went into hemorrhagic shock, a hematoma was present in the groin, four (4) units of blood was given, one (1) unit of fresh frozen plasma (ffp), and hematocrit was around 0.3 at all time, and the patient expired due to cardiogenic shock (cgs) and systemic organ failure hours later. |  |  |  |  | The investigation for the reported access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the result of the clinical review states, "the clinicians stated that the bleeding occurred after the introducer was removed. a massive hematoma developed at the site which caused a gap between the skin and the femoral artery. resultantly, the instrument was initially unable to properly seal the access site. the bleeding was unable to be effectively stopped by the clinicians despite the administering of blood products, and the patient became unstable and eventually fell into hemorrhagic shock. the team was eventually able to stop the bleeding by making a small incision through the hematoma to access the arteriotomy and inserting the blue suture hub of the instrument to seal the incision. the clinicians asserted that best practices were followed during impella setup and insertion, including with introducer removal. ultimately, the patient expired several hours later due to cardiogenic shock and multiple organ failure. all impella products were discarded. the patient's condition was previously unstable during pci, requiring the need of resuscitation prior to impella use. the results of the investigation state the following, "after reviewing the clinical information, the cause of the access site bleeding was not determined since product was not returned". as noted in the impella ifu: impella cp with smart assist system. impella rp with smart assist system with automated impella controller: section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound." ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36915 | 1220648-2021-00905 | 12022100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12022100&pc=OZD | 05/27/2021 | 07/21/2021 |  | 06/17/2021 |  |  | 05/27/2021 | 06/17/2021 |  |  | 11/30/2021 |  | 06/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 IMC PUMP SET CANADA |  |  | ABIOMED, INC. |  |  | US |  | 1475168 |  | 005064 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 39 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 05/27/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 5.0 pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had a hematoma that was evacuated and then closed with a drain put into place. | [5 empty] | Upon further communication with the rep, we found out that the hematoma was unrelated to the impella, but rather a patient vascular condition. the product was returned and evaluated. the root cause of the blood leaking from the pump plug was determined to be a hole in the catheter, which was likely caused by a suture needle prick. | The impella 5.0 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36916 | 1220648-2021-00906 | 12022101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=12022101&pc=OZD | 05/24/2021 | 06/17/2021 |  | 06/17/2021 |  |  | 05/24/2021 | 06/17/2021 |  |  | 04/30/2023 |  | 06/17/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022069880 |  | 0048-0045 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | USER FACILITY | [5 empty] | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 05/24/2021 | * | Invalid/Incomplete |  |  | The medical center placed an impella cp for hemodynamic support for a patient admitted in cardiogenic shock. the team found the patient to have multivessel coronary artery disease and had angioplasty performed with the support of the impella. during the support the team found the impella to alarm with suction. the patient had many units of blood and fluid delivered for the suction and groin site bleed. the bleed at the access site was also treated with surgical repair. when sent to the ct department for imaging the team found no retroperitoneal bleed. of note, the sheath was left in til the patient was admitted to the icu and the blue suture pads had been advanced into the tract, which is not as trained and intended, and could have contributed to the groin site bleed. | [5 empty] | The medical center did return the product for analysis and evaluation. the cause of the access site bleed was related to the management of the access site, which included the placement of the butterfly into the arteriotomy. the failure will be monitored and trended.
36917 | 1220648-2021-00901 | 11999309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11999309&pc=OZD | 05/19/2021 | 06/22/2021 |  | 06/15/2021 |  |  | 05/19/2021 | 06/15/2021 |  |  | 11/30/2022 |  | 06/15/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021020214 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 05/19/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical. approximately 2 days into support, the patient developed leg ischemia in the impella leg. as treatment, the device was prematurely removed. | The complainant reported a 67 year old female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical. approximately 2 days into support, the patient developed arm ischemia in the impella arm. as treatment, the device was prematurely removed. |  |  |  |  | The impella cp optical was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | Section b5 states, "the patient developed leg ischemia in the impella leg." this was erroneous and should read, "the patient developed arm ischemia in the impella arm."
36918 | 1220648-2021-00902 | 11999355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11999355&pc=OZD | 05/21/2021 | 06/30/2021 |  | 06/15/2021 |  |  | 05/21/2021 | 06/15/2021 |  |  | 01/31/2022 |  | 06/15/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | CANADA 5.0 AIC |  |  | ABIOMED INC. |  |  | US |  | 1466192 |  | 005064-RO |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 62 YR | [5 empty] | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 05/21/2021 | * | Invalid/Incomplete |  |  | The medical team in (b)(6) had an impella 5.0 supporting a patient admitted in cardiogenic shock suffering from an ami. the hemodynamic support was enough to allow for the iabp to be removed and patient improvement. after just under 24 hours of support the purge flow was low and purge pressure rose. the team attempted troubleshooting but, the pump stopped and was explanted. there was clot observed at the pump at explant. after this pump stop the patient was seen to be stable and the impella was not replaced. | [5 empty] | To date the medical center in (b)(6) has not yet returned for analysis the pump product or data logs. without product a root cause for the pump stop, within the indicated use timeframe, is not possible. upon completion of the investigation, a final report will be filed. | The medical center returned the pump product for analysis. the medical center did not return the data logs however. when the pump was analyzed the pump was found to have blood ingress. the blood ingress cause however could not be determined as no data logs were returned, as such the root cause is not determined. the failure mode will be monitored and trended.
36919 | 1220648-2021-00896 | 11973460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11973460&pc=OZD | 05/15/2021 | 06/10/2021 |  | 06/10/2021 |  |  | 05/15/2021 | 06/10/2021 |  |  | 10/31/2022 |  | 06/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD |  |  | IMPELLA CP PUMP SET, |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021011903 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/17/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6)-year-old (b)(6) female patient had impella cp pump inserted in the left femoral for high-risk percutaneous coronary interventions (hrpci). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the day after removal, thrombus was surgically removed for thrombus obstruction in the lower extremities. the thrombotic obstruction occurred on the impella inserted side. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36920 | 1220648-2021-00897 | 11973464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11973464&pc=OZD | 05/14/2021 | 06/10/2021 |  | 06/10/2021 |  |  | 05/17/2021 | 06/10/2021 |  |  | 03/31/2022 |  | 06/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021017072 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | HEALTH PROFESSIONAL |  |  |  |  | 69 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 05/17/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in an unknown location for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that cerebral hemorrhage and cerebral herniation occurred during impella support which was confirmed via ct. lifesaving was severe (with cerebral herniation), and the treatment policy was changed to palliative treatment. due to palliative treatment unnecessary devices were removed including impella. there was no additional treatment for cva and the physical noted that the pump worked as intended. date of death is unknown. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36921 | 1220648-2021-00898 | 11973466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11973466&pc=OZD | 05/07/2021 | 06/10/2021 |  | 06/10/2021 |  |  | 05/18/2021 | 06/10/2021 |  |  | 02/28/2023 |  | 06/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022043371 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | HEALTH PROFESSIONAL |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/18/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the femoral artery (side not provided for myocarditis. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that thrombocytopenia occurred during impella support. as a result, an unknown number of platelets were transfused. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36922 | 1220648-2021-00899 | 11979166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11979166&pc=OZD | 05/14/2021 | 06/10/2021 |  | 06/10/2021 |  |  | 05/19/2021 | 06/10/2021 |  |  | 02/28/2023 |  | 06/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022052352 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 49 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/19/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old asian male patient had impella cp pump inserted in the left femoral for acute myocardial infarction (ami) with shock. the patient had access site bleeding, an unknown amount of red blood cells (rbcs) was administered, and as a result hemostasis was performed by surgical vascular suture. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36923 | 1220648-2021-00900 | 11979214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11979214&pc=OZD | 05/21/2021 | 06/10/2021 |  | 06/10/2021 |  |  | 05/21/2021 | 06/10/2021 |  |  | 12/31/2022 |  | 06/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022023830 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 41 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/21/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male patient had impella cp pump inserted in the right femoral for acute myocardial infarction (ami) with shock. the patient had cerebral infarction during pump support and the treatment for cerebral infarction was performed as usual. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36924 | 1220648-2021-00894 | 11931551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11931551&pc=OZD | 02/26/2021 | 06/03/2021 |  | 06/03/2021 |  |  | 05/07/2021 | 06/03/2021 |  |  | 10/31/2022 |  | 06/03/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021017027 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/07/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the femoral artery for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that while on impella support the patient developed an ischemic cerebrovascular accident (cva). it is unknown if other additional treatment was performed. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36925 | 1220648-2021-00895 | 11931595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11931595&pc=OZD | 03/16/2021 | 06/03/2021 |  | 06/03/2021 |  |  | 05/07/2021 | 06/03/2021 |  |  | 12/31/2022 |  | 06/03/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022023946 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/07/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old female patient had impella 2.5 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient developed access site bleed during impella support. as a result, the patient was administered 2 units of blood products. no further information was provided. | [5 empty] | The impella 2.5 was not return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36926 | 1220648-2021-00892 | 11915507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11915507&pc=OZD | 06/29/2020 | 06/01/2021 |  | 06/01/2021 |  |  | 05/07/2021 | 06/01/2021 |  |  | 01/31/2022 |  | 06/01/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1473816 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 86 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/07/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old asian female patient had impella cp pump inserted in the right femoral for acute myocardial infarction (ami) with shock. the patient had cerebrovascular accident (cva). | [5 empty] | Device not returned by customer. the impella cp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
36927 | 1220648-2021-00893 | 11915539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11915539&pc=OZD | 05/09/2021 | 07/09/2021 |  | 06/01/2021 |  |  | 05/09/2021 | 06/01/2021 |  |  | 02/28/2023 |  | 06/01/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022043425 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 43 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 05/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella 5.5 pump inserted via axillary approach. the pump stopped, explanted and another 5.5 pump was inserted. | [5 empty] | The device was returned for evaluation and did not have any visual damage, however there was blood ingress in the purge lumen, indicating blood in the motor. without the data logs, the events leading up to the blood ingress is unknown. the root cause of the pump stop was blood ingress, however it is unknown what allowed the blood to enter without the data logs. | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36928 | 1220648-2021-00891 | 11900603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11900603&pc=OZD | 11/11/2020 | 05/27/2021 |  | 05/28/2021 |  |  | 05/06/2021 | 05/28/2021 |  |  | 12/31/2021 |  | 05/28/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1461608 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/06/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2020 and explanted on (b)(6) 2020. it was reported that the patient developed access site bleed during impella support. as a result, the patient was administered blood products and pressure hemostasis was applied. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36929 | 1220648-2021-00890 | 11900606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11900606&pc=OZD | 05/02/2021 | 05/27/2021 |  | 05/28/2021 |  |  | 05/02/2021 | 05/28/2021 |  |  | 07/31/2022 |  | 05/28/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC |  |  | US |  | 1492166 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/02/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that the patient developed signs and symptoms if cerebral hemorrhage. initially it was noted as ischemic hemorrhage. the physician was unable to determine if the event was related to impella or ecmo since both were being used at the same time. as a result, the patient underwent a craniotomy thrombectomy. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36930 | 1220648-2021-00889 | 11900607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11900607&pc=PYX | 05/06/2021 | 07/23/2021 |  | 05/28/2021 |  |  | 05/06/2021 | 05/28/2021 |  |  | 02/28/2023 |  | 05/28/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1484873 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Death | Yes | No | Material Separation | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 05/06/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction (ami/cgs) and cardiogenic shock. during delivery, the guidewire fell out of place and the pump was pulled back in order to re-attempt delivery. in the process of pulling back the device, the pump broke off within the 23 fr introducer and required a surgical snare for removal. the event resulted in blood loss that required 2 units of blood product delivered. the patient was inserted with a new rp, but ultimately expired. the cause of death was noted as cgs, however, it's alleged the blood lost with the original rp contributed to this patient's outcome. | [5 empty] | The device was returned and an investigation was performed. visual inspection found the device's outflow cage bent and confirmed the reported catheter detachment. the failure was reproduced with excessive force. not including the bent outflow cage, there was no other evidence of device failure. therefore, we believe the cause of the reported issue was a deformed outflow cage that led to the physician using excessive force, leading to the breakage. | The impella rp was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36931 | MW5101587 | 11900443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11900443&pc=OJE | 05/06/2021 | 05/25/2021 |  | 05/27/2021 |  |  |  | 05/27/2021 |  |  | 05/31/2022 |  | 05/27/2021 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | ABIOMED IMPELLA RP |  |  | OJE |  |  | 004334 |  |  | ABIOMED INC. |  |  | US |  |  |  | 004334 |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 60 YR | [5 empty] | Death | Yes | No | Break | [6 empty] | Insufficient Information | [12 empty] | Death | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Patient was diagnosed with multi-vessel coronary artery disease and was referred for cabg surgery. patient was declining and surgeon requested impella and impella rp (rt heart) be placed prior to surgery. devices were placed, but there was difficulty placing the catheter for the impella rp due to patient's anatomy. during insertion the catheter broke. another catheter was successfully placed, but the patient continued to decline and expired in the cath lab. fda safety report id# (b)(4).
36932 | 1220648-2021-00888 | 11881480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11881480&pc=OZD | 05/02/2021 | 07/23/2021 |  | 05/25/2021 |  |  | 05/03/2021 | 05/25/2021 |  |  | 11/30/2022 |  | 05/25/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2022028682 |  | 0048-0014 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 05/03/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with "deep" cardiogenic shock, receiving cardiopulmonary resuscitation prior to being taken to the hospital. the impella cp optical was selected for hemodynamic support while percutaneous coronary intervention was performed. during support, the impella had perforated the left ventricle and the patient developed a pericardial tamponade. surgical intervention was performed and the perforation was patched successfully. | [5 empty] | The impella cp optical was returned and an investigation was conducted. the device was free of any visual deficiencies. the returned data logs could not confirm a relationship between the perforation and the operation of the device. | The impella cp optical has been obtained and an analysis is underway. a supplemental report will be filed upon completion of the investigation.
36933 | 1220648-2021-00886 | 11872454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11872454&pc=OZD | 01/16/2021 | 05/24/2021 |  | 05/24/2021 |  |  | 04/27/2021 | 05/24/2021 |  |  | 10/31/2022 |  | 05/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021011325 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/27/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock. the impella cp device was selected for hemodynamic support. during use of the device, the patient endured an ischemic stroke (cerebrovascular accident). the relationship to the impella could not be established. | [5 empty] | The impella cp was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36934 | 1220648-2021-00887 | 11872468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11872468&pc=OZD | 07/28/2020 | 05/24/2021 |  | 05/24/2021 |  |  | 04/27/2021 | 05/24/2021 |  |  | 01/31/2022 |  | 05/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1470941 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | HEALTH PROFESSIONAL | USER FACI |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/27/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female impella patient enduring an access site bleed during support. unknown amount of blood transfusion was required. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36935 | 1220648-2021-00882 | 11865428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11865428&pc=OZD | 03/02/2020 | 05/21/2021 |  | 05/21/2021 |  |  | 04/26/2021 | 05/21/2021 |  |  | 10/31/2021 |  | 05/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC |  |  | US |  | 1453002 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed an access site bleed during impella support. as a result, blood products were administered, amount unknown. no further information was reported. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36936 | 1220648-2021-00881 | 11865450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11865450&pc=OZD | 06/25/2020 | 05/21/2021 |  | 05/21/2021 |  |  | 04/26/2021 | 05/21/2021 |  |  | 12/31/2020 |  | 05/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1391292 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report | [8 empty] | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted into the patients femoral (unknow side) for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that the patient developed hemolysis during impella support. an unknown amount of blood product was administered as the patient had systemic bleeding other than that caused by impella. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36937 | 1220648-2021-00883 | 11865491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11865491&pc=OZD | 04/23/2021 | 05/21/2021 |  | 05/21/2021 |  |  | 04/26/2021 | 05/21/2021 |  |  | 12/31/2022 |  | 05/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET, |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022055480 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella 2.5 pump inserted in an unknown insertion for an unknown indication. it was reported that the patient developed hemolysis during impella support. lowering axillary level did not improve the issue. dialysis was introduced due to deterioration of renal function; however, it is unknown whether renal function deteriorated due to hemolysis using impella or whether the patient developed renal failure due to heart failure. the patient was also administered paptoglobulin. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36938 | 1220648-2021-00884 | 11865518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11865518&pc=OZD | 04/27/2021 | 07/08/2021 |  | 05/21/2021 |  |  | 04/28/2021 | 05/21/2021 |  |  | 11/30/2021 |  | 05/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET, |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1452994 |  |  |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 63 YR | [5 empty] | Malfunction | Yes | Yes | Pumping Stopped | [6 empty] | Insufficient Information | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 04/29/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) -year-old male patient had impella 5.0 pump inserted in the right axillary for post-cardiotomy cardiogenic shock (pccs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported there was an impella defective failure after unplugging the pump to another console. the site attempted to restart the device; however, the connect cable was not recognized. the site found that it was impossible to continue to use impella. the cause of the issue could not be found as there were no abnormalities in the connection. impella support was discontinued and ecmo was managed independently. no further information was provided. | [5 empty] | The impella 5.0 was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion. | The impella 5.0 was returned for review and found to have a large amount of dried blood on the sidearm and the grey connectors. the grey connectors were inspected and found to have blood inside the connection and one of the leads corroded. the pump would not connect to the aic. thus the root cause of the pump stop was blood ingress into the lemo connectors leading to pin corrosion. the data logs were returned for review and confirm the placement signal issue and the pump disconnection. the product was released with a completed and reviewed device history record under the abiomed quality system. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event.
36939 | 1220648-2021-00885 | 11865521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11865521&pc=OZD | 01/10/2021 | 05/21/2021 |  | 05/21/2021 |  |  | 05/06/2021 | 05/21/2021 |  |  | 08/31/2022 |  | 05/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 259693 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 42 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/06/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old female patient had impella cp pump inserted in an unknown location for an unknown indication. the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the patient developed access site bleed during impella support. as a result, the patient was administered 10 or more units of blood products. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36940 | 1220648-2021-00880 | 11856561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11856561&pc=OZD | 05/11/2021 | 06/22/2021 |  | 05/20/2021 |  |  | 05/11/2021 | 05/20/2021 |  |  | 03/31/2023 |  | 05/20/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED INC. |  |  | US |  | 2022053284 |  | 0048-0002 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hemorrhagic Stroke | Thrombosis/Thrombus | [9 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 05/11/2021 | * | Invalid/Incomplete |  |  | The (b)(6) medical team had placed an impella cp for an ami/cgs patient at a community hospital. they then transferred the patient for care escalation to a tertiary center. the support was ended due to hemolysis concerns on the second day of support, at the tertiary center. when the impella cp was explanted there was biomaterial/clot observed adhered to the pump. at the access site at the first medical center the arteriotomy was in a poor position, as it was noted to be high, and this prompted vascular repair and many units of rbcs to be infused. the blood loss had lead to a diagnosis of hemorrhagic shock. the patient survived these adverse events. | [5 empty] | Since the initial report was filed, the investigation into the presumed hemolysis and access site bleed has concluded. the medical center in switzerland returned the pump product as well as data logs, and an investigation was performed. the cause of the access site bleed and vascular damage was poor technique at initial insertion and when placing the repositioning sheath. the pump was run through benchtop hemolysis testing and was found have ingested biomaterial that caused the hemolysis. | The investigation into the presumed hemolysis and access site bleeding is on-going at this time. the data logs have been returned for analysis, but no impella pump product has as of yet been returned. upon completion of the investigation, a final report will be filed.
36941 | 1220648-2021-00876 | 11845288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11845288&pc=PYX | 04/21/2021 | 05/18/2021 |  | 05/18/2021 |  |  | 04/21/2021 | 05/18/2021 |  |  | 10/31/2022 |  | 05/18/2021 |  | TEMPORARY NONROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED INC. |  |  | US |  | 2021009799 |  | 004334 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 04/21/2021 | * | Invalid/Incomplete |  |  | An impella rp was placed via the femoral vein for a patient in biventricular failure. when the introducer sheath was removed, the pump came out of right ventricular position. in the attempts made to re-advance to optimal position the patient was decompensating. the hemodynamics dropped and blood was lost. the blood loss was estimated to be 2 units. this blood was lost as the pump and sheath were replaced. the team infused back 3 units of blood to the patient and placed a new, second impella rp. | [5 empty] | The medical center team discarded the impella product. as such, no investigation of the product is possible. to date, no root cause of the blood loss and pump loosing position has been determined. should new clinical information come to light that leads to root cause and a definitive product malfunction, a final report will be filed.
36942 | 1220648-2021-00877 | 11845355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11845355&pc=OZD | 02/06/2021 | 05/18/2021 |  | 05/18/2021 |  |  | 04/20/2021 | 05/18/2021 |  |  | 04/30/2022 |  | 05/18/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1488338 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/20/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old female patient had impella cp pump inserted in the right femoral for post-cardiotomy cardiogenic shock (pccs). it was reported the patient developed lower limb ischemia during impella support. as a result, blood was sent from the side branch of the ecmo circuit to the site where lower limb ischemia was occurring. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36943 | 1220648-2021-00878 | 11845361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11845361&pc=OZD | 01/22/2021 | 05/18/2021 |  | 05/18/2021 |  |  | 04/21/2021 | 05/18/2021 |  |  | 11/30/2021 |  | 05/18/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1452229 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/21/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old male patient had impella 5.0 pump inserted in the right axillary for post-cardiotomy cardiogenic shock (pccs). the patient was implanted on (b)(6) 2020 and explanted on (b)(6) 2021. it was reported the patient develop sepsis while on impella support. as a result, the patient was administered antibiotic agents and the mcs device was replaced. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36944 | 1220648-2021-00879 | 11845378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11845378&pc=OZD | 04/22/2021 | 05/18/2021 |  | 05/18/2021 |  |  | 04/22/2021 | 05/18/2021 |  |  | 12/31/2022 |  | 05/18/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022025696 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | OTHER |  |  |  |  | 89 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [40 empty] | Initial submission |  |  |  |  | 04/22/2021 | * | Invalid/Incomplete |  |  | An impella cp was placed via single access for hemodynamic support of a patient suffering from cardiogenic shock and an acute myocardial infarction. the support allowed for the team to perform a multivessel coronary angioplasty. the medical team some weeks later filed a user medwatch pertaining to a vascular injury at the impella cp access site. the field representative had not been aware of the extent of the harm or the need for intervention. the medwatch documented the need for thrombectomy and angioplasty after impella explanted. | [5 empty] | The medical center did not return the impella pump or introducer for analysis. therefore, the root cause of the harm and necessary intervention can not be determined. the failure mode will be monitored and trended.
36945 | 1220648-2021-00872 | 11830111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11830111&pc=OZD | 04/05/2021 | 06/30/2021 |  | 05/15/2021 |  |  | 04/21/2021 | 05/15/2021 |  |  | 07/31/2021 |  | 05/15/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1430738 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 44 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 04/21/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old asian male patient had impella 5.0 pump inserted in the right axillary. the patient had bleeding from impella insertion site and as a result ten (10) units or more blood transfusion was given. | [5 empty] | Review of the provided data logs confirmed several spikes in motor current occurring on april 5, 2021. at 18:45 on april 5th, a motor current spike triggered a pump stop and an 'impella stopped - motor current high' alarm. the pump restarted shortly after; however, additional motor current spikes were observed along with unstable pump flow. evaluation of the returned product found biomaterial around the impeller. the root cause of the low pump flow was ingested biomaterial. the root cause of the access site bleeding was determined to be unrelated to the impella device since it was reported that the patient was bleeding in multiple places. | The impella 5.0 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
36946 | 1220648-2021-00873 | 11830112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11830112&pc=OZD | 04/20/2021 | 05/15/2021 |  | 05/15/2021 |  |  | 04/22/2021 | 05/15/2021 |  |  | 08/31/2021 |  | 05/15/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1453454 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/22/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian female patient had impella 2.5 pump inserted in the right femoral. the patient was bleeding from the access site and as a result four (4) units of red blood cells (rbcs) was administered and surgical repair was done. | [5 empty] | Device discarded by customer. the impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36947 | 1220648-2021-00874 | 11830113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11830113&pc=OZD | 02/03/2021 | 09/15/2024 |  | 05/15/2021 |  |  | 04/20/2021 | 05/15/2021 |  |  | 06/30/2021 |  | 05/15/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1425997 |  | 005048-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2021 |  |  |  |  | The procedural outcome is patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported an asian patient had impella 2.5 pump inserted for cardiogenic shock (cgs). the patient had cardiac tamponade. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2021-00874 was provided. | Device discarded by customer. the impella 2.5 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
36948 | 1220648-2021-00875 | 11830114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11830114&pc=OZD | 03/09/2020 | 05/18/2021 |  | 05/15/2021 |  |  | 05/06/2021 | 05/15/2021 |  |  | 09/30/2021 |  | 05/15/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1446479 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 05/06/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old asian male patient had impella cp pump inserted for acute myocardial infarction (ami) with shock. the patient had access site bleeding and as a result blood transfusion was administered. | [5 empty] | We received information, that there was no blood transfusion administered. | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36949 | 1220648-2021-00870 | 11819771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11819771&pc=OZD | 04/17/2021 | 05/13/2021 |  | 05/13/2021 |  |  | 04/17/2021 | 05/13/2021 |  |  | 10/31/2022 |  | 05/13/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2021013216 |  | 0048-0014 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Positioning Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/17/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp smart assist pump. during support, the patient became "agitated" resulting in the pump falling out of the ventricle. during an attempted removal of the pump, it became kinked and a dissection formed in the patient's aorta. the patient was taken to the operating room for dissection repair and pump removal. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36950 | 1220648-2021-00868 | 11819772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11819772&pc=OZD | 04/20/2021 | 05/13/2021 |  | 05/13/2021 |  |  | 04/20/2021 | 05/13/2021 |  |  | 02/28/2023 |  | 05/13/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022055277 |  | 0048-0045 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 04/20/2021 | * | Invalid/Incomplete |  |  | A patient was admitted for a high risk coronary angioplasty. the patient had previous had a bypass (cabg) procedure and now all grafts were occluded. an impella cp was placed for hemodynamic support for the angioplasty procedure. after the 14fr introducer was removed and the impella pump's repositioning sheath was placed, the team observed bleeding at the access site. the team had done all due diligence with tightening of perclose sutures, angle matching, and purse string sutures. the team made the decision to remove the impella pump and replace it. they also infuse 2 units of red blood cells. | [5 empty] | The cause of the access site bleeding could not be determined as the user facility did not return the product for analysis. the failure mode will be monitored and trended.
36951 | 1220648-2021-00869 | 11819789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11819789&pc=OZD | 04/21/2021 | 05/13/2021 |  | 05/13/2021 |  |  | 04/21/2021 | 05/13/2021 |  |  | 10/31/2022 |  | 05/13/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2021013770 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | USER FACILITY | [5 empty] | 61 YR | [5 empty] | Injury | Yes | No | Use of Device Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 04/21/2021 | * | Invalid/Incomplete |  |  | A patient was deteriorating in the icu after suffering from a stemi. the patient was to have a high risk angioplasty to treat the coronary artery that had caused the myocardial infarction. the pump supported the patient for 5 days when the team made the decision to transfer the patient to a tertiary medical center for escalation of care. at the tertiary center the impella was explanted due to concerns of hemolysis. when explanted the pump had biomaterial adhered to it. the team made the choice to place an iabp and vv ecmo rather than another impella pump. | [5 empty] | The user facility returned the pump product for analysis and testing. the data logs were not returned. upon visual inspection and testing biomaterial was seen to be adhered to the pump. the root cause for the hemolysis is the failure to properly position the pump. this was determined via clinical case review. the failure mode will be monitored and trended.
36952 | 1220648-2021-00871 | 11819813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11819813&pc=OZD | 04/20/2021 | 07/09/2021 |  | 05/13/2021 |  |  | 04/20/2021 | 05/13/2021 |  |  | 02/28/2023 |  | 05/13/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022049682 |  | 0048-0045 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 51 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 04/20/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock. after insertion of the 14 fr introducer into the right femoral artery, the patient exhibited ischemia, even after removal of the introducer and advancement of the impella repositioning sheath. lower limb blood feeding was done as treatment via an antegrade sheath. during support, hemolysis was reported via plasma-free hemoglobin testing results of 300 mg/dl. as treatment, the pump was removed and an impella 5.5 was inserted into the patient's right axillary artery. | [5 empty] | The impella cp optical pump was returned, however, the introducer was not returned. simulated use testing was performed for hemolysis and passed without issue. we are unable to determine the cause of the hemolysis as we could not reproduce the reported issue. we also cannot determine the cause of the ischemia as the introducer was not received. | The impella cp optical was returned and an investigation of device is underway. a supplemental mdr will be filed at the completion of our analysis.
36953 | 1220648-2021-00863 | 11805720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11805720&pc=OZD | 04/15/2021 | 05/07/2021 |  | 05/11/2021 |  |  | 04/15/2021 | 05/11/2021 |  |  | 03/31/2019 |  | 05/11/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, US |  |  | ABIOMED INC. |  |  | US |  | 1287481 |  | 0048-0003 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Hospitalization | Required Intervention | [26 empty] | Initial submission |  |  |  |  | 04/15/2021 | * | Invalid/Incomplete |  |  | There was an impella cp utilized for a high risk coronary angioplasty in 2017. the patient was a part of the (b)(6) registry. the access was a transcaval one via the right femoral vein. the (b)(6) team reviewed the case in 2021 and found the patient had expired due to a bleed that was not resolved at the transcaval access. the transcaval access is not an ifu approved/indicated access site. the abiomed representative in 2017 was unaware of the adverse event. the representative had noted the access site was successfully closed via an amplatzer device to the aorta and two perclose to the femoral vein. of note, the amplatzer is not approved for transcaval use. | [5 empty] | The root cause of the access site bleed and subsequent death can not tied definitively to the use of impella. all product was discarded in 2017 at the time of registry enrollment and pump use. no product investigation is possible. the most likely cause is the use of the amplatzer device as a closure to the transcaval access. the device is not indicated for this use. in addition, the transcaval access of the cp is not indicated. the failure mode will be monitored and trended.
36954 | 1220648-2021-00862 | 11805732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11805732&pc=OZD | 04/14/2021 | 05/07/2021 |  | 05/11/2021 |  |  | 04/14/2021 | 05/11/2021 |  |  | 01/31/2023 |  | 05/11/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022040210 |  | 0048-0045 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | USER F |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | ECMO | [22 empty] | Initial submission |  |  |  |  | 04/14/2021 | * | Invalid/Incomplete |  |  | The customer placed an impella cp in a pccs patient via the 10mm graft to the axillary artery. when the patient was unable to come off-pump during the coronary artery bypass and mitral valve replacement surgery, the team placed va ecmo and inserted the impella pump via the femoral artery after removal from the previous graft site. the patient was transferred to the tertiary center for further monitoring and care. at the tertiary center the patient's limb was noted to be ischemic. the team placed a perfusion catheter. when this failed to restore the limb, the team performed a fasciotomy. the impella was explanted after 26 hours of support. | [5 empty] | There was limb ischemia when the impella cp was placed via the femoral artery. the patient was on high doses of pressors at the time of impella support. the cause of the ischemia was most likely the condition and size of the native vessel. the failure mode will be monitored and trended.
36955 | 1220648-2021-00864 | 11805733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11805733&pc=OZD | 04/10/2021 | 05/07/2021 |  | 05/11/2021 |  |  | 04/13/2021 | 05/11/2021 |  |  | 10/31/2022 |  | 05/11/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021020010 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 89 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Myocardial Contusion | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/13/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6)-year-old (b)(6) female patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021, explant date was not provided. it was reported that the insertion process was difficult due to vascular meandering and severe calcification. va-ecmo had to be introduced due to deterioration of hemodynamics. after left ventricular angiography, additional treatment was discontinued due to worsening cardiac rupture. it is unknown whether the exacerbation of myocardial rupture was due to impella insertion or accidental. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36956 | 1220648-2021-00866 | 11807965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11807965&pc=OZD | 04/17/2021 | 05/11/2021 |  | 05/11/2021 |  |  | 04/17/2021 | 05/11/2021 |  |  | 12/31/2022 |  | 05/11/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022022204 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Rupture | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/17/2021 | * | Invalid/Incomplete |  |  | The complainant reported a 76 white female patient had impella cp pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). the pump was entangled in the mitral chordae, the papillary muscle was ruptured resulting in surgical removal of impella and mitral valve replacement. | [5 empty] | Device not returned by customer. the impella cp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
36957 | 1220648-2021-00865 | 11807975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11807975&pc=OZD | 04/19/2021 | 05/11/2021 |  | 05/11/2021 |  |  | 04/19/2021 | 05/11/2021 |  |  | 08/31/2022 |  | 05/11/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1500625 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/19/2021 | * | Invalid/Incomplete |  |  | The complainant reported a 76 white female patient had impella cp pump inserted in the left femoral artery (lfa). the patient had left leg limb ischemia, the pump was surgically removed by a vascular surgeon via cut down/repair and a new device was inserted in the right axillary. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36958 | 1220648-2021-00867 | 11809043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11809043&pc=OZD | 04/01/2021 | 05/11/2021 |  | 05/11/2021 |  |  | 04/12/2021 | 05/11/2021 |  |  | 12/31/2022 |  | 05/11/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022035770 |  | 0048-0045 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 87 YR | [5 empty] | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission |  |  |  |  | 04/12/2021 | * | Invalid/Incomplete |  |  | The complaint reported an (b)(6) male patient presenting for high risk percutaneous coronary intervention. during the procedure, a j-wire accidentally engaged the impella through the outflow cage and the pump stopped. the pump restarted upon removing the j-wire, however, with low flows. upon receipt of the device on (b)(6) 2021, a broken impeller blade was discovered. | [5 empty] | The device was returned and a visual inspection identified a broken impeller blade. data log review confirmed a pump stop had occurred during use of the device. based on the information provided from the customer, it was determined that the pump stop and subsequent impeller fracture was caused by interaction with the j-wire device.
36959 | MW5101291 | 11811941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11811941&pc=OZD | 04/19/2021 | 05/07/2021 |  | 05/11/2021 |  |  |  | 05/11/2021 |  |  |  |  | 05/11/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA CP WITH SMARTASSIST |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 20220039120 |  |  |  | No | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 58 YR |  |  | 126 |  |  | Malfunction | No | Yes | High impedance | High Sensing Threshold | Connection Problem | Unintended Movement |  |  |  | Insufficient Information | [12 empty] | Other | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Impella heart pump catheter malfunction during initial placement of the device. impella monitor detected and prompt error message of impella device stated "high impedance" "high threshold". impella catheter removed and replaced with new working impella. upon inspection by rep., who came in after the fact stated they believed it was possible that the orange wire connected to the disposable 'flag' did not all come out when pulled to discard. they believed that upon preparing the impella device, it was taken out of the package and the small paper flag connected to the pump that shows the user which small hole to use for the tubing to be fed through the pump that is attached to an orange wire which is supposed to be pulled off. upon pulling the orange wire, some of the orange wire may have possibly been left behind with out knowledge of the users. the wire caused the pump to stall and not work after it entered the patient. the manager stated this flag/wire is new to the packaged impella device. the impella device was totally removed and another was immediately placed in the patient without incident. this product was given to the rep after the fact. fda safety report id# (b)(4).
36960 | 1220648-2021-00860 | 11783360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11783360&pc=OZD | 04/09/2021 | 11/26/2024 |  | 05/06/2021 |  |  | 04/09/2021 | 05/06/2021 |  |  | 06/30/2022 |  | 05/06/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1493297 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup |  |  | 04/09/2021 |  |  |  |  | Additional information was provided indicating that the patient also developed an infection at the impella access site after the device was explanted. the patient was administered antibacterial agents to treat the infection. | The complainant reported a (b)(6) male patient had impella 5.0 pump inserted in the right axillary for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported hemolysis occurred, which was presumed to be caused by the improper position of the pump catheter. as a result, continuous hemofiltration/hemodiafiltration (chdf) was added to protect the patient's kidneys. it was noted that hemolysis did settled down. no further information was provided. | Additional information was provided indicating that the patient also developed bleeding including hematoma after removal which needed blood transfusion. |  |  |  | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | This report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr (b)(4) was provided in sections b2, b5, d1, d3, d4, d6, e1, g2, and h6.
36961 | 1220648-2021-00861 | 11783370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11783370&pc=OZD | 04/09/2021 | 05/06/2021 |  | 05/06/2021 |  |  | 04/12/2021 | 05/06/2021 |  |  | 10/31/2022 |  | 05/06/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022028391 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [41 empty] | Initial submission |  |  |  |  | 04/12/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient was implanted on (b)(6) 2021 and explanted on (b)(6) 2021. it was reported that the impella device was explanted due to bleeding at the access site. the device was replaced by aibp. it was noted that there was patient movement. additionally, blood products were administered. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36962 | MW5101146 | 11776322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11776322&pc=OZD | 04/18/2021 | 04/30/2021 |  | 05/04/2021 |  |  |  | 05/04/2021 |  |  | 02/28/2023 |  | 05/04/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA HEART PUMP |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 67 YR |  |  | 86 |  |  | Death | No | Yes | Pumping Stopped | [6 empty] | Insufficient Information | [12 empty] | Death | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Alarm went off on impella device due to pump failure. issue was addressed immediately, provider notified, and impella representative contacted who recommended restart. after multiple attempts and console changes, pump could not be restarted. patient subsequently expired, pump was pulled by representative and provider, and device sequestered. fda safety report id#: (b)(4).
36963 | 1220648-2021-00858 | 11761661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11761661&pc=OZD | 04/08/2021 | 05/15/2021 |  | 05/03/2021 |  |  | 04/08/2021 | 05/03/2021 |  |  | 04/30/2022 |  | 05/03/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1481044 |  | 0550-0008 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | LITERA |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 04/08/2021 | * | Invalid/Incomplete |  |  | The patient was in the operating suite (in 2020) for a cardiopulmonary bypass. the team had 4 vessels bypassed and due to a need for hemodynamic support, the impella 5.5 was placed for support via direct access to the aorta. the impella supported the patient for 8 days successfully and then was weaned and explanted. the physician later published upon this patient and noted that at explant the patient went into cardiogenic shock attributed to the damage aortic valve leaflet. the leaflet injury was treated by surgery. the aortic valve was replaced. the publication attributes this harm to the use of impella. | [5 empty] | The publication states that the aortic valve injury occurred during impella support but it was unknown if the pump caused the issue. the product was no longer available for analysis or testing, and nor were the data logs. the pump was inserted without a guidewire via the direct route. the presumed root cause of the aortic valve damage was excessive force during implant. the damaged valve was replaced. the failure mode will be monitored and trended. | The impella support was in (b)(6) 2020. all product was subsequently disposed of. when the literature piece was discovered, the investigation began into the event. the root cause of the aortic valve damage has not yet been determined. at conclusion of the investigation, a final report will be filed.
36964 | 1220648-2021-00859 | 11765771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11765771&pc=OZD | 04/06/2021 | 05/03/2021 |  | 05/03/2021 |  |  | 04/06/2021 | 05/03/2021 |  |  | 12/31/2022 |  | 05/03/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022023443 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 23 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 04/06/2021 | * | Invalid/Incomplete |  |  | A younger patient with history of being covid+ was admitted with cardiomyopathy and other cardiac comorbidities of chf, hypertension, myocarditis, asthma, and left ventricular thrombosis. when upon admission his ef was noted to be 15% the impella cp was placed for hemodynamic support. the cp was to support him while decisions made for either heart transplant or lvad placement. the patient began to experience signs of hemolysis while on impella support. the urine color was changed and the team attempted to rectify the hemolysis with pump repositioning. when measures failed, the team made choice to remove the impella and replace it with another pump. the second cp pump was placed and his support continued and hemolysis/hematuria was cleared. | [5 empty] | The medical facility returned the impella catheter for analysis. the analysis revealed that the root cause of the hemolysis was ingestion of biomaterial. the biomaterial was wrapped around the impeller and therefore cause of the hemolysis. the failure mode will be monitored and trended.
36965 | 1220648-2021-00857 | 11757205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11757205&pc=OZD | 04/07/2021 | 04/30/2021 |  | 04/30/2021 |  |  | 04/07/2021 | 04/30/2021 |  |  | 09/30/2022 |  | 04/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1502523 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | HEALTH PROFESSIONAL |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/07/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old male patient had impella cp pump inserted in the right femoral for high-risk percutaneous coronary interventions (hrpci). it was reported that during impella support bleeding from the access site was observed and was noted to be severe. as a result, considerable amount of red blood cells, fresh frozen plasma and albumin were administered. it was noted that the long sheath was peeled away without being pulled out of the body during the peel away. the bleeding was caused because the blood vessel became damage caused by peeling away without completely removing the 25 cm long sheath (peel away) from the body. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36966 | 1220648-2021-00856 | 11757227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11757227&pc=OZD | 06/20/2020 | 04/30/2021 |  | 04/30/2021 |  |  | 04/06/2021 | 04/30/2021 |  |  | 12/31/2021 |  | 04/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1461608 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/06/2021 | * | Invalid/Incomplete |  |  | The complainant reported a patient whose age and gender are unknown had impella cp pump inserted in an unknown insertion site for post-cardiotomy cardiogenic shock (pccs). it was reported the patient developed lower limb ischemia during impella support. as a result, antegrade puncture was performed, which improved the issue no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36967 | 1220648-2021-00853 | 11746979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11746979&pc=PYX | 04/07/2021 | 04/29/2021 |  | 04/29/2021 |  |  | 04/07/2021 | 04/29/2021 |  |  | 12/31/2022 |  | 04/29/2021 |  | TEMPORARY NONROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED INC. |  |  | US |  | 2022027402 |  | 004334 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 51 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [41 empty] | Initial submission |  |  |  |  | 04/07/2021 | * | Invalid/Incomplete |  |  | A patient was admitted in cardiogenic shock and had and impella rp placed for right heart hemodynamic support. after over one day of support the team made the choice to explant the pump due to concerns of hemolysis. the patient had labs hemolyzing and despite repositioning of the pump, the hemolysis concerns were not resolved. the team infused 4 units of blood products, explanted the rp, and placed the patient on ecmo support. | [5 empty] | The medical center did not return the impella rp pump for analysis and bench top hemolysis testing. a root cause can not be determined at this time with information known and shared. should more information be shared that leads to a definitive root cause, a final report will be filed.
36968 | 1220648-2021-00854 | 11749848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11749848&pc=OZD | 04/07/2021 | 04/29/2021 |  | 04/29/2021 |  |  | 04/07/2021 | 04/29/2021 |  |  | 11/30/2022 |  | 04/29/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2021009526 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Obstruction of Flow | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 04/07/2021 | * | Invalid/Incomplete |  |  | A female patient was admitted in cardiogenic shock and suffering from an acute myocardial infarction. the medical team placed an impella cp for hemodynamic support. after 10 hours there was access site bleeding and the team noted urine color change. they explanted the pump for the concerns of hemolysis and the access site bleeding. at the explant the team observed signs of a cerebral event. due to concerns of a stroke, the patient was taken to the interventional radiology suite for clot removal from the carotid artery. the physician attributed the clot to the use of impella and it's explant. | [5 empty] | The investigation of the pump is not possible. the product was discarded and as such, no analysis or bench top testing for the presumed hemolysis is possible. without further clinical details, the root cause of the hemolysis, access site bleeding, and cva after pump explant is not possible. should more details be shared that lead to definitive root cause, a final report will be filed.
36969 | 1220648-2021-00855 | 11749851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11749851&pc=OZD | 04/06/2021 | 04/29/2021 |  | 04/29/2021 |  |  | 04/06/2021 | 04/29/2021 |  |  | 01/31/2023 |  | 04/29/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022036312 |  | 0048-0045 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/06/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting for hemodynamic support, pre-operative cabg. the patient was suspected to have hemolysis as evident by dark red urine output. it's unclear if and plasma-free hemoglobin testing was completed to confirm hemolysis, however. as treatment, the pump was premature removed. | [5 empty] | The impella device was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36970 | 1220648-2021-00852 | 11739147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11739147&pc=OZD | 04/04/2021 | 07/09/2021 |  | 04/28/2021 |  |  | 04/04/2021 | 04/28/2021 |  |  | 10/31/2022 |  | 04/28/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2021017459 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Malfunction | Yes | Yes | No Flow | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup | Followup |  |  | 04/04/2021 | * | Invalid/Incomplete |  |  | A patient was admitted with pneumonia and shortness of breath. when in the cardiac cath lab the study found him to be suffering from cardiomyopathy and an ef of 20%. the medical team placed an impella 5.5 for hemodynamic support while pending decisions made on heart transplant workup. the impella 5.5 had a pump stop after 4 days. the team explanted the pump and observed biomaterial to be adhered to the pump. | [5 empty] | The final pump and data log analysis has completed. the data logs show purge pressure changes indicative of a kink in the purge line. the root cause of the pump stoppage was due to the kink. the purge sidearm was also observed to be damaged due to excessive force by the operator. the failure mode will be monitored and trended. | Since the initial report was filed, the investigation has completed. the pump and data logs were analyzed. the purge sidearm of the pump was found to be broken, due to excessive force. this breakage allowed for blood ingress into the pump. the ingress caused the pump stop. the failure mode will be monitored and trended. | The investigation into the impella 5.5 pump stoppage is on-going at this time. upon completion of the investigation, a final report will be filed.
36971 | MW5101070 | 11744572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11744572&pc=PBL | 04/27/2021 | 04/27/2021 |  | 04/28/2021 |  |  |  | 04/28/2021 |  |  | 01/31/2023 |  | 04/28/2021 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | PBL |  |  | 0048-0003 |  |  | ABIOMED EUROPE GMBH |  |  |  |  | 306176 |  |  |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 65 YR | [5 empty] | Injury | No | Yes | Material Twisted/Bent | Therapeutic or Diagnostic Output Failure | [5 empty] | Appropriate Term / Code Not Available | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Device was kinked in the body, worked on suction and failed to support the left ventricle during the procedure. fda safety report id # (b)(4).
36972 | 11723687 | 11723687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11723687&pc=OZD | 03/26/2021 | 04/09/2021 | 04/09/2021 | 04/26/2021 | 04/09/2021 | 04/26/2021 |  | 04/26/2021 |  |  |  |  | 04/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 21462313-34-36 |  | 0550-0008 |  | No answer provided | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 20440 DA |  |  | 87 |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Appropriate Term / Code Not Available | [41 empty] | Initial submission | Followup | Followup |  |  |  | Y | HOSPITAL |  |  | A potential heart transplant patient being maintained with the impella 5.5 left ventricular assist device (lvad). rn had found a clear collection of fluid on patient abdomen from an unknown source. further investigation lead to the discovery that the fluid had been leaking from the impella specifically from the "red impella plug" component proximal to the patient. the fluid on abdomen likely was the heparinized impella purge solution. the cardiothoracic intensive care unit (cticu) team, ventricular assist device (vad) team, and impella device representative had all been contacted regarding this finding. the impella device representative, explained how the leaking fluid would eventually lead to the total failure of the pump. this finding significantly alters the plan for patient care as the focus shifted to evaluating patient tolerance of imminent device removal. furthermore, patient will likely go to or and other therapies will have to be considered. the most concerning detail of this event is how there had been tegaderm (transparent film dressing) wrapped around the "red impella plug" component which was partially containing the leak from the device. it is not known who had initially placed the tegaderm but the leak from the device has likely been occurring for an indiscriminate amount of time. | A potential heart transplant patient being maintained with the impella 5.5 left ventricular assist device (lvad). rn had found a clear collection of fluid on patient abdomen from an unknown source. further investigation lead to the discovery that the fluid had been leaking from the impella specifically from the "red impella plug" component proximal to the patient. the fluid on abdomen likely was the heparinized impella purge solution. the cardiothoracic intensive care unit (cticu) team, ventricular assist device (vad) team, and impella device representative had all been contacted regarding this finding. the impella device representative, explained how the leaking fluid would eventually lead to the total failure of the pump. this finding significantly alters the plan for patient care as the focus shifted to evaluating patient tolerance of imminent device removal. furthermore, patient will likely go to or and other therapies will have to be considered. the most concerning detail of this event is how there had been tegaderm (transparent film dressing) wrapped around the "red impella plug" component which was partially containing the leak from the device. it is not known who had initially placed the tegaderm but the leak from the device has likely been occurring for an indiscriminate amount of time. | A potential heart transplant patient being maintained with the impella 5.5 left ventricular assist device (lvad). rn had found a clear collection of fluid on patient abdomen from an unknown source. further investigation lead to the discovery that the fluid had been leaking from the impella specifically from the "red impella plug" component proximal to the patient. the fluid on abdomen likely was the heparinized impella purge solution. the cardiothoracic intensive care unit (cticu) team, ventricular assist device (vad) team, and impella device representative had all been contacted regarding this finding. the impella device representative, explained how the leaking fluid would eventually lead to the total failure of the pump. this finding significantly alters the plan for patient care as the focus shifted to evaluating patient tolerance of imminent device removal. furthermore, patient will likely go to or and other therapies will have to be considered. the most concerning detail of this event is how there had been tegaderm (transparent film dressing) wrapped around the "red impella plug" component which was partially containing the leak from the device. it is not known who had initially placed the tegaderm but the leak from the device has likely been occurring for an indiscriminate amount of time.
36973 | 1220648-2021-00849 | 11725292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11725292&pc=OZD | 04/01/2021 | 04/26/2021 |  | 04/26/2021 |  |  | 04/01/2021 | 04/26/2021 |  |  | 10/31/2022 |  | 04/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2022024719 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/01/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient had impella 5.0 pump inserted. the patient had aortic valve damage. | [5 empty] | Device discarded by customer. the impella 5.0 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36974 | 1220648-2021-00850 | 11725293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11725293&pc=OZD | 04/03/2021 | 04/26/2021 |  | 04/26/2021 |  |  | 04/03/2021 | 04/26/2021 |  |  | 01/31/2023 |  | 04/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022042968 |  | 0048-0045 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/03/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old asian female patient had impella cp pump inserted in the left femoral artery (lfa) for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had hemolysis determined from hematuria and elevated lactate dehydrogenase (ldh) and as a result four (4) units of packed red blood cells (prbcs), a total of 4 units of platelets and fresh frozen plasma (ffp) was administered. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36975 | 1220648-2021-00851 | 11725305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11725305&pc=OZD | 03/29/2021 | 04/26/2021 |  | 04/26/2021 |  |  | 03/29/2021 | 04/26/2021 |  |  | 12/31/2022 |  | 04/26/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022027596 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 03/29/2021 | * | Invalid/Incomplete |  |  | A patient was admitted in cardiogenic shock and had an impella cp placed via the left femoral artery. during the two days of support the pump came out of optimal left ventricle position and was replaced. when the second pump was being placed the team observed reduced blood flow and the limb was ischemic. the patient was taken to the operating suite for thrombectomy and fasciotomy. afterwards, the limb condition improved. of note the patient was not therapeutically anticoagulated while on impella support due to bleeding from another site. | [5 empty] | The medical center did not return the product for analysis. imaging was not provided for analysis. the team did note that due to unrelated bleeding, the patient was not therapeutically anticoagulated. the cause of the ischemia is patient condition. the failure mode will be monitored and trended.
36976 | 1220648-2021-00846 | 11699414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11699414&pc=OZD | 02/14/2021 | 04/21/2021 |  | 04/21/2021 |  |  | 03/29/2021 | 04/21/2021 |  |  | 02/28/2021 |  | 04/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1405009 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/29/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella 2.5 pump inserted in an unknown insertion site for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that the patient developed an access site bleed. as a result, blood products were administered, amount was not provided. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36977 | 1220648-2021-00847 | 11704444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11704444&pc=OZD | 02/10/2021 | 04/21/2021 |  | 04/21/2021 |  |  | 03/30/2021 | 04/21/2021 |  |  | 08/31/2022 |  | 04/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1496671 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/30/2021 | * | Invalid/Incomplete |  |  | The complainant reported an asian male patient had impella cp pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had hemolysis and as a result blood transfusion was administered. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36978 | 1220648-2021-00848 | 11704448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11704448&pc=OZD | 03/24/2021 | 09/15/2024 |  | 04/21/2021 |  |  | 03/24/2021 | 04/21/2021 |  |  |  |  | 04/21/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, US |  |  | ABIOMED INC. |  |  | US |  |  |  | 0048-0003 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility |  |  |  |  | 60 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | Pericardial Effusion | [11 empty] | Hospitalization | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/24/2021 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | In 2017 a patient was enrolled in an abiomed registry. the patient had an impella cp placed for hemodynamic support of an acute myocardial infarction and cardiogenic shock. the patient's condition continued to deteriorate and she was observed in an echo to have a tamponade. there was a perforation to the left ventricle. the patient was taken to the operating suite for repair. unfortunately, the patient expired. the perforation was discussed with the abiomed field representative afterwards and was told the pump was repositioned many times during cpr and the patient was in very poor condition prior to impella implant. no complaint was made against the impella pump. years later, in march of 2021, the committee reviewed the case and the team determined the perforation may have been caused/contributed to by the use of impella. |  |  |  |  | Since the initial report was filed, the investigation into the left ventricle perforation has concluded. the cause of the perforation was most likely the improper pump positioning. the pump was positioned during cpr being performed. the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2021-00848 was provided in sections b2, b5, d6a, d6b, h1, and h6. | The investigation into the left ventricular perforation has not completed. the product was discarded. further clinical details on how the pump-to-patient interaction occurred are being sought. should more details come that lead to root cause of the perforation, a final report will be filed.
36979 | 1220648-2021-00840 | 11692451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11692451&pc=OZD | 06/16/2020 | 04/20/2021 |  | 04/20/2021 |  |  | 03/23/2021 | 04/20/2021 |  |  | 08/31/2021 |  | 04/20/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1436348 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/23/2021 | * | Invalid/Incomplete |  |  | The complainant reported a patient whose age and gender were not provided had an impella cp pump inserted in an unknown insertion site an unknown indication. it was reported the patient developed impella access site bleed. as a result, blood products were administered, with amount administered unknown. patient's act values were 160 per seconds. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36980 | 1220648-2021-00842 | 11692481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11692481&pc=OZD | 02/17/2021 | 04/20/2021 |  | 04/20/2021 |  |  | 03/25/2021 | 04/20/2021 |  |  | 07/31/2022 |  | 04/20/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1490580 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/25/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) male patient had impella cp pump inserted in the right femoral for high-risk percutaneous coronary interventions (hrpci) it was reported the patient developed hemolysis during impella support. a s a result, blood products were administered, amount unknown and p-level was lowered. hemolysis resolved. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36981 | 1220648-2021-00841 | 11692486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11692486&pc=OZD | 10/01/2020 | 04/20/2021 |  | 04/20/2021 |  |  | 03/23/2021 | 04/20/2021 |  |  | 04/30/2022 |  | 04/20/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1488338 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/23/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed impella access site bleed. as a result, blood products were administered, amount unknown and 1 stitch suture was added. patient's act's were 240 per seconds and then switched to 150 per seconds no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36982 | 1220648-2021-00843 | 11692491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11692491&pc=OZD | 05/12/2020 | 09/16/2024 |  | 04/20/2021 |  |  | 03/30/2021 | 04/20/2021 |  |  | 08/31/2020 |  | 04/20/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1364775 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 68 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 03/30/2021 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a (b)(6) year-old (b)(6) male patient had impella 5.0 pump inserted in the right axillary for post-cardiotomy cardiogenic shock (pccs). the complainant reported the patient developed an ischemic stroke during impella support. ultimately, the patient expired. date of death is unknown. no further information was provided. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2021-00843 was provided. | The impella 5.0 was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36983 | 1220648-2021-00844 | 11692492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11692492&pc=OZD | 03/25/2021 | 04/30/2021 |  | 04/20/2021 |  |  | 03/25/2021 | 04/20/2021 |  |  | 01/31/2023 |  | 04/20/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED INC. |  |  | US |  | 2022039046 |  | 0048-0045 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 03/25/2021 | * | Invalid/Incomplete |  |  | A patient was admitted for a high risk angioplasty of the coronary artery. he was known to have left main coronary disease and so an impella cp was placed for hemodynamic support during the angioplasty. the team accessed via the right femoral artery and proceeded with the case. at the case end the team attempted to close the arteriotomy, when the closure failed and the patient suffered an access site bleed. the team treated this with a balloon tamponade, blood replacement, and a viabahn stent. | [5 empty] | The investigation has completed since the initial report was filed. the returned introducer and pump product was analyzed, as were the data logs. there was no damage or abnormalities observed in the returned product. the bleeding was due to the high anticoagulation and not due to the impella product. the act was noted to be supra-therapeutic after the patient was placed on bivalirudin. the failure mode will be monitored and trended. | The investigation into the patient injury is on-going. at the conclusion of the investigation, a final report will be filed.
36984 | 1220648-2021-00845 | 11692506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11692506&pc=OZD | 03/22/2021 | 05/11/2021 |  | 04/20/2021 |  |  | 03/22/2021 | 04/20/2021 |  |  | 02/28/2022 |  | 04/20/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1469258 |  | 0550-0008 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Malfunction | No | Yes | Unexpected Shutdown | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | ECMO | [21 empty] | Initial submission | Followup |  |  |  | 03/22/2021 | * | Invalid/Incomplete |  |  | A patient was admitted to the operating suite (or) for mitral/tricuspid valve repair. while in the or she suffered from ventricular fibrillation storm and was shocked multiple times by her icd. due to this deteriorating condition she had an impella 5.5 placed via a graft to axillary artery. when her condition still was poor, the team placed ecmo for right ventricular failure. the patient stabilized and was released from the or to the icu for continued monitoring. after 4 days of support she returned to the or for washout and while on the or table the impella pump stopped. the pump was able to be restarted. the pump's purge pressure parameters were not optimal, and so tpa was added to the purge. the pump remained on for support and after 15 days of total impella support the pump was explanted when care was withdrawn and she expired. | [5 empty] | The medical center did not return the impella pump product. the logs however were returned and they were analyzed. the analysis determined that the root cause was biomaterial ingestion. this was determined after analysis of the purge pressure, flows, and effect the infusion of the tpa made to the pump performance. the failure mode will be monitored and trended. | The investigation is on-going at this time into the product malfunction. upon completion of the investigation, a final report will be filed.
36985 | 1220648-2021-00836 | 11685432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11685432&pc=OZD | 03/30/2021 | 06/01/2021 |  | 04/19/2021 |  |  | 03/30/2021 | 04/19/2021 |  |  | 12/31/2022 |  | 04/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022025802 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 03/30/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella cp pump inserted in the left femoral for acute myocardial infarction (ami) with shock. the repositioning sheath advanced into the artery, the patient was bleeding around the insertion site that was uncontrollable. the pump was removed and patient was taken to operating room (or) for surgical closure. | [5 empty] | Product evaluation confirmed that the sheath was not peeled along the scorelines and had been cut. imaging of the valve revealed an uneven tear along the scoreline. | The impella cp pump was not returned by the customer, only the introducer was returned, and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation if the information is related to this event.
36986 | 1220648-2021-00837 | 11685465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11685465&pc=OZD | 02/17/2020 | 04/19/2021 |  | 04/19/2021 |  |  | 03/23/2021 | 04/19/2021 |  |  | 09/30/2021 |  | 04/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1447303 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Paralysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/31/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient was diagnosed with femoral nerve paralysis due to loss of sensation and muscle weakness in the left femoral nerve area after the impella was removed from the left femoral artery. the symptoms disappeared and the patient was discharged home. relationship with impella is unknown. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36987 | 1220648-2021-00839 | 11685471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11685471&pc=OZD | 03/24/2021 | 04/19/2021 |  | 04/19/2021 |  |  | 03/24/2021 | 04/19/2021 |  |  | 02/28/2023 |  | 04/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022047459 |  | 0048-0045 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/24/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella cp pump inserted in the right femoral artery (rfa) for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had distal leg thrombosis that resulted in ischemia and the pump was removed. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
36988 | 1220648-2021-00838 | 11685478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11685478&pc=OZD | 03/23/2021 | 04/19/2021 |  | 04/19/2021 |  |  | 03/23/2021 | 04/19/2021 |  |  | 01/31/2022 |  | 04/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 1459786 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 0 YR | [5 empty] | Death | Yes | No | Insufficient Information | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 03/23/2021 | * | Invalid/Incomplete |  |  | The complainant reported a patient of unknown age and gender presenting with acute myocardial infarction and cardiogenic shock. the impella cp was selected for hemodynamic support and inserted independent of abiomed on-site support. the patient was transferred to another hospital where they expired due to hemorrhagic shock. an autopsy is being performed to identify what, if any, relationship the impella had with the patient's injury and/or outcome. | [5 empty] | The impella cp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. should the device or any additional information be obtained, a supplemental report will be filed.
36989 | 1220648-2021-00835 | 11680490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11680490&pc=OZD | 03/25/2021 | 04/16/2021 |  | 04/16/2021 |  |  | 03/25/2021 | 04/16/2021 |  |  | 12/31/2022 |  | 04/16/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022027535 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 92 YR | [5 empty] | Injury | Yes | No | Obstruction of Flow | Patient-Device Incompatibility | [5 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 03/25/2021 | * | Invalid/Incomplete |  |  | An elderly patient was admitted in cardiogenic shock and had a coronary angioplasty in the cardiac catheterization lab. the patient had a guidewire perforation and his condition deteriorated and so the impella cp was placed for hemodynamic support. unfortunately the elderly patient suffered limb ischemia from the impella sheath insertion. to treat the limb ischemia the patient required thrombectomy. in retrospect, the patient was known to have peripheral arterial disease. | [5 empty] | The medical facility did not return the product for analysis and investigation. without further data shared regarding degree of underlying peripheral arterial disease and imaging, the root cause of the ischemia is believed to be patient condition and underlying size and condition of the patient's vasculature. the failure mode of limb ischemia will be monitored and trended.
36990 | 1220648-2021-00830 | 11663568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11663568&pc=OZD | 03/03/2021 | 04/14/2021 |  | 04/14/2021 |  |  | 03/17/2021 | 04/14/2021 |  |  | 11/30/2022 |  | 04/14/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021009690 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 19 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/17/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female patient enduring flash pulmonary edema and aortic insufficiency that ultimately required a replacement valve. according to the surgeon, mild inflammation was found on the valve that was believed to have been caused by impella use. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36991 | 1220648-2021-00831 | 11663571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11663571&pc=OZD | 03/19/2021 | 06/10/2021 |  | 04/14/2021 |  |  | 03/20/2021 | 04/14/2021 |  |  | 05/31/2022 |  | 04/14/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1488045 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Material Twisted/Bent | [6 empty] | Renal Impairment | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 03/20/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old patient endure a delay in support resulting from a kinked cannula causing lower than expected pump flow. reduced cardiac output and reduced mean arterial pressure were exhibited. the impella was removed and replaced, but due to reduced/delayed support, dialysis was required. the physician attributed the impella as a contributing factor in the patient's need for dialysis. | [5 empty] | The impella cp optical was returned for investigation. upon completion of our analysis, a supplemental mdr will be filed. | The device was returned and a kink in the cannula was confirmed. data logs also confirmed numerous suction alarms with a decrease in pump flow. simulated use testing found the pump to run within specification. evidence suggests the cannula kink was caused by the patient's condition (small ventricle).
36992 | 1220648-2021-00832 | 11663575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11663575&pc=OZD | 03/19/2021 | 04/14/2021 |  | 04/14/2021 |  |  | 03/19/2021 | 04/14/2021 |  |  | 07/31/2022 |  | 04/14/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED INC. |  |  | US |  | 1502237 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Device Dislodged or Dislocated | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | Cardiac Perforation | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 03/19/2021 | * | Invalid/Incomplete |  |  | A patient was admitted in cardiogenic shock suffering from an acute myocardial infarction. the team placed an impella 5.0 for support via access from the axillary artery and associated graft. the 5.0 remained on for a support of 7 days when the team made choice to perform bypass (cabg) surgery. the cabg was being performed and the heart was lifted for the surgical procedure. while able to observe the heart, the team noted that the impella had perforated the left ventricle (lv). the perforation was repaired while the patient was on the bypass pump. the physician noted that lifting of the heart may have allowed the impella to advance into and perforate the lv wall at the apex, which due to diseased and infarcted left anterior descending coronary artery, was friable tissue. after the impella was removed, the team performed thrombectomy to treat limb ischemia. the arm was thought to be ischemic due to clot from the impella axillary accessed limb. | [5 empty] | The medical center returned product and data logs for analysis. the cause of the ventricular perforation was the lifting of the heart in the surgical suite, which allowed for the impella pump to move too deep within the ventricle and cause the perforation. the perforation was at the area of the already friable tissue which was weakened. the limb ischemia was due to the patient's native condition, and associated thrombosis. the failure mode will be monitored and trended.
36993 | 11663606 | 11663606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11663606&pc=PBL | 02/20/2021 | 03/02/2021 | 03/30/2021 | 04/14/2021 | 03/30/2021 | 04/14/2021 |  | 04/14/2021 |  |  |  |  | 04/14/2021 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA LR 0.018 PTFE GUIDEWIRE |  |  | PBL |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  | 0048-0003 |  | 0048-0003 |  | No | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 32485 DA | [5 empty] | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | Thrombosis/Thrombus | [11 empty] | Other | [28 empty] | Initial submission | [5 empty] | Y | Invalid/Incomplete |  |  | Impella 14fr cp placement, right femoral artery. switched from d5w to d5w with heparin later. later, impella removal; found to have thrombus and dissection at site of insertion (right femoral) requiring thrombectomy and angioplasty.
36994 | 1220648-2021-00833 | 11668420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11668420&pc=OZD | 02/09/2021 | 04/14/2021 |  | 04/14/2021 |  |  | 03/22/2021 | 04/14/2021 |  |  | 09/30/2022 |  | 04/14/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022023192 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | HEALTH PROFESSIONAL |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/22/2021 | * | Invalid/Incomplete |  |  | 'the complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed hemolysis on the day of insertion. suction occurred due to volume loss and hemolysis was observed; however, it improved the next day by lowering the infusion load and auxiliary level. on (b)(6) 2021, the patient again showed signs of hemolysis. hemolysis was analyzed by urine color and lactate dehydrogenase levels. the patient was administered haptoglobin and hchdf was added. the physician stated that impella was likely to be pulled out due to patient anatomy. no further information was provided. | [5 empty] | The impella cp is expected to be returned. a supplemental mdr will be filed with the results of the investigation, upon completion.
36995 | 1220648-2021-00834 | 11668495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11668495&pc=OZD | 03/22/2021 | 04/14/2021 |  | 04/14/2021 |  |  | 03/22/2021 | 04/14/2021 |  |  | 09/30/2022 |  | 04/14/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021006101 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 41 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/22/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient's urine was wine red and hemolysis was suspected. as a result, red blood cells were administered, amount unknown. pcps were also rotating at a high flow rate, and the patient's general condition was poor. the physician stated that he was unable to determine if the adverse event was cause by impella. lower limb ischemia was also noted for which a reverse sheeth was inserted to treat the limb. no further information was provided. | [5 empty] | The impella cp was has not been returned. therefore an investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36996 | 1220648-2021-00828 | 11635063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11635063&pc=OZD | 03/15/2021 | 04/08/2021 |  | 04/08/2021 |  |  | 03/15/2021 | 04/08/2021 |  |  | 09/30/2022 |  | 04/08/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1503048 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 03/15/2021 | * | Invalid/Incomplete |  |  | A patient was admitted with an st elevation myocardial infarction (stemi). the hospital made the choice to transfer her to a tertiary medical center for escalation of care. the team made the choice to intervene upon her coronary artery that was occluded. when her condition deteriorated post intervention they placed and impella cp for support. unfortunately the access site for the impella developed a hematoma. the team attempted to treat it and reduce by placing pressure with a femstop. when this failed, the team determined that blood replacement was necessary. multiple units of blood products were infused, being both red blood cells and platelets. later the family made the choice to withdraw care and she expired after 2 days of impella support. | [5 empty] | The medical center discarded the impella product upon the family withdrawing care and the patient expiring. the definitive cause of the access site hematoma is not possible to be determined without product return. the cause was most likely patient management of the access site. the failure mode will be monitored and trended.
36997 | 1220648-2021-00829 | 11635071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11635071&pc=OZD | 03/15/2021 | 09/15/2024 |  | 04/08/2021 |  |  | 03/16/2021 | 04/08/2021 |  |  | 12/31/2022 |  | 04/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022029751 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 90 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 03/16/2021 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a (b)(6) female patient had impella cp pump inserted in the right femoral for high-risk percutaneous coronary interventions (hrpci). it was reported the patient developed bleeding at the impella access site. as a result, the patient was administered 600ml of albumin. it was noted that no sutures were added. the patient subsequently died of sepsis. no further information was provided. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2021-00829 was provided. | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
36998 | 1220648-2021-00825 | 11618330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11618330&pc=OZD | 03/09/2021 | 04/06/2021 |  | 04/06/2021 |  |  | 03/09/2021 | 04/06/2021 |  |  | 07/31/2022 |  | 04/06/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1488085 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | USER FACILITY |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [22 empty] | Initial submission |  |  |  |  | 03/09/2021 | * | Invalid/Incomplete |  |  | In (b)(6) 2020 an impella cp was placed in (b)(6) to support a patient admitted in cardiogenic shock. the cp was placed via the femoral artery and supported the patient for 8 days in combination with ecmo. after 8 days of successful support the pump was explanted. months later as part of a registry review the abiomed team was notified that during the support there were concerns for bleeding as well as limb ischemia. the blood loss allowed for hematoma development, which was treated by blood product infusion. the limb ischemia occurred in the impella accessed limb on day of explant. the access site had been sutured at explant. further details are unknown and not documented within the registry documentation. | [5 empty] | The complainant in (b)(6) did not return the product from the patient support in 2020. the registry data did not elaborate on the bleeding and limb ischemia. without the product and patient information regarding the hematoma and limb ischemia root cause could not be determined. in addition as the limb ischemia was reported after explant, the relationship of the ischemia to the use of impella was deemed unrelated. the failure modes will be monitored and trended.
36999 | 1220648-2021-00826 | 11618343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11618343&pc=OZD | 03/10/2021 | 04/06/2021 |  | 04/06/2021 |  |  | 03/10/2021 | 04/06/2021 |  |  | 01/31/2023 |  | 04/06/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022032057 |  | 0550-0008 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission |  |  |  |  | 03/10/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting with cardiomyopathy. during support, the patient pump unexpectedly fell out of the ventricle and prolapsing on itself. a snare was required to straighten the pump and remove in one piece. a new impella was placed without further issue. | [5 empty] | The impella 5.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37000 | 1220648-2021-00827 | 11618356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11618356&pc=OZD | 09/02/2020 | 04/06/2021 |  | 04/06/2021 |  |  | 03/09/2021 | 04/06/2021 |  |  | 04/30/2022 |  | 04/06/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1479906 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [40 empty] | Initial submission |  |  |  |  | 03/12/2021 | * | Invalid/Incomplete |  |  | The complainant reported an access site bleed and thrombocytopenia that prompted delivery of multiple units of blood product (platelets). | [5 empty] | The impella cp was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37001 | 1220648-2021-00820 | 11614707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11614707&pc=OZD | 03/14/2021 | 05/20/2021 |  | 04/05/2021 |  |  | 03/14/2021 | 04/05/2021 |  |  | 12/31/2022 |  | 04/05/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022022126 |  | 0048-0045 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 03/14/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella cp pump inserted in the left groin. the patient had hemolysis with plasma free hemoglobin resulted in 244, 280, 323, 921 and the pump was removed. | [5 empty] | The pump and data logs was returned for evaluation. review of the provided data logs showed multiple positioning alarms. including 'impella position unknown', 'impella in aorta', and 'impella in ventricle'. analysis of the returned product found no damage, that would have contributed to the reported hemolysis. the root cause of the hemolysis was improper positioning of the device. | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37002 | 1220648-2021-00821 | 11614713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11614713&pc=OZD | 03/09/2021 | 04/12/2021 |  | 04/05/2021 |  |  | 03/09/2021 | 04/05/2021 |  |  | 01/31/2023 |  | 04/05/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022037372 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | Followup |  |  |  | 03/09/2021 | * | Invalid/Incomplete |  |  | A patient was admitted with multiple cardiac comorbidities. the team observed multivessel coronary artery disease as well as an ef of only 20% and aortic valve insufficiency (ai). the plan was made for coronary artery bypass surgery (cabg) and aortic valve replacement with an impella 5.5 providing for hemodynamic support. the surgery was performed and when the patient was taken off the cabg pump, the impella stopped. the team attempted to troubleshoot the pump stop by restarting multiple times. when this failed, the team explanted the 5.5 pump. the patient was monitored and found to be stable. no second impella pump was placed. the pump had run for almost 5 hours of support before the pump stop. | [5 empty] | Since the initial report was filed. the medical center forwarded to abiomed the impella product. the product went through analysis. and the root cause of the pump stop was determined. the pump stopped, due to biomaterial ingestion. the failure mode will be monitored and trended. | The full investigation into the pump stop of the impella 5.5 was not possible. the team did not return the pump for investigation. the team did communicate that upon explant the pump had biomaterial, presumed to be plaque, adhered to the pump. this adhered material prevented the impeller from operating as necessary. without the pump, this can not be investigated further to determine the root cause of the stop and biomaterial composition. the failure mode will be monitored and trended.
37003 | 1220648-2021-00822 | 11616194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11616194&pc=OZD | 02/25/2021 | 04/05/2021 |  | 04/05/2021 |  |  | 03/09/2021 | 04/05/2021 |  |  | 10/31/2022 |  | 04/05/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2021012013 |  | 0048-0024-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 86 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Rupture | Aortic Valve Insufficiency/ Regurgitation | [11 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 03/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old female patient presenting with acute myocardial infarction and cardiogenic shock. the impella cp device was used for hemodynamic support. during support, the patient exhibited aortic regurgitation. a ruptured chordae tendineae was found. the patient ultimately expired. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37004 | 1220648-2021-00823 | 11616197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11616197&pc=OZD | 03/10/2021 | 04/05/2021 |  | 04/05/2021 |  |  | 03/10/2021 | 04/05/2021 |  |  | 01/31/2023 |  | 04/05/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022038087 |  | 0550-0008 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/10/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) patient presenting for hemodynamic support using the impella 5.5 pump. during support, a pocket hematoma formed at the axillary insertion site. as treatment, an evacuation was performed. | [5 empty] | The impella 5.5 was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37005 | 1220648-2021-00824 | 11616200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11616200&pc=OZD | 10/06/2020 | 04/05/2021 |  | 04/05/2021 |  |  | 03/09/2021 | 04/05/2021 |  |  | 04/30/2020 |  | 04/05/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1488338 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for use with the impella cp. during support, the patient developed an access site bleed. a blood transfusion was administered as treatment. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37006 | 1220648-2021-00811 | 11607264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11607264&pc=OZD | 03/11/2021 | 05/04/2021 |  | 04/02/2021 |  |  | 03/11/2021 | 04/02/2021 |  |  | 11/30/2022 |  | 04/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022022674 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 03/11/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella 5.5 pump inserted in the right axillary without issues. the patient had hemolysis, this pump was removed, and another impella 5.5 pump was used. | [5 empty] | The pump was returned for evaluation, and found to have a large thrombus in the outflow cage. there were no signs of damage seen. the device was cleaned and hemolysis tested, receiving an mih of 2.86. thus, this confirms the pump was not the root cause of the hemolysis. the root cause of hemolysis was ingested biomaterial. | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37007 | 1220648-2021-00810 | 11607306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11607306&pc=OZD | 03/09/2021 | 05/04/2021 |  | 04/02/2021 |  |  | 03/09/2021 | 04/02/2021 |  |  | 10/31/2022 |  | 04/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021018370 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 03/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) white male patient had impella 5.5 pump inserted in the right axillary artery for acute myocardial infarction (ami)/cardiogenic shock (cgs). during surgery, the physician lifted the heart and impella perforated the ventricle and this was repaired. | [5 empty] | The device was returned for evaluation. and showed no signs of damage, that could've contributed to the left ventricle (lv) perforation. it was likely, that the patient had a pre-existing complication in the myocardium. that allowed the pump to perforate, when the physician lifted the heart. | Device not returned by customer. the impella 5.5 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37008 | 1220648-2021-00813 | 11610584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11610584&pc=OZD | 03/06/2021 | 03/02/2022 |  | 04/02/2021 |  |  | 03/06/2021 | 04/02/2021 |  |  | 09/30/2022 |  | 04/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021006152 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 73 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | A |  |  | No | Initial submission | Followup |  |  |  | 03/06/2021 |  |  |  |  | The complainant reported a (b)(6) year-old (b)(6) male patient had impella cp pump inserted in the left femoral for high-risk percutaneous coronary interventions (hrpci). it was reported that suction persisted for several days after the device was turned on. volume loading was continued; however, since ecmo was also taking flow, filling became difficult to maintain. hemolysis was treated by introducing continuous hemodiafiltration (chdf. no further information was provided. | [5 empty] | Supplemental filed to include additional information received from the customer: the patient also exhibited renal failure. there was no additional treatment beyond what was originally reported. the original investigation stands as-is. the impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37009 | 1220648-2021-00812 | 11610591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11610591&pc=OZD | 03/07/2021 | 05/25/2021 |  | 04/02/2021 |  |  | 03/07/2021 | 04/02/2021 |  |  | 01/31/2023 |  | 04/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022038050 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 03/07/2021 | * | Invalid/Incomplete |  |  | The complainant reported a 77 year old male patient presenting with post cardiotomy cardiogenic shock for hemodynamic support using the impella 5.5 device. after approximately 2 days of support, the patient exhibited hemolysis that was confirmed by multiple draws of plasma-free hemoglobin values exceeding 40mg/dl within a 24 hour period. the pump was prematurely removed as a result. the patient remained in stable condition, thereafter. | [5 empty] | The impella 5.5 was received for investigation. a final mdr will be filed with results of the analysis. | Both the device and data logs were returned for review. visual inspection of the device found no damage. and data log analysis found no gross positioning issues, during support. simulated use testing for hemolysis, found the pump operated within specification. we are unable to determine a definitive root cause. however, we believe the reported issue was related to the patient's condition.
37010 | 1220648-2021-00817 | 11610700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11610700&pc=OZD | 08/31/2020 | 09/15/2024 |  | 04/02/2021 |  |  | 03/09/2021 | 04/02/2021 |  |  | 04/30/2022 |  | 04/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1480165 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 58 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 03/09/2021 |  |  |  |  | Additional information was received noting that the family made the decision to withdraw care, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral. indication was not provided. it was reported the patient developed an ischemic cerebrovascular accident (cva). no further information was provided. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2021-00817 was provided. | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37011 | 1220648-2021-00814 | 11610708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11610708&pc=OZD | 03/09/2021 | 04/02/2021 |  | 04/02/2021 |  |  | 03/09/2021 | 04/02/2021 |  |  | 10/31/2021 |  | 04/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1453032 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) -year-old (b)(6) male patient had impella cp pump inserted in an unknown insertion site for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed bleeding at the impella access site. as a result, blood products were administered, amount was not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37012 | 1220648-2021-00815 | 11610709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11610709&pc=OZD | 03/05/2021 | 01/14/2022 |  | 04/02/2021 |  |  | 03/08/2021 | 04/02/2021 |  |  | 10/31/2022 |  | 04/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2022041171 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | White | Malfunction | No | Yes | Material Fragmentation | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/08/2021 |  |  |  |  | A publication on this event was received, which clarified the events that actually transpired: the 0.018" guide wire endured resistance when attempting to remove from the impella. the wire snapped and the catheter folded onto itself requiring straightening/removal with a surgical snare. no fragments of the device were left in the patient and they remained in stable condition, thereafter. | The complainant reported a (b)(6) male patient presenting for high risk percutaneous coronary intervention using the impella cp smart assist for hemodynamic support. a piece of the impella device broke off into the patient. a surgical snare was required to remove the fragmented piece. the patient was noted in stable condition. |  |  |  |  | The pump and data logs were returned for analysis. visual inspection of the pump found the 0.018" guidewire was wrapped around the impeller blades and damaged. a portion of the guidewire was missing and not returned with the device. damage to the catheter, proximal to the pump, was found; indicative of an interaction with a snare. lastly, the catheter was kinked. an analysis of the data logs found a spike in motor current with the pump failing to start. based on our analysis, evidence suggests the reported issue was caused by the use failing to remove the guidewire prior to starting the pump. | The investigation results provided in the supplemental investigation have not changed, in light of this new information received via received publication. h6 coding and b5 narrative have been updated to clarify event originally reported. h(6): changed medical device problem code from 1261 to 2981, added 4641 and 4624 to health effect impact code. b(5): include additional information received via publication, clarifying the event originally reported. | We are currently requesting the device be returned to abiomed for an investigation. should we receive the device, or any new information is obtained, a supplemental mdr will be filed.
37013 | 1220648-2021-00819 | 11610710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11610710&pc=OZD | 03/12/2021 | 04/02/2021 |  | 04/02/2021 |  |  | 03/12/2021 | 04/02/2021 |  |  | 12/31/2022 |  | 04/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022027551 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/12/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient had impella cp pump inserted in the right leg for cardiogenic shock (cgs). the patient had ischemic leg distal of impella, required intubation and vasopressors, a 5.0 pump was inserted, and this pump was removed. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37014 | 1220648-2021-00818 | 11610720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11610720&pc=OZD | 03/09/2021 | 06/07/2021 |  | 04/02/2021 |  |  | 03/09/2021 | 04/02/2021 |  |  | 09/30/2022 |  | 04/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021003999 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 03/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting with post cardiotomy cardiogenic shock for hemodynamic support using the impella 5.5 device. upon removal of device and preparing for a new 5.5 (due to lower than expected flows), the ventricle was found perforated. as treatment, the ventricle was repaired with sutures. | [5 empty] | The device and data logs were returned for analysis. the data logs found multiple improper positioning alarms. the device tested and no alarms were triggered. based on further follow up discussion with the abiomed on-site representative, the physician chose to forgo using a guidewire for placement and had encountered some difficulty when crossing the valve. based on our analysis and the information on hand, we feel the cause of the ventricle perforation was use related, since a guide wire was not used during insertion. | The impella 5.5 was requested back from the account. should the device be returned, or any new information be received, a supplemental mdr will be filed.
37015 | 1220648-2021-00808 | 11596602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11596602&pc=OZD | 02/09/2021 | 03/31/2021 |  | 03/31/2021 |  |  | 03/09/2021 | 03/31/2021 |  |  | 06/30/2022 |  | 03/31/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1484370 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 86 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) asian female patient had impella cp optical pump inserted in the right femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had bleeding from impella insertion site that required blood transfusion. | [5 empty] | Device discarded by customer. the impella cp optical pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37016 | 1220648-2021-00809 | 11598869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11598869&pc=OZD | 03/04/2021 | 04/06/2021 |  | 03/31/2021 |  |  | 03/04/2021 | 03/31/2021 |  |  | 11/30/2022 |  | 03/31/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2021020255 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Extravasation | Vascular Dissection | [11 empty] | Hospitalization | Required Intervention |  |  |  |  | MEDTRONIC TAVR | [22 empty] | Initial submission | Followup |  |  |  | 03/04/2021 | * | Invalid/Incomplete |  |  | A patient was admitted for a high risk angioplasty of the coronary arteries. the patient has known left main coronary artery disease along with aortic valve stenosis. he was being prepped for a possible tavr. to give the patient hemodynamic support during the tavr an impella cp was placed. the patient tolerated the tavr but, when the impella was explanted the patient had an access site complication observed by the physician. the femoral artery had a perforation. the team placed a covered stent to remedy the issue. despite hemoglobin levels dropping, no replacement blood products were infused. | [5 empty] | Since the original report was filed, the investigation has concluded. the pump product, data logs, and one dilator for introducer sheath were returned for analysis. no imaging was returned despite request made. when the pump was run on the bench the pump performed to specification. the introducer's 12fr dilator was returned with tip damage. the root cause, of the presumed impella caused perforation, was deemed to be use related. the failure mode will be monitored and trended. | The investigation is on-going at this time. upon conclusion of the investigation, a final report will be filed.
37017 | 1220648-2021-00804 | 11588671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11588671&pc=OZD | 03/04/2021 | 03/30/2021 |  | 03/30/2021 |  |  | 03/05/2021 | 03/30/2021 |  |  | 08/31/2021 |  | 03/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1436980 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/05/2021 | * | Invalid/Incomplete |  |  | The complainant reported a female patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). patient's age was not provided. it was reported that at the time of removal a large amount of thrombus was removed with a fogarty catheter. it was noted there were no signs of limb ischemia. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37018 | 1220648-2021-00803 | 11588672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11588672&pc=OZD | 03/06/2021 | 03/30/2021 |  | 03/30/2021 |  |  | 03/07/2021 | 03/30/2021 |  |  | 12/31/2022 |  | 03/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022028932 |  | 0048-0024-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Death | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 03/07/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male patient had impella cp optical pump inserted in the right femoral. the patient had aortic dissection and patient was on do not attempt resuscitation (dnar) and expired. the pump had perforated the left thoracic cavity at aortic arch. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37019 | 1220648-2021-00805 | 11588676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11588676&pc=OZD | 03/04/2021 | 03/30/2021 |  | 03/30/2021 |  |  | 03/04/2021 | 03/30/2021 |  |  | 01/31/2023 |  | 03/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022036516 |  | 0550-0008 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/04/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella 5.5 pump inserted in the right axillary artery. the pump stopped after three (3) days, the patient was bleeding from jackson-pratt (jp) drain at axillary site, red blood cells (rbc), fresh frozen plasma, (ffp), and platelets was administered. | [5 empty] | Device not returned by customer. the impella 5.5 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37020 | 1220648-2021-00807 | 11588682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11588682&pc=OZD | 02/09/2021 | 03/30/2021 |  | 03/30/2021 |  |  | 03/09/2021 | 03/30/2021 |  |  | 01/31/2022 |  | 03/30/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1470969 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) asian female patient had impella cp optical pump inserted in the left femoral. the patient had bleeding from impella insertion site that required blood transfusion and thrombocytopenia that required platelet (plt) administration. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37021 | 1220648-2021-00806 | 11588685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11588685&pc=OZD | 03/03/2021 | 03/30/2021 |  | 03/30/2021 |  |  | 03/03/2021 | 03/30/2021 |  |  | 10/31/2022 |  | 03/30/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 2021017587 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 26 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pseudoaneurysm | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 03/03/2021 | * | Invalid/Incomplete |  |  | In (b)(6) a patient was admitted with myocarditis and had an impella cp placed via the femoral artery for hemodynamic support. va ecmo was also added for right heart support. impella support continued for approximately 5 days while the plan was being developed for either vad or heart transplant. as the patient's condition improved, the impella was explanted. upon explant the patient developed a groin site complication. the team observed a pseudoaneurysm at the access site of the impella and surgical hemostasis was necessary. further details of what the surgeon did to the patient's anatomy were not possible, due to covid restrictions within the medical center. no imaging or product return was possible. | [5 empty] | The complainant in (b)(6) did not return the product or imaging for analysis. the team did report that the patient was young with no history of arteriosclerosis and believe the access site complication to be due to the explanting physician's technique. without product return the root cause is presumed to be patient condition. what surgical technique was done to achieve hemostasis is unknown. the failure mode will be monitored and trended.
37022 | 1220648-2021-00795 | 11576264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11576264&pc=OZD | 02/26/2021 | 03/26/2021 |  | 03/26/2021 |  |  | 02/27/2021 | 03/26/2021 |  |  | 10/31/2022 |  | 03/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021018096 |  | 0550-0008 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 29 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/27/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with cardiomyopathy. after returning to the icu from having a swan-ganz catheter placed, the patient developed excessive bleeding from the impella insertion site. the impella was removed, the patient was delivered several units of blood products, and surgery to repair the damaged artery was performed. it should be noted that the physician feels the placement of the swan-ganz catheter may have nicked a suture holding the graft to the artery in which the impella was placed. this sequence of events was felt to have put excess strain on the artery resulting in the reported tear. | [5 empty] | The impella 5.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37023 | 1220648-2021-00796 | 11576335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11576335&pc=OZD | 02/19/2020 | 03/29/2021 |  | 03/26/2021 |  |  | 03/02/2021 | 03/26/2021 |  |  | 10/31/2021 |  | 03/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1453032 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 03/02/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed access site bleed. as a result, the patient was administered blood products, amount was not provided. no further information was provided. | [5 empty] | The impella was returned for analysis. the cause of the access site bleeding was not determined since no information regarding the bleed was provided. the impella tempubunsho recommends that during impella support the patient should be anticoagulated to where the patient¿s activated clotting time (act) is between 160 and 180 seconds. if the patient¿s act is above this range then there is an increased risk of bleeding at the impella access site. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event. | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37024 | 1220648-2021-00797 | 11576577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11576577&pc=OZD | 01/13/2021 | 03/26/2021 |  | 03/26/2021 |  |  | 03/02/2021 | 03/26/2021 |  |  | 10/31/2022 |  | 03/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021017584 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/02/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) female patient had impella cp pump inserted in the right axillary for high-risk percutaneous coronary interventions (hrpci). it was reported that the patient developed access site bleed. as a result, the patient received blood products, amount was not provided. additionally, the surgical suturing was also performed. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37025 | 1220648-2021-00798 | 11576645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11576645&pc=OZD | 03/01/2021 | 03/26/2021 |  | 03/26/2021 |  |  | 03/01/2021 | 03/26/2021 |  |  | 11/30/2022 |  | 03/26/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2021020183 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 03/01/2021 | * | Invalid/Incomplete |  |  | A patient was admitted in cardiogenic shock. she was suffering from vt (ventricular tachycardia) storm and after being cardioverted multiple times she had an impella cp placed for hemodynamic support. the impella was placed via the right femoral artery. as her care needed escalation, she was transferred to the tertiary center. when admitted to the tertiary center the impella limb was observed to be ischemic. the limb was seen to have become pulseless. when the team at the tertiary hospital was unable to successfully perfuse the limb, after failing to place a distal perfusion line, the impella was explanted and an iabp (intra-aortic balloon pump) was placed instead. | [5 empty] | The medical center discarded the impella product and as such no investigation was conducted. the cause of the limb ischemia that prompted pump explant is presumed to be patient condition. the failure mode will be monitored and trended.
37026 | 1220648-2021-00799 | 11576663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11576663&pc=OZD | 03/01/2021 | 05/12/2021 |  | 03/26/2021 |  |  | 03/01/2021 | 03/26/2021 |  |  | 12/31/2022 |  | 03/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022023697 |  | 005060 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 66 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | Followup |  |  |  | 03/01/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting post coronary artery bypass graft and cardiotomy. the impella 2.5 was inserted as a "weaning strategy" from ecmo. less than 1 hour into pump use, the device stopped. a successful restart was made, however, the pump stopped again. the pump was removed and replaced and a new impella 5.0 with no further issues. | [5 empty] | The pump was returned and an investigation was completed. a large mass of biomaterial was found on the flow separator. once the mass was removed, the pump was able to run within specification. we believe the root cause for the reported pump stop was biomaterial ingestion due to patient's condition. | The impella cp was returned by the customer and an investigation is underway. a supplemental mdr will be filed at the completion of our evaluation.
37027 | 1220648-2021-00800 | 11576673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11576673&pc=OZD | 03/01/2021 | 05/12/2021 |  | 03/26/2021 |  |  | 03/02/2021 | 03/26/2021 |  |  | 07/31/2022 |  | 03/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1492166 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 03/02/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that at explant it was noted that the patient's lower limb showed signs of leg ischemia. thrombotic obstruction was confirmed via fogati. blood vessel incision was performed to remove the thrombus. it was noted that prior to impella insertion the patient had very poor blood vessel properties; however, the placement of impella may have promoted thrombus formation. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The device was returned for analysis. the cause of the ischemia was most likely, the condition/size of the patient vessel. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event.
37028 | 1220648-2021-00801 | 11576680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11576680&pc=OZD | 03/02/2021 | 07/27/2021 |  | 03/26/2021 |  |  | 03/02/2021 | 03/26/2021 |  |  | 10/31/2022 |  | 03/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022028450 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 86 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombocytopenia | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup | Followup |  |  | 03/02/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old asian male patient had impella cp optical pump inserted in the left femoral for acute myocardial infarction (ami)- cardiogenic shock (cgs). the patient had subcutaneous hematoma, as a result the hematoma was surgically removed, ten (10) units of fresh frozen plasma (ffp), and ten (10) units of red blood cells (rbc) was administered. | [5 empty] | We received additional information on (b)(6) 2021 that the patient had thrombocytopenia and platelet transfusion was administered. based on the clinical account of high activated clotting time (acts), the root cause of the access site bleeding was due to act management. | Abiomed received information that the patient also had access site infection and antibiotics was given. | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37029 | 1220648-2021-00802 | 11576697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11576697&pc=OZD | 06/20/2020 | 03/26/2021 |  | 03/26/2021 |  |  | 03/06/2021 | 03/26/2021 |  |  | 03/31/2022 |  | 03/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 1474462 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/06/2021 | * | Invalid/Incomplete |  |  | We received a publication from cardiology journal, 2021, vol. 28, no. 1, 185-186 that a (b)(6) year old female patient had impella cp pump inserted in the left femoral artery for hemodynamic support. injection of the left femoral artery through contralateral access demonstrated a thrombosis as a result a mechanical thrombectomy using rotarex was performed allowing immediate antegrade blood flow restoration in the left common femoral artery. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37030 | 1220648-2021-00794 | 11555392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11555392&pc=OZD | 02/22/2021 | 03/23/2021 |  | 03/23/2021 |  |  | 02/22/2021 | 03/23/2021 |  |  | 07/31/2022 |  | 03/23/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1492973 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 02/22/2021 | * | Invalid/Incomplete |  |  | A patient with multiple cardiac and oncologic comorbidities was admitted in cardiogenic shock. she had an impella cp placed for hemodynamic support. the cp allowed for cardiac angioplasty to occur. at conclusion of the intervention, her leg began to become mottled. the color change was indicative of limb ischemia. the team made choice after 10 hours of support to explant the cp from the ischemic leg and place another impella via the axillary artery insertion. vascular surgery intervened and restored flow to the limb. | [5 empty] | The medical center did not return for analysis the impella pump product. no imaging of the ischemic limb nor data logs from the impella console were returned for analysis and investigation. the patient had known peripheral arterial disease. the cause of the limb ischemia is patient condition.
37031 | 1220648-2021-00784 | 11540288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11540288&pc=OZD | 02/26/2021 | 03/22/2021 |  | 03/22/2021 |  |  | 02/26/2021 | 03/22/2021 |  |  | 12/31/2022 |  | 03/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022030529 |  | 0550-0008 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old black or african american male patient had impella 5.5 pump placed surgically via anterior direct aortic approach. there was an aortic valve dissection that resulted in surgical repair with three (3) units of platelets, two (2) units of cryo, and two (2) units of fresh frozen plasma (ffp) being administered. | [5 empty] | Device not returned by customer. the impella 5.5 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37032 | 1220648-2021-00785 | 11540316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11540316&pc=OZD | 02/22/2021 | 03/25/2021 |  | 03/22/2021 |  |  | 02/22/2021 | 03/22/2021 |  |  | 12/31/2022 |  | 03/22/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022026890 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 54 YR | [5 empty] | Malfunction | No | Yes | Failure to Power Up | [6 empty] | Low Cardiac Output | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 02/22/2021 | * | Invalid/Incomplete |  |  | A patient with cardiomyopathy was admitted and placed on impella 5.5 support. after 5 days of support the impella began to exhibit purge alarms which prompted the team to add tpa to the purge, however this did not remedy the purge alarms and the pump stopped. when they were unable to re-start the impella 5.5 the pump was explanted and replaced. during the interruption in the support the team escalated his vasoactive support due to his condition deteriorating while pump exchange taking place. | [5 empty] | The product has been returned for analysis. the analysis of the impella 5.5, and it's stop is on-going at this time. when the investigation is completed a final report will be filed. | Since the original report was filed, the investigation has concluded. the cause of the impella 5.5 stopping was determined after review and analysis of the returned data logs and pump product. the cause was blood ingress into the motor of the impella pump. the failure mode will be monitored and trended.
37033 | 1220648-2021-00786 | 11546325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11546325&pc=OZD | 02/25/2021 | 03/22/2021 |  | 03/22/2021 |  |  | 02/25/2021 | 03/22/2021 |  |  | 07/31/2022 |  | 03/22/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD AIC PUMP SET US |  |  | ABIOMED INC. |  |  | US |  | 1492909 |  | 005082 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 74 YR | [5 empty] | Malfunction | No | Yes | Loss of Power | [6 empty] | Appropriate Term / Code Not Available | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 02/25/2021 | * | Invalid/Incomplete |  |  | A patient was admitted with multiple cardiac and renal comorbidites for a coronary artery bypass surgery. post cardiotomy he exhibited shock and had an impella ld placed for hemodynamic support. after just over 19 hours of ld support the pump stopped. the pump stop was unable to be rectified and so the pump was explanted and an iabp was placed for support. the patient survived. | [5 empty] | The medical center returned to abiomed the impella ld product for investigation and analysis, as well as the pump data logs. when the pump was opened and broken down upon the benchtop there was an observation made of blood ingress into the motor. the failure mode will be monitored and trended.
37034 | 1220648-2021-00787 | 11546359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11546359&pc=OZD | 02/24/2021 | 03/22/2021 |  | 03/22/2021 |  |  | 02/24/2021 | 03/22/2021 |  |  | 07/31/2021 |  | 03/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1430738 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/24/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 5.0 pump inserted in the right axillary for cardiomyopathy. it was reported that the patient developed an access site bleed. as a result, blood products four units of red blood cells were administered. heparin was also removed from the temporary purge solution. a ct showed no re-bleeding, and heparin was added again. no surgical procedure was performed. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37035 | 1220648-2021-00788 | 11546367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11546367&pc=OZD | 02/24/2021 | 03/22/2021 |  | 03/22/2021 |  |  | 02/24/2021 | 03/22/2021 |  |  | 11/30/2022 |  | 03/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022032133 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/24/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) male patient had impella cp pump inserted in the left femoral for high-risk percutaneous coronary interventions (hrpci). it was reported that the patient developed a hematoma at the impella access site. as a result, the patient was administered blood products, number of units were not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37036 | 1220648-2021-00792 | 11547595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11547595&pc=OZD | 02/26/2021 | 03/22/2021 |  | 03/22/2021 |  |  | 02/26/2021 | 03/22/2021 |  |  | 10/31/2022 |  | 03/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021017567 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for myocarditis. it was reported that after replacing the peel-away sheath with the indwelling sheath bleeding occurred. the bleeding did not stop. the angle was maintained, and pressure was fixed from above to stop bleeding. the patient was administered mannitol, adenine, and phosphate (map). no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37037 | 1220648-2021-00790 | 11547597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11547597&pc=OZD | 02/24/2021 | 03/22/2021 |  | 03/22/2021 |  |  | 02/24/2021 | 03/22/2021 |  |  | 01/31/2022 |  | 03/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1465351 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/24/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right axillary for cardiomyopathy. it was reported that bleeding from below the right clavicle of the impella insertion site was noted. as a result surgical treatment was rendered. details of the treatment was not provided. no blood transfusion was required. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37038 | 1220648-2021-00791 | 11547614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11547614&pc=OZD | 03/25/2020 | 03/22/2021 |  | 03/22/2021 |  |  | 02/24/2021 | 03/22/2021 |  |  | 04/30/2020 |  | 03/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1347824 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/24/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 5.0 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that aortic regurgitation (ar) occurred during support with impella. after impella was explanted ar improved, therefore, no additional treatment was provided; however, the device was removed due to ar. the length from the a valve to the suction part is unknown. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37039 | 1220648-2021-00789 | 11547640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11547640&pc=OZD | 11/27/2020 | 03/22/2021 |  | 03/22/2021 |  |  | 02/24/2021 | 03/22/2021 |  |  | 01/31/2022 |  | 03/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1470953 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 46 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [41 empty] | Initial submission |  |  |  |  | 02/24/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) female patient had impella cp pump inserted in the right femoral for myocarditis. it was reported that the patient developed a cerebrovascular accident during impella support. patient's outcome was reported as sequelae. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37040 | 1220648-2021-00793 | 11547661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11547661&pc=OZD | 02/24/2021 | 03/22/2021 |  | 03/22/2021 |  |  | 02/26/2021 | 03/22/2021 |  |  | 07/31/2022 |  | 03/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1492166 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 83 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed limb ischemia during impella support. as a result, femoral to femoral bypass was implemented. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37041 | 1220648-2021-00781 | 11529927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11529927&pc=OZD | 02/26/2021 | 03/19/2021 |  | 03/19/2021 |  |  | 02/26/2021 | 03/19/2021 |  |  | 12/31/2022 |  | 03/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022026816 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 59 YR | [5 empty] | Malfunction | No | Yes | Fracture | Separation Problem | [5 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 02/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient had impella cp optical pump inserted in the left femoral artery (lfa). the inlet cage and pigtail broke off inside the patient and as a result cardiopulmonary resuscitation (cpr) was initiated but the patient expired. | [5 empty] | Device not returned by customer. the impella cp optical pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37042 | 1220648-2021-00782 | 11529940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11529940&pc=OZD | 02/25/2021 | 04/14/2021 |  | 03/19/2021 |  |  | 02/25/2021 | 03/19/2021 |  |  | 09/30/2022 |  | 03/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1501812 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 73 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 02/25/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white female patient had impella 5.5 pump inserted in the right axillary. the pump stopped during the intended support time according to the instructions for use (ifu). a cp pump was inserted after the 5.5 pump was removed. | [5 empty] | The data logs were returned. and confirmed, that the pump was stopped, due to a motor current mismatch. when the pump was returned to head quarters, there was no damage found to the motor. however, the electrical resistances of 2 phases were 6 ohms higher. this was likely, due to the long duration of blood/fluids in the red pump plug. a suture hole at the 30 cm mark was found confirming, the root cause of the blood ingress into the red pump plug. and ultimately the pump stop. | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37043 | 1220648-2021-00783 | 11529954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11529954&pc=OZD | 02/27/2021 | 04/14/2021 |  | 03/19/2021 |  |  | 02/27/2021 | 03/19/2021 |  |  | 03/31/2021 |  | 03/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 1411909 |  | 005040 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | CONSUMER | FOREIGN | HEALTH PROFE |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 02/27/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella 2.5 pump inserted in the left femoral. the patient had hemolysis and as a result ten (10) units of erythrocyte concentrates (ecc) was administered. | [5 empty] | The product was returned for evaluation and no delaminated epotek or biomaterial was found during analysis. data log review shows many suction events occurring throughout the case. since the transesophageal echocardiography (tte) revealed the pump was moved to good positioning with no inflow/outflow blockages, the root cause of the suspected hemolysis was likely due to the patient's condition. | The impella 2.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37044 | MW5100119 | 11529137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11529137&pc=PYX | 03/10/2021 | 03/16/2021 |  | 03/18/2021 |  |  |  | 03/18/2021 |  |  |  |  | 03/18/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA PURGE LINE |  |  | PYX | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 30 YR |  |  | 63 |  |  | Injury | No | Yes | Break | [6 empty] | Insufficient Information | [12 empty] | Hospitalization | Life Threatening | Other | [26 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | A (b)(6) female who presents with postpartum cardiomyopathy who arrested on the floor x2. rosc obtained on maximum drips. taken emergently for left axillary impella. she has been recovering in the icu and was ready to transfer to (b)(6). while on the helicopter pad for transfer to (b)(6), the impella purge line broke thereby necessitating emergent replacement. she was taken from the helicopter pad directly to the hybrid room for emergency replacement. given the events, it was not prudent for her to transfer that evening in the helicopter, thus delaying her transfer to higher level of care for heart transplant list. i have additional documents that will not allow me to upload because they are not a picture. the documents show that abiomed knows that this is an issue and has created a "workaround" for this problem. please contact me via email so that i may share those documents. thank you. fda safety report id# (b)(4).
37045 | 1220648-2021-00777 | 11504542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11504542&pc=OZD | 02/18/2021 | 05/06/2021 |  | 03/17/2021 |  |  | 02/18/2021 | 03/17/2021 |  |  | 10/31/2022 |  | 03/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021018093 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 49 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | Followup |  |  |  | 02/18/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella 5.5 pump. after less than 24 hours of support, the pump stopped. attempts to restart the pump were not successful, therefore, the device was removed and replaced with a new impella. no further issues were reported nor any patient harm was reported. | [5 empty] | The impella 5.5 was received and an investigation is underway. upon completion, a supplemental mdr will be filed with investigation results. | The device and data logs were received. upon physical inspection of the pump, there was no sign of damage or obstructions. however, blood was seen under the impeller. data logs found a sudden rise in purge pressure with a delayed purge blocked alarm. the cause of the pump stop was determined, to be the result of blood ingress.
37046 | 1220648-2021-00778 | 11508231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11508231&pc=OZD | 02/19/2021 | 03/30/2021 |  | 03/17/2021 |  |  | 02/19/2021 | 03/17/2021 |  |  | 11/30/2022 |  | 03/17/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022021112 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | VA ECMO | [22 empty] | Initial submission | Followup |  |  |  | 02/19/2021 | * | Invalid/Incomplete |  |  | A patient was admitted in cardiogenic shock suffering from an acute myocardial infarction. he had presented to a community hospital after 2 weeks of symptoms, and was transferred to the tertiary center for cardiac care. when assessed in the cardiac catheterization lab, multivessel coronary artery disease was diagnosed, and the team placed an impella cp for hemodynamic support. the patient was intubated and continued to deteriorate. the team placed va ecmo. when an echo was performed a visualization was made of free wall rupture of the left ventricle (lv) at the apex, which was attributed to the use of, and placement of, the impella. the patient was taken to the operating suite and the lv was repaired. in addition he received 13 units of blood product. later the family made choice to only allow care and comfort. he later expired. | [5 empty] | Since the original report was filed, the investigation into the patient injury had concluded. the ventricular perforation was due to underlying condition. the previous myocardial infarction had caused the tissue to be necrotic and fragile. the perforation was due to the pre-existing condition. the failure mode will be monitored and trended. | The investigation into the ventricular perforation is on-going at this time. upon completion of the investigation a final report will be filed.
37047 | 1220648-2021-00779 | 11508243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11508243&pc=OZD | 02/19/2021 | 03/17/2021 |  | 03/17/2021 |  |  | 02/19/2021 | 03/17/2021 |  |  | 08/31/2022 |  | 03/17/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED INC. |  |  | US |  | 1503737 |  | 005062 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 39 YR | [5 empty] | Injury | Yes | No | Positioning Problem | [6 empty] | Appropriate Term / Code Not Available | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/19/2021 | * | Invalid/Incomplete |  |  | A (b)(6) male was admitted to the hospital suffering from cardiomyopathy. for hemodynamic support the team placed an impella 5.0. the 5.0 was placed as the patient awaited an lvad in surgery. after 31 days of successful support the 5.0 was explanted. the abiomed rep was notified that they had observed that the pump was adhered to the to the aortic valve leaflet. the team surgically removed the pump and placed va ecmo for 2 days while they prepared for the surgical lvad. the pump had been used beyond indication when the valve was compromised. the patient survived and an lvad was placed. | [5 empty] | The investigation into the pump migrating to, and adhering to, the aortic valve has not completed. the product, logs, and imaging have not yet been returned for analysis. the cause of the damage to this valve is unknown at this time. upon completion of the investigation, a final report will be filed.
37048 | 1220648-2021-00780 | 11508247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11508247&pc=OZD | 02/21/2021 | 03/17/2021 |  | 03/17/2021 |  |  | 02/21/2021 | 03/17/2021 |  |  | 10/31/2022 |  | 03/17/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021011934 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/21/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that the patient experienced bleeding and a hematoma formed at the impella access site. the patient was administered six units of red blood cells with subsequent transfusion for which administration is unknown. it was noted that there was no patient movement throughout support. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37049 | 1220648-2021-00775 | 11485537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11485537&pc=OZD | 11/07/2020 | 03/15/2021 |  | 03/15/2021 |  |  | 02/15/2021 | 03/15/2021 |  |  | 12/31/2021 |  | 03/15/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1463499 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/15/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed an infection at the impella access site. as a result, an antibacterial agent was administered, details were not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37050 | 1220648-2021-00776 | 11490238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11490238&pc=OZD | 02/16/2021 | 03/15/2021 |  | 03/15/2021 |  |  | 02/18/2021 | 03/15/2021 |  |  | 10/31/2022 |  | 03/15/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021017099 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 81 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | [12 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 02/18/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient experienced cardiac arrest and cardiac tamponade occurred. act's were extended to operate impella. it was noted that blood gradually exuded into the heart sac. no additional treatment was provided. the patient ultimately expired. the patient became cpa for which pcps was inserted. at that time, hypoxemia occurred and caused hypoxic encephalopathy, resulting in death. impella was working until the patient died. relationship to impella: unknown, because cpa might have caused pericardial effusion due to the use of impella. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37051 | 11462143 | 11462143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11462143&pc=PBL | 01/25/2021 | 02/09/2021 | 02/25/2021 | 03/11/2021 | 02/25/2021 | 03/11/2021 |  | 03/11/2021 |  |  |  |  | 03/11/2021 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | PBL |  |  | IMPELLA LR 0.018¿ PTFE GUIDEWIRE |  |  | ABIOMED |  |  | US |  |  |  | 0048-0003 |  | No | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 29200 DA |  |  | 99 |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [28 empty] | Initial submission | [5 empty] | Y | Invalid/Incomplete |  |  | After the impella cp was removed from the right common femoral artery there was persistent bleeding, despite femo-stop and manual compression and purse-string suture. this required emergent balloon angioplasty of the femoral artery to stop bleeding.
37052 | 1220648-2021-00770 | 11455533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11455533&pc=OZD | 02/24/2021 | 03/10/2021 |  | 03/10/2021 |  |  | 02/24/2021 | 03/10/2021 |  |  | 09/30/2022 |  | 03/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022023202 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/24/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) female patient had impella cp pump inserted in the left femoral for post-cardiotomy cardiogenic shock (pccs). it was reported that while attempting to explant the impella the discharge port part was caught in perceval (bioprosthetic valve designed). due to this the physician was unable to pull or push the impella. troubleshooting such as using a guiding wire to push impella to shift the position or using a snare from another site approach to hook it on the pigtail of impella and pushing it forward to catch it; however, this was not successful. the patient¿s chest was opened and impella was able to be removed. the patient remained stable. the physician stated that the insertion method may have caused the issue. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device has not been possible. should any new information be returned, a supplemental mdr will be filed.
37053 | 1220648-2021-00772 | 11455535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11455535&pc=OZD | 02/09/2021 | 03/10/2021 |  | 03/10/2021 |  |  | 02/09/2021 | 03/10/2021 |  |  | 12/31/2021 |  | 03/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1459421 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed access site bleed. as a result, the patient required blood administration. the number of units is unknown. no further information was provided. | [5 empty] | The impella cp has not been returnedby the customer and therefore, investigation of device is not possible. should any new information be returned, a supplemental mdr will be filed.
37054 | 1220648-2021-00771 | 11455536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11455536&pc=OZD | 08/18/2020 | 03/10/2021 |  | 03/10/2021 |  |  | 02/09/2021 | 03/10/2021 |  |  | 01/31/2022 |  | 03/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1470941 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed access site bleed. as a result, the patient required blood administration. the number of units is unknown. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37055 | 1220648-2021-00773 | 11455544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11455544&pc=OZD | 12/01/2020 | 03/10/2021 |  | 03/10/2021 |  |  | 02/09/2021 | 03/10/2021 |  |  | 08/31/2022 |  | 03/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1496671 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in an unknown location for myocarditis. it was reported the patient developed access site bleed. as a result, the patient required blood administration. the number of units is unknown. successful hemostasis was achieved by pressure hemostasis. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37056 | 1220648-2021-00774 | 11455554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11455554&pc=OZD | 01/30/2021 | 03/10/2021 |  | 03/10/2021 |  |  | 02/09/2021 | 03/10/2021 |  |  | 09/30/2022 |  | 03/10/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1503408 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in an unknown location for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed access site bleed. as a result, a red blood cells were administered, number of units is unknown. no further information was provided. | [5 empty] | The device was returned for analysis. the cause of the access site bleeding was not determined. the impella tempubunsho recommends that during impella support the patient should be anticoagulated to where the patient¿s activated clotting time (act) is between 160 and 180 seconds. if the patient¿s act is above this range then there is an increased risk of bleeding at the impella access site.
37057 | 1220648-2021-00767 | 11448633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11448633&pc=OZD | 02/09/2021 | 03/09/2021 |  | 03/09/2021 |  |  | 02/09/2021 | 03/09/2021 |  |  | 08/31/2021 |  | 03/09/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1444442 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hypovolemic Shock | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 02/09/2021 | * | Invalid/Incomplete |  |  | A patient was admitted via the emergency room in (b)(6) suffering from an acute myocardial infarction and cardiogenic shock. he had ecmo placed and this was followed by coronary angioplasty with an impella cp placed for hemodynamic support. the patient's condition was poor and after 2 days on support the patient expired from hypovolemic shock. after approximately 1 year there was a registry shared inclusive of this patient and abiomed was notified that during the support the patient experienced signs of hemolysis that prompted the team to infuse blood products. the exact amount of blood products infused was not shared. the blood products also benefitted the patient who was suffering from bleeding at the impella access site. there was thought that perhaps the patient was suffering from dic (disseminated intravascular coagulation) which could have contributed to the blood loss. | [5 empty] | The complainant in (b)(6) did not return for investigation the impella product. the information shared via registry did not lead to a root cause that ties the presumed hemolysis or bleeding to the use of impella. the cause of the access bleeding was the high act values and improper anticoagulation while the injury, of presumed hemolysis, was due to improper pump positioning. the failure mode will be monitored and trended.
37058 | 1220648-2021-00769 | 11448647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11448647&pc=OZD | 02/12/2021 | 03/09/2021 |  | 03/09/2021 |  |  | 02/13/2021 | 03/09/2021 |  |  | 03/31/2022 |  | 03/09/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1479287 |  | 0550-0008 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/13/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old hispanic/latino female patient had impella 5.5 pump inserted in the right axillary. while in the operating room (or) the pump perforated the left ventricle (lv), a surgical repair was done, and device was removed. | [5 empty] | Device discarded by customer. the impella 5.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37059 | 1220648-2021-00768 | 11448649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11448649&pc=OZD | 03/03/2021 | 03/09/2021 |  | 03/09/2021 |  |  | 03/03/2021 | 03/09/2021 |  |  | 10/31/2022 |  | 03/09/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021005215 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 75 YR | [5 empty] | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission |  |  |  |  | 03/03/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male patient had impella cp optical pump inserted. the pump was returned and upon evaluation, the impeller blades had been broken off. | [5 empty] | The returned product was visually inspected, and no abnormalities were found on the cannula. inspection of the outflow cage revealed damage to the impeller. the pump was then run, and the failure mode was reproduced (2.8 l/min at p-9), the cannula was then detached; this confirmed the impeller damage, as the lower half of one the impeller blades had been broken off.
37060 | 1220648-2021-00762 | 11437133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11437133&pc=OZD | 02/11/2020 | 03/08/2021 |  | 03/08/2021 |  |  | 02/09/2021 | 03/08/2021 |  |  | 08/31/2021 |  | 03/08/2021 |  | : TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1434959 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed access site bleed. as a result, red blood cells and fresh frozen plasma were administered. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37061 | 1220648-2021-00763 | 11437135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11437135&pc=OZD | 01/06/2021 | 03/08/2021 |  | 03/08/2021 |  |  | 02/09/2021 | 03/08/2021 |  |  | 07/31/2022 |  | 03/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1488085 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that while on impella support the patient experienced access site bleed. as a result, blood products were administered, amount unknown. patient's act's were managed between 160-180s. subsequently, the patient also developed hemolysis for which haptoglobin was administered, the amount is also unknown. no further information was reported. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37062 | 1220648-2021-00764 | 11439431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11439431&pc=OZD | 03/22/2020 | 03/08/2021 |  | 03/08/2021 |  |  | 02/09/2021 | 03/08/2021 |  |  | 10/31/2021 |  | 03/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1453055 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for myocarditis. patient¿s age was not provided. it was reported that while on impella support the patient experienced access site bleed. as a result, red blood cells were administered, amount unknown. patient's act's were managed between 160-180s. subsequently, the patient also developed hemolysis for which haptoglobin was administered, the amount is also unknown. no further information was reported. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37063 | 1220648-2021-00766 | 11440338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11440338&pc=OZD | 10/23/2020 | 03/08/2021 |  | 03/08/2021 |  |  | 02/09/2021 | 03/08/2021 |  |  | 07/31/2022 |  | 03/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1493301 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting with a history of a coronary artery bypass graft and ventricular failure. the impella cp was selected for hemodynamic support and inserted without any reported issues. during support, the patient developed an access site bleed that prompted multiple units of blood product delivered. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37064 | 1220648-2021-00765 | 11440392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11440392&pc=OZD | 11/28/2020 | 03/08/2021 |  | 03/08/2021 |  |  | 02/09/2021 | 03/08/2021 |  |  | 08/31/2022 |  | 03/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1496670 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp device. during support, the patient exhibited signs of hemolysis via red urine and lowered hemoglobin levels, however, no plasma-free hemoglobin test to confirm hemolysis was completed. as treatment, haptoglobin was delivered. the patient also developed an access site bleed that prompted multiple units of blood product delivered. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37065 | 1220648-2021-00757 | 11423853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11423853&pc=OZD | 09/29/2020 | 03/05/2021 |  | 03/05/2021 |  |  | 02/09/2021 | 03/05/2021 |  |  | 01/31/2022 |  | 03/05/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1476728 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 46 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male patient had impella cp pump inserted in the left femoral for acute myocardial infarction (ami) with shock. the patient was bleeding from the access site and an unknown amount of blood transfusion was administered. | [5 empty] | Device not returned by customer. the impella cp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37066 | 1220648-2021-00758 | 11423887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11423887&pc=OZD | 02/09/2021 | 03/05/2021 |  | 03/05/2021 |  |  | 02/09/2021 | 03/05/2021 |  |  | 07/31/2022 |  | 03/05/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1490681 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 41 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/09/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male patient had impella cp pump inserted in the right femoral. the patient had hemolysis and access site bleeding and as a result an unknown amount of red blood cells (rbc) transfusion was administered. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37067 | 1220648-2021-00759 | 11423945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11423945&pc=OZD | 02/10/2021 | 03/05/2021 |  | 03/05/2021 |  |  | 02/10/2021 | 03/05/2021 |  |  | 12/31/2022 |  | 03/05/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022023327 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 78 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 02/10/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported developed hemorrhagic cerebral infarction along with celiac artery embolism, renal artery embolism, and lower limb embolism. furthermore, a large amount of mural thrombus was seen in the aorta. the patient ultimately expired. cause of death was reported as thrombosis. hit was also suspected. date of patients death was not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37068 | 1220648-2021-00760 | 11424014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11424014&pc=OZD | 02/12/2021 | 03/05/2021 |  | 03/05/2021 |  |  | 02/12/2021 | 03/05/2021 |  |  | 06/30/2022 |  | 03/05/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1491982 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/12/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient had impella cp optical pump inserted in the right femoral artery. the patient was bleeding profusely when the repositioning sheath was inserted and as a result three (3) units of red blood cells (rbcs) was administered. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37069 | 1220648-2021-00761 | 11424164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11424164&pc=OZD | 01/28/2020 | 03/05/2021 |  | 03/05/2021 |  |  | 02/10/2021 | 03/05/2021 |  |  | 10/31/2021 |  | 03/05/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 1446055 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 38 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/10/2021 | * | Invalid/Incomplete |  |  | We received a publication from esc european society of cardiology, european heart journal- case reports (2020) 4, 1-4, ecmella: successful rescue cardiopulmonary support in post-coronary artery bypass graft cardiogenic shock with cardiac arrest¿case report that a 38 year old white male patient had impella cp pump inserted in the left femoral. during the procedure, there was an accidental tear of the left femoral artery, with bleeding, that was controlled with surgical repair. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37070 | 11412817 | 11412817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11412817&pc=OZD | 01/18/2021 | 02/05/2021 | 02/05/2021 | 03/04/2021 | 02/05/2021 | 03/04/2021 |  | 03/04/2021 |  |  |  |  | 03/04/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 21535 DA |  |  | 98 |  |  | Malfunction | No | Yes | Backflow | Output Problem | [5 empty] | Left Ventricular Dysfunction | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Patient has cardiogenic shock which required an axillary impella 5.5. this was done as a bridge to transplant. his device acutely failed. we pulled it back into the aorta. on the way to ct, the impella began alarming "impella stopped" and "impella stopped, retrograde flow" p level dropped to 0. attempted to reset p level. did not work. began alarming again. returned to room, changed consoled, experienced the same alarms, impella @ 45. echocardiogram paged. ct surgeon at bedside and pulled impella out of the heart into ascending aorta during echo. pt transported emergently to or for replacement. patient was brought to the or for exchange. intra-operative echocardiogram demonstrated a depressed and dilated lv. there was no lv thrombus. th old device was removed uneventfully. the graft was completely de-clotted. the valve was crossed uneventful. the impella was inserted uneventfully. the impella was directed away from the mitral valve and measured about 6.0cm from the aortic valve. the device was functioning well. his septum was midline. it was flowing about 4.5l at p8. his drips were adjusted. the axillary cutdown site was hemostatic. overall, he tolerated the procedure well and was transferred back to the icu in stable condition. defective device was given to on-site device rep per direction from manager and charge nurse for further diagnosis. impella devices need special packaging requirement and are time sensitive for patient safety so manger gave approval to release to vendor from the or.
37071 | 1220648-2021-00752 | 11406207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11406207&pc=OZD | 02/02/2021 | 03/03/2021 |  | 03/03/2021 |  |  | 02/02/2021 | 03/03/2021 |  |  | 09/30/2022 |  | 03/03/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1503408 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/03/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for myocarditis. it was reported the patient developed access site bleeding during re-positioning. as a result, the patient was administered blood products (red blood products, fresh frozen plasma). the patient bleed for a third time for which a muscle relaxant was administered. the physician was unable to control the bleeding and the impella device was explanted. the physician believed that the patient movement during support was likely to led to bleeding due to lack of proper sedation control. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37072 | 1220648-2021-00754 | 11406219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11406219&pc=OZD | 02/05/2021 | 03/03/2021 |  | 03/03/2021 |  |  | 02/08/2021 | 03/03/2021 |  |  | 09/30/2022 |  | 03/03/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2021006074 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 39 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/08/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the left femoral for myocarditis. it was reported the patient developed massive bleeding from the impella access site. although surgical treatment was performed, hemostasis was difficult, and impella was removed. no further information was provided. | [5 empty] | The impella was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37073 | 1220648-2021-00755 | 11406220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11406220&pc=OZD | 02/10/2021 | 03/03/2021 |  | 03/03/2021 |  |  | 02/10/2021 | 03/03/2021 |  |  | 01/31/2022 |  | 03/03/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1471967 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 50 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [22 empty] | Initial submission |  |  |  |  | 02/10/2021 | * | Invalid/Incomplete |  |  | A patient was admitted to a medical center in (b)(6) suffering from and acute myocardial infarction and vf storm. she had an impella cp and ecmo inserted for support. the ecmo had been placed first and impella followed. the patient was successfully supported for over 5 days when the pump was explanted. months later the patient support was reviewed per a registry enrollment. the registry in japan noted that there was an impella access site bleed that was treated by a manual hold and blood product infusion in excess of 2 units of blood replacement. | [5 empty] | The investigation into the reported, via registry, access site bleeding has completed. the root cause is access management. the pressure applied was not sufficient to achieve hemostasis. due to this, the blood product replacement was necessary. the failure mode will be monitored and trended.
37074 | 1220648-2021-00753 | 11406252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11406252&pc=OZD | 02/05/2021 | 03/03/2021 |  | 03/03/2021 |  |  | 02/05/2021 | 03/03/2021 |  |  | 12/31/2021 |  | 03/03/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1459421 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 87 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/05/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that mild hemolysis occurred after impella placement. attempts were made to resolve hemolysis, but tee did not show the position clearly. the impella was removed based on the overall situation. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37075 | 1220648-2021-00756 | 11406310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11406310&pc=OZD | 02/10/2021 | 03/03/2021 |  | 03/03/2021 |  |  | 02/10/2021 | 03/03/2021 |  |  | 10/31/2020 |  | 03/03/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED INC. |  |  | US |  | 1373741 |  | 005040 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pseudoaneurysm | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 02/10/2021 | * | Invalid/Incomplete |  |  | A publication was found in periodic literature review conducted by abiomed. the publication was of an italian impella support case in which the team accessed the brachial artery percutaneously for impella access. this is not a site noted for access in the ifu, however the team in italy found it to be of adequate diameter and one of the only access site available to them. the femoral arteries were not an option due to pre-existing condition. the impella 2.5 was placed and supported the patient hemodynamically for over 2 hours for the angioplasty. after explant the team observed a brachial pseudoaneurysm. the surgeon treated the brachial-radial artery pseudoaneurysm with a covered stent and the patient survived the procedure. | [5 empty] | The literature piece is all the information that is available for investigation. no product was returned for analysis. no imaging of the anatomy or further details are forthcoming. the event occurred in 2019. the root cause of the pseudoaneurysm is patient condition.
37076 | 1220648-2021-00751 | 11398667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11398667&pc=OZD | 09/18/2020 | 03/02/2021 |  | 03/02/2021 |  |  | 01/26/2021 | 03/02/2021 |  |  | 01/31/2022 |  | 03/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1470969 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed cardiac tamponade during support. as a result, pericardial fluid became stored and required draining via chest tube. there were no issues noted with the patient's right side of the heart. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37077 | 1035166-2021-00026 | 11398854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11398854&pc=DYB | 01/29/2021 | 07/23/2021 |  | 03/02/2021 |  |  | 02/03/2021 | 03/02/2021 |  |  | 10/01/2022 |  | 03/02/2021 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | CR-02148 |  | 0052-3006 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 61 YR | [5 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported there was a 23 fr introducer that was leaking at the scoring next to the orange bezel around the valve. the problem was found in less than 60 seconds during 5.0 axillary insertion procedure on a (b)(6) male patient. another axillary kit was used. the procedure was completed in normal fashion and the rest of the case was unremarkable. no additional medical or surgical intervention was required. the device was discarded. | [5 empty] | The device was used for treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. the investigation will focus on review of product documentation. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. per qa procedure abiomed introducer sheath in-process and final inspections:vacuum leak test: perform pressure and vacuum leak test per procedure (latest revision). this procedure is performed by qa 100%. per instructions for use (ifu): when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete.
37078 | 1220648-2021-00749 | 11387900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11387900&pc=OZD | 02/03/2021 | 03/17/2021 |  | 02/26/2021 |  |  | 02/03/2021 | 02/26/2021 |  |  | 07/31/2022 |  | 02/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1494180 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 02/03/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient had impella cp optical pump inserted in the right femoral artery (rfa). the patient had hematoma on the right groin and as a result less than ten (10) units of packed red blood cells (prbcs) and one (1) unit of platelet was given over four to five days of support. | [5 empty] | The device was returned for evaluation and the purge gap was large confirming the bearing had failed at the end of support on the 13th day. this is beyond the indication of use for the impella cp pump. the repositioning unit was inspected and no damage was found that could have contributed to the access site bleeding. the clinical suggests the patient may have been experiencing disseminated intravascular coagulation (dic) during the case. the root cause of the access site bleeding was unable to be determined, but was likely due to a patient's condition. | The impella cp optical pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37079 | 1220648-2021-00750 | 11387979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11387979&pc=OZD | 02/05/2021 | 03/24/2021 |  | 02/26/2021 |  |  | 02/05/2021 | 02/26/2021 |  |  | 11/30/2022 |  | 02/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021018364 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 53 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 02/05/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old black or african american male patient had impella 5.5 pump inserted through the axillary artery. the pump stopped while patient was on support, patient had to be taken to the operating room (or) for pump removal and replacement through axillary graft. | [5 empty] | The pump was returned for evaluation and no exterior damage was seen. the catheter did not have any kinks, nor did the purge sidearm have any damage. the catheter jacket was opened over the motor end cap to visualize the purge lumen and blood was confirmed. the data logs were reviewed and showed that there were two kink situations that occurred, with the last causing blood to ingress. in each of these events the purge blocked alarm was severely delayed. thus the root cause of the pump stop was blood ingress, exacerbated by a delayed purge blocked alarm. | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37080 | 1220648-2021-00746 | 11378068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11378068&pc=OZD | 05/14/2020 | 02/25/2021 |  | 02/25/2021 |  |  | 01/26/2021 | 02/25/2021 |  |  | 12/31/2021 |  | 02/25/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1459421 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male (b)(6) had impella cp pump inserted in the left femoral for aortic valvuloplasty. it was reported that the patient developed access site bleed. as a result, blood products were administered. amount is unknown. the issue improved after removal. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37081 | 1220648-2021-00747 | 11378079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11378079&pc=OZD | 02/17/2020 | 02/25/2021 |  | 02/25/2021 |  |  | 01/26/2021 | 02/25/2021 |  |  | 09/30/2021 |  | 02/25/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | 22 CHERRY HILL DR |  |  | US |  | 1446706 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that during impella support the patient developed access site bleed. as a result, the patient was administered blood product, unknown amount and surgical sutures were also added. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37082 | 1220648-2021-00748 | 11379440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11379440&pc=OZD | 07/04/2020 | 02/25/2021 |  | 02/25/2021 |  |  | 01/26/2021 | 02/25/2021 |  |  | 04/30/2022 |  | 02/25/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1479524 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a female (b)(6) patient whose age was not provided had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was mentioned that the patient developed bleeding at insertion site during removal the pump for which blood products were administered. no further information was provided. | [5 empty] | The cause of the access site bleeding was not determined. the impella tempubunsho recommends that during impella support the patient should be anticoagulated to where the patient¿s activated clotting time (act) is between 160 and 180 seconds. if the patient¿s act is above this range then there is an increased risk of bleeding at the impella access site. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event.
37083 | 1220648-2021-00731 | 11372600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11372600&pc=OZD | 08/05/2020 | 02/24/2021 |  | 02/24/2021 |  |  | 01/26/2021 | 02/24/2021 |  |  | 04/30/2022 |  | 02/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1479524 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp. during support, the patient developed an access site bleed that required multiple units of blood product delivered. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37084 | 1220648-2021-00732 | 11374098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11374098&pc=OZD | 12/01/2020 | 02/24/2021 |  | 02/24/2021 |  |  | 01/26/2021 | 02/24/2021 |  |  | 01/31/2022 |  | 02/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1471967 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 42 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that the patient developed lower limb ischemia. impella was successfully weaned and explanted. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37085 | 1220648-2021-00733 | 11374124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11374124&pc=OZD | 08/22/2020 | 02/24/2021 |  | 02/24/2021 |  |  | 01/26/2021 | 02/24/2021 |  |  | 12/31/2021 |  | 02/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1459421 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/04/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed access site bleed. the physician stated the bleeding most likely related to the acces site being expanded when the first impella was replaced from the same side. surgical suturing and blood transfusion were performed. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37086 | 1220648-2021-00734 | 11374259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11374259&pc=OZD | 06/23/2020 | 02/24/2021 |  | 02/24/2021 |  |  | 01/26/2021 | 02/24/2021 |  |  | 12/31/2021 |  | 02/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1461513 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral as an aid to the mvr operation. it was reported that the patient developed lower limb ischemia. as a result, blood transfer from pcps blood transfer tube was performed. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37087 | 1220648-2021-00735 | 11374280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11374280&pc=OZD | 05/09/2020 | 02/24/2021 |  | 02/24/2021 |  |  | 01/26/2021 | 02/24/2021 |  |  | 07/31/2020 |  | 02/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1359038 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | HEALTH PROFESSIONAL |  |  |  |  | 41 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 5.0 pump inserted in the right axillary for cardiomyopathy. it was reported the patient developed access site bleed. as a result, blood products were administered. further details were not provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37088 | 1220648-2021-00736 | 11374302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11374302&pc=OZD | 04/04/2020 | 02/24/2021 |  | 02/24/2021 |  |  | 01/26/2021 | 02/24/2021 |  |  | 12/31/2021 |  | 02/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1461513 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 40 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for an unknown indication. it was reported the patient developed an infection at the access site. as a result, antibiotics were administered to treat the infection. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37089 | 1220648-2021-00738 | 11374330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11374330&pc=OZD | 01/27/2021 | 02/24/2021 |  | 02/24/2021 |  |  | 01/27/2021 | 02/24/2021 |  |  | 10/31/2022 |  | 02/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021011959 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | HEALTH PROFESSIONAL |  |  |  |  | 83 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/27/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that after insertion of the impella device, the patient showed signs of lower extremity ischemia. bypass was performed from the healthy side with a reverse sheath. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37090 | 1220648-2021-00739 | 11374340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11374340&pc=OZD | 12/07/2020 | 02/24/2021 |  | 02/24/2021 |  |  | 01/26/2021 | 02/24/2021 |  |  | 04/30/2022 |  | 02/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1480165 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed lower limb ischemia. pcps were also inserted. no additional treatment was provided for lower limb ischemia. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37091 | 1220648-2021-00737 | 11374350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11374350&pc=OZD | 07/13/2020 | 02/24/2021 |  | 02/24/2021 |  |  | 01/26/2021 | 02/24/2021 |  |  | 01/31/2022 |  | 02/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1470941 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old japanese male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that the patient developed access site bleed during support and while explanting the device. as a result, the patient was administered blood products. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37092 | 1220648-2021-00740 | 11374437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11374437&pc=OZD | 01/15/2021 | 02/24/2021 |  | 02/24/2021 |  |  | 01/26/2021 | 02/24/2021 |  |  | 06/01/2022 |  | 02/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1493502 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 5.0 pump inserted (insertion site was not provided) for the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that the patient experienced a papillary muscle rupture due to mr onset during impella 5.0 support. as a result, impella was removed, mvr surgery was performed. the patient is currently in the icu. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37093 | 1220648-2021-00741 | 11374460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11374460&pc=OZD | 09/05/2020 | 02/24/2021 |  | 02/24/2021 |  |  | 01/26/2021 | 02/24/2021 |  |  | 01/31/2022 |  | 02/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1473816 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that the patient developed lower limb ischemia. treatment details were not provide; however, it was noted that the patient limb ischmia improved. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37094 | 1220648-2021-00743 | 11374472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11374472&pc=OZD | 10/05/2020 | 02/24/2021 |  | 02/24/2021 |  |  | 01/26/2021 | 02/24/2021 |  |  | 12/31/2020 |  | 02/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1397692 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted in the right axillary for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that when the impella was inserted from the right femoral artery; the patient developed the patient developed access site bleed. it was difficult to control bleeding at the right femoral artery. as a result, two units of red blood cells were administered, surgical sutures (cigarette sutures), and fixing adjustments. the impella was also as caught in the aortic arch and could not be inserted. as a result, insertion of the impella was switched to the right axillary. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37095 | 1220648-2021-00744 | 11374492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11374492&pc=OZD | 02/02/2021 | 03/12/2021 |  | 02/24/2021 |  |  | 02/02/2021 | 02/24/2021 |  |  | 08/31/2022 |  | 02/24/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1501290 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 02/02/2021 | * | Invalid/Incomplete |  |  | The patient was admitted for a cardioversion/ep study. after cardioverted he exhibited atrial fibrillation with rvr and was found to be in cardiogenic shock. an impella cp was placed for hemodynamic support. the cp supported him well, though on the 2nd day of support the patient was exhibiting signs of hemolysis in labs hemolyzing and hemoglobin levels dropping. the team made choice to infuse 2 units of blood product and leave the impella pump support on. the pump remained for 7 days of support til care was withdrawn and pump was explanted. | [5 empty] | The medical center has forwarded the impella pump for investigation and analysis. the analysis and hemolysis testing is on-going at this time. upon completion of the investigation, a final report will be filed. | Since the original report was filed, the investigation has come to a conclusion. the root cause of the hemolysis was determined to be patient condition. the pump was run on the benchtop and found to have passed abiomed run hemolysis testing. the failure mode will be monitored and trended.
37096 | 1220648-2021-00745 | 11374500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11374500&pc=OZD | 04/16/2020 | 02/24/2021 |  | 02/24/2021 |  |  | 01/26/2021 | 02/24/2021 |  |  | 07/31/2021 |  | 02/24/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1430942 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 84 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) female patient had impella 2.5 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that the patient developed lower limb ischemia. as a result, impella was replaced from the right axillary. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37097 | 1220648-2021-00742 | 11374522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11374522&pc=OZD | 10/05/2020 | 02/24/2021 |  | 02/24/2021 |  |  | 01/26/2021 | 02/24/2021 |  |  | 01/31/2022 |  | 02/24/2021 |  | : TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1476728 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed lower limb ischemia. as a countermeasure, blood was sent to the lower limbs. subsequently, the patient also developed hemolysis. as a result, haptoglobin was administered. the patient also experienced access site bleed for which blood products were administered. impella was successfully weaned and explanted. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37098 | 1220648-2021-00716 | 11366956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11366956&pc=OZD | 08/13/2020 | 02/23/2021 |  | 02/23/2021 |  |  | 01/26/2021 | 02/23/2021 |  |  | 07/31/2020 |  | 02/23/2021 |  | : TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1359038 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 85 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old (b)(6) patient had impella 5.0 pump inserted in the right axillary. the patient had access site bleeding and as a result an unknown amount of blood transfusion was administered. | [5 empty] | Device discarded by customer. the impella 5.0 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37099 | 1220648-2021-00715 | 11366979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11366979&pc=OZD | 09/12/2020 | 02/23/2021 |  | 02/23/2021 |  |  | 01/26/2021 | 02/23/2021 |  |  | 01/31/2022 |  | 02/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1470969 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction (ami). the patient had hemolysis and as a result an unknown amount of red blood cells (rbc) was given. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37100 | 1220648-2021-00718 | 11366991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11366991&pc=OZD | 08/25/2020 | 02/23/2021 |  | 02/23/2021 |  |  | 01/26/2021 | 02/23/2021 |  |  | 08/31/2021 |  | 02/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1436980 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella 2.5 pump. a "persistent" access site bleed developed during support. beyond adjusting the fixation angle, the patient required a vascular suture and 2 units of red bloods cell delivered. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37101 | 1220648-2021-00717 | 11366993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11366993&pc=OZD | 01/26/2021 | 02/23/2021 |  | 02/23/2021 |  |  | 01/26/2021 | 02/23/2021 |  |  | 01/31/2022 |  | 02/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1470968 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report | [8 empty] | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | An (b)(6) year old female patient presented with acute myocardial infarction and cardiogenic shock. the impella cp was selected for hemodynamic support. during support, the patient developed an access site bleed and suspected hemolysis that required an unknown quantity of blood products. the device was removed and the patient was stable. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37102 | 1220648-2021-00719 | 11367091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11367091&pc=OZD | 09/30/2020 | 02/23/2021 |  | 02/23/2021 |  |  | 01/26/2021 | 02/23/2021 |  |  | 07/31/2022 |  | 02/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1489399 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 41 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/05/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp pump. during support, the patient developed an access site. as treatment, 4 units of blood product were delivered and the fixation angle was adjusted. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37103 | 1220648-2021-00720 | 11367100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11367100&pc=OZD | 04/20/2020 | 02/23/2021 |  | 02/23/2021 |  |  | 01/26/2021 | 02/23/2021 |  |  | 12/31/2021 |  | 02/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1463499 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 57 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old patient of, unknown gender, presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp pump. during support, the patient developed an access site bleed. as treatment, multiple units of red blood cells. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37104 | 1220648-2021-00721 | 11367177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11367177&pc=OZD | 05/22/2020 | 02/23/2021 |  | 02/23/2021 |  |  | 01/26/2021 | 02/23/2021 |  |  | 07/31/2021 |  | 02/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1430942 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 49 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella 2.5 pump inserted in the right femoral for myocarditis. it was reported lower limb ischemia was suspected after impella placement. the issue was resolved after removing impella. there was a risk of lower limb ischemia due to the small blood vessel diameter; however, there were no abnormalities in blood flow confirmation, palpation, or visual observation by echo. the sheath was inserted from the opposite side as a precaution. no further information was reported. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37105 | 1220648-2021-00722 | 11367189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11367189&pc=OZD | 05/30/2020 | 02/23/2021 |  | 02/23/2021 |  |  | 01/26/2021 | 02/23/2021 |  |  | 12/31/2021 |  | 02/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1461512 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/05/2021 | * | Invalid/Incomplete |  |  | The complainant reported a patient of unknown age and gender presenting with cardiomyopathy. during support, the patient exhibiting signs of hemolysis via urine color, however, no plasma-free hemoglobin testing was conducted to confirm this. as treatment, haptoglobin was administered. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37106 | 1220648-2021-00724 | 11369359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11369359&pc=OZD | 07/08/2020 | 02/23/2021 |  | 02/23/2021 |  |  | 01/26/2021 | 02/23/2021 |  |  | 12/31/2021 |  | 02/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1461512 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed hemolysis during impella support. as a result, haptoglobin was administered. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37107 | 1220648-2021-00723 | 11369360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11369360&pc=OZD | 11/24/2020 | 02/23/2021 |  | 02/23/2021 |  |  | 01/26/2021 | 02/23/2021 |  |  | 08/31/2022 |  | 02/23/2021 |  | : TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1502522 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/03/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that the patient developed access site bleeding which required 30 units of blood products. subsequently, the patient developed an infection at the access site for which antibacterial agents were administered for sepsis. no further information was provided. | [5 empty] | The impella cp was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37108 | 1220648-2021-00725 | 11369362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11369362&pc=OZD | 06/02/2020 | 02/23/2021 |  | 02/23/2021 |  |  | 01/26/2021 | 02/23/2021 |  |  | 06/30/2021 |  | 02/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1429014 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/05/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting with myocarditis for hemodynamic support using an impella 5.0 pump. during support, the patient developed an access site bleed that required multiple units of a blood product delivered. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37109 | 1220648-2021-00727 | 11369366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11369366&pc=OZD | 07/07/2020 | 02/23/2021 |  | 02/23/2021 |  |  | 01/26/2021 | 02/23/2021 |  |  | 01/31/2022 |  | 02/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1473816 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 80 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [41 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | An (b)(6) year old female patient presented with acute myocardial infarction and cardiogenic shock. the impella cp was selected for hemodynamic support. during support, the patient developed an infection at the access site. the pump was weaned and removed, however, the patient ultimately expired. the impella was alleged to have been a causal or contributory factor in this patient's outcome. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37110 | 1220648-2021-00728 | 11369367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11369367&pc=OZD | 03/22/2020 | 02/23/2021 |  | 02/23/2021 |  |  | 01/26/2021 | 02/23/2021 |  |  | 06/30/2020 |  | 02/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1349665 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 57 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female patient presenting with myocarditis for hemodynamic support using the impella 5.0 pump. during support, the patient developed an access site bleed that required blood product delivery. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37111 | 1220648-2021-00726 | 11369368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11369368&pc=OZD | 05/28/2020 | 02/23/2021 |  | 02/23/2021 |  |  | 01/26/2021 | 02/23/2021 |  |  | 01/31/2022 |  | 02/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1465638 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting for high risk percutaneous coronary intervention. the impella cp was selected for hemodynamic support. during support, the patient endured an access site bleed. multiple units of blood products were delivered as treatment. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37112 | 1220648-2021-00729 | 11369373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11369373&pc=OZD | 10/28/2020 | 02/23/2021 |  | 02/23/2021 |  |  | 01/26/2021 | 02/23/2021 |  |  | 03/31/2022 |  | 02/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1476730 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 84 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/02/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year-old japanese male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that impella was inserted into the lower limbs without angiography. the pump was removed the next day and it was not possible to hear with doppler before the removal. therefore, the lower limb became ischemic. ischemia resolved and no further issues were reported. | [5 empty] | The impella cp pump was discarded by the customer and data logs were not received. an evaluation will be performed and a final yoshiki will be filed.
37113 | 1220648-2021-00730 | 11369377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11369377&pc=OZD | 02/09/2021 | 04/08/2022 |  | 02/23/2021 |  |  | 02/10/2021 | 02/23/2021 |  |  | 01/31/2021 |  | 02/23/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1400797 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemothorax | Cardiac Tamponade | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/10/2021 |  |  |  |  | The complainant reported a 68 year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella 5.0 device. during the removal of the pump, a hemothorax formed. a blood transfusion was provided, however, the patient ultimately expired after a dnr was enacted. the physician alleged that the impella may have caused or contributed to this injury and outcome. | A supplemental mdr is being filed as new information was provided revealing the patient to have endured a cardiac tamponade in relationship to the device. |  |  |  |  | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | This supplemental mdr is being filed to include the report of cardiac tamponade. the device has not been received, therefore, no physician investigation of the product could be completed. if new information is received, a supplemental mdr will be filed.
37114 | 1220648-2021-00712 | 11358181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11358181&pc=OZD | 01/25/2021 | 02/22/2021 |  | 02/22/2021 |  |  | 01/25/2021 | 02/22/2021 |  |  | 12/31/2020 |  | 02/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1400793 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 22 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/25/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 5.0 pump inserted in the right axillary for cardiomyopathy. it was reported that immediately after impella explant, a transesophageal echocardiography revealed mild aortic regurgitation. additionally, blood pressure during extracorporeal circulation was very low. aortic valve replacement using a bioprosthetic valve was performed due to the patient's aortic regurgitation which was expected to worsen after the lvad transplant which the patient was undergoing. observation of the patient's aortic valve after replacement surgery revealed damage that was not inflammatory and suggested that it was caused by mechanical damage from impella. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37115 | 1220648-2021-00713 | 11358255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11358255&pc=OZD | 07/16/2020 | 02/22/2021 |  | 02/22/2021 |  |  | 01/26/2021 | 02/22/2021 |  |  | 01/31/2022 |  | 02/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1465559 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) female patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed limb ischemia when impella was explanted. the patient was able to recover with no issues. treatment details were not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37116 | 1220648-2021-00714 | 11359069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11359069&pc=OZD | 10/17/2020 | 02/22/2021 |  | 02/22/2021 |  |  | 01/26/2021 | 02/22/2021 |  |  | 06/30/2022 |  | 02/22/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP. |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1484369 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction (ami) w/ shock. the patient had hemolysis and as a result an unknown amount of haptoglobin was administered. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37117 | 1220648-2021-00706 | 11347814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11347814&pc=OZD | 02/23/2020 | 02/19/2021 |  | 02/19/2021 |  |  | 01/26/2021 | 02/19/2021 |  |  | 09/30/2021 |  | 02/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1446739 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old (b)(6) male patient had impella cp pump inserted in the left femoral. the patient was bleeding and an unknown amount of red blood cells (rbc) and fresh frozen plasma (ffp) was administered. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37118 | 1220648-2021-00707 | 11347822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11347822&pc=OZD | 01/27/2021 | 02/19/2021 |  | 02/19/2021 |  |  | 01/27/2021 | 02/19/2021 |  |  | 10/31/2022 |  | 02/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2021019945 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 57 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Appropriate Term / Code Not Available | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/27/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient had impella 5.0 pump inserted on the right side. the pump stopped, patients hemodynamics decompensated, and chest compressions were initiated, and decision was made to replace with a new pump. | [5 empty] | Device discarded by customer. the impella 5.0 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37119 | 1220648-2021-00708 | 11347840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11347840&pc=OZD | 01/27/2021 | 03/17/2021 |  | 02/19/2021 |  |  | 01/27/2021 | 02/19/2021 |  |  | 09/30/2022 |  | 02/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2021008603 |  | 0048-0045 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/27/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient had impella cp pump inserted for acute myocardial infarction (ami)/ cardiogenic shock (cgs). the patient had hemolysis that was 55mg/dl and 99mg/dl that was measured within a 24 hour period and the pump was removed to resolve hemolysis. | [5 empty] | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The product was returned for evaluation. product analysis revealed delaminated epotek on the sleeve bearing. data log analysis did not reveal any major positioning or suction issues, although there was a brief period at the beginning of the case where the outflow was close to the valve due to the ventricular placement signal (ps) but still pulsatile motor current. there are no positioning issues present in the logs for the remainder of the case, but it is possible the outlet remained close to the valve. the pump passed hemolysis testing per astm f1841. therefore, the root cause of the hemolysis was due to improper positioning.
37120 | 1220648-2021-00705 | 11349999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11349999&pc=OZD | 01/21/2021 | 04/02/2021 |  | 02/19/2021 |  |  | 01/22/2021 | 02/19/2021 |  |  | 03/31/2021 |  | 02/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1418897 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 49 YR | [5 empty] | Injury | No | Yes | Pumping Stopped | [6 empty] | Insufficient Information | [41 empty] | Initial submission | Followup |  |  |  | 01/22/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 5.0 pump inserted in the ascending aorta. indication was not provided. it was reported the pump stopped. the physician stated that this was due to long term use of the device. no attempts were made to restart the pump because of the patient's do not resuscitate (dnr) policy. | [5 empty] | The impella was returned for analysis. the cause of the pump stop was ingested biomaterial. potential pump stop is listed as a potential effect associated with the use of the impella system. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event. | The impella 5.0 was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37121 | 1220648-2021-00710 | 11350822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11350822&pc=OZD | 01/28/2021 | 02/19/2021 |  | 02/19/2021 |  |  | 01/28/2021 | 02/19/2021 |  |  | 09/30/2022 |  | 02/19/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1504709 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | USER F |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Retroperitoneal Hemorrhage | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 01/28/2021 | * | Invalid/Incomplete |  |  | A patient was admitted for a high risk coronary angioplasty. she had disease in all 3 coronary arteries that would be treated with the hemodynamic support of an impella cp. the cp was seen to be out of optimal left ventricle position. the pump was replaced. during the pump replacement a retroperitoneal bleed was suspected originating at the impella access site. the bleed was treated with 2 units of blood replacement. | [5 empty] | The medical center returned for analysis the impella pump product. the pump was inspected and found to have no abnormalities. the root cause of the patient's bleed was therefore due to patient management. the failure mode will be monitored and trended.
37122 | 1220648-2021-00709 | 11350835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11350835&pc=OZD | 06/16/2020 | 02/19/2021 |  | 02/19/2021 |  |  | 01/26/2021 | 02/19/2021 |  |  | 12/31/2021 |  | 02/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1461513 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old asian male patient had impella cp pump inserted in the right femoral for acute myocardial infarction (ami) with shock. the patient had access site bleeding, hemolysis with haptoglobin administration, and leg ischemia. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37123 | 1220648-2021-00711 | 11350911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11350911&pc=OZD | 01/27/2021 | 03/08/2021 |  | 02/19/2021 |  |  | 01/27/2021 | 02/19/2021 |  |  | 11/30/2022 |  | 02/19/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2021020206 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | USER F |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Retroperitoneal Hemorrhage | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 01/27/2021 | * | Invalid/Incomplete |  |  | A (b)(6) year old patient with multiple cardiac comorbidities was brought to the medical center for a high risk coronary angioplasty, as he was not a repeat surgical candidate. he had an impella cp placed for hemodynamic support during the angioplasty. at the conclusion of the angioplasty the impella was explanted and the access site closed by the manta device. shortly after the device was placed the patient began to complain of discomfort and was found to have a retroperitoneal bleed, thought to have originated from the impella access site. the patient was surgically treated for the bleed. | [5 empty] | Since the original report was filed, the investigation into the access site bleeding from vascular damage has concluded. the complainant did return the impella product for evaluation. no defect to the pump was noted that could have caused the damage. the pump was run and was within normal limits of performance. the cause of bleeding was not determined. further details were not shared to lead to a root cause. | The medical center did not return the impella pump product or introducer sheath for analysis. we are awaiting the product return. should more information be shared that leads to a definitive root cause of the retroperitoneal bleed, a follow-up report will be filed. the failure mode will be monitored and trended.
37124 | 1220648-2021-00704 | 11342791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11342791&pc=OZD | 01/26/2021 | 02/18/2021 |  | 02/18/2021 |  |  | 01/26/2021 | 02/18/2021 |  |  | 04/30/2022 |  | 02/18/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1479524 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [22 empty] | Initial submission |  |  |  |  | 01/26/2021 | * | Invalid/Incomplete |  |  | A (b)(6) year old patient in (b)(6) was admitted in cardiogenic shock suffering from an acute myocardial infarction. he was taken for cardiac catheterization and coronary angioplasty. multiple vessels were intervened upon and for hemodynamic support an impella cp and ecmo were inserted. the impella supported the patient for 5 days when it was explanted due to concerns of thrombus attaching to the impella pump. the patient had suffered a stroke on the 4th day of support. the impella was removed first and the ecmo circuit subsequently. the patient expired. the stroke/cva and outcome of the patient's death were shared with abiomed months after the support and explant via a registry. | [5 empty] | The investigation is on-going at this time. the impella product was discarded in (b)(6) 2020. whether more details are forthcoming is not known at this time. nor are we aware if data log analysis will be possible. the covid conditions have prevented further follow-up. should more details be share that lead to a root cause of the cva/stroke, a final report will be filed.
37125 | 1220648-2021-00700 | 11337658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11337658&pc=OZD | 01/27/2021 | 02/17/2021 |  | 02/17/2021 |  |  | 01/27/2021 | 02/17/2021 |  |  | 12/31/2022 |  |  |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2021019656 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH | [10 empty] | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission |  |  |  |  | 01/27/2021 | * | Invalid/Incomplete |  |  | The medical center admitted a patient for a tavr procedure. while having the valve procedure, the patient decompensated and had an impella cp placed while undergoing cpr. the impella pump supported the patient successfully with some decreased pump flows til explanted approximately 3 hours after insertion. the impella pump was returned for analysis of the low flows. the pump was found to have a damaged impeller. | [5 empty] | When the impella was returned for the analysis of the low flows, the impeller was found to be damaged. the medical center was not aware of the malfunction. the failure mode will be monitored and trended.
37126 | 1220648-2021-00701 | 11338267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11338267&pc=OZD | 01/21/2021 | 02/17/2021 |  | 02/17/2021 |  |  | 01/22/2021 | 02/17/2021 |  |  | 10/31/2022 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021019864 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [41 empty] | Initial submission |  |  |  |  | 01/22/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that at time of impella insertion, an attempt was made to place a 14fr long peel away sheath. the patient's blood iliac vessels were meandering, stenotic, and calcified, and it was difficult to pass through the 7fr sheath. the peel away sheath was inserted after pta was performed. the iliac blood vessel ruptured when the peel away sheath was inserted, it was recovered by placing a viabahn. the impella was inserted again using a long sheath. there is no problem with the operation of the pump and there was no significant change in patient status. the patient's blood vessel condition was poor prior to inserting impella. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37127 | 1220648-2021-00702 | 11338837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11338837&pc=OZD | 12/18/2020 | 02/17/2021 |  | 02/17/2021 |  |  | 01/21/2021 | 02/17/2021 |  |  | 09/30/2022 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1503409 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [41 empty] | Initial submission |  |  |  |  | 01/21/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) female patient had impella cp pump inserted for bts purposes. site of insertion was not provided. it was reported that after the impella device was removed, the patient's left common iliac artery was dissected. the cause is unknown; however, there was calcification of ao and at the dissociation site. after removal, vascular sutures were performed surgically. the patient remained stable after the repair. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37128 | 1220648-2021-00698 | 11331698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11331698&pc=OZD | 01/12/2021 | 02/16/2021 |  | 02/16/2021 |  |  | 01/17/2021 | 02/16/2021 |  |  | 09/30/2022 |  | 02/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2021006029 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/19/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that after the impella device was explanted, there was poor blood flow in the lower limbs. the physician believes that a thrombus may have flown after the removal of impella. as a result, the patient received treatment for lower limb ischemia to restore blood flow to the lower limbs. details were not provided. no further information was provided. | [5 empty] | The impella cp was has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37129 | 1220648-2021-00699 | 11334525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11334525&pc=OZD | 01/22/2021 | 02/16/2021 |  | 02/16/2021 |  |  | 01/22/2021 | 02/16/2021 |  |  | 12/31/2022 |  | 02/16/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2021019745 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 01/22/2021 | * | Invalid/Incomplete |  |  | A patient was admitted in cardiogenic shock suffering from an acute myocardial infarction. when admitted to the cardiac catheterization lab he was found to by hypoxic and in need of cpr. after 30 minutes of cpr the team was able to place an impella cp via the left femoral artery. the cp pump was placed despite observed iliac artery disease and a prosthetic valve, both of which made pump placement a challenge. after team was able to placed the cp to the patient's left ventricle, the team did observe a loss of the pedal pulse in the impella limb. the team placed a femoral-to-femoral external bypass to perfuse the left limb. the pump remained on for support for roughly 24 hours and was then explanted and escalated to an impella 5.5 placed via the axillary artery. | [5 empty] | The medical center discarded the impella product. the investigation into the reported limb ischemia is not possible without product return and analysis. imaging was shared which does reveal existing peripheral arterial disease. patient condition contributed to the limb ischemia that prompted an external bypass to perfuse the limb. the failure mode will be monitored and trended.
37130 | 1220648-2021-00694 | 11321125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11321125&pc=OZD | 01/15/2021 | 03/30/2021 |  | 02/12/2021 |  |  | 01/15/2021 | 02/12/2021 |  |  | 10/31/2022 |  | 02/12/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2021012666 |  | 0048-0014 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 69 YR | [5 empty] | Death | Yes | Yes | Material Separation | Patient Device Interaction Problem | [5 empty] | Mitral Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/15/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting with cardiomyopathy, reduced left ventricular ejection fraction, a history of mitral valve reconstruction, and a history of a cabg. the impella cp smart assist pump was selected for hemodynamic support and inserted. after the pump was initiated, however, repositioning was required due to the device's catheter becoming "kinked" and the motor exhibiting elevated current. upon an attempted pull back of the device, for purposes of repositioning, the physician endured resistance. the pump's pigtail ultimately separated from the device and remained in the patient's anatomy. the physician attributed the resistance to the pigtail becoming stuck on either the neo-chordae or mitral valve ring. the impella was removed, pigtail was retrieved with a surgical snare, and a second impella was inserted. the patient expired shortly after new device insertion. the cause of death was noted as "hemodynamic breakdown" resulting from trauma endured by the aforementioned issue with this pump. | [5 empty] | The impella cp optical device and data logs were obtained from the customer and an investigation was performed. visual inspection found the inflow cage separated from the cannula. analysis of the returned data logs found a drop in motor current and a spike in pump flow, followed by "impella in ventricle" alarms that is believed to be when the damage occurred. based on the information reported and analysis of this device/data logs, we believe the cause of the detachment was related to excessive force resulting from improper positioning. | The impella cp optical was received and an investigation is underway. a supplemental mdr will be filed upon completion of the evaluation.
37131 | 1220648-2021-00695 | 11324081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11324081&pc=OZD | 01/15/2021 | 02/12/2021 |  | 02/12/2021 |  |  | 01/15/2021 | 02/12/2021 |  |  | 10/31/2022 |  | 02/12/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021006672 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | Yes | Material Twisted/Bent | Failure to Eject | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/15/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock. the impella cp optical was selected for support. during an attempted delivery of the device, the catheter kinked and stuck in the patient's vasculature. surgical intervention was required to remove the device via cutdown of the artery. the patient developed an access site bleed and required 4 unit blood transfusion. the physician opted to refrain from inserting a new impella in order to let the patient rest through the night and re-assess the following day. unfortunately, the patient expired the following day, however, this was related to multi-organ failure and not related to the complication endured with this device. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37132 | 1220648-2021-00696 | 11324114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11324114&pc=OZD | 01/10/2021 | 02/12/2021 |  | 02/12/2021 |  |  | 01/18/2021 | 02/12/2021 |  |  | 10/31/2022 |  | 02/12/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST JP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2021012000 |  | 0048-0024-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 01/18/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted laty for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that the patient developed cerebral infarction during impella support and was discontinued. aggressive treatment was not desired by the family and on (b)(6) 2021, the patient expired due to cerebral infarction. per the physician, there may have been a blood clot in the ascending aorta during impella insertion. causal relationship between the device and the adverse event is unknown; however, the impella was operating fine until the patient expired. the physician does not have any negative impression of impella as he believes the patient would have suffered a stroke with or without inserting impella. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37133 | 1220648-2021-00697 | 11324760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11324760&pc=OZD | 01/18/2021 | 02/12/2021 |  | 02/12/2021 |  |  | 01/18/2021 | 02/12/2021 |  |  | 10/31/2022 |  | 02/12/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2021013113 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 59 YR | [5 empty] | Malfunction | No | Yes | Unexpected Shutdown | [6 empty] | Insufficient Information | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 01/18/2021 | * | Invalid/Incomplete |  |  | A (b)(6) year old male with cardiomyopathy was admitted while awaiting heart transplant. he had an extensive cardiac history inclusive of ventricular tachycardia and mitral regurgitation, while awaiting transplant he was placed on an impella 5.5 for hemodynamic support. on the 2nd day of support the impella 5.5 stopped while he was being transported to a ct scan. the pump was replaced urgently in the operating suite. the new 5.5 supported him for 17 days til it was removed for heart transplant. | [5 empty] | The medical center did return the impella 5.5 pump product for analysis. the pump stopped due to a coil short. the coil short will be investigated fully under a capa.
37134 | 1220648-2021-00693 | 11315085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11315085&pc=OZD | 01/18/2021 | 02/11/2021 |  | 02/11/2021 |  |  | 01/18/2021 | 02/11/2021 |  |  | 10/31/2022 |  | 02/11/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021009008 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/18/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old hispanic/latino male patient had impella cp optical pump inserted. the patient had pericardiocentesis in the catheterization laboratory, then taken to operating room (or) for patch of left ventricle (lv) perforation and coronary artery bypass graft (cabg) for the left main (lm) dissection. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37135 | 11301115 | 11301115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11301115&pc=PBL | 01/08/2021 | 01/14/2021 | 01/14/2021 | 02/09/2021 | 01/14/2021 | 02/09/2021 |  | 02/09/2021 |  |  |  |  | 02/09/2021 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003.0814 |  | Yes | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 23360 DA |  |  | 92 |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | During an interventional cardiac procedure, a temporary ventricular assist device (impella) was inserted in the left ventricle of the heart. while in the ventricle, the device was having issues with it's motor and eventually seized. patient was dependent on this device for procedure. broken device was removed and a new one was immediately inserted.
37136 | 1220648-2021-00689 | 11293676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11293676&pc=OZD | 05/19/2020 | 02/08/2021 |  | 02/08/2021 |  |  | 01/13/2021 | 02/08/2021 |  |  |  |  | 02/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/13/2021 | * | Invalid/Incomplete |  |  | We received a publication from said a, sayed l. percutaneous thrombectomy of impella-associated iliac artery thrombosis using the flowtriever system. clin case rep. 2020;8:2645¿2649. a 65 years old male patient had developed acute hypoxemic respiratory failure, coronary angiography revealed total occlusion of the left main artery (lm). the impella cp pump was inserted in the right common femoral artery (cfa). over the following three days of support, the right iliac artery angiogram showed a thrombus located in the external iliac artery just above the tip of the impella sheath, flowtriever system was used to disrupt the thrombus, balloon angioplasty of the right common iliac artery was performed to achieve tamponade and facilitate the removal of the impella. the patient was noted to have signs of limb ischemia. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37137 | 1220648-2021-00690 | 11294733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11294733&pc=OZD | 01/11/2021 | 02/08/2021 |  | 02/08/2021 |  |  | 01/11/2021 | 02/08/2021 |  |  | 09/30/2022 |  | 02/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021007626 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/11/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock. the impella cp optical was selected for hemodynamic support and inserted without issue. during support, the patient endured an access site bleed at the site of the 14 fr introducer. as a note, the physician elected to leave the introducer in place during support, against ifu recommendations. treatment required manual pressure, use of a sandbag, and 6 units of blood products over the course of a weekend. the pump was used until the patient was deemed stable enough for removal. | [5 empty] | The impella cp optical and 14 fr introducer were discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37138 | 1220648-2021-00692 | 11298122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11298122&pc=OZD | 01/11/2021 | 02/08/2021 |  | 02/08/2021 |  |  | 01/15/2021 | 02/08/2021 |  |  | 09/30/2022 |  | 02/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021006133 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | HEALTH PROFESSIONAL |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/15/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs) it was reported impella cp was inserted due to three-vessel lesions from recent mi. the patient's arteriosclerosis was initially strong; however, the pulse of the lower limbs did not touch. as a result, reverse insertion was performed due to lower limb ischemia. it was noted there was blistering due to reperfusion. additionally, the left arm that was used for pci also developed blisters due to hemostasis which indicates the issue may be patient related. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37139 | 1220648-2021-00691 | 11298278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11298278&pc=OZD | 12/24/2020 | 02/08/2021 |  | 02/08/2021 |  |  | 01/15/2021 | 02/08/2021 |  |  | 12/31/2021 |  | 02/08/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1466157 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | HEALTH PROFESSIONAL |  |  |  |  | 17 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/15/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella 5.0 pump inserted in the right axillary for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patients blood vessel became damaged due to strong pull. as a report, re-construction vessel surgery was performed. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37140 | 11285068 | 11285068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11285068&pc=PBL | 12/26/2019 | 12/29/2020 | 12/31/2020 | 02/05/2021 | 12/31/2020 | 02/05/2021 |  | 02/05/2021 |  |  |  |  | 02/05/2021 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 12775 DA |  |  | 102 |  |  | Malfunction | No | Yes | Increase in Pressure | Material Twisted/Bent | [5 empty] | Insufficient Information | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | During a cath lab procedure- patient and cardiogenic shock and impella was used. during placement of the impella into the aorta, the catheter folded and possibly kinked. impella console stated purge pressure high. impella cp later exchanged for new catheter brought by rep. no harm to the patient noted.
37141 | 1220648-2021-00687 | 11279083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11279083&pc=OZD | 01/12/2021 | 02/04/2021 |  | 02/04/2021 |  |  | 01/12/2021 | 02/04/2021 |  |  | 11/30/2022 |  | 02/04/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2021019148 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 44 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 01/12/2021 | * | Invalid/Incomplete |  |  | A (b)(6) female with multiple underlying comorbidities was admitted suffering from an acute myocardial infarction. she had a dissected left main coronary artery that was intervened upon while under support of the impella cp pump. during the intervention she began to deteriorate and the team discovered that she had bleeding to her abdomen due to a bleed from the inferior vena cava. a laparotomy was performed and she had more than 10 units of blood infused. while in the icu for continued monitoring her lower limb exhibited signs of ischemia. due to her previous bleed, the team did not utilize anticoagulation in the impella per ifu recommendations. external bypass was placed to perfuse her limbs. after 26 hours of impella support the pump was explanted. when explanted the team performed balloon angioplasty to the impella site, specifically the common femoral artery, and the limb was improved. | [5 empty] | The investigation into the limb ischemia and subsequent intervention has completed. when the impella was inserted to the femoral artery there was limb ischemia. the ischemia was found in the artery that had been noted to be <4mm on angiogram. the root cause of the ischemia was the condition/size of the native vessel. the failure mode will be monitored and trended.
37142 | 1220648-2021-00688 | 11282268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11282268&pc=OZD | 01/13/2021 | 02/04/2021 |  | 02/04/2021 |  |  | 01/13/2021 | 02/04/2021 |  |  |  |  | 02/04/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | UNKNOWN |  |  |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | LITERATURE | [5 empty] | 55 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [22 empty] | Initial submission |  |  |  |  | 01/13/2021 | * | Invalid/Incomplete |  |  | While conducting a literature review, and abiomed employee discovered a publication of a patient in (b)(6) who diagnosed as covid 19+ while hospitalized for his angina and myocardial infarction (ami). during his hospitalization and treatment for his ami he had an impella cp placed for care escalation. while on impella cp and ecmo support, patient experienced signs of hemolysis which prompted the team to explant the impella. no further details regarding the hemolysis while on impella support are known. at the time of the support, in july 2020, abiomed field personnel were not present for case support due to covid restrictions. if attempted troubleshooting of the impella and/or ecmo was attempted is unknown. | [5 empty] | The complainant's publication is all the details available for analysis and investigation. the product was discarded by the customer and no abiomed personnel were present at the time of support, thus root cause is unable to be determined. the failure mode will be monitored and trended.
37143 | MW5099225 | 11285253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11285253&pc=OZD | 01/18/2021 | 02/02/2021 |  | 02/04/2021 |  |  |  | 02/04/2021 |  |  |  |  | 02/04/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  |  |  | 264771 |  | 0048-003 |  | No | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 56 YR |  |  | 84 |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Death | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Left ventricular apex perforation findings post impella device placement. fda safety report id# (b)(4).
37144 | 1220648-2021-00686 | 11271983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11271983&pc=OZD | 01/08/2021 | 03/02/2021 |  | 02/03/2021 |  |  | 01/08/2021 | 02/03/2021 |  |  | 11/30/2022 |  | 02/03/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021019164 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/08/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old asian male patient had impella cp optical pump inserted. the patient had severe hemolysis and family decided to make do not resuscitate (dnr). | [5 empty] | The device was returned for evaluation and no biomaterial was found. the pump was hemolysis tested and past per astm f1841. thus the root cause of hemolysis was likely due to the patient's condition. | The impella cp optical pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37145 | 1220648-2021-00685 | 11267917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11267917&pc=OZD | 01/08/2021 | 02/02/2021 |  | 02/02/2021 |  |  | 01/08/2021 | 02/02/2021 |  |  | 07/31/2022 |  | 02/02/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1499665 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/08/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white female patient had impella cp optical pump inserted. after implanting the 14fr x 13 cm sheath the flow as occluded to the distal extremity, the patient had lower leg limb ischemia and surgical exploration and repair of the common femoral artery was done by a cardio thoracic surgeon. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37146 | 1220648-2021-00679 | 11258518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11258518&pc=OZD | 12/27/2020 | 02/01/2021 |  | 02/01/2021 |  |  | 01/05/2021 | 02/01/2021 |  |  | 06/30/2022 |  | 02/01/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1486359 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/05/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year-old japanese male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed hemolysis during impella support. the patient presented with signs of red urine. additionally, bilirubin and renal function deteriorated. as a result, the impella device was explanted. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37147 | 1220648-2021-00681 | 11258918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11258918&pc=OZD | 01/06/2021 | 02/01/2021 |  | 02/01/2021 |  |  | 01/06/2021 | 02/01/2021 |  |  | 10/31/2022 |  | 02/01/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021013414 |  | 0048-0024 |  | No | [5 empty] | No | Lay User/Patient |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Perforation of Vessels | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/06/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp optical. shortly after placement of the 14 fr abiomed introducer, a hematoma formed. the pump was inserted, the peel away introducer was removed, and the impella repositioning sheath was advanced. at this point, a vessel perforation was found. as treatment, the impella was removed, a balloon tamponade was placed, and the patient received 2 units of packed red blood cells. an intra aortic balloon pump was later inserted for continued hemodynamic support. the physician attributed the hematoma and vessel perforation to a combination of the patient's body habitus and removal of the peel away sheath. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37148 | 1220648-2021-00680 | 11258925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11258925&pc=OZD | 01/03/2021 | 02/01/2021 |  | 02/01/2021 |  |  | 01/03/2021 | 02/01/2021 |  |  | 07/31/2022 |  | 02/01/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1498942 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | USER F |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/04/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp optical. during support, the patient developed a hematoma and oozing at the impella access site. as treatment, manual pressure and multiple units of blood product were delivered. additionally, the patient exhibited leg ischemia in the impella leg. as treatment, a "fem-fem" bypass was performed. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37149 | 1220648-2021-00682 | 11259768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11259768&pc=OZD | 01/04/2021 | 02/01/2021 |  | 02/01/2021 |  |  | 01/04/2021 | 02/01/2021 |  |  | 07/31/2022 |  | 02/01/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1498322 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Malfunction | Yes | No | Coagulation in Device or Device Ingredient | Use of Device Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 01/04/2021 | * | Invalid/Incomplete |  |  | There was a patient admitted with cardiomyopathy and placed on an impella 5.5 pump for hemodynamic support, while awaiting a lvad for his end stage heart failure. after 22 days of support the pump was exhibiting alarms for high purge pressure. the medical team had accidentally cracked the purge side arm, and attempted to troubleshoot it with a blunt tip needle. the team was aware that this was a temporary fix and plans were made for a pump exchange. the pump stopped while in the operating suite, 4 hours after the damage was done to the purge side arm. the pump was removed and with the pump removal thrombus was found to be within the access/arm. the patient underwent thrombectomy and a new impella was placed. | [5 empty] | The medical facility returned the impella pump products for analysis. the pump had been used beyond indicated use of 14 days per ifu, when the pump stopped on day 23 after the purge sidearm was damaged. there was some thrombotic material observed at explant and thrombectomy were necessary to preserve the limb. the pump stop and thrombotic event will be monitored and trended.
37150 | 1220648-2021-00683 | 11259880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11259880&pc=OZD | 01/06/2021 | 03/02/2021 |  | 02/01/2021 |  |  | 01/06/2021 | 02/01/2021 |  |  | 09/30/2022 |  | 02/01/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021012476 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/06/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male patient had impella cp optical pump inserted in the right leg. the patient had oozing from the insertion site, continued to ooze even after mattress suture was applied to a subcutaneous tissue, patient continued to decline and leg become ischemic, decision was made to escalate to axillary 5.0 pump. the cp optical pump was surgically removed, thrombectomy and fasciotomy was performed to the right lower leg. | [5 empty] | The impella cp optical pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The product was returned for evaluation did not reveal any defects in the guide wire repositioning unit (gwru) valve that would have contributed to the bleeding. the root cause of the access site bleeding was not determined but likely due to a pump patient interaction. the root cause of the limb ischemia was likely due to the patient's condition as the patient was on a high dose of pressors.
37151 | 1220648-2021-00684 | 11262045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11262045&pc=OZD | 01/10/2021 | 02/01/2021 |  | 02/01/2021 |  |  | 01/15/2021 | 02/01/2021 |  |  | 01/31/2021 |  | 02/01/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1404640 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 42 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 01/15/2021 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year-old (b)(6) male patient had impella 5.0 pump inserted in the left axillaty for cardiomyopathy. it was reported that the patients condition was stable; however, hemodynamics deteriorated. a ct scan confirmed a widespread stroke which resulted in patient death due to cerebral infarction. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37152 | 1220648-2021-00678 | 11250555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11250555&pc=OZD | 12/12/2020 | 01/29/2021 |  | 01/29/2021 |  |  | 12/12/2020 | 01/29/2021 |  |  | 10/30/2022 |  | 01/29/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 2021003262 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 83 YR | [5 empty] | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Hospitalization | [28 empty] | Initial submission |  |  |  |  | 12/12/2020 | * | Invalid/Incomplete |  |  | The medical center admitted an (b)(6) year old male suffering from an acute myocardial infarction and cardiogenic shock. when performing the angioplasty the patient began to deteriorate and the team placed an impella cp for hemodynamic support. the patient in addition was found to need right sided heart support and so the team placed an impella rp. during the angioplasty procedure the impella cp was found to exhibit low flows and so the pump was replaced by a new impella cp. when the first impella pump was returned to abiomed for investigation the impeller blade was found to be damaged, and as such a report is filed for the product malfunction. | [5 empty] | The medical center returned the impella cp product for investigation of the low flows. the impeller blade was found to be fractured upon analysis. what interacted with the impeller and caused the damage is not known. the failure mode will be monitored and trended.
37153 | 1220648-2021-00677 | 11244113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11244113&pc=OZD | 01/01/2021 | 01/28/2021 |  | 01/28/2021 |  |  | 01/01/2021 | 01/28/2021 |  |  | 09/30/2022 |  | 01/28/2021 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1501813 |  | 0550-0008 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Difficult to Remove | Difficult or Delayed Separation | [5 empty] | Unspecified Heart Problem | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 01/01/2021 | * | Invalid/Incomplete |  |  | A patient with cardiomyopathy had an impella 5.5 supporting him for 39 days while awaiting heart transplant. when the team attempted to explant the pump in the operating suite in preparation for the heart transplant, they found the pump was unable to be pulled back or advanced. the team manipulated the graft and still found that the pump was unable to be explanted. the team made the choice to cut the impella pump and remove it while still contained within the heart's ventricle. part of the pump was removed along with the heart, while the remaining portion was removed via the graft at the axillary artery. the entire pump was removed and the patient received a new heart. | [5 empty] | The medical center discarded the impella pump product. without product return, a root cause of the pump's inability to be explanted in the normal manner is not determined. the failure mode will be monitored and trended.
37154 | 1220648-2021-00674 | 11237096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11237096&pc=OZD | 01/04/2021 | 01/27/2021 |  | 01/27/2021 |  |  | 01/05/2021 | 01/27/2021 |  |  | 12/31/2021 |  | 01/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1467618 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 20 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/05/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old black or african american female patient had impella cp optical pump inserted. the patient developed thrombus in the left ventricle (lv), the pump was removed, the next day the patient lost pulses and thrombectomy via percutaneous was performed by an interventional radiology. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37155 | 1220648-2021-00675 | 11237859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11237859&pc=OZD | 01/04/2021 | 02/26/2021 |  | 01/27/2021 |  |  | 01/04/2021 | 01/27/2021 |  |  | 01/31/2022 |  | 01/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1468627 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/04/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient had impella cp optical pump inserted for acute myocardial infarction (ami)/ cardiogenic shock (cgs). the patient developed severe hemolysis. as a result, the pump was removed, two (2) units of red blood cells (rbc) was given, and another cp pump was inserted. | [5 empty] | Data logs was reviewed and did not reveal any major positioning issues or suction during the case. hemolysis testing was performed and the pump passed hemolysis testing per fda astm standard f1841. the root cause of the elevated ldh was likely due to poor patient condition. | The impella cp optical pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37156 | 1220648-2021-00676 | 11238079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11238079&pc=OZD | 01/04/2021 | 01/27/2021 |  | 01/27/2021 |  |  | 01/04/2021 | 01/27/2021 |  |  | 07/31/2022 |  | 01/27/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1489510 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 47 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/04/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient had impella cp pump inserted for acute myocardial infarction (ami) with shock. when the repositioning sheath manipulated arterial bleeding was seen coming from the access site, two (2) units of blood products was given, manual pressure was held until a balloon was placed from the opposite side to tamponade the artery from the inside. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37157 | 11228739 | 11228739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11228739&pc=OZD | 01/08/2021 | 01/20/2021 | 01/20/2021 | 01/26/2021 | 01/20/2021 | 01/26/2021 |  | 01/26/2021 |  |  |  |  | 01/26/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | V8.2 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No answer provided | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 16790 DA | [5 empty] | Malfunction | No | Yes | Output Problem | Therapeutic or Diagnostic Output Failure | [5 empty] | Insufficient Information | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The rp impella stopped reporting flows on friday. there was a motor current, however it was not appropriately reporting. sensor was not working. unable to monitor settings. reported to abiomed rep. reported by doctor.
37158 | 1220648-2021-00672 | 11207277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11207277&pc=OZD | 12/25/2020 | 01/20/2021 |  | 01/21/2021 |  |  | 12/25/2020 | 01/21/2021 |  |  | 10/31/2022 |  | 01/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021004172 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/25/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp optical device. the pump was inserted without any reported difficulties. after the patient's procedure, however, his leg became ischemic and required premature impella removal. a new impella was placed in the patient's axillary artery and support continued without issue. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37159 | 1220648-2021-00673 | 11208410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11208410&pc=OZD | 12/11/2020 | 01/21/2021 |  | 01/21/2021 |  |  | 12/28/2020 | 01/21/2021 |  |  | 03/31/2022 |  | 01/21/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1476730 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 80 YR | [5 empty] | Injury | No | Yes | Labelling, Instructions for Use or Training Problem | Failure to Pump | [5 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission |  |  |  |  | 12/28/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that during aortic blockage, the impella sheath was also clamped with a soft jaw. the pump was then restarted before the aortic block was released; however, it did not work. at that time, the "current consumption abnormality alarm" was sounding. when the pump was operated several times, the "current consumption abnormality alarm" disappeared and it started to operate. there was no adverse impact to the patient. after the device was explanted, it was placed in a water basin and it worked without issues. the clinical team later confirmed that the device was turned off while inside the patient. no further information was provided. | [5 empty] | The returned device showed no damage or biomaterial and was able to run without issues. data logs indicate that the clinical team shut the pump off for about 2.5 hours while it was inside the patient. when the clinical team attempted to restart the pump, it would not start. data logs also showed the water basin run that occurred in the field. the pump was able to run, but with erratic motor current, spikes, and periods of complete suction. correlating this with the fact that the pump was off for 2.5 hours while inside the patient, highly suggests the root cause of the pump not able to start was ingested biomaterial from stagnant blood flow. impella ifu warns that the device should never be turned off while inside the patient. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event.
37160 | 1220648-2021-00670 | 11201533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11201533&pc=OZD | 12/21/2020 | 01/20/2021 |  | 01/20/2021 |  |  | 12/21/2020 | 01/20/2021 |  |  | 08/31/2022 |  | 01/20/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1500949 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 88 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Cardiac Output | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/21/2020 | * | Invalid/Incomplete |  |  | The complainant reported an 88 year old male patient presenting with acute myocardial infarction and cardiogenic shock with "severe" aortic stenosis. a balloon aortic valvuloplasty was performed and the patient was placed on the impella cp optical for hemodynamic support. the physician alleged the patient to have decreased cardiac output as the result of having the impella in place. the impella was ultimately removed and the patient's hemodynamics improved. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37161 | 1220648-2021-00671 | 11202100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11202100&pc=OZD | 12/22/2020 | 01/20/2021 |  | 01/20/2021 |  |  | 12/22/2020 | 01/20/2021 |  |  | 08/31/2022 |  | 01/20/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1504139 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/22/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female patient presented as a surgical turn down for high risk percutaneous coronary intervention. the impella cp optical was selected for hemodynamic support and inserted without any reported issue. during support, however, the patient endured a loss of distal pulses. the pump was removed and manual pressure was held, however, pulses did not return. a thrombectomy as performed and blood flow was restored to the leg. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37162 | 1220648-2021-00669 | 11195456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11195456&pc=OZD | 12/23/2020 | 01/19/2021 |  | 01/19/2021 |  |  | 12/23/2020 | 01/19/2021 |  |  | 10/31/2022 |  | 01/19/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021003274 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/23/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient had impella cp optical pump inserted in the right femoral artery (rfa). the pump was removed and the patient had an above the knee amputation. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37163 | 1035166-2021-00006 | 11189294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11189294&pc=DYB | 12/14/2020 | 04/13/2021 |  | 01/18/2021 |  |  | 12/22/2020 | 01/18/2021 |  |  | 07/01/2021 |  | 01/18/2021 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-14501 |  | 0052-3021 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that the patient presented post triple coronary artery bypass and hemodynamics were declining. a severe st elevation was noted, prompting the selection of the impella rp for use in this patient. during acute myocardial infarction/cardiogenic shock procedure, 23 fr introducer was inserted in the patient and a kink was identified midway up the device. the impella rp was inserted through the introducer, but could not successfully cross into the ventricle for proper placement. the impella device was found damaged upon removal and the physician did not feel comfortable removing the introducer, so vascular surgery was consulted. the vascular surgeon removed the introducer and closed the site, however, a perforation was found in the patient's vein was found. the perforation was alleged to have been caused by the kinked introducer. no additional information is available and no other adverse event reported. | [5 empty] | The device was used for treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. however, vascular perforation was repaired and impella was delivered. potential contributing factors to the kinks could be if the sheath was advanced through an area of resistance or patient anatomy could have contributed to the kink.the investigation was focus on a review of product documentation.the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. following controls are in place to mitigate the reported product issue. per procedure abiomed introducer sheath in-process and final inspection: measure dimensions: measure tip i.d. of sheath using appropriate pin gauges to be in range of 0.305" - 0.307" for 23f. visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. per ifu : never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress, follow up report will be submitted if we find any further additional information.
37164 | 1220648-2021-00668 | 11182585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11182585&pc=OZD | 12/18/2020 | 01/15/2021 |  | 01/16/2021 |  |  | 12/18/2020 | 01/16/2021 |  |  | 09/30/2022 |  | 01/16/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021010073 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/18/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting for high risk percutaneous coronary intervention. the impella cp optical was selected for support. shortly after inserting the 14 fr introducer, angiography found an occlusion of blood flow to the patient's lower limb of the leg inserted upon. the physician was concerned for possible damage to the patient's transplanted kidney, thus, the procedure was aborted. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37165 | 1220648-2021-00664 | 11178482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11178482&pc=OZD | 12/07/2020 | 01/15/2021 |  | 01/15/2021 |  |  | 12/18/2020 | 01/15/2021 |  |  | 01/31/2022 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC |  |  | US |  | 1465351 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [41 empty] | Initial submission |  |  |  |  | 12/18/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old japanese male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that during impella support, the patient developed bleeding at the access site. as a result, the patient was administered red blood products, amount was not provided. the bleeding did not subside, and withdrawal was prioritized over further transfusions to maintain impella. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37166 | 1220648-2021-00665 | 11178491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11178491&pc=OZD | 12/18/2020 | 01/15/2021 |  | 01/15/2021 |  |  | 12/18/2020 | 01/15/2021 |  |  | 10/31/2022 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2021011915 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [41 empty] | Initial submission |  |  |  |  | 12/18/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for high-risk percutaneous coronary intervention (hrpci). it was reported the patient developed lower limb ischemia during pump support. the impella device was explanted due to hemodynamics and percutaneous coronary intervention (pci) could be performed safety. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37167 | 1220648-2021-00666 | 11178955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11178955&pc=OZD | 12/18/2020 | 01/15/2021 |  | 01/15/2021 |  |  | 12/18/2020 | 01/15/2021 |  |  | 12/31/2022 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1493300 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [41 empty] | Initial submission |  |  |  |  | 12/18/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) female patient had impella cp pump inserted in the left femoral for post-cardiotomy cardiogenic shock (pccs). it was reported that patient experienced hemolysis during pump support. as a result, the impella device was explanted. per the physician, hemolysis is likely related to impella since bilirubin level decreased from 14 to 7 before and after withdrawal. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37168 | 1220648-2021-00667 | 11178998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11178998&pc=OZD | 12/23/2020 | 01/15/2021 |  | 01/15/2021 |  |  | 12/24/2020 | 01/15/2021 |  |  | 09/30/2022 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET, |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1503408 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [41 empty] | Initial submission |  |  |  |  | 12/27/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old japanese male patient had impella cp pump inserted in the left femoral for high-risk percutaneous coronary intervention (hrpci). it was reported bleeding was noted at the access site when impella was fixed. tobacco anastomosis was performed which subsidized the bleed; however, several hours later the bleeding started again. the patient was administered 2 units of red blood cells. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37169 | 1220648-2021-00663 | 11170597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11170597&pc=OZD | 12/21/2020 | 01/14/2021 |  | 01/14/2021 |  |  | 12/21/2020 | 01/14/2021 |  |  | 09/30/2022 |  | 01/14/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2021007577 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 49 YR | [5 empty] | Malfunction | No | Yes | Unexpected Shutdown | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 12/21/2020 | * | Invalid/Incomplete |  |  | A patient with cardiomyopathy was admitted to the medical center with end stage heart failure. he had a iabp placed but, after a few days discovered that the care had to be escalated to an impella 5.5 while awaiting heart transplant. unfortunately after just 2 hours of support, the pump stopped. the team had seen purge flow and purge pressure alarms prior to the pump stop, as well as motor current changes. the pump was replaced. the second 5.5 pump is still supporting the patient successfully for over 22 days to date. | [5 empty] | The medical center returned the pump product for analysis. the pump had prematurely stopped due to a coil short. the failure mode will be monitored and trended.
37170 | 1220648-2021-00662 | 11166998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11166998&pc=OZD | 12/18/2020 | 03/02/2021 |  | 01/13/2021 |  |  | 12/18/2020 | 01/13/2021 |  |  | 09/30/2021 |  | 01/13/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1452370 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 50 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/18/2020 | * | Invalid/Incomplete |  |  | The complaint reported a (b)(6) year old male patient presenting for high risk percutaneous coronary intervention. the impella cp was selected for support. approximately 2 days later, the decision was made to escalate the patient to greater support with the impella 5.5. when the impella cp was removed, a thrombus causing limb ischemia was identified at the access site (right common femoral artery). a surgical cutdown was performed to remove the clot and the site was closed. | [5 empty] | The impella cp optical was not returned, however, we were sent the 14 fr introducer that was involved in this event. an investigation of device is underway and a supplemental mdr will be filed with the results of our analysis. | The impella cp optical pump was not returned, however, we received the introducer involved in the event. historical and trending documentation identifies ischemia as a patient condition that is not caused by a malfunction with the impella device.
37171 | 1220648-2021-00661 | 11154319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11154319&pc=PYX | 12/14/2020 | 03/03/2021 |  | 01/12/2021 |  |  | 12/14/2020 | 01/12/2021 |  |  | 01/31/2021 |  | 01/12/2021 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1410106 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | Yes | Material Twisted/Bent | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/14/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting post triple coronary artery bypass procedure. the decision was made to insert an impella rp in response to patient's decreased hemodynamics while recovering in the cvicu. during the attempted delivery of the impella, the 23 fr introducer became "severely kinked" midway up the device. the pump was attempted for delivery, but despite multiple attempts, could not successfully place the device past the patient's tricuspid heart valve. the pump was removed, however, surgical removal of the introducer was requested as the physician did not feel comfortable doing so. the surgeon was called and successfully removed the introducer and closed the access site, however, a perforation in the patient's vein was noted as a result of the kink. | [5 empty] | The impella rp was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The introducer was returned. a physical analysis confirmed multiple kinks in the introducer. the cause of the "severely kinked" introducer was related to the patient's condition and improper technique by the user.
37172 | 1220648-2021-00658 | 11147483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11147483&pc=OZD | 12/16/2020 | 02/01/2021 |  | 01/11/2021 |  |  | 12/16/2020 | 01/11/2021 |  |  | 09/30/2022 |  | 01/11/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1497768 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 12/16/2020 | * | Invalid/Incomplete |  |  | A (b)(6) year old patient was transferred to a tertiary medical center in cardiogenic shock suffering from an acute myocardial infarction. an impella cp was placed for hemodynamic support. the team was considering escalating her care to an impella 5.5 or centrimag after 7 days of support. she was exhibiting signs of hemolysis, in the form of dark red urine and hemolyzing labs. the impella cp was removed and a centrimag was placed surgically. in addition the presumed hemolysis was treated with 4 units of blood product replacement. | [5 empty] | The impella pump product has not yet been returned for analysis. the root cause of the hemolysis is not determined. upon investigation being completed, the final report will be filed. | Since the initial report was filed, the investigation has closed. the root cause of the presumed hemolysis, that prompted pump explant, was not determined. the medical center did not return the pump product or the aic data logs for analysis. the failure mode will be monitored and trended.
37173 | 1220648-2021-00659 | 11147504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11147504&pc=OZD | 12/14/2020 | 01/11/2021 |  | 01/11/2021 |  |  | 12/14/2020 | 01/11/2021 |  |  | 09/30/2022 |  | 01/11/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1503426 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | PROTEK DUO LVAD | [22 empty] | Initial submission |  |  |  |  | 12/14/2020 | * | Invalid/Incomplete |  |  | A (b)(6) female patient with a protek duo lvad was admitted with myocarditis. the team placed an impella cp for hemodynamic support in combination with the lvad. after several hours the team observed limb ischemia and removed the impella from the femoral artery, which was thought to be small at just 6mm in diameter. the ischemia was diagnosed by lost pedal pulses and the limb becoming cooler over 1-2 hours. they replaced the pump with an access in the axillary artery. the physician did confirm low native pulsatility, but did not attribute the ischemia to this, but rather to the impella placement. | [5 empty] | The medical facility discarded the impella pump after removal. the request for imaging to evaluate the iliofemoral anatomy was not fulfilled, as no imaging was shared. the root cause of the ischemia is not from the use of impella, but rather due to underling condition and vessel size/diameter. the team had noted there were baseline vascular issues for this patient. the failure mode will be monitored and trended.
37174 | 1220648-2021-00660 | 11147761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11147761&pc=OZD | 01/04/2021 | 01/11/2021 |  | 01/11/2021 |  |  | 01/04/2021 | 01/11/2021 |  |  | 08/31/2022 |  | 01/11/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 1501539 |  | 0048-0014 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 90 YR | [5 empty] | Malfunction | No | Yes | Fracture | No Apparent Adverse Event | [5 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission |  |  |  |  | 01/04/2021 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male patient had impella cp pump inserted for cardiogenic shock (cgs). the pump was returned, visual evaluation found both of the impeller blades were broken. | [5 empty] | Visual inspection found both impeller blades were broken. a percutaneous coronary intervention (pci) guidewire was used to approach coronary arteries, the wire must pass the outflow cage, next to impeller blades. the root cause of the impeller blades breaking is due to the blades having interacted with an object during pci procedure. the guidewire interaction with impella could be the probable root cause but since the guidewire was not returned for evaluation, root cause couldn't be determined.
37175 | 1220648-2021-00656 | 11132650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11132650&pc=OZD | 12/15/2020 | 02/04/2021 |  | 01/07/2021 |  |  | 12/15/2020 | 01/07/2021 |  |  | 09/30/2021 |  | 01/07/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1446479 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/15/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) patient had impella cp pump inserted in the left femoral for myocarditis. it was reported the patient experienced bleeding from impella access site. as a result, the patient was administered blood products, amount unknown. hemostasis was attempted by adding sutures to the skin. it was noted that the patient also experienced bleeding from pcps site and stomach, unrelated to impella. no further information was provided. | [5 empty] | The impella cp was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion. | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. product was not returned.
37176 | 1220648-2021-00655 | 11127000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11127000&pc=OZD | 12/14/2020 | 01/06/2021 |  | 01/06/2021 |  |  | 12/14/2020 | 01/06/2021 |  |  | 08/31/2022 |  | 01/06/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1502456 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/14/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that the patient experienced bleeding from impella access site. as a result, the patient was administered 4 units of red blood cells. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37177 | 1220648-2021-00649 | 11116322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11116322&pc=OZD | 12/02/2020 | 01/05/2021 |  | 01/05/2021 |  |  | 12/11/2020 | 01/05/2021 |  |  | 01/31/2022 |  | 01/05/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1470968 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/11/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for myocarditis. it was reported the patient had oozing and swelling. the following day the physician added sutures to the blood vessel and unknown amount of red count cell, fresh frozen plasma and plasma count were administered, amount unknown. it was noted that the patient was bleeding from both impella access site and extracorporeal membrane oxygenation (ecmo). both impella and ecmo were ultimately explanted and the patient was transferred to another hospital with stable circulation. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37178 | 1220648-2021-00650 | 11116325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11116325&pc=OZD | 12/10/2020 | 01/05/2021 |  | 01/05/2021 |  |  | 12/11/2020 | 01/05/2021 |  |  | 03/31/2022 |  | 01/05/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1477975 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 34 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/11/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) male patient had impella cp pump inserted in the left femoral for myocarditis. it was reported that after replacing the peel-away sheath with a repositioning sheath, there was subcutaneous swelling around the impella access site which became prominent. the physician surgeon sutured the blood vessels in the catheterization room; however, oozing continued. as a result, the patient was administered red count cell and fresh frozen plasmas, amount was not provided. additional sutures will be added if necessary. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37179 | 1220648-2021-00651 | 11116876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11116876&pc=OZD | 11/06/2020 | 01/05/2021 |  | 01/05/2021 |  |  | 11/06/2020 | 01/05/2021 |  |  | 06/30/2022 |  | 01/05/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1492264 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/06/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient had impella cp optical pump inserted in the left femoral. there was persistent groin bleeding, two (2) units of packed red blood cells (prbcs) was administered, and the pump was removed. | [5 empty] | Device discarded by customer. the cause of the access site bleeding was not determined since product was not returned.
37180 | 1220648-2021-00652 | 11116964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11116964&pc=OZD | 11/19/2020 | 01/05/2021 |  | 01/05/2021 |  |  | 12/11/2020 | 01/05/2021 |  |  | 09/30/2021 |  | 01/05/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1444438 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Infarction, Cerebral | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/11/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 5.0 pump inserted in the right axillary for unloading purposes. it was reported that cerebral infarction was confirmed after removal of the impella device. additionally, there was paralysis of the left upper and lower limbs. the doctor diagnosed that the thrombus adhered to the pigtail flew to the cerebral blood vessel from the right inner neck, which may have caused the stroke. furthermore, it was reported that the patient is still in the hcu and is waiting for a transplant with paralysis remaining. however, due to the wide range of infarcts, it is difficult to apply for transplantation. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37181 | 1220648-2021-00653 | 11117005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11117005&pc=OZD | 12/11/2020 | 01/05/2021 |  | 01/05/2021 |  |  | 12/11/2020 | 01/05/2021 |  |  | 09/30/2022 |  | 01/05/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1503408 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 39 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/14/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral artillery for left ventricle unloading without coronary problem on cag. it was reported that after inserting the impella, bleeding did not subside after replacing it with the indwelling sheath. the physician added surgical sutures to the blood vessels. the patient was also administered 500ml of red blood count. the patient also developed hemolysis. the patient was placed on continuous hemofiltration/hemodiafiltration (chdf) to treat the hemolysis. the cause of the hemolysis is unknown. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37182 | 1220648-2021-00654 | 11117023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11117023&pc=OZD | 12/12/2020 | 01/05/2021 |  | 01/05/2021 |  |  | 12/12/2020 | 01/05/2021 |  |  | 07/31/2022 |  | 01/05/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1492164 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 85 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/12/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) patient had impella cp pump for acute myocardial infarction and cardiogenic shock (ami/cgs). insertion site not provided. it was reported the patient developed lower limb ischemia after surgical removal of impella. percutaneous transluminal angioplasty (pta) was performed to release the ischemia. it was noted that the patient had pre-existing antistreptolysin o (aso) disease; therefore, the causal relationship between impella and lower limb ischemia is unknown. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37183 | 1220648-2021-00647 | 11108744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11108744&pc=OZD | 12/09/2020 | 01/04/2021 |  | 01/04/2021 |  |  | 12/09/2020 | 01/04/2021 |  |  | 09/30/2022 |  | 01/04/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1505139 |  | 0550-0008 |  | Yes | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 12/09/2020 | * | Invalid/Incomplete |  |  | The cardiac team escalated the care for a patient with many cardiac comorbidities from an iabp to the impella 5.5. the patient had a surgical cutdown and graft placed to allow for the placement of the 5.5. after 4 days of support the patient's impella 5.5 lost appropriate position and was replaced. at the pump replacement a hematoma evacuation was also done. the hematoma was at the impella site. the second pump successfully supported the patient for another 20 days. no further hematoma care was necessary. | [5 empty] | The complainant did not return for analysis the introducer product or impella pump for analysis. without further clinical details or the product analysis the root cause of the access site bleed and hematoma was not possible to be determined. if more information is shared or product analysis is possible, a final report will be filed.
37184 | 1220648-2021-00648 | 11112010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11112010&pc=OZD | 12/10/2020 | 01/04/2021 |  | 01/04/2021 |  |  | 12/10/2020 | 01/04/2021 |  |  | 08/31/2022 |  | 01/04/2021 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1500627 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Insufficiency | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 12/10/2020 | * | Invalid/Incomplete |  |  | A male patient with a history of aortic valve disease was admitted for a valve replacement. the team found the patient to be in shock post-cardiotomy cardiogenic shock, and so the plan was to place an impella cp for hemodynamic support. the found the that the patient experienced aortic valve insufficiency (ai) when the impella was in position. the echo showed severe ai. the valve was termed "wide open". the team troubleshot by pump positioning changes and found that the ai was not resolvable. the pump was explanted. | [5 empty] | The customer did not return the impella pump for analysis nor did they share any of the imaging that would allow for investigation of the reported ai, such as the echo imaging performed when the pump was in position. to date no root cause has been determined. should more information be shared that leads to a root cause, a final report will be filed.
37185 | 1220648-2020-00644 | 11086982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11086982&pc=OZD | 12/04/2020 | 12/29/2020 |  | 12/29/2020 |  |  | 12/04/2020 | 12/29/2020 |  |  | 12/31/2021 |  | 12/29/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1462414 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Aortic Dissection | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/04/2020 | * | Invalid/Incomplete |  |  | The complainant reported a patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella device. the pump was prepared and inserted without any reported issues. during support, the patient's act was continuously monitored so as to keep within the impella instructions for use recommended range (160-180 seconds). during pump removal, an aortic dissection and bi-lateral ischemia developed. the allegation indicated the impella's pigtail may have been caught on plaque during removal of the device, thus, resulting in the dissection and subsequent occlusion of blood flow within the left and right legs. vascular surgery was required to treat these injuries. | [5 empty] | The device was received and a failure investigation was completed. upon examination of the pump there were no abnormalities observed that could have contributed to the reported vascular damage. we believe the cause of the injury was due to poor patient condition.
37186 | 1220648-2020-00645 | 11087001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11087001&pc=OZD | 12/05/2020 | 12/29/2020 |  | 12/29/2020 |  |  | 12/05/2020 | 12/29/2020 |  |  | 08/31/2022 |  | 12/29/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1496670 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 30 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/05/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) female patient had impella cp pump inserted in the left femoral. indication was not provided. it was reported that there was swelling around the insertion site for which pressure was applied. the patient was administered 6 units of red count cells, 10 units of fresh frozen plasma and 10 units of plasma count. the patient also required surgical repair; details were not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37187 | 1220648-2020-00646 | 11091363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11091363&pc=OZD | 12/07/2020 | 01/28/2021 |  | 12/29/2020 |  |  | 12/07/2020 | 12/29/2020 |  |  | 10/31/2021 |  | 12/29/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1468735 |  | 005042 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 90 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | Perforation | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/07/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient had impella 2.5 pump inserted in the left femoral artery. the patient decompensated quickly with internal bleeding and cardiopulmonary resuscitation (cpr) was initiated, rapid transfusion totaling to four (4) units of packed red bloods cells (prbcs) was given and an unknown about of fresh frozen plasma (ffp) were given. | [5 empty] | The device was returned for evaluation and found to be severely damaged at the outflow cage. there were multiple scratch marks along the yoke and endcap indicating interaction with the guidewire. the issue occurred with an abbott guidewire and the wire ended up getting caught on a contralateral sheath from the impella access site. thus the root cause of the vascular damage was likely due to the sheath manipulation due to improper technique. | Device not returned by customer. the impella 2.5 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37188 | 1220648-2020-00643 | 11083461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11083461&pc=OZD | 12/04/2020 | 01/28/2021 |  | 12/28/2020 |  |  | 12/05/2020 | 12/28/2020 |  |  | 09/30/2022 |  | 12/28/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA RP PUMP |  |  | OZD |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1502413 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/05/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white female patient had impella rp pump inserted in the right femoral vein (rfv). the patient had suspected severe hemolysis and as a result the pump was removed. | [5 empty] | The returned pump had a cut catheter and a detached cannula-outflow likely from the removal process and shipping. the inflow was visually inspected and biomaterial was observed. the root cause of the hemolysis was likely ingested biomaterial. | The impella rp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37189 | 11071799 | 11071799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11071799&pc=OZD | 11/20/2020 | 12/10/2020 | 12/10/2020 | 12/24/2020 | 12/10/2020 | 12/24/2020 |  | 12/24/2020 |  |  |  |  | 12/24/2020 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No | [5 empty] | Invalid/Incomplete | No Information |  |  | User Facility report | [8 empty] | 31025 DA |  |  | 88 |  |  | Injury | Yes | Yes | No Device Output | [6 empty] | Cardiopulmonary Arrest | [12 empty] | Life Threatening | Required Intervention | [27 empty] | Initial submission | Followup | Followup |  |  |  | Y | HOSPITAL |  |  | This patient was on an impella pump. at night, the impella controller sounded an alarm and displayed two error messages: "impella stopped: controller failure" and "impella stopped: retrograde flow", and subsequently continually displayed the message "controller error". as the controller was malfunctioning the p-level automatically dropped to p-0. we confirmed the cm placement of the impella catheter, at felt for a pulse, but without the impella support the patient had lost his pulse and we began a code. during the code and during the display of error messages we spoke with the impella representative on the phone (she was able to remotely see the impella screen to monitor the device in real-time), and at early morning, i subsequently spoke with another person from impella to update her as well. we sequestered the controller and will note for it to be assessed and serviced. | This patient was on an impella pump. at night, the impella controller sounded an alarm and displayed two error messages: "impella stopped: controller failure" and "impella stopped: retrograde flow", and subsequently continually displayed the message "controller error". as the controller was malfunctioning the p-level automatically dropped to p-0. we confirmed the cm placement of the impella catheter, at felt for a pulse, but without the impella support the patient had lost his pulse and we began a code. during the code and during the display of error messages we spoke with the impella representative on the phone (she was able to remotely see the impella screen to monitor the device in real-time), and at early morning, i subsequently spoke with another person from impella to update her as well. we sequestered the controller and will note for it to be assessed and serviced. | This patient was on an impella pump. at night, the impella controller sounded an alarm and displayed two error messages: "impella stopped: controller failure" and "impella stopped: retrograde flow", and subsequently continually displayed the message "controller error". as the controller was malfunctioning the p-level automatically dropped to p-0. we confirmed the cm placement of the impella catheter, at felt for a pulse, but without the impella support the patient had lost his pulse and we began a code. during the code and during the display of error messages we spoke with the impella representative on the phone (she was able to remotely see the impella screen to monitor the device in real-time), and at early morning, i subsequently spoke with another person from impella to update her as well. we sequestered the controller and will note for it to be assessed and serviced.
37190 | 1220648-2020-00635 | 11072398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11072398&pc=OZD | 11/16/2020 | 12/21/2020 |  | 12/24/2020 |  |  | 11/16/2020 | 12/24/2020 |  |  | 01/31/2022 |  | 12/24/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1469503 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Renal Failure | Blood Loss | [10 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 11/16/2020 | * | Invalid/Incomplete |  |  | A patient was admitted in cardiogenic shock suffering from an acute myocardial infarction. when in the emergency room she arrested with ventricular fibrillation and was resuscitated. in the cardiac catheterization suite she was found to have multivessel coronary artery disease and was placed on an impella cp for hemodynamic support during the multivessel angioplasty of her left sided coronary arteries. she was then prepped for transfer to a tertiary center. at this time she was observed to have a small access site hematoma that was successfully expressed and controlled. when transferred to the tertiary center the access site again began to bleed and hemostasis was not achieved without impella pump explant, 2 units of blood products, and the placement of perclose. some weeks later the physician notified abiomed medical affairs that this blood loss contributed to an acute kidney injury. | [5 empty] | The complainant did not return the product for analysis. should more data be shared with abiomed by the customer that leads to the root cause of the bleed, hematoma, and acute kidney injury, a final report will be filed.
37191 | 1220648-2020-00636 | 11065564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11065564&pc=OZD | 11/23/2020 | 08/12/2021 |  | 12/23/2020 |  |  | 11/24/2020 | 12/23/2020 |  |  | 01/31/2022 |  | 12/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1476728 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 51 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 11/24/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old asian male patient had impella cp pump inserted in the right femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had hemolysis worsened after insertion of the pump and as a result six (6) units of red blood cells (rbcs) was administered. | [5 empty] | Abiomed received information that the patient had ischemic cerebrovascular accident (cva). the data logs show periods of low motor current pulsatility and suction and position issue alarms. ultimately the root cause of hemolysis was unable to be determined because the pump was not returned for review. there was a cva reported, but its occurrence in relation to impella is unknown thus the patient injury is deemed unrelated. | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed. if the customer does return the product back, then a follow-up supplemental will be filed.
37192 | 1220648-2020-00637 | 11065659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11065659&pc=OZD | 11/25/2020 | 12/23/2020 |  | 12/23/2020 |  |  | 11/27/2020 | 12/23/2020 |  |  | 04/30/2022 |  | 12/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1479524 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 90 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/27/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old asian male patient had impella cp pump inserted in the right femoral for acute myocardial infarction with shock (ami). the patient had lower limb ischemia and as a result the pump was explanted. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed. if the product is returned back a supplemental will be filed.
37193 | 1220648-2020-00638 | 11067057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11067057&pc=OZD | 12/03/2020 | 01/27/2021 |  | 12/23/2020 |  |  | 12/03/2020 | 12/23/2020 |  |  | 06/30/2022 |  | 12/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1494055 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 42 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/03/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old black or african american female patient had impella 5.0 pump inserted. the pump stopped during use. | [5 empty] | The returned product was visually inspected and no abnormalities were observed. the pump was run in a basin of water and a high motor current pump stop was reproduced. the pump was restarted and the pump activated and ran with normal flows until an abrupt high motor current pump stop occurred. the pump electrical resistances were all within specification. the ct scan revealed that the pump bearing retainer had wear and that the balls were out of position. | The impella 5.0 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37194 | 1220648-2020-00639 | 11067134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11067134&pc=OZD | 12/02/2020 | 01/13/2021 |  | 12/23/2020 |  |  | 12/02/2020 | 12/23/2020 |  |  | 10/31/2022 |  | 12/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2021004536 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/02/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella cp optical pump inserted for high risk percutaneous coronary intervention (hrpci). the red impella plug was leaking blood and purge fluid, a continuous slow leak that resulted in a drop of the patient¿s hemoglobin and hematocrit (h&h) and as a result the patient received one (1) unit of packed red blood cells (prbcs) and one (1) unit of fresh frozen plasma (ffp). | [5 empty] | The impella cp optical pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The cp optical pump was returned for evaluation. a hole was found on the catheter, which is the cause of blood leaked from pump plug. a micro-puncture by closing suture could have damaged the catheter resulted in blood ingress into the catheter and leaked out from pump plug.
37195 | 1220648-2020-00642 | 11068724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11068724&pc=OZD | 12/02/2020 | 12/23/2020 |  | 12/23/2020 |  |  | 12/02/2020 | 12/23/2020 |  |  | 07/31/2022 |  | 12/23/2020 |  | : TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1492164 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/02/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old japanese male patient had impella cp pump inserted in the left ventricle for myocarditis. it was reported that after the peel-away sheath was replaced with an indwelling sheath there was no bleeding at first; however, bleeding began to increase gradually at the impella access site. the bleeding subsided to some extent while the angle was maintained. the patient was administered red blood cells, amount unknown. no further information was provided. | [5 empty] | The impella cp has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37196 | 1220648-2020-00640 | 11068785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11068785&pc=OZD | 12/01/2020 | 02/04/2021 |  | 12/23/2020 |  |  | 12/02/2020 | 12/23/2020 |  |  | 06/30/2022 |  | 12/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1484369 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 83 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/02/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral for high risk percutaneous coronary interventions (hrpci). it was reported that after 24 hours of impella support on p8, a suction alarm suddenly occurred, and the pumps flow rate dropped. an echo confirmed there was no issue with the pumps position or volume. the left and right ventricle conditions were checked, and no issues were noted. the physician suspected a blood clot and the impella was removed and replaced with intra-aortic balloon pump (iabp). a white thrombus was found in the cannula after removal. the patient subsequently experienced hemolysis and a decreased blood pressure due thrombus formation. the patient showed signs and symptoms of dark red urine with a rise in lactate dehydrogenase and potassium after the suction alarm occurred. the physician adjusted the pumps position and ultimately explanted impella. cardiac stimulant was added when the pump flow rate and patients' blood pressure dropped. no further information was provided. | [5 empty] | The impella cp was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion. | The impella cp was returned for analysis. the root cause of the event was ingested biomaterial. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event.
37197 | 1220648-2020-00641 | 11068817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11068817&pc=OZD | 12/01/2020 | 12/23/2020 |  | 12/23/2020 |  |  | 12/01/2020 | 12/23/2020 |  |  | 08/31/2022 |  | 12/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1502456 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 84 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/02/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6)-year-old (b)(6) female patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed lower limb ischemia while on impella support. the physician was unable to hear the blood flow in the lower limbs with doppler and was unable to feel a pause in the limb. a reverse sheath was inserted and femoral to femoral bypass was performed from the contralateral side to treat the ischemia. subsequently, the patient experienced hemolysis. while confirming the pumps position via echo, it was noted that the inlet was located at the a valve. the pumps position was adjusted to correct the issue. however, due to hemodynamic instability, pcps were inserted and the impella device was explanted due to hemolysis. no further information was provided. | [5 empty] | The impella cp has not been received and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37198 | 1220648-2020-00633 | 11047560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11047560&pc=OZD | 11/27/2020 | 12/21/2020 |  | 12/21/2020 |  |  | 11/30/2020 | 12/21/2020 |  |  | 09/30/2022 |  | 12/21/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1503408 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/30/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old asian male patient had impella cp pump inserted in the left femoral for acute myocardial infarction with shock (ami). the patient had access site bleeding and as a result 10 units of red blood cells (rbcs), 6 units of fresh frozen plasma (ffp), and 20 units of platelets concentrate (pc) was administered. | [5 empty] | Device not returned by customer. the impella cp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37199 | 1220648-2020-00634 | 11047731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11047731&pc=OZD | 11/19/2020 | 01/21/2021 |  | 12/21/2020 |  |  | 11/30/2020 | 12/21/2020 |  |  | 05/31/2022 |  | 12/21/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1488353 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 11/30/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella cp optical pump inserted. the patient had massive hemolysis and as a result the pump was removed. | [5 empty] | The returned pump was visually inspected and no biomaterial, delaminated epotek, or other pump damage was observed. the pump was run in a basin of water and operated to specification. the cause of the hemolysis was likely due to the patient's condition since the hemolysis was unable to be reproduced in our evaluation. no blood lab results were provided to confirm the hemolysis. the pump was tested for hemolysis and passed. | The impella cp optical pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37200 | 1220648-2020-00631 | 11028871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11028871&pc=OZD | 11/25/2020 | 12/17/2020 |  | 12/17/2020 |  |  | 11/25/2020 | 12/17/2020 |  |  | 08/31/2022 |  | 12/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1500618 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 11/25/2020 | * | Invalid/Incomplete |  |  | A patient was admitted in cardiogenic shock suffering from an acute myocardial infarction. the patient had arrested multiple times, and so the cardiac catheterization staff performed coronary angioplasty with support of an impella cp. upon transfer to the holding area, prior to icu admission, the team observed a bleed at the impella access site. the team was unable to control the bleed and so the patient was sent to the operating room for vascular repair. in addition, 2 units of blood product were infused. | [5 empty] | The root cause of the vascular damage could not be determined as no product was returned for analysis. the failure mode will be monitored and trended.
37201 | 1220648-2020-00632 | 11029679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11029679&pc=OZD | 11/30/2020 | 12/17/2020 |  | 12/17/2020 |  |  | 11/30/2020 | 12/17/2020 |  |  | 04/30/2021 |  | 12/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1419522 |  | 005048-JP |  | Yes | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 17 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/30/2020 | * | Invalid/Incomplete |  |  | A pediatric patient ((b)(6) yo) in (b)(6) was running a marathon and suffered from ventricular fibrillation. upon admission to the medical center she was diagnosed with cardiomyopathy and placed on ecmo and an impella 2.5 pump for hemodynamic support. while on support the impella access site was found to be bleeding. the bleeding prompted the team to infuse multiple units of red blood cells, fresh frozen plasma, and platelets. after 4 days of support on the impella 2.5 the team decided to escalate her care to the impella 5.0. the patient also had exhibited signs of hemolysis while on the 2.5 pump support. the blood products infused for the access site bleed also benefited the hemolysis. | [5 empty] | The complainant in (b)(6) did not return the impella product for analysis. without hemolysis testing on the benchtop and investigation of the pump, no root cause can be determined for either the hemolysis or access site bleeding. the failure mode will be monitored and trended.
37202 | 1220648-2020-00628 | 11020908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11020908&pc=PYX | 12/05/2020 | 01/13/2021 |  | 12/16/2020 |  |  | 12/05/2020 | 12/16/2020 |  |  | 01/31/2021 |  | 12/16/2020 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED INC. |  |  | US |  | 1402793 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Positioning Problem | Insufficient Information | [5 empty] | Perforation | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 12/05/2020 | * | Invalid/Incomplete |  |  | A patient was admitted with cardiomyopathy and had a vad placed. when the patient continued to decompensate, the team placed an impella rp for right sided heart support. the placement of the rp was done via the 23fr sheath to the femoral vein. placement was complicated. the rp came out of position and the team replaced the sheath to reinsert to the right ventricle. during the pump/sheath replacement the team observed that the sheath was extravascular. the vein had been perforated and required vascular repair. | [5 empty] | The medical center did not return the impella product for analysis. without the product and further clinical details, the root cause of the vascualr damage was not able to be determined. the failure mode will be monitored and trended. | The us complainant has returned the impella pump but, not the introducer sheath. the investigation is ongoing at this time. the vessel perforation investigation to root cause is not completed. when the investigation is done, a final report will be filed.
37203 | 1220648-2020-00629 | 11022844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11022844&pc=OZD | 12/08/2020 | 01/15/2021 |  | 12/16/2020 |  |  | 12/08/2020 | 12/16/2020 |  |  | 08/31/2022 |  | 12/16/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1501563 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/08/2020 | * | Invalid/Incomplete |  |  | The patient was admitted in cardiogenic shock, suffering from an acute myocardial infarction. the medical team placed an impella cp for hemodynamic support. on the second day of support the labs were hemolyzing and the urine color had changed, both were signs of hemolysis, and as such the team attempted to troubleshoot the hemolysis. they repositioned the impella pump multiple times. when the patient ceased to make urine, they removed the impella pump and also began dialysis. | [5 empty] | Since the initial report was filed, the investigation has completed. the pump was returned for analysis as were the data logs. the data logs confirm that the pump was repeatedly out of position, which was the cause of the hemolysis. the pump had biomaterial around the impeller as well. the data logs show the team was attempting to reposition the pump repeatedly, and eventually removed the pump to reduce the hemolysis. the failure mode will be monitored and trended. | The complainant has returned the impella pump and data logs. pump positioning images have been requested but, not received yet for analysis. the analysis is on-going at this time. upon completion of the investigation, a final report will be filed.
37204 | 1220648-2020-00630 | 11022895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11022895&pc=OZD | 12/07/2020 | 12/16/2020 |  | 12/16/2020 |  |  | 12/07/2020 | 12/16/2020 |  |  | 07/31/2022 |  | 12/16/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1497094 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 49 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Great Vessel Perforation | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [22 empty] | Initial submission |  |  |  |  | 12/07/2020 | * | Invalid/Incomplete |  |  | A patient was transferred to a tertiary medical center after being diagnosed with scad (spontaneous coronary artery dissection) at a community hospital. at the community hospital she had an iabp placed for support. upon the admission to the tertiary center the team chose to remove the balloon pump and place an impella cp for escalated care. the pump was placed and position evaluated by echo. when the pump position was seen to be suboptimal, the team attempted to reposition. when the repositioning failed, the pump was removed and replaced. during the pump replacement, the patient began to complain of back pain and difficulty breathing. the team sedated her, placed ecmo for respiratory support, and sent her to the operating suite (or) for further assessment. while in the or an aortic dissection was observed in the descending aorta. the aorta damage was not known to be due to the use of impella, or other devices, and patient condition. the patient expired in the or. | [5 empty] | The complainant discarded the impella pump product. no analysis was possible. the root cause of the aortic perforation is not known at this time. should more data or imaging be shared, that leads to root cause, another report will be filed.
37205 | 1220648-2020-00626 | 11011755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11011755&pc=OZD | 12/07/2020 | 01/19/2021 |  | 12/15/2020 |  |  | 12/07/2020 | 12/15/2020 |  |  | 09/30/2022 |  | 12/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED INC. |  |  | US |  | 1502754 |  | 005040 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Material Twisted/Bent | Positioning Problem | [5 empty] | Perforation of Vessels | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 12/07/2020 | * | Invalid/Incomplete |  |  | A patient in (b)(6) was admitted for a high risk coronary angioplasty with planned impella 2.5 hemodynamic support. the medical team attempted to place the impella to the patient's left ventricle when the pump became kinked while in the aorta, and was unable to pass through the valve. the team was unable to retract the pump and had to call for a surgical removal of the impella pump. the impella pump placement was aborted. | [5 empty] | Since the original report was filed, the investigation has concluded. when the pump was being advanced over the aortic arch and became kinked, and it was unable to be un-kinked. the pump was surgically removed from the patient. the product was not returned for analysis from the eu. the patient had no underlying anatomy that prompted the kink or surgical removal. the event was thought to be due the excessive force asserted by the operating physician. the force caused the cannula to kink. the failure mode will be monitored and trended. | The customer in the (b)(6) has not yet returned the product for analysis. clinical details have not yet been shared regarding the anatomy and what was done to the vessel in the surgical removal of the impella pump. a final report will be filed when more clinical details and product analysis is completed.
37206 | 1220648-2020-00627 | 11013156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=11013156&pc=OZD | 11/20/2020 | 01/05/2021 |  | 12/15/2020 |  |  | 11/20/2020 | 12/15/2020 |  |  | 09/30/2022 |  | 12/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1504947 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Material Puncture/Hole | Material Integrity Problem | [5 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 11/20/2020 | * | Invalid/Incomplete |  |  | A patient was admitted for a high risk coronary angioplasty. the patient had multivessel coronary artery disease of the left main, left anterior descending and circumflex arteries. the team made the choice to place an impella cp for hemodynamic support during the angioplasty. the team observed an access site hematoma at the impella sheath site when completing the interventional case. the team performed contralateral leg balloon angioplasty and a manual compression to the bleeding access site. the access site bleed was closed with these measures and placement of groin closure devices, of angioseal and perclose. when the impella introducer sheath was examined a hole was noted in the middle of the introducer. | [5 empty] | Since the original report was filed, the investigation into the hematoma has concluded. the product was returned and the introducer sheath was found to be kinked. the kind caused the vascular damage. the failure mode will be monitored and trended. | The complainant has not yet returned the impella introducer product for analysis. upon completion of the investigation, a final report will be filed.
37207 | 1220648-2020-00623 | 10994717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10994717&pc=OZD | 11/14/2020 | 12/11/2020 |  | 12/12/2020 |  |  | 11/17/2020 | 12/12/2020 |  |  | 07/31/2022 |  | 12/12/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC |  |  | US |  | 1493300 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Mitral Valve Stenosis | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/17/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted for cardiomyopathy. it was reported that prior to impella insertion the patient had moderate aortic stenosis (as) and calcification of the valve for which the patient did not receive treatment. per the physician, calcium was blown off at the time of insertion or removal and the patient developed cerebral infarction. the presence or absence of a thrombus in the left chamber was confirmed by echo, therefore the possibility was low. treatment details for cerebral infarction were not provided. an echo showed that impella was facing the posterior wall side, and it was suspected that the valve leaflets were suppressed. the patient developed mitral valve stenosis and the impella device was changed to iabp. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37208 | 1220648-2020-00624 | 10994736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10994736&pc=OZD | 11/26/2020 | 12/11/2020 |  | 12/12/2020 |  |  | 11/26/2020 | 12/12/2020 |  |  | 08/31/2022 |  | 12/12/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1502522 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/26/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old (b)(6) male patient had impella cp pump inserted in the left femoral for high risk percutaneous coronary intervention (hrpci). there was a lot of bleeding at the impella insertion site, suture to the vessel was done by a vascular surgeon at the insertion site. hemostasis did not occur so the pump was removed after percutaneous coronary intervention (pci) was completed. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37209 | 1220648-2020-00625 | 10994742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10994742&pc=OZD | 11/24/2020 | 12/11/2020 |  | 12/12/2020 |  |  | 11/25/2020 | 12/12/2020 |  |  | 02/28/2021 |  | 12/12/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA RP PUMP |  |  | OZD |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1412303 |  | 004334 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/25/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old (b)(6) female patient had impella rp pump inserted. the patient had internal bleeding, there was presumed swelling of the leg, abdomen and pelvis area due to impella rp insertion site, and as a result two (2) units of packed red blood cells (prbcs) was administered. | [5 empty] | Device not returned by customer. the impella rp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37210 | 1220648-2020-00622 | 10992164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10992164&pc=OZD | 11/20/2020 | 02/12/2021 |  | 12/11/2020 |  |  | 11/20/2020 | 12/11/2020 |  |  |  |  | 12/11/2020 |  | AUTOMATED IMPELLA CONTROLLER |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1494422 |  | 0042-0040-US |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 85 YR | [5 empty] | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | [12 empty] | Required Intervention | [5 empty] | IMPELLA CP OPTICAL PUMP | IMPELLA CP OPTICAL PUMP | [21 empty] | Initial submission | Followup |  |  |  | 11/20/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) patient had impella cp optical pump inserted for high risk percutaneous coronary intervention (hrpci). the aic failed causing the pump to stop. it took about 5-7 minutes to obtain another aic, during this time, the patient was intubated cardiopulmonary resuscitation (cpr) was performed, once patient was back on support, return of spontaneous circulation (rosc) was obtained. | [5 empty] | After the pump stopped, it was unable to be restarted due to persistence of alarms #1024. after console was powered down and restarted 30 minutes later, alarm #1024 presented immediately after boot-up. analysis on console logs prior to the reported event revealed occurrence of alarm #1024 upon console boot-up on 2020-10-16, which was not reported to abiomed. the console was powered on several times and on occasion had a demo/test pump attached but not ran. after console was brought to danvers for testing, the issue was reproduced: during an overnight run with pump simulator, the simulator stopped after console presented with alarms #6 and #1024 recorded in logs (same sequence as during clinical use on 2020-11-20). after further troubleshooting of the console's components it's been determined that the issue was caused by defective impellatronic board, and specifically a breakdown in communication between two of its microcontrollers (cpld and dsp) located on the impellatronic assembly. after the impellatronic assembly had been replaced, console was tested for over 40 hours, and alarms #6 or #1024 did not reoccur, and console with pump worked within specification. independent investigation of the pump 250544 used during this clinical case, confirmed that the pump operated within specification, and did not cause console malfunction. analysis of the device history review (dhr) documentation of console ic8190 did not indicate any issues related to the field failure. in the course of investigation console logs from the burn-in test in emfg were analyzed and found to contain alarms #6 and #1024. | The aic was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37211 | 1220648-2020-00621 | 10978549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10978549&pc=OZD | 12/07/2020 | 12/09/2020 |  | 12/09/2020 |  |  | 12/07/2020 | 12/09/2020 |  |  | 07/31/2022 |  | 12/09/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED INC. |  |  | US |  | 1484643 |  | 0048-0002 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Blood Loss | [11 empty] | Death | Hospitalization | Required Intervention | [26 empty] | Initial submission |  |  |  |  | 12/07/2020 | * | Invalid/Incomplete |  |  | A patient was admitted to the hospital in cardiogenic shock in (b)(6). an impella cp was placed via the left femoral artery for hemodynamic support during a cardiac angioplasty. prior to transfer to the icu, for continued monitoring, the medical team assessed the patient due to declining arterial pressure. the impella access site was visualized and found to have an arterial damage and associated retroperitoneal bleed. the patient showed signs of hemorrhagic shock. the team attempted to treat the issue with contralateral leg access and stent grafting as well as blood product infusion. unfortunately, the patient expired from the hemorrhagic shock. | [5 empty] | The complainant in the eu did not return for investigation any product nor was imaging shared for analysis. the vessel damage and associated hemorrhagic shock is due to underlying patient condition. the physician had commented that the patient had sclerotic vessels. there was a chance for this complication to occur during any intervention and with any introducer placement to the femoral artery. the failure mode will be monitored and trended.
37212 | 1220648-2020-00620 | 10957974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10957974&pc=OZD | 11/09/2020 | 12/07/2020 |  | 12/07/2020 |  |  | 11/09/2020 | 12/07/2020 |  |  | 07/31/2022 |  | 12/07/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1492164 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/09/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) patient had impella cp pump inserted for acute myocardial infarction and cardiogenic shock (ami/cgs). patient's age and gender were not provided. it was reported the patient experienced access site bleeding which stopped when the catheter was maintained at a certain angle. it was noted that to prevent atelectasis, there was patient movement during support. the patient received sutures at the insertion site and the patient was administered eight units of red blood cells. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37213 | 1220648-2020-00619 | 10929366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10929366&pc=OZD | 11/04/2020 | 12/01/2020 |  | 12/01/2020 |  |  | 11/04/2020 | 12/01/2020 |  |  | 08/31/2022 |  | 12/01/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPOR BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1503542 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Positioning Problem | [6 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 11/04/2020 | * | Invalid/Incomplete |  |  | An elderly patient was admitted with severe multivessel coronary artery disease. the cardiothoracic team turned him down for bypass surgery and so the team determined to do a high risk angioplasty with impella cp support for hemodynamic monitoring. the team placed the cp via the right femoral artery and after angioplasty sent him to the icu for monitoring. while in the icu the patient was rolled for a chest xray and with the movement the introducer sheath was pulled out. the patient had a hematoma develop as blood was lost. the introducer sheath had not been removed prior to the patient's admission to the icu due to vessel anatomy and disease. the team was unable to control the bleed with manual pressure or an occlusion balloon. the team infused 3 units of blood and sent him the operating department, but he expired before hemostasis could be achieved. | [5 empty] | The complainant has not returned any product to date for analysis. should product or data logs be returned and an investigation performed that leads to a root cause of the bleed, a final report will be filed.
37214 | 1220648-2020-00618 | 10919377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10919377&pc=OZD | 10/30/2018 | 11/30/2020 |  | 11/30/2020 |  |  | 11/04/2020 | 11/30/2020 |  |  | 03/31/2020 |  | 11/30/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 1337591 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 85 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Regurgitation | Rupture | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/04/2020 | * | Invalid/Incomplete |  |  | We received a publication from elsevier, 12 august 2020, ¿cardiovascular revascularization medicine¿, regarding a (b)(6) male patient had impella cp pump inserted for support. despite severe aortic stenosis, which is a contraindication for impella cp pump placement; patient experienced mitral chordae tendineae rupture leading to severe mitral regurgitation and treatment was done. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37215 | 1220648-2020-00617 | 10900159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10900159&pc=OZD | 11/02/2020 | 11/24/2020 |  | 11/25/2020 |  |  | 11/02/2020 | 11/25/2020 |  |  | 07/31/2022 |  | 11/25/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1496179 |  | 0550-0008 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 29 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/02/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old black or african american male patient had impella 5.5 pump inserted for myocarditis. the patient had two (2) measurements of pfhg that exceeded 40mg/dl recorded within a 24 hour period that was 398 and 402 and as a result the pump was removed and hemolysis began to resolve immediately following pump replacement. | [5 empty] | Device not returned by customer. the impella 5.5 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37216 | 1220648-2020-00616 | 10900188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10900188&pc=OZD | 11/02/2020 | 11/24/2020 |  | 11/25/2020 |  |  | 11/02/2020 | 11/25/2020 |  |  | 08/31/2022 |  | 11/25/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1496978 |  | 0550-0008 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Coagulation Disorder | Hemostasis | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/02/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male patient had impella 5.5 pump inserted. the patient had coagulopathy after insertion, the major bleeding was in the proximal graft sutures located on the aorta, after three hours and a total of ten (10) units of packed red blood cells (prbcs), fresh frozen plasma (ffp), platelets, cryo, kcentra, and factor vii was given. | [5 empty] | Device discarded by customer. the impella 5.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37217 | 1220648-2020-00615 | 10885162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10885162&pc=OZD | 11/02/2020 | 11/23/2020 |  | 11/23/2020 |  |  | 11/02/2020 | 11/23/2020 |  |  | 07/31/2022 |  | 11/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | UNKNOWN |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 11/02/2020 | * | Invalid/Incomplete |  |  | An impella cp was placed for a patient suffering from an acute myocardial infarction in (b)(6). the patient was seen to have multivessel coronary artery disease and coronary angioplasty was begun. during the time in the icu the patient exhibited signs of hemolysis. the team was unsure of the cause, as it was either an effect from the shock, contrast medium, or impella use. a decision was made to initiate chdf and infuse blood products. after just over 120 hours the team explanted the impella. | [5 empty] | The complainant in (b)(6) did not return the pump for analysis. the data logs were returned and not found to have found to have functioned improperly. there were some suction alarms and some position alarms but, nothing to determine the root cause to be the use of impella. the root cause can not be definitively determined. the failure mode will be monitored and trended.
37218 | 1220648-2020-00611 | 10874900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10874900&pc=OZD | 11/03/2020 | 11/20/2020 |  | 11/20/2020 |  |  | 11/06/2020 | 11/20/2020 |  |  | 06/30/2022 |  | 11/20/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1484369 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/06/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute heart failure. it was reported the patient developed hemolysis shortly after impella was implanted. the patient was administered haptoglobin and transfused blood, exact amount was not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37219 | 1220648-2020-00612 | 10874910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10874910&pc=OZD | 10/31/2020 | 12/29/2020 |  | 11/20/2020 |  |  | 11/03/2020 | 11/20/2020 |  |  | 01/31/2022 |  | 11/20/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1471967 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 41 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 11/03/2020 | * | Invalid/Incomplete |  |  | Additional information was received indicating hemolysis did not improve after additional treatment; however, improvedment was noted after removal of the impella device. | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral because the physician was unable to withdraw the patient from shock state. it was reported the patient developed hemolytic in the urine the day after impella was inserted. this was due to insufficient volume and increase peripheral vascular resistance. the pumps position was adjusted to help alleviate hemolysis. the distance from the a valve to the suction part was 50mm and was adjusted to 38 mm. this did not improve the issue. the patient was placed on chdf because the hemolysis persisted, renal function had deteriorated, and the patient's urine excretion had decreased. although bilirubin was still high. chdf was withdrawn because volume control was possible only by excretion of self-urine. myocardial recovery was noted, and the patient was successfully weaned. haptoglobin administration was also performed. no further information was provided. |  |  |  |  | The impella cp was returned for investigation. delaminated epotek was found around the sleeve bearing, which is the most likely root cause of the hemolysis. data logs did not show any evidence of major positioning or suction alarms that may have contributed to the hemolysis. detail clinical patient information was not provided; therefore, abiomed is unable to determine the exact cause of the hemolysis. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event. | The impella cp was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37220 | 1220648-2020-00613 | 10874937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10874937&pc=OZD | 11/02/2020 | 11/20/2020 |  | 11/20/2020 |  |  | 11/02/2020 | 11/20/2020 |  |  | 12/31/2021 |  | 11/20/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1461512 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 84 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/10/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) female patient had impella cp pump inserted in the left femoral for hemodynamic instability. it was reported that prior to impella placement a transcatheter aortic valve implantation (tavi) was performed. due to the tavi procedure the patient experienced a mitral valve perforation and the papillary muscle ruptured. the patient was transferred to (b)(6) hospital for intubation management. although a mitra clip was implanted, the mitral regurgitation (mr) did not release. due to the patient¿s hemodynamic instability, impella cp was inserted. after returning to the icu, the impella device was explanted due to post-operative bleeding; however, the patient¿s hemodynamics were stable. on the morning of (b)(6) 2020, a ct show extensive cerebral hemorrhage and an infarct lesion. the physician stated that it is unknown whether it started at the time of tavi or where at some other point. information regarding treatment was requested but not provided. | [5 empty] | Hemorrhagic strokes are typically related to uncontrolled hypertension, excessive anticoagulation and antithrombotic regimens, trauma or intra-cranial vascular anomalies, and not the device or the procedure. abiomed does not manufacture anticoagulants. the treating physician is responsible for the total dose of anticoagulants administered to the patient. the event is being reported due to the infarct lesion. the impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37221 | 1220648-2020-00614 | 10874947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10874947&pc=OZD | 11/01/2020 | 12/30/2020 |  | 11/20/2020 |  |  | 11/01/2020 | 11/20/2020 |  |  | 08/31/2022 |  | 11/20/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1496978 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 57 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 11/01/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male patient had impella 5.5 pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs). the pump stopped after several days of support, the patient was stable when this occurred so the pump was removed. | [5 empty] | The data logs were returned and confirmed that there was a sudden rise in purge pressure. the device was returned for evaluation and a severe kink was found directly after the pump plug. the pump ultimately stopped due to blood ingress. | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37222 | 10875173 | 10875173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10875173&pc=OZD | 11/02/2020 | 11/04/2020 | 11/04/2020 | 11/20/2020 | 11/04/2020 | 11/20/2020 |  | 11/20/2020 |  |  |  |  | 11/20/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0550-0008 |  | Yes | [5 empty] | No | No Information |  |  | User Facility report | [14 empty] | Malfunction | No | Yes | Appropriate Term/Code Not Available | [6 empty] | Hemolysis | [12 empty] | Other | [28 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella implanted, multiple suction alarms and significant hemolysis. patient underwent impella replacement in or with significant improvement to hemolysis and impella function.
37223 | 1220648-2020-00608 | 10868088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10868088&pc=OZD | 10/25/2020 | 12/29/2020 |  | 11/19/2020 |  |  | 10/26/2020 | 11/19/2020 |  |  | 04/30/2022 |  | 11/19/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1484055 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 10/26/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient had impella cp optical pump inserted in the right axillary. the patient¿s insertion site was bleeding and as a result a total of five (5) units of packed red blood cells (prbcs) was administered. | [5 empty] | Further communications with the representative was conducted, the physician stated to the rep that he was unsure of the location of the bleeding. following pump explant, it was confirmed that there was no bleeding at the access site, therefore there was no patient injury related to impella. this is not considered a complaint as there was no product malfunction or adverse event related to impella device. | The impella cp optical pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37224 | 1220648-2020-00609 | 10868090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10868090&pc=OZD | 10/25/2020 | 11/19/2020 |  | 11/19/2020 |  |  | 10/25/2020 | 11/19/2020 |  |  | 03/31/2022 |  | 11/19/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1482913 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/25/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white female patient had impella cp optical pump inserted in the left femoral artery (lfa). the patient was being transferred from another hospital, upon arrival to the intensive care unit (icu) the patient had significant oozing at the groin site, a vascular surgeon had to repair the artery, 8-10 units of red blood cells (rbcs) was given, and 3-4 units of platelets was given. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37225 | 1220648-2020-00610 | 10868188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10868188&pc=OZD | 10/15/2020 | 11/19/2020 |  | 11/19/2020 |  |  | 11/18/2020 | 11/19/2020 |  |  | 03/31/2022 |  | 11/19/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1478771 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 71 YR | [5 empty] | Malfunction | No | Yes | Fracture | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission |  |  |  |  | 10/22/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white female patient had impella cp optical pump inserted. the pump was getting lower than expected flows right from the beginning. volume was given but the issue did not resolve. the pump was able to restart three (3) times but there was a continuous suction and flows that would not go over 0.8l. the physician decided to remove the pump. | [5 empty] | The data logs was reviewed and confirmed that there was a spike in motor current, followed by a few seconds later another motor current spiked and a pump stop occurred. the pump was able to be restarted on the third attempt but flow dropped to 0.8l/min and this indicated a potential interaction with the impeller blades. the pump was returned and visual inspection found both impeller blades were broken. this broken blade pattern has been reproduced by inserting a guidewire in the outflow and running the pump. the pump was able to run in the lab with low flows. the root cause of the pump stop is likely due to the single access guidewire or catheter getting caught in the outflow during the single access procedure. this likely stopped the pump temporarily and broke the impeller blades.
37226 | 1220648-2020-00605 | 10853393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10853393&pc=OZD | 10/23/2020 | 12/03/2020 |  | 11/17/2020 |  |  | 10/23/2020 | 11/17/2020 |  |  | 05/31/2022 |  | 11/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1488263 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 10/23/2020 | * | Invalid/Incomplete |  |  | A patient was admitted to the hospital via the emergency room and was assessed. the patient was found to have multivessel coronary artery disease. the team made the choice to place an impella cp for hemodynamic support during the coronary angioplasty procedure. the team found that the access site for the impella pump, which was the right femoral artery, had persistent bleeding. the team attempted to troubleshoot by a purse string suture and pressure dressing, but his lab values of hematocrit and hemoglobin were depressed and so the team explanted the pump. in addition the pump was explanted due to concerns for hemolysis. the labs were hemolyzing and the urine color had changed. | [5 empty] | Since the initial report was filed the investigation into the access site bleeding and presumed hemolysis has concluded. the introducer product was not returned for analysis, and as such the root cause of the bleeding could not be determined. the data logs were returned and revealed good pump positioning with minor suction alarms. the root cause of the hemolysis could not be determined with solely the data logs, and as such no root cause could be determined. the failure mode will be monitored and trended. | The medical center has returned the data logs but, not yet the pump. the pump analysis into the root cause of the access site bleeding and presumed hemolysis has not yet begun. upon completion of the investigation, a final report will be filed.
37227 | 1220648-2020-00606 | 10853402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10853402&pc=OZD | 10/16/2020 | 02/24/2021 |  | 11/17/2020 |  |  | 10/23/2020 | 11/17/2020 |  |  | 07/31/2022 |  | 11/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 1486706 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 10/23/2020 | * | Invalid/Incomplete |  |  | The patient had very difficult and calcified access in the femoral artery. a previous stented iliac artery and restenosis stent was in place. the patient had impella cp pump inserted in the femoral, manta closure device was used, but then late in the night leg amputation was required as the patient had ischemia. | [5 empty] | Section d4 was updated as further information was provided about the pump. the cause of the ischemia was likely the condition or size of the patient's vessel. | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37228 | 1220648-2020-00607 | 10854892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10854892&pc=OZD | 10/26/2020 | 11/17/2020 |  | 11/17/2020 |  |  | 10/26/2020 | 11/17/2020 |  |  | 08/31/2022 |  | 11/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1496978 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 62 YR | [5 empty] | Malfunction | No | Yes | No Apparent Adverse Event | [6 empty] | No Consequences Or Impact To Patient | [41 empty] | Initial submission |  |  |  |  | 10/26/2020 | * | Invalid/Incomplete |  |  | A patient was admitted in cardiogenic shock and had an impella cp placed for hemodynamic support. after one day the team made the decision to escalate his care due to biventricular failure. the team took the patient to the operating department for the placement of the impella rp to support the right heart and to escalate the left side to the impella 5.5. unfortunately, the 5.5 stopped and would not start once placed to the left ventricle. the pump was replaced and support proceeded. | [5 empty] | The medical center returned the pump for analysis. the pump ran on the bench. upon inspection of the pump there was a mark observed indicative of the pump being operated with the guidewire retained in the pump. the root cause of the pump unable to start was the user starting the pump with the gw still inside. this goes against ifu and user training. the failure mode will be monitored and trended.
37229 | 1220648-2020-00602 | 10844896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10844896&pc=OZD | 10/19/2020 | 12/30/2020 |  | 11/16/2020 |  |  | 10/19/2020 | 11/16/2020 |  |  | 11/30/2021 |  | 11/16/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1452994 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 10/19/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 5.0 pump inserted in the right axillary for cardiomyopathy. it was reported that the patient developed a cerebral infarction approximately two 2 days after the insertion of impella. when the device was explanted the patient¿s, condition had improved to the extend where that patient could rehabilitate. the physician noted that at the time of insertion, the bleeding was uncontrollable, and a thrombus had adhered in the inlet at the time of explant. the relationship between impella and cerebral infarction is unknown. | [5 empty] | The impella 5.0 was returned for analysis. there is no evidence to indicate any relationship between the cerebral infarction and the use of impella. the cause of the patient injury was patient condition and the relation to impella was determined to be unrelated. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event. | The impella 5.0 was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37230 | 1220648-2020-00603 | 10845076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10845076&pc=OZD | 10/16/2020 | 11/16/2020 |  | 11/16/2020 |  |  | 10/21/2020 | 11/16/2020 |  |  | 07/31/2022 |  | 11/16/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1493301 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/21/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that bleeding was observed from both inguinal regions of the body during ecpella support. bleeding was noted from both impella and ecmo access site. two stitches were added and blood products were administered, amount was not provided. | [5 empty] | The impella cp has not been return by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37231 | 1220648-2020-00604 | 10845906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10845906&pc=OZD | 10/24/2020 | 11/16/2020 |  | 11/16/2020 |  |  | 10/24/2020 | 11/16/2020 |  |  | 07/31/2022 |  | 11/16/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1493301 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/24/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute heart failure. it was reported the patient developed hemolysis for which chdf was performed. hemolysis was confirmed by blood samples; however, pfhg measurements were not provided. after confirming hemolysis, impella was removed because the condition was stable. the patient had renal failure and is being managed by chdf. the patient was also administered red blood cell, fresh frozen plasma and fluid replacement. the physician stated that the hemolysis may have developed due to the location of the impella. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37232 | 1220648-2020-00600 | 10836247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10836247&pc=OZD | 10/22/2020 | 11/13/2020 |  | 11/13/2020 |  |  | 10/22/2020 | 11/13/2020 |  |  | 09/30/2022 |  | 11/13/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1499125 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 10/22/2020 | * | Invalid/Incomplete |  |  | A patient with postcardiotomy cardiogenic shock (pccs) had an impella 5.5 placed for hemodynamic support. there was difficulty in placement of the 5.5 pump across the valve due to low volume and intermittent valve opening. once placed there was a concern for inadequate volume within the left ventricle, and so volume was delivered. once delivered across the valve and into position there was an observation made of a posterior wall of the left ventricle perforation. the perforation was surgically repaired. the patient survived. | [5 empty] | The medical center returned for analysis the impella pump product. no imaging was returned for analysis, despite request. the root cause of the ventricular perforation was traced to the patient condition and underlying infarction that resulted in ventricle wall rupture.
37233 | 1220648-2020-00601 | 10837573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10837573&pc=OZD | 10/08/2020 | 11/13/2020 |  | 11/13/2020 |  |  | 10/23/2020 | 11/13/2020 |  |  | 02/28/2022 |  | 11/13/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, IN. |  |  | US |  | 1474275 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 81 YR | [5 empty] | Malfunction | No | Yes | No Apparent Adverse Event | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission |  |  |  |  | 10/23/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) years old hispanic/latino male patient had impella cp optical pump inserted independently. the pump had flows between 0.3 l/min at p2 with suction, the pump was returned for investigation. | [5 empty] | The data shows the pump experienced spikes in motor current, drops in motor speed and gradual decrease in flows. the device was returned for evaluation and the impeller was found to be damaged. considering the patient had a transcatheter aortic valve replacement (tavr), the root cause of the low flows was interaction with the device during use.
37234 | 1035166-2020-00074 | 10828657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10828657&pc=DYB | 10/15/2020 | 04/12/2021 |  | 11/12/2020 |  |  | 10/15/2020 | 11/12/2020 |  |  | 04/01/2022 |  | 11/12/2020 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-17047 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN |  |  |  |  | 52 YR | [5 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Myocardial Infarction | Cardiogenic Shock | [40 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that a 14fr introducer appeared not to be sealed (leak). the problem was found during preparation for an acute myocardial infarction (ami)/ cardiogenic shock (cgs) procedure on a (b)(6)- year old male patient. the physicians decided quickly to open another impella set for a new introducer as the patient was in very poor condition. the patient expired at some point during, or after the procedure was completed. | [5 empty] | The device used in treatmeant. one 14f 13 cm long abiomed introducer sheath was returned from the customer with the dilator. there were no other accessories. visible blood was found on the sheath and inside the sideport tubing. upon initial inspection, the hemostatic valve was observed to be damaged. the valve was inspected in more detail under 10x magnification. the damage was found to be a tear that stemmed off of the valve split line. the introducer was manually tested by introducing water through sideport of the sheath using a syringe. the introducer leaked through the valve. the tear in the hemostatic valve is the cause for the leak through the valve. the probable root cause of the tear could be off center dilator/device insertion during use. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspection. no manufacturing defects were found. per procedure adelante s2s introducer sheath in-process and final inspection: vacuum leak test ,-perform pressure and vacuum leak test per procedure.per procedure introducer set, silicone seal slitting once the cutting cycle is complete, the part will be placed on the machine tray, remove the seal from the tray and perform a visual inspection using a 10x microscope. at the start of each work order the first 5 acceptable seals must be provided to qa for acceptance. split seals: if the work order is for split seals prior to providing qa with the 5 samples, visually inspect and pull test the first 3 seals per section seal pull test. if the pull tests are within specification, continue to run the first 5 samples for qa acceptance, visually inspect and submit to qa. production may continue and seals can be "run at risk" while qa inspection is in progress. per ifu precuations: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. directions for use: slowly retract the guidewire and dilator, leaving the sheath in position. the hemostasis valve will reduce the loss of blood and the inadvertent aspiration of air through the sheath. introduce the catheter through the hemostasis valve/sheath and advance it into position. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete.
37235 | 1220648-2020-00599 | 10822191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10822191&pc=OZD | 10/08/2020 | 12/11/2020 |  | 11/11/2020 |  |  | 10/12/2020 | 11/11/2020 |  |  | 08/21/2021 |  | 11/11/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1444442 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | Thrombus | [11 empty] | Death | [28 empty] | Initial submission | Followup |  |  |  | 10/12/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impelled cp pump inserted in the right femoral for myocarditis. it was reported there were no issues during the case, however, when the impelled was explanted it was noted that a thrombus had formed on the impelled side. on (b)(6) 2020, after removal of the impelled, a doppler signal could not be heard from both dorsalis pedis arteries, right after manual compression. cyanosis suddenly appeared from the waist down. a computed tomography (ct) revealed complete occlusion below the iliac artery, there were no findings of occlusion at admission. a fogarty embolectomy catheter was used to remove a blood clot of 10 cm or more. the thrombectomy caused the patient¿s condition to deteriorate with ck rising from the 60s' to 54000s, chdf was introduced to improve renal function. pulse therapy was also performed for myocarditis. the physician stated that patient factors may have attributed to the event. however, since the event occurred immediately after removal of impela, it cannot be said that there is no causal relationship with the impelled. | [5 empty] | The impella cp was returned. no thrombus or defects were found with the pump upon inspection. there is no evidence to suggest the thrombus is related to the impella as it was discovered post impella removal. the cause of the patient injury was patient condition and the relation to impella was determined to be unrelated. no further information was provided. the manufacturer is closing the file on this event. | The impella cpx was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37236 | 1220648-2020-00595 | 10816101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10816101&pc=OZD | 10/18/2020 | 11/10/2020 |  | 11/10/2020 |  |  | 10/19/2020 | 11/10/2020 |  |  | 06/30/2021 |  | 11/10/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1428735 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 74 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | Cardiac Tamponade | [11 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 10/19/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the impella 2.5 was inserted with no issues. there were no vital problems; however, the patient¿s condition worsened around 7:30 on (b)(6) 2020. the patient's blood pressure gradually decreased, and the patient expired. the cause of death was reported as cardiac tamponade. per the physician, the causal relationship is unknown. an echocardiogram showed that there was water effusion the cardiac cavity. the physician performed cpr for cpa in the initial care room at the time of the emergency transport. the physician commented that since there was a pericardial effusion at the time, the patient may have had a cardiac tamponade. there was a dissection which may have widened when the impella was inserted. it was noted that the impella worked without any issues. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37237 | 1220648-2020-00596 | 10816118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10816118&pc=OZD | 10/16/2020 | 11/10/2020 |  | 11/10/2020 |  |  | 10/16/2020 | 11/10/2020 |  |  | 07/31/2022 |  | 11/10/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1490580 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/16/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that after removal of the impella device, there was lower limb ischemia due to thrombus which was prominent on the insertion side. a thrombectomy was performed with a fogarty catheter which improved the patient's condition. the physician noted that prior to impella insertion there was stenosis in the iliac artery, and a plain old balloon angioplasty (poba) was inserted prior to impella insertion. the further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37238 | 1220648-2020-00597 | 10816151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10816151&pc=OZD | 10/15/2020 | 02/25/2021 |  | 11/10/2020 |  |  | 10/15/2020 | 11/10/2020 |  |  | 07/31/2022 |  | 11/10/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1493301 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup | Followup |  |  | 10/15/2020 | * | Invalid/Incomplete |  |  | Additional information was received indicating the hemolysis led to renal failure which required dialysis. | The complainant reported a (b)(6) year-old (b)(6) female patient had impella cp pump inserted in the right femoral for myocarditis. it was reported that the patient developed hemolysis during impella support. as a result, the patient was administered haptoglobin, amount was not provided. there was also worsening kidney function due to hemolysis. |  |  |  |  | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella cp was returned. the cause of the hemolysis was patient condition. the most likely cause of the hemolysis is caused by the underlying patient condition and unrelated to the use of impella. it is rare that hemolysis is caused by improper impella positioning. it is extremely rare that an impella malfunction leads to hemolysis. no further information was provided. the manufacturer is closing the file on this event. | The impella cp was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37239 | 1220648-2020-00598 | 10817958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10817958&pc=PYX | 10/16/2020 | 12/01/2020 |  | 11/10/2020 |  |  | 10/16/2020 | 11/10/2020 |  |  | 10/31/2021 |  | 11/10/2020 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED INC. |  |  | US |  | 1460452 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Perforation of Vessels | [12 empty] | Hospitalization | [5 empty] | IABP | IABP | [21 empty] | Initial submission | Followup |  |  |  | 10/16/2020 | * | Invalid/Incomplete |  |  | A patient was admitted after a cardiac arrest in the community. upon admission to the hospital, the medical team evaluated him and performed coronary angioplasty with stenting in 2 vessels. when his condition continued to deteriorate the team made choice to place an impella rp for right sided heart support, in combination with the intra-aortic balloon pump (iabp) and ventilator, that was supporting him. after 4 days of support the team decided to explant the impella rp. upon explant the team observed what they thought to be venous tissue of approximately 2cm adhered to the pump. the tissue was sent for pathology. to date no vessel repair has been performed to the knowledge of abiomed personnel, nor has the pathology report been shared yet to confirm that the pump caused venous vessel damage. | [5 empty] | Since the initial report was filed, the investigation into the vascular damage has concluded. the pump product and data logs were returned for analysis. the data logs revealed no issues during the support. the pump had a slight amount of damage to the outflow cage, but not enough damage to have caused vessel damage or tear. as such, the root cause of the vessel damage could not be determined. the failure mode will be monitored and trended. | The product has been returned for analysis. the pump data logs have also been returned for analysis. the analysis is on-going at this time. in addition, further clinical details inclusive of the pathology report are not yet relayed to abiomed. upon completion of the investigation, a final report will be filed.
37240 | 1220648-2020-00593 | 10804462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10804462&pc=OZD | 10/16/2020 | 11/06/2020 |  | 11/06/2020 |  |  | 10/16/2020 | 11/06/2020 |  |  | 01/31/2022 |  | 11/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1467005 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/16/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male patient, impella 5.0 was attempted to be inserted via axillary approach, resistance was met, the physician used a 23fr sheath but upon removal a kink and crack at the valve caused about one (1) liter of blood loss. as a result, five (5) units of packed red blood cells (prbcs) was given and axillary insertion was aborted. | [5 empty] | Device discarded by customer. the impella 5.0 pump was not returned by the customer therefore a failure analysis investigation cannot be completed.
37241 | 1220648-2020-00594 | 10804488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10804488&pc=OZD | 10/11/2020 | 11/06/2020 |  | 11/06/2020 |  |  | 10/12/2020 | 11/06/2020 |  |  | 07/31/2022 |  | 11/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1489399 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/12/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) male patient had impella cp pump inserted in the right femoral for post-cardiotomy cardiogenic shock (pccs). it was reported that prior the impella insertion the patient had lower extremity arterial blockage. impella was inserted and soon after cyanosis appeared due to ischemia of the lower limbs. an additional sheath was inserted into both lower limbs to bypass and have blood flow from the healthy side. at the time of the event the patient was high dose of pressor, and act's were managed at 180 seconds. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37242 | 1220648-2020-00592 | 10796088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10796088&pc=OZD | 10/13/2020 | 11/05/2020 |  | 11/05/2020 |  |  | 10/13/2020 | 11/05/2020 |  |  | 06/30/2022 |  | 11/05/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1489494 |  | 0550-0008 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Off-Label Use | Malposition of Device | [5 empty] | Cardiac Perforation | [12 empty] | Hospitalization | Other | Required Intervention | [26 empty] | Initial submission |  |  |  |  | 10/13/2020 | * | Invalid/Incomplete |  |  | A patient was admitted to the medical center 2 weeks prior and had an emergent repair of a ascending aortic aneurysm. the patient was placed on ecmo and an intra aortic balloon pump. her recovery was not possible on those 2 devices and so the team placed an impella 5.5 via the pulmonary artery to unload the right ventricle and hopefully explant the ecmo. the 5.5 is not indicated for right sided support. when attempting to position the 5.5 pump the patient suffered from ventricular perforation. the patient was sent for vascular repair but, after all measures were done, inclusive of tevar (thoracic endovascular aortic repair) and blood product infusion (8 units of red blood cells, 4 ffp, as well as 8 liters of normal saline and 2 liters of plasmalyte plus cellsaver) she was sent to the icu for care withdrawal with the family. | [5 empty] | The medical team did not return the product for analysis. data logs were returned for analysis. according to the medical team, during repositioning of the impella 5.5, the nose cone of the pump perforated the rv. there were several other areas of injury, attributed by the surgeon, due to cardiac massage. the root cause of the ventricle perforation was most likely a rare complication of acute myocardial infarction (ami) resulting in a ventricular wall rupture during the ongoing cardiac massage.
37243 | 1220648-2020-00590 | 10771694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10771694&pc=OZD | 10/09/2020 | 11/02/2020 |  | 11/02/2020 |  |  | 10/09/2020 | 11/02/2020 |  |  | 07/31/2022 |  | 11/02/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1497531 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/09/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) white female patient had impella cp optical pump inserted via the right femoral artery (rfa). the patient computerized tomography (ct) scan revealed bleeding at the insertion site, the pump was removed, the vascular surgeon repaired the site, and an unknown about of packed red blood cells (prbcs) was given along with intravenous fluids (ivf). another pump was inserted for support. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37244 | 1220648-2020-00591 | 10771695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10771695&pc=OZD | 10/08/2020 | 11/16/2020 |  | 11/02/2020 |  |  | 10/08/2020 | 11/02/2020 |  |  | 12/31/2020 |  | 11/02/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA RP PUMP |  |  | OZD |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1401532 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 10/08/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male patient had impella rp pump inserted. the patient had significant bleeding from the insertion site, a clot was possible, so the decision was made to remove the pump, and mattress suture to the vessel was placed by a cardiac surgery physician assistant. | [5 empty] | Correction to b5. describe event or problem: there was no medical intervention or blood products given for the access site bleeding. the suture was done to the access site surface not the vessel. the field lt log was reviewed and the high purge pressure alarms was noticed throughout the run. there was no increases in motor current or pump stops that was observed but pump flow did drop at the end of the run. the returned pump was visually inspected and a deformed outflow cage and a kinked catheter were observed. there was no biomaterial observed. the kinked catheter was cut open and a kinked purge line was observed. the purge line was then cut, and fluid flowed freely through the purge, confirming the kinked purge was the cause of the blockage. there was no issue observed with the repo unit that would relate to bleeding. | The impella cp optical pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37245 | 1220648-2020-00588 | 10758019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10758019&pc=OZD | 09/30/2020 | 10/29/2020 |  | 10/29/2020 |  |  | 10/06/2020 | 10/29/2020 |  |  | 01/31/2022 |  | 10/29/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1474083 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/06/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 5.0 pump inserted in the left axillary for myocarditis. it was reported the patient developed access site bleed while on impella support. the patient was administered 14 units of fresh frozen plasma and 10 units of platelets due to the event and other causes which were not provided. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37246 | 1220648-2020-00589 | 10758061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10758061&pc=OZD | 09/29/2020 | 10/29/2020 |  | 10/29/2020 |  |  | 10/06/2020 | 10/29/2020 |  |  | 03/31/2022 |  | 10/29/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1477975 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 46 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/06/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed hemolysis during impella support. the position of the pump was evaluated by echo and no issues were noted. pfhg measurements were not provided. the patient was administered haptoglobin to resolve the hemolysis. the was improvement in effluent from dialysis treatment; however, there was no improvement in hb reduction. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37247 | 1220648-2020-00587 | 10748504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10748504&pc=OZD | 10/07/2020 | 10/28/2020 |  | 10/28/2020 |  |  | 10/07/2020 | 10/28/2020 |  |  | 01/31/2017 |  | 10/28/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, US |  |  | ABIOMED INC. |  |  | US |  | 1183017 |  | 0048-0003 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | LITERA |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Regurgitation | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 10/07/2020 | * | Invalid/Incomplete |  |  | A publication was discovered that outlined an impella cp support from 2015. the publication detailed a patient in which a male patient admitted in cardiogenic shock and suffering from an acute myocardial infarction was found to have a 100% blocked left anterior descending coronary artery that could not be treated due to poor left ventricle function. the impella cp was placed for hemodynamic support. after 5 days the patient was transferred to a tertiary center for escalated care. the impella was removed at the tertiary center as they planned on either a bypass surgery or a durable vad. shortly after the pump was removed the patient condition deteriorated and the team observed new mitral valve regurgitation (mr) under tee echo. the new mr was attribute to the use of impella. the patient expired during the bypass and mitral valve surgery, due to intrathoracic bleeding. | [5 empty] | The case event took place in 2015. no product or data logs are available for review. no investigation into the mitral regurgitation attributed to the use of impella is possible. the cause is unable to be determined. the failure mode will be trended.
37248 | 1220648-2020-00585 | 10726227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10726227&pc=OZD | 09/28/2020 | 10/23/2020 |  | 10/23/2020 |  |  | 09/28/2020 | 10/23/2020 |  |  | 06/30/2022 |  | 10/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1486359 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | HEALTH PROFESSIONAL |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/28/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old japanese male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that after moving the patient to the intensive care unit (icu) access site bleed at the impella access site was noted. the impella device was weaned and explanted and the patient's hemodynamics remained stable. the access site was settled by pressure hemostasis and did not require sutures. additionally, the patient received 16 units of red cell count and 6 units of fresh frozen plasma. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37249 | 1220648-2020-00586 | 10726270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10726270&pc=OZD | 09/29/2020 | 02/24/2021 |  | 10/23/2020 |  |  | 09/30/2020 | 10/23/2020 |  |  | 07/31/2021 |  | 10/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1432800 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Sepsis | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 09/30/2020 | * | Invalid/Incomplete |  |  | Additional information was provided indicating the patient also developed sepsis which required antibacterial drug administration. | The complainant reported a (b)(6) year-old japanese male patient had impella 2.5 pump inserted in the left femoral for high risk percutaneous coronary intervention (hrpci). it was reported the patient¿s potassium levels became high during impella insertion, as a result continuous hemodiafiltration (chdf was performed. hemolysis was diagnosis based on blood sampling results during pci after insertion of impella. the following day potassium levels dropped; however, hemolytic reaction was remarkable and impella was explanted early. it was noted the patient was on dialysis three times a week and was scheduled for dialysis the date after impella was implanted. no further information was provided. |  |  |  |  | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37250 | 1220648-2020-00584 | 10717882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10717882&pc=OZD | 10/02/2020 | 11/13/2020 |  | 10/22/2020 |  |  | 10/02/2020 | 10/22/2020 |  |  | 07/31/2022 |  | 10/22/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1495607 |  | 0550-0008 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 48 YR | [5 empty] | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Consequences Or Impact To Patient | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 10/02/2020 | * | Invalid/Incomplete |  |  | A patient with a history of cardiogenic shock and coronary angioplasty had an impella 5.5 placed for hemodynamic support. the patient had been on support in weeks prior, and then his condition deteriorated. the impella 5.5 was placed via axillary access, and the patient was sent to the operating room for bypass surgery. after the surgery, the impella pump had a purge fluid leak and stopped. the pump was unable to be re-started and was replaced, as the patient needed continued hemodynamic support. | [5 empty] | To date, no product has been returned for analysis. the data logs are also not yet returned. without product and logs, the root cause of the purge fluid leak and pump stop can not be determined. upon completion of the investigation, a final report will be filed. | Since the initial report was filed the investigation has concluded. the cause of the purge fluid leak was sidearm damage that has an associated capa. the root cause of the pump stopping after this damage was ingested biomaterial. the biomaterial was observed after the pump was explanted, and prior to the explant the patient was not properly anticoagulated. the act was reported to be only 92 at the time of the pump stop, which goes against ifu recommendations.
37251 | 1220648-2020-00582 | 10697375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10697375&pc=OZD | 09/22/2020 | 10/19/2020 |  | 10/19/2020 |  |  | 09/22/2020 | 10/19/2020 |  |  | 04/30/2022 |  | 10/19/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1484961 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | USER F |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | Blood Loss | [11 empty] | Death | Hospitalization | [27 empty] | Initial submission |  |  |  |  | 09/22/2020 | * | Invalid/Incomplete |  |  | A patient with coronary artery disease was transferred into the tertiary center for care escalation and intervention. an impella cp was placed for hemodynamic support via femoral artery access. this placement was done despite knowledge of peripheral arterial disease. the case continued for the angioplasty with coronary atherectomy. after the angioplasty the patient complained of discomfort and began to deteriorate. the team observed a perforation to the right iliac artery and aorta. the impella was removed, a lucas device placed for cpr support and chest compressions, and the team prepped a covered stent to treat the perforation. unfortunately, the patient expired before hemostasis was achieved. | [5 empty] | The medical center discarded the impella product. no analysis is possible. should more details be shared that lead to a root cause of the vessel damage and outcome, a final report will be filed.
37252 | 1220648-2020-00583 | 10697416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10697416&pc=OZD | 09/25/2020 | 10/19/2020 |  | 10/19/2020 |  |  | 09/25/2020 | 10/19/2020 |  |  | 03/31/2022 |  | 10/19/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 1471046 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Mitral Perforation | [11 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 09/25/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old female patient had impella cp pump inserted in the femoral. on the third day of support, the patient experienced leg ischemia and decision was made to explant the pump. after a few minutes, the patient appeared to become very unstable and the mitral valve seemed damaged from the pump removal. the patient expired due critical condition of cardiogenic shock under cardiopulmonary resuscitation (cpr) and severe mitral valve (mv) regurgitation. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37253 | 10689422 | 10689422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10689422&pc=OZD | 09/24/2020 | 10/02/2020 | 10/02/2020 | 10/16/2020 | 10/02/2020 | 10/16/2020 |  | 10/16/2020 |  |  |  |  | 10/16/2020 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA PURGE CASSETTE |  |  | OZD |  |  | 0043-0002 |  |  | ABIOMED |  |  | US | [5 empty] | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | User Facility report | [8 empty] | 26645 DA |  |  | 93 |  |  | Malfunction |  | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | Followup |  |  |  |  | Y | HOSPITAL |  |  | From staff: i was assisting rn in changing the purge cassette on the impella. the purge cassette was changed without difficulty and i was making sure the tubings were secure with adequate slack and the tubing cracked at the yellow connection. rn immediately occluded tubing while i ran to get another purge cassette. i set up the next purge cassette without incident. no harm to the patient. rep from abiomed notified. manufacturer response for impella purge cassette, (brand not provided) (per site reporter). rep from abiomed will pick up the cassette next week. this is a disposable piece of the impella. | From staff: i was assisting rn in changing the purge cassette on the impella. the purge cassette was changed without difficulty and i was making sure the tubings were secure with adequate slack and the tubing cracked at the yellow connection. rn immediately occluded tubing while i ran to get another purge cassette. i set up the next purge cassette without incident. no harm to the patient. rep from abiomed notified. manufacturer response for impella purge cassette, (brand not provided) (per site reporter). rep from abiomed will pick up the cassette next week. this is a disposable piece of the impella.
37254 | 1035166-2020-00061 | 10686905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10686905&pc=DYB | 09/10/2020 | 02/17/2021 |  | 10/15/2020 |  |  | 09/17/2020 | 10/15/2020 |  |  | 09/01/2021 |  | 10/15/2020 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-15000 |  | 0052-3021 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [11 empty] | Malfunction | Yes | Yes | Deformation Due to Compressive Stress | No Apparent Adverse Event | [5 empty] | Myocardial Infarction | Cardiogenic Shock | [40 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The patient was presented to the physician with acute myocardial infraction and cardiogenic shock. md attempting to place rp, tried to get rp thru sheath but then it kinked and wouldn't go further. new rp pulled and attempted unsuccessfully. pt coded and expired while trying to place rp. the rep did return impella pump product, but no sheath. cause of the patient death is unknown. no additional information is available. | [5 empty] | The following sections were updated in follow-up 1: h11: h11: corrected data - h6 medical device problem code corrected to 3189. original code (2889) as inadvertently reported in error, correct code is 3189. the device was not returned for analysis. no adverse event has been reported. this medwatch report has been filed for administration purposes because patient expired during procedure. the investigation was focused on the review of product documentation. inspection procedures require any oscor product to pass all in-process and qa final inspection before shipping to the customer. per procedure abiomed introducer sheath in-process and final inspection, the devices in this lot were inspected as follows: 9.2.2 visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. this is performed by qa on the first 5 consecutive properly tipped parts. per section 12.2, visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. this is performed by qa on 100% of the sheaths. the instructions for use (ifu) informs the user: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete.
37255 | 1220648-2020-00580 | 10673220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10673220&pc=OZD | 09/18/2020 | 10/13/2020 |  | 10/13/2020 |  |  | 09/18/2020 | 10/13/2020 |  |  | 07/31/2022 |  | 10/13/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1495130 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | USER F |  |  |  |  | 59 YR | [5 empty] | Malfunction | No | Yes | Insufficient Information | [6 empty] | Not Applicable | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 09/18/2020 | * | Invalid/Incomplete |  |  | A patient with cardiomyopathy was to have an impella 5.5 placed in the operating room. the pump immediately stopped upon insertion. this stoppage prompted the team to replace the pump. the team placed a second 5.5 pump which supported the patient successfully until a lvad was placed on the 7th day of support. | [5 empty] | The medical center returned the impella 5.5 pump for analysis. the pump was run on the benchtop. the pump was able to start and there were no defects or malfunctions found. the pump stop could not be reproduced. the medical center has yet to return the pump data logs for review. should the log analysis help determine a root cause of the pump stop, a supplemental report will follow. to date the root cause is unable to be determined.
37256 | 1220648-2020-00581 | 10673230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10673230&pc=OZD | 09/19/2020 | 10/13/2020 |  | 10/13/2020 |  |  | 09/19/2020 | 10/13/2020 |  |  | 07/31/2022 |  | 10/13/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP 550 SET, EU |  |  | ABIOMED INC. |  |  | US |  | 1495512 |  | 0550-0008 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Use of Device Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 09/19/2020 | * | Invalid/Incomplete |  |  | A patient transferred to the tertiary medical center in germany in cardiogenic shock. the patient had depressed right and left sided heart function with diminished ejection fractions. an impella 5.5 was placed for hemodynamic support. on the second day of support the patient was exhibiting signs of hemolysis with elevated plasma free hemoglobin levels. the pump position was assessed and the pump was chosen to be replaced. a second 5.5 pump was placed. the second impella supported the patient successfully til the patient expired in multi-organ failure 10 days later. | [5 empty] | The investigation is ongoing at this time. the product was discarded, however requests for data logs and further clinical details are expected. upon completion of the investigation, should a root cause of the hemolysis be determined, a final report will be filed.
37257 | 1220648-2020-00579 | 10641681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10641681&pc=OZD | 09/10/2020 | 10/07/2020 |  | 10/07/2020 |  |  | 09/10/2020 | 10/07/2020 |  |  | 06/30/2022 |  | 10/07/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1491183 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | HEALTH PROFESSIONAL | [5 empty] | 46 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 09/10/2020 | * | Invalid/Incomplete |  |  | A combative and noncompliant patient was admitted suffering from an acute myocardial infarction. the patient had an ejection fraction of 10% and left anterior descending coronary artery disease. the team placed an impella cp for support via the left femoral artery for hemodynamic support during the cardiac catheterization and associated coronary angioplasty. after placement of the impella the limb showed no distal flow. to reduce the limb ischemia the team placed and antegrade sheath for external bypass to the limb. immediately flow restored and pulse returned to the distal anatomy. five days later the impella pump was successfully weaned and explanted. no lasting harm or injury to the limb from the brief limb ischemia. | [5 empty] | The medical facility discarded the impella pump product at explant. no root cause can as of yet be determined. should the facility share more clinical details regarding the event that leads to a cause of the limb ischemia, a supplemental report will be filed.
37258 | 1220648-2020-00572 | 10634039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10634039&pc=OZD | 09/08/2020 | 10/06/2020 |  | 10/06/2020 |  |  | 09/09/2020 | 10/06/2020 |  |  | 01/31/2022 |  | 10/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1465559 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 86 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/09/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that while on impella support there was side leak bleeding at the impella access site. sutures were added to stop the bleeding and red blood cells were administered; exact amount was not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37259 | 1220648-2020-00573 | 10634046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10634046&pc=OZD | 09/07/2020 | 10/06/2020 |  | 10/06/2020 |  |  | 09/07/2020 | 10/06/2020 |  |  | 04/30/2022 |  | 10/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1480165 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/07/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that while the patient was being supported with v-a ecmo+impella there was bleeding from the impella access site. as a result, the patient was administered red blood cells, fresh frozen plasma and palette count, amount was not provided. the patient's condition improved after blood products were administered. cardiac function significantly improved, and no further support was required. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37260 | 1220648-2020-00574 | 10634053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10634053&pc=OZD | 09/12/2020 | 10/06/2020 |  | 10/06/2020 |  |  | 09/12/2020 | 10/06/2020 |  |  | 04/30/2022 |  | 10/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED INC. |  |  | US |  | 1475900 |  | 005062 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 09/12/2020 | * | Invalid/Incomplete |  |  | A patient had a cardiac catheterization done as an outpatient and was found to have multivessel disease. the patient was a surgical turndown and so a high risk angioplasty was planned with impella support. an impella 5.0 was placed via the axillary. the impella access site was bleeding and the patient required 5 units of blood replacement during the 3 days of support. the patient was also noted to have inflammation and there was ecchymosis into the torso and right arm. after the 3 days of support the pump was explanted and the patient recovered. | [5 empty] | The investigation into the access site bleed is ongoing at this time. upon completion of the investigation, a final report will be drafted.
37261 | 1220648-2020-00575 | 10634063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10634063&pc=OZD | 09/12/2020 | 10/22/2020 |  | 10/06/2020 |  |  | 09/12/2020 | 10/06/2020 |  |  | 05/31/2020 |  | 10/06/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1487075 |  | 0048-0024 |  | Yes | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 41 YR | [5 empty] | Injury | Yes | No | Migration or Expulsion of Device | Malposition of Device | [5 empty] | Cardiac Tamponade | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 09/12/2020 | * | Invalid/Incomplete |  |  | A female patient of (b)(6) was admitted to a (b)(6) hospital and transferred to a tertiary center with plan for an intra-aortic balloon pump (iabp) and possible mitral valve replacement (mvr). when admitted to the tertiary center the patient was sent to the operating room for the iabp placement. when her condition deteriorated, the team made the choice to instead place an impella 5.5 for hemodynamic support. when her vasculature was found to be of too small a diameter for the 5.5, the team decided to place an impella cp instead. the cp was successfully supporting the patient when on the second day of support, after coughing, the impella waveforms were alarming and the pump position was assessed. the team observed a pericardial effusion. the impella was seen to be protruding from the myocardium when she was in the or for a pericardial window. the surgical repair was done and the patient received blood products. the impella remained on for support for another 8 days. it was then explanted. the patient survived. | [5 empty] | The complainant did discard the product after impella explant. more clinical data has been requested. should more information be shared that leads to a root cause of the perforation and impella migration, a final report will be filed. | Since the initial report was filed, the investigation has completed. the pump and data logs were not returned for review. it was confirmed by the clinical account that the implanting physician modified the pump prior to implant by removing the pigtail curl. this would make the tip of the pump unsafe to the myocardium and does not properly follow the ifu. the root cause of the ventricle perforation was most likely improper modification of the pump by the operating physician.
37262 | 1220648-2020-00576 | 10634068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10634068&pc=OZD | 09/07/2020 | 10/06/2020 |  | 10/06/2020 |  |  | 09/07/2020 | 10/06/2020 |  |  | 03/31/2022 |  | 10/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1482913 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/07/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) female patient had impella cp optical pump inserted in the left groin with no issues. the patient was diagnosed with hemolysis by hemolyzed labs. as treatment the pump was removed. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37263 | 1220648-2020-00577 | 10634389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10634389&pc=OZD | 09/07/2020 | 10/06/2020 |  | 10/06/2020 |  |  | 09/07/2020 | 10/06/2020 |  |  | 06/30/2022 |  | 10/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1484369 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 87 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/07/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) female patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that while on impella support, the patient developed lower limb ischemia at the right femoral. the patient's coronary had severe calcification and required a a rota. it is believed there was peripheral vascular disorders and calcification along with tortuous vasculature. the vessel size was evaluated and angiography was performed. there were no other devices on the patient's lower limbs. the lower limb ischemia was treated with thrombectomy. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37264 | 1220648-2020-00578 | 10635139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10635139&pc=OZD | 09/18/2020 | 10/06/2020 |  | 10/06/2020 |  |  | 09/18/2020 | 10/06/2020 |  |  | 04/30/2022 |  | 10/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED INC. |  |  | US |  | 1473864 |  | 005060 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN | HEALTH PR |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Unintended Movement | [6 empty] | Cardiac Tamponade | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 09/18/2020 | * | Invalid/Incomplete |  |  | A male patient was admitted for a coronary artery bypass (cabg) surgical procedure. an impella 5.0 was placed prior to the surgical case for hemodynamic support. of note there was an observation made of a known aneurysm in the left ventricle. the impella pump came in contact with the pre-existing aneurysm when the pump was not appropriately positioned, as it had migrated towards the ventricular apex. the aneurysm was punctured and surgical repair and blood products were needed. once the repair was done, the impella pump support proceeded. the pump remained on for support til wean and explant. | [5 empty] | The product was discarded. further clinical details and data logs are requested. the distributor has been contacted. the investigation has concluded without further clinical details from user in italy. the root cause of the perforation was the due to the pump being in improper position. the impella was reported to be too deep in the ventricle, causing the pigtail to damage a pre-existing aneurysm.
37265 | 1220648-2020-00571 | 10603183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10603183&pc=PYX | 09/10/2020 | 10/22/2020 |  | 09/29/2020 |  |  | 09/10/2020 | 09/29/2020 |  |  | 02/28/2021 |  | 09/29/2020 |  | TEMPORARY RIGHT SIDED HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED INC. |  |  | US |  | 1415968 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 35 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 09/10/2020 | * | Invalid/Incomplete |  |  | A (b)(6) year old male with cardiomyopathy had a lvad for 1 year. the patient condition deteriorated and right sided heart support was needed. an impella rp was placed for support. during the support the team observed presumed hemolysis. the patient had a change in the color of the urine output, as well as elevated labs. after just 24 hours the pump was removed and a protek duo was placed due to the hemolysis. | [5 empty] | The complainant has returned the impella rp product. an analysis is on-going at this time. upon completion, a final report will be filed. | Since the initial report was filed the investigation has concluded. the medical center did return the pump for hemolysis testing and the data logs for analysis. the data logs did not indicate there were any issues with the pump such as suction or pump malposition that could have caused the presumed hemolysis. the pump did not show signs of damage, but there was biomaterial around the impeller. the root cause of the issue was most likely the patient's own biomaterial.
37266 | 1035166-2020-00030 | 10583189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10583189&pc=DYB | 08/14/2020 | 02/19/2021 |  | 09/24/2020 |  |  | 08/27/2020 | 09/24/2020 |  |  | 03/01/2022 |  | 09/24/2020 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-16668 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | Yes | Break | Fluid/Blood Leak | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup | Followup |  |  |  | * | Invalid/Incomplete |  |  | It was reported that during placement of the impella cp to the right femoral artery, user was moving the pump through the peel-away introducer and the introducer has delaminated in several spots. both parts of the peel-away sheath had detached from each other on the distal end causing massive blood leak. the pump has been introduced into the heart and the procedure was completed while keeping hard pressure on the leak spot. delaminated introducer caused a massive hemorrhage on the outside of the device. the medical staff has decided to continue the procedure. the hemorrhage was limited by maintaining strong pressure. during surgical operation 3 units of red blood cells were transfused. the patient had symptoms of hemorrhagic shock after the procedure. no additional information is available. | [5 empty] | H5: product is label as single use. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | The following sections were updated in follow-up. the device used in treatment. one 14f abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. blood was found on and inside the sheath and sideport tubing. an 11cm split was observed on the sheath shaft. pon evaluation of the returned sheath, it was found that there were splits in the sheath shaft approximately 11cm from the distal tip and an indentation under the hub. the splits in the sheath shaft started approximately 2cm under the hub on the scoreline and further down the sheath shaft confirms the splits are on the scorelines. the distal tip of the sheath is split and part of the split was curled under. the hemostatic valve looked normal and intact. returned device analysis reveals the 14f abiomed introducer sheath is within manufacturing specifications. the splitting of the sheath shaft might have been caused by manipulating the device during use or by ancillary device used with the sheath. according to the device history record, the introducer sheath passed all in-process and qa final inspection steps before shipping to the customer including visual, dimensional and mechanical testing. no manufacturing defects were found. per procedure abiomed introducer sheath in-process and final inspection: first 5 consecutive properly tipped parts, inspect 100%. visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. this is performed by qa on the first 5 consecutive properly tipped parts. visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. this is performed by qa on 100% of the sheaths. per ifu precautions:aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress, follow up report will be submitted if we find any further additional information.
37267 | 1220648-2020-00568 | 10572746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10572746&pc=OZD | 08/26/2020 | 10/13/2020 |  | 09/23/2020 |  |  | 08/26/2020 | 09/23/2020 |  |  | 05/31/2022 |  | 09/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1487310 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 08/26/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) patient had impella cp optical pump inserted. the pump was removed due to a pump stop, upon removal clots was noted in the femoral artery. percutaneous transluminal coronary angioplasty (ptca) and thrombectomy was done. the patient¿s leg was cold so a cp pump was placed on the opposite leg. | [5 empty] | The impella cp optical pump was returned for evaluation. the analysis of the pump and data logs revealed that the root cause of the pump stop was due to bearing wear. the pump stopped after eight (8) days of use which is outside the pump's five (5) day indication for use. the root cause of the ischemia is likely due to the patient's condition or size of the patient's vessel. | The impella cp optical pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37268 | 1220648-2020-00569 | 10576544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10576544&pc=OZD | 09/14/2020 | 09/23/2020 |  | 09/23/2020 |  |  | 09/14/2020 | 09/23/2020 |  |  | 04/30/2020 |  | 09/23/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED INC. |  |  |  |  | 1349052 |  | 0048-0002 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 84 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 09/14/2020 | * | Invalid/Incomplete |  |  | In 2018 a male patient was admitted suffering from an acute myocardial infarction and cardiogenic shock. as the patient was admitted undergoing cpr and in critical condition an impella cp was placed for hemodynamic support. this was done without onsite support of any abiomed personnel. after just over 10 hours of support the patient expired with organ failure and underlying comorbidities. in 2020 the team in germany published upon the event. the team noted in the publication that there was limb ischemia in the impella accessed leg that was treated with ipsilateral leg access at the superficial femoral artery and placement of a bypass. the limb had peripheral perfusion restored. | [5 empty] | The product was discarded in 2018. no analysis of the product is possible. the review of the publication and device history review is all that is possible. the root cause of the limb ischemia was most likely due to patient condition, but without product this can not be proven.
37269 | 1220648-2020-00570 | 10576597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10576597&pc=OZD | 08/28/2020 | 09/23/2020 |  | 09/23/2020 |  |  | 08/28/2020 | 09/23/2020 |  |  | 01/31/2022 |  | 09/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 1468126 |  | 005040 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/28/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old caucasian female patient had impella 2.5 pump inserted in the femoral. the patient had air embolism and the left ventricle (lv)/right ventricle (rv) failure with need for resuscitation for about five (5) minutes. | [5 empty] | Device discarded by customer. the impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37270 | 1220648-2020-00567 | 10568831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10568831&pc=OZD | 08/27/2020 | 09/22/2020 |  | 09/22/2020 |  |  | 08/27/2020 | 09/22/2020 |  |  | 10/31/2021 |  | 09/22/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED INC. |  |  | US |  | 1474606 |  | 005040 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 08/27/2020 | * | Invalid/Incomplete |  |  | A (b)(6) year old female patient in (b)(6) was admitted for a high risk coronary artery angioplasty under hemodynamic support of the impella 2.5 pump. access to place the 2.5 was complicated by the patient's adipose tissue, but the pump was successfully placed and the case proceeded. the pump remained of for support of just under one day. the explant of the pump was complicated by bleeding and the development of a hematoma at the access site. the patient did received blood replacement. | [5 empty] | The medical center in (b)(6) did return the impella product for analysis. upon examination of the impella pump there was an observation made of damage to the wound closure tip. this damage allowed for the blood loss and hematoma development. the failure mode will be monitored and trended.
37271 | 1220648-2020-00563 | 10554856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10554856&pc=OZD | 08/26/2020 | 09/18/2020 |  | 09/18/2020 |  |  | 08/26/2020 | 09/18/2020 |  |  | 12/31/2021 |  | 09/18/2020 |  | : TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1461512 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/26/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for symptoms of heart failure. it was reported the patient developed hemolysis about 12 hours after impella was inserted. it was noted that the patient had hematuria prior to impella inswetion and therefore it is belived that impella was not the only cause of the hemolysis. hematuria improved by position and volume adjustment. pfhg measurements were not obtained. the patient was administrated six units of fresh frozen platelets and catecholamines. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37272 | 1220648-2020-00564 | 10554927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10554927&pc=OZD | 08/24/2020 | 09/18/2020 |  | 09/18/2020 |  |  | 08/24/2020 | 09/18/2020 |  |  | 06/30/2022 |  | 09/18/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1484370 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 47 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/24/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the right femoral for acute exacerbation of as. it was reported the patient serious ischemia on the lower limbs below the impella access site. limb ischemia was due to the insertion of the repo sheath into the femoral artery. there was no significant vessel meandering. the patient underwent a fasciotomy procedure to treat the compartment syndrome associated with lower limb ischemia. it was noted the by the physician that the cause of the compartment syndrome was the use of pcps for cardiac arrest associated with hyperkalemia. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37273 | 1220648-2020-00565 | 10555353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10555353&pc=OZD | 09/11/2019 | 09/18/2020 |  | 09/18/2020 |  |  | 09/17/2020 | 09/18/2020 |  |  | 11/30/2020 |  | 09/18/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1385057 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Code Available | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/17/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that patient developed thrombocytopenia. the reported information indicated the event was related to the impella device, pcps device and the procedure. the patient was administered 30 units of concentrated platelets. the patient soon recovered and impella was successfully explanted. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37274 | 1220648-2020-00566 | 10556469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10556469&pc=OZD | 08/15/2020 | 09/18/2020 |  | 09/18/2020 |  |  | 08/21/2020 | 09/18/2020 |  |  | 03/31/2022 |  | 09/18/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1481675 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/21/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock. the impella cp optical was selected for hemodynamic support and inserted without issue. during use of the device, the patient exhibited red urine that could not be resolved with normal troubleshooting attempts. after approximately 6 days, the device was prematurely removed. an impella 5.5 was then placed without further issues reported. | [5 empty] | The impella cp optical was received and a failure investigation was completed. other than a possible purge gap blockage, no physical abnormalities were noted. a simulated use test was performed and the device had operated as intended. upon examination of the pump, after the test run, biomaterial was found in the outflow cage. investigational findings suggest the patient's reported hemolysis was caused by ingested biomaterial within the pump.
37275 | 1220648-2020-00561 | 10546214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10546214&pc=OZD | 08/27/2020 | 09/17/2020 |  | 09/17/2020 |  |  | 08/27/2020 | 09/17/2020 |  |  | 01/31/2022 |  | 09/17/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED INC. |  |  | US |  | 1466423 |  | 0048-0032 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aneurysm | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 08/27/2020 | * | Invalid/Incomplete |  |  | A patient was admitted for a high risk coronary angioplasty of the left main artery, left anterior descending artery, and right coronary artery. to have hemodynamic support during the angioplasty, the team chose to place an impella cp via the femoral artery. the angioplasty was accomplished successfully. at the case end the team observed the impella access site to have developed a pseudo-aneurysm. the vessel was treated with ballooning for 30 minutes. the patient recovered and was discharged stable. | [5 empty] | The root cause of the vessel damage could not be determined as no product was returned for analysis. the failure mode will be monitored and trended.
37276 | 1220648-2020-00562 | 10546238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10546238&pc=OZD | 08/26/2020 | 09/17/2020 |  | 09/17/2020 |  |  | 08/26/2020 | 09/17/2020 |  |  | 01/31/2022 |  | 09/17/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1468630 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 87 YR | [5 empty] | Injury | Yes | No | Inadequate or Insufficient Training | Improper or Incorrect Procedure or Method | [5 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 08/26/2020 | * | Invalid/Incomplete |  |  | An elderly male patient was admitted in cardiogenic shock and was taken the cardiac catheterization lab for assessment of the coronary arteries. the patient had multivessel coronary artery disease and for hemodynamic support the team made a choice to initiate impella support. in a user error the incorrect product was opened and the team placed a 23 fr introducer meant for right sided support into the arterial/left sided support artery. the artery was not of sufficient size/diameter to accommodate such a larger introducer sheath. the artery was damaged and limb ischemia developed. the patient was taken to the operating suite for sheath removal and arterial repair. a new 14fr introducer was placed and impella support initiated with the impella cp. | [5 empty] | The team did not return for analysis the improperly placed introducer sheath. the team did return the impella pump for analysis and investigation. the introducer was improper for arterial placement. the cause of the limb ischemia is the condition/size of the patient vessel. the failure mode will be monitored and trended.
37277 | 1220648-2020-00555 | 10534021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10534021&pc=OZD | 08/18/2020 | 09/15/2020 |  | 09/15/2020 |  |  | 08/18/2020 | 09/15/2020 |  |  | 12/31/2020 |  | 09/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 1382217 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 14 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/18/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old male patient had impella cp pump inserted in the left axillary. the patient had hemolysis, the pump was removed to resolve hemolysis, and a left ventricular assist device (lvad) was placed. | [5 empty] | Device not returned by customer. the impella cp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37278 | 1220648-2020-00556 | 10534027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10534027&pc=OZD | 08/18/2020 | 09/15/2020 |  | 09/15/2020 |  |  | 08/18/2020 | 09/15/2020 |  |  | 10/31/2021 |  | 09/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 005040 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 81 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | Cardiac Perforation | [11 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 08/18/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male presenting for an elective high risk percutaneous coronary intervention. at the end of the procedure, a dissection to the patient's left main (unrelated to the impella device) was identified. the decision was made to transfer the patient by helicopter to an outside hospital for intervention to repair the dissection. once the patient arrived at the new hospital, the operating physician found the impella had punctured completely through the patient's circumflex. the patient expired as it was clear no intervention would save the patient. it was unclear to the physician as to how the impella had perforated the circumflex, but it was believed to have occurred during transport. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37279 | 1220648-2020-00557 | 10534043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10534043&pc=OZD | 08/18/2020 | 11/11/2020 |  | 09/15/2020 |  |  | 08/20/2020 | 09/15/2020 |  |  | 05/31/2022 |  | 09/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 1478824 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 14 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 08/20/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting with cardiomypathy for hemodynamic support using the impella cp pump. the device was inserted in the left axillary artery without issue, however, the patient exhibited hemolysis during support. as a result, the pump was prematurely removed and a left ventricular assist device was placed. | [5 empty] | Both the device and the pump's data logs were returned for analysis. the data logs found no significant positioning issues endured during the case. the pump was inspected and no damaged or biomaterial was identified. we believe the root cause of the suspected hemolysis was due to the patient's original condition. | The impella cp was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37280 | 1220648-2020-00558 | 10534048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10534048&pc=OZD | 08/19/2020 | 09/15/2020 |  | 09/15/2020 |  |  | 08/19/2020 | 09/15/2020 |  |  | 02/28/2021 |  | 09/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1404060 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/19/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock. the impella cp optical was selected for hemodynamic support, was prepped, inserted, and initiated without issue. the percutaneous coronary intervention procedure was successful. upon removal of the impella device, a thrombus was identified on the tip of the impella within the iliac artery. intervention in the form of a thrombectomy was performed to remove the thrombus. the patient was monitored for 3 additional days and discharged without further issue. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37281 | 1220648-2020-00559 | 10534065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10534065&pc=OZD | 02/25/2019 | 09/15/2020 |  | 09/15/2020 |  |  | 09/14/2020 | 09/15/2020 |  |  | 09/30/2020 |  | 09/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1374528 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/14/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 5.0 pump inserted in the right axillary for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed an infection during pump support. the reported information stated the event was cause by impella and the patient's status. the patient was administered drug treatment to treat the infection. the impella was weaned and explanted after 12 days of support. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37282 | 1220648-2020-00560 | 10534079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10534079&pc=OZD | 08/17/2020 | 09/15/2020 |  | 09/15/2020 |  |  | 08/20/2020 | 09/15/2020 |  |  | 12/31/2021 |  | 09/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1461513 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 47 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/20/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for cardiomyopathy. it was reported that four days after impella insertion, the patient experienced bleeding at the access site. while troubleshooting the issue the physician noted the puncture angle was loose. the puncture site was firmly fixed to stop the bleeding, and the patient received 2 units of red cell count for anemia. the further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37283 | 1220648-2020-00552 | 10527330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10527330&pc=OZD | 08/19/2020 | 09/14/2020 |  | 09/14/2020 |  |  | 08/19/2020 | 09/14/2020 |  |  | 09/30/2021 |  | 09/14/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1453344 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 82 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/19/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year- old (b)(6) male patient had impella 5.0 pump inserted. there was a pump stop after six (6) days of support, and another 5.0 pump was used. | [5 empty] | The impella 5.0 pump was not returned by the customer, therefore, a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37284 | 1220648-2020-00553 | 10527375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10527375&pc=OZD | 08/20/2020 | 11/02/2020 |  | 09/14/2020 |  |  | 08/20/2020 | 09/14/2020 |  |  | 06/30/2022 |  | 09/14/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1485743 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 71 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup | Followup |  |  | 08/20/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male patient had impella 5.5 pump inserted smoothly. there was a pump stop after six (6) days of support and another 5.5 pump was used. | [5 empty] | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The pump was returned for evaluation and confirmed to have a loose catheter due to inadequate glue adhesion to the kink protector. | The pump was returned for evaluation. the root cause of the open line was due to the catheter being forcefully pulled out of the pump handle.
37285 | 1220648-2020-00554 | 10527394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10527394&pc=OZD | 08/18/2020 | 10/13/2020 |  | 09/14/2020 |  |  | 08/18/2020 | 09/14/2020 |  |  | 04/30/2022 |  | 09/14/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1484688 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Injury | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 08/18/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male patient had impella cp optical pump inserted in the right common femoral. there was some wire resistance around the right iliac, surgical repair of the right iliac artery and common femoral artery was performed and impella 5.0 pump was added for support. the patient expired due to cardiogenic shock (cgs). | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed. | The returned pump was examined in the lab and no biomaterial was found in the inflow or outflow. a large kink was found in the catheter just distal to the motor housing, which may have occurred as the pump encountered resistance on explant. the guidewire was not returned. imaging of the patient's iliac arteries was not taken prior to insertion to assess or calcification or tortuosity, and thus the root cause of the patient's vessel damage cannot be determined.
37286 | 1220648-2020-00551 | 10508280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10508280&pc=OZD | 01/31/2018 | 09/09/2020 |  | 09/09/2020 |  |  | 08/11/2020 | 09/09/2020 |  |  | 01/31/2018 |  | 09/09/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  |  |  | 1232171 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 42 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/11/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with non-ischemia cardiomyopathy. the patient was also in cardiogenic shock, acute liver failure, and renal failure. the impella 5.0 was selected for hemodynamic support. the device was prepped, inserted, and initiated without any reported issue. during support, the patient exhibited an altered mental status and imaging found a clot in the left ventricle; at the tip of the catheter. the device was removed and a ct scan revealed infarcts in the frontal lobe and left corona radiata. the physician deemed the infarcts were caused by the thrombus. the patient, however, successfully went on to receive a heart transplant. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37287 | 1220648-2020-00547 | 10494768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10494768&pc=OZD | 07/23/2018 | 09/04/2020 |  | 09/04/2020 |  |  | 08/11/2020 | 09/04/2020 |  |  | 04/30/2020 |  | 09/04/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1341542 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Infarction, Cerebral | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/11/2020 | * | Invalid/Incomplete |  |  | We received a publication from asaio journal 2020, ¿ischemic stroke and intracranial hemorrhages during impella cardiac support¿, case no. 5: acute right occipital and left frontal infarcts in regards to a (b)(6) years old caucasian male patient had impella 5.0 pump inserted in the right axillary was very torturous and difficult. three days post-implant, computerized tomography (ct) imaging of brain obtained for increased lethargy showed two ischemic infarcts in the right occipital and left frontal lobes. impella was removed after eight (8) days of support after the patient¿s hemodynamics tolerated the device to be weaned. the patient's condition stabilized but developed pulseless electrical activity cardiac arrest on day 26 of hospitalization. | [5 empty] | Device discarded by customer. the impella 5.0 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37288 | 1220648-2020-00548 | 10494778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10494778&pc=OZD | 08/11/2020 | 09/04/2020 |  | 09/04/2020 |  |  | 08/11/2020 | 09/04/2020 |  |  | 02/28/2022 |  | 09/04/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED, INC. |  |  | US |  | 1463935 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/11/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white female patient had impella cp pump inserted in the right femoral without issues. a few hours later in the intensive care unit (icu) the patient had cold leg, no foot pulses palpable so the decision was made to explant the impella to improve blood circulation. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37289 | 1220648-2020-00549 | 10494813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10494813&pc=OZD | 08/10/2020 | 10/20/2020 |  | 09/04/2020 |  |  | 08/10/2020 | 09/04/2020 |  |  | 04/30/2022 |  | 09/04/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1481616 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 08/10/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old japanese male patient had impella cp pump inserted in the right axillary for pre-op coronary artery bypass graft (cabg). it was reported that two-three hours after the impella device was inserted, the patient¿s lactate dehydrogenase and bilirubin levels increased along with bloody urine. the patient was diagnosed with hemolysis. repositioning of the device was unsuccessful, and it was ultimately explanted. at the time of the event, the patient was not on ecmo. no further information was provided. | [5 empty] | The device was returned for investigation. the clinical team confirmed that the hemolysis had no causal relationship with impella. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event. | The impella cp was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37290 | 1220648-2020-00550 | 10494912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10494912&pc=OZD | 12/17/2018 | 09/04/2020 |  | 09/04/2020 |  |  | 04/11/2019 | 09/04/2020 |  |  | 08/31/2020 |  | 09/04/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1363700 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/11/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old japanese male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed hemolysis during pump support. per the physician the onset of the hemolysis occurred after the impella was inserted. it was noted the cause of the hemolysis was due to suction. as a result of the event, the patient was administered 400 units of haptoglobin. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37291 | 1220648-2020-00545 | 10487241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10487241&pc=OZD | 07/31/2020 | 09/03/2020 |  | 09/03/2020 |  |  | 08/07/2020 | 09/03/2020 |  |  | 10/31/2021 |  | 09/03/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1453032 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 83 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/07/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old male patient presenting with post cardiotomy cardiogenic shock. the impella cp was chosen for support, inserted, and initiated without any reported issues. the device was used until planned removal, however, the patient exhibited limb ischemia. a dissection and thrombus was found at the access site. the impella's pigtail was alleged to have caused or contributed to these injuries, and may have occurred when removing the device. intervention in the form of "emergency evt" and stent placement was administered. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37292 | 1220648-2020-00546 | 10487247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10487247&pc=OZD | 08/05/2020 | 10/22/2020 |  | 09/03/2020 |  |  | 08/05/2020 | 09/03/2020 |  |  | 03/31/2022 |  | 09/03/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1471575 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 68 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Not Applicable | [41 empty] | Initial submission | Followup |  |  |  | 08/05/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella 5.5 device. during support, the device stopped and was unable to be restarted. the patient was taken to the operating room and the device as removed. there was no 2nd device implanted. | [5 empty] | The device was received to investigate the cause of the reported pump stop. the purge side arm was not returned. the device's data logs confirmed a loss of purge pressure resulting from a damaged purge side arm. extended use without adequate purge pressure resulted in blood ingress, causing the reported pump stop. | The impella 5.5 was received and an investigation of device is underway. a supplemental mdr will be filed at the conclusion of our investigation.
37293 | MW5096417 | 10493490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10493490&pc=OZD | 08/20/2020 | 09/01/2020 |  | 09/03/2020 |  |  |  | 09/03/2020 |  |  | 09/30/2021 |  | 09/03/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | CATHETER IMPELLA AIC 5.0 VENTRICULAR ASSIST DEVICE |  |  | OZD | [5 empty] | ABIOMED INC. |  |  | US |  |  |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 82 YR |  |  | 93 |  |  | Malfunction | No | Yes | Failure to Power Up | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | The impella device shut off and would not turn back on and after troubleshooting. it was determined to reimplant the device.
37294 | 1220648-2020-00543 | 10483984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10483984&pc=OZD | 08/06/2020 | 09/29/2020 |  | 09/02/2020 |  |  | 08/06/2020 | 09/02/2020 |  |  | 09/30/2021 |  | 09/02/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1446479 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 08/06/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral for myocarditis. it was reported that after the impella was inserted, the pumps position became deep in the ventricle in the middle of the night. the following morning the physician adjusted the position under echocardiography. continuous hemodiafiltration (chdf) was red, bilirubin and lactate dehydrogenase (ldh) levels were high and impella¿s position was confirmed once again. there were no issues noted with positioning. the physician suspected that since patient¿s septum was being stretched and the myocardium was partially thickened, hemolysis was likely to occur due to water removal of chdf. the following day, the ldh levels were 20,000 and the t-bill increased to two digits. plasma free hemoglobin (pfhg) test showed hemolysis levels were high, as a result, impella was removed. after removal there was no drop in hemoglobin blood products were not provided. the patient¿s liver function was originally poor; therefore, the cause of the hemolysis is unknown. | [5 empty] | The impella cp was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion. | The returned pump was visually inspected and no abnormalities were observed. the data logs were reviewed and an impella position wrong alarm triggered in the first hour of use. no position alarms triggered after this point. suction briefly occurred the next day and p-level was later lowered to p2. the cause of the hemolysis was most likely patient condition because the hemolysis did not resolve after impella was removed. pfhb blood test was not performed. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event.
37295 | 1220648-2020-00544 | 10484011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10484011&pc=OZD | 04/26/2019 | 09/02/2020 |  | 09/02/2020 |  |  | 04/26/2019 | 09/02/2020 |  |  | 06/30/2020 |  | 09/02/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1352792 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 49 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/26/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella 2.5 pump. during support, the patient exhibited signs of hemolysis. it is unclear as to what specific indication led the physician to believe hemolysis was present. as treatment, haptoglobin was administered. | [5 empty] | This impella 2.5 was involved in a complaint from 2018 that was assessed as not reportable for an unrelated failure mode (hematoma without any intervention required). new information (hemolysis requiring haptoglobin) was received, however, since this complaint took place almost 2 years ago, the product has since been discarded/destroyed. the original investigation consisted of a device history record, which found no manufacturing issues with this device.
37296 | 1220648-2020-00542 | 10477739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10477739&pc=OZD | 08/03/2020 | 09/01/2020 |  | 09/01/2020 |  |  | 08/03/2020 | 09/01/2020 |  |  | 11/30/2021 |  | 09/01/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1469722 |  | 005042 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/03/2020 | * | Invalid/Incomplete |  |  | The complainant reported a(b)(6) year old female patient presenting for high risk percutaneous coronary intervention for hemodynamic support using the impella 2.5 device. both femoral arteries were noted as "tight" and patient had severe calcific disease. the pump was inserted after ballooning the iliac and femoral arteries and the patient's procedure was completed successfully. the impella 2.5 was removed, however, the decision was made to leave the 13 fr introducer in the patient's access site and have vascular surgery remove. while waiting for vascular surgery, the patient developed limb ischemia and required surgical repair. during the repair, both the introducer (and clots that had formed) were removed. | [5 empty] | The impella 2.5 and 13 fr introducer were not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37297 | 1220648-2020-00541 | 10468399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10468399&pc=OZD | 08/06/2020 | 08/28/2020 |  | 08/28/2020 |  |  | 08/06/2020 | 08/28/2020 |  |  | 01/31/2022 |  | 08/28/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1468627 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 25 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Regurgitation | Rupture | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/06/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female patient had impella cp optical pump inserted in the right femoral. the decision was made to remove the pump, while in the operating room (or) a transesophageal echocardiography (tee) was done to evaluate the left ventricle (lv) prior to explant and they noticed severe mitral regurgitation (mr) and papillary rupture. the surgeon opened the patient¿s chest, a3 leaflet was repaired and papillary was repaired. the patient was on veno arterial venous (vav) extracorporeal membrane oxygenation (ecmo) for three (3) days. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37298 | 1220648-2020-00540 | 10453136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10453136&pc=OZD | 07/01/2020 | 09/09/2020 |  | 08/26/2020 |  |  | 07/01/2020 | 08/26/2020 |  |  | 11/30/2021 |  | 08/26/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1465244 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 07/01/2020 | * | Invalid/Incomplete |  |  | A patient was transferred to a tertiary medical center with cardiogenic shock, acute kidney injury, and acute liver injury. at the tertiary center the heart failure team made the choice to place an intra-aortic balloon pump (iabp) and give medical therapy in the form of multiple cardiac med drips. the choice was then made to escalate care to an impella pump. the team was unable to placed a 5.0 pump due to vessel diameter, and instead placed an impella cp and admitted the patient to the icu for hemodynamic monitoring after explant of the iabp. on the 6th day of impella support the team explanted the impella with concerns for hemolysis. a plasma free hemoglobin was drawn and was elevated, as were ldh levels, and labs were hemolyzing. the patient had been receiving crrt dialysis and blood products. the impella was replaced by another iabp for hemodynamic support. | [5 empty] | Since the initial report was filed the investigation into the reported hemolysis has concluded. the medical center returned part of the pump for analysis. no biomaterial or pump damage/scratches were noted. the pump handle had been cut off and no further testing was possible. the root cause of the hemolysis was not determined. the failure mode will be monitored and trended. | The investigation into the alleged hemolysis caused by impella use is on-going. upon investigation completion a final report will be submitted.
37299 | 10453327 | 10453327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10453327&pc=PBL | 08/12/2020 | 08/18/2020 | 08/18/2020 | 08/26/2020 | 08/18/2020 | 08/26/2020 |  | 08/26/2020 |  |  |  |  | 08/26/2020 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PBL |  |  | IMPELLA CP PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 240879 |  | 0048-0003.0814 |  | No answer provided | [5 empty] | Invalid/Incomplete | Health Professional |  |  | User Facility report | [8 empty] | 21535 DA |  |  | 86 |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Elderly with history of coronary artery disease and resp failure and chronic systolic heart failure. upon going into carcinogenic shock, ventricular tachycardia/ventricular fibrillation - patient went into cardiac arrest and cardiopulmonary resuscitation (cpr) initiated. an impella was inserted and patient taken to helipad for transfer to another hospital. when on the helipad, the circuit/cassette failed. another circuit was obtained and replaced prior to transfer.
37300 | 1220648-2020-00539 | 10451652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10451652&pc=OZD | 07/30/2020 | 08/25/2020 |  | 08/25/2020 |  |  | 07/30/2020 | 08/25/2020 |  |  | 12/31/2020 |  | 08/25/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1400793 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/30/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) female patient had impella 5.0 pump inserted in the right femoral for an unknown indication. it was reported the patient had mild ar before impella was inserted. after insertion of the impella device, ar became severe. since impella was 42mm deep, position was adjusted; however, this did not improve the ar. the physician stated that of the three leaflets of the aortic valve, two leaflets were joined due to calcification and movement was poor. the rest of the leaflets could not be sealed with impella due to calcification. it was noted that ar was exacerbated. as a result of the event, impella was explanted and the patient was managed by intra-aortic balloon pump (iabp) and extracorporeal membrane oxygenation (ecmo). no further information provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37301 | 1220648-2020-00535 | 10441935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10441935&pc=OZD | 07/29/2020 | 10/21/2020 |  | 08/24/2020 |  |  | 07/29/2020 | 08/24/2020 |  |  | 03/31/2022 |  | 08/24/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1480003 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 07/29/2020 | * | Invalid/Incomplete |  |  | The complainant reported a patient presenting with acute myocardial infarction and cardiogenic shock. the impella cp optical was selected for hemodynamic support and inserted without issue. during support, the patient developed red urine, thus, hemolysis was suspected. device repositioning was attempted, however, the hemolysis could not be resolved. as treatment, the device was prematurely removed and support was continued with an intra-aortic balloon pump. | [5 empty] | The impella cp optical was received from the customer and an investigation of device is underway. upon completion, a supplemental mdr will be filed. | The device was returned and an investigation was conducted. no defects or biomaterial that could have caused or contributed to the suspected hemolysis was found. data logs found positioning issues had occurred early in the case. we believe the root cause of the suspected hemolysis was improper positioning.
37302 | 1220648-2020-00534 | 10441937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10441937&pc=OZD | 07/25/2020 | 08/24/2020 |  | 08/24/2020 |  |  | 07/25/2020 | 08/24/2020 |  |  | 11/30/2021 |  | 08/24/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1466419 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 12 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/25/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old hispanic/latino patient presenting with dilated cardiomyopathy. the patient required hemodynamic support until a heart transplant was received. the device was inserted without issue, however, a pocket hematoma developed at the access site during support. multiple units of blood product were required over the course of support and act was purposely kept on the "lower side" at <160 seconds. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37303 | 1220648-2020-00536 | 10444621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10444621&pc=OZD | 07/21/2020 | 08/24/2020 |  | 08/24/2020 |  |  | 07/29/2020 | 08/24/2020 |  |  | 09/30/2021 |  | 08/24/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1449375 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | HEALTH PROFESSIONAL |  |  |  |  | 50 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/29/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella 5.0 pump inserted in the left femoral for cardiomyopathy. it was reported that soon after impella insertion and echo confirmed aortic regurgitation. impella's position was adjusted; however, this did not improve the patient¿s condition. the impella was removed, and central extracorporeal membrane oxygenation (ecmo) was introduced. after removal of impella, the patient underwent aortic valve replacement (avr) to treat the moderate ar. the physician stated the patient had been transferred from another hospital for which the situation was unknown. it was also noted that it cannot be concluded the impella device cause the ar, since at the other hospital the patient's ar was mild and became moderate after the impella was inserted. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37304 | 1220648-2020-00537 | 10444663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10444663&pc=OZD | 07/14/2020 | 08/24/2020 |  | 08/24/2020 |  |  | 07/30/2020 | 08/24/2020 |  |  | 01/31/2022 |  | 08/24/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC |  |  | US |  | 1473816 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | HEALTH PROFESSIONAL |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/30/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) female patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that when the impella was explanted a thrombus was found in the patient's left iliac artery. a fogarty embolectomy catheter was used to performed thrombectomy. per the physician, the patient had the tendency to experience ischemia in the lower extremities and had thrombus formation (site unknown). | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37305 | 1220648-2020-00538 | 10444684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10444684&pc=OZD | 07/31/2020 | 10/27/2020 |  | 08/24/2020 |  |  | 07/31/2020 | 08/24/2020 |  |  | 03/31/2022 |  | 08/24/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1479162 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 07/31/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient had impella cp optical pump inserted. the patient¿s groin was oozing and as a result four (4) units of blood was administered and a femstop was placed to achieve hemostasis. | [5 empty] | Device not returned by customer. the impella cp optical pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed. | Correction to b5.: the complainant reported a 52 year old white male patient had impella cp optical pump inserted. the physician suspected a thrombus on the pump inlet during echocardiogram (echo) and wanted to exchange the pump for a new one. the patient was reported to have had pink urine during support and four (4) units of blood were given. correction to h6 as on initial report patient code was not filled out. the pump was returned for evaluation and biomaterial was confirmed to be wrapped around the impeller. only urine was used to determine hemolysis, no plasma free hemoglobin was taken. the root cause of hemolysis was likely biomaterial.
37306 | 1220648-2020-00533 | 10414587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10414587&pc=OZD | 07/20/2020 | 10/16/2020 |  | 08/17/2020 |  |  | 07/20/2020 | 08/17/2020 |  |  | 02/28/2022 |  | 08/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1471434 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 07/20/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient presenting with cardiomyopathy. the impella 5.5 pump was selected for hemodynamic support and inserted without issue. approximately 3 days into support, a pocket hematoma formed at the patient's access site the patient's access site. this was discovered after repositioning of the device. intervention in the form of a surgical drain was required to resolve. | [5 empty] | The device was returned and an investigation was performed. no damage to the blue butterfly suture was found that could lead to an access site hematoma. we believe, based on investigational findings and the clinical account, the cause of the pocket hematoma was caused by manipulation of the access graft during repositioning. | The impella 5.5 was received and an investigation of the device is underway. upon completion of our analysis, a supplemental mdr will be submitted.
37307 | 1220648-2020-00532 | 10410937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10410937&pc=OZD | 07/21/2020 | 08/14/2020 |  | 08/14/2020 |  |  | 07/21/2020 | 08/14/2020 |  |  | 04/30/2022 |  | 08/14/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1479906 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 69 YR | [5 empty] | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/21/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) female patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed access site bleeding. per the physician, anticoagulation regimen associated with heparin administration using impella became difficult which caused bleeding to become poorly controlled. eia dissociation and thrombotic obstruction was also observed. the physician stated that it is highly possible that the blood vessel was damaged when the introducer was inserted. the patient was administered 10 units of red blood cells and 10 units of fresh frozen plasma. additionally, blood clots were collected by vascular surgery and stents were placed. as support was no longer needed, impella was successfully weaned and explanted. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37308 | 1220648-2020-00531 | 10397362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10397362&pc=OZD | 07/17/2020 | 08/12/2020 |  | 08/12/2020 |  |  | 07/17/2020 | 08/12/2020 |  |  | 02/28/2022 |  | 08/12/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1471852 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 45 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/17/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old black or african american male patient presenting with acute myocardial infarction and cardiogenic shock. the impella cp optical was selected for hemodynamic support and was initiated without issue. after approximately 3 days, the patient endured hemolysis that was confirmed via plasma-free hemoglobin testing. device repositioning was attempted in order to resolve the hemolysis, however, was no successful. the device was prematurely removed and replaced with a non-impella pump. | [5 empty] | The impella cp optical was not received and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37309 | 1220648-2020-00528 | 10386106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10386106&pc=OZD | 07/15/2020 | 08/24/2020 |  | 08/10/2020 |  |  | 07/15/2020 | 08/10/2020 |  |  | 01/31/2022 |  | 08/10/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1463871 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 32 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 07/15/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old black or african american male patient had impella 5.5 pump inserted in the right axillary without issues. the pump stopped during support so another pump was inserted. | [5 empty] | The device and data logs were returned for evaluation. the data logs showed that although the pump was primed outside of the body for a long time, the purge pressure was high from the start. the logs confirmed that suddenly the purge pressure "improved" at the end of the case, but what was thought to be a fix was likely damage caused by excess pressure applied from a manual syringe. the motor current began to rise after the drop in purge pressure. the pump was confirmed to have blood ingress which was the root cause of the pump stop. | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37310 | 1220648-2020-00529 | 10386167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10386167&pc=OZD | 07/17/2020 | 08/10/2020 |  | 08/10/2020 |  |  | 07/17/2020 | 08/10/2020 |  |  | 10/31/2021 |  | 08/10/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1453164 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/17/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that prior to impella insertion both sides of the femoral artery (fa) had severe stenosis. endovascular therapy (evt) was performed on the left femoral artery (lfa). after inserting impella, the device seemed to be a little deep on the echocardiography (ucg). impella¿s position was adjusted, and the shaft was pulled about 1 cm when aortic regurgitation (ar) suddenly occurred. cardiopulmonary arrest (cpa) took place for less than 1 minute, and pa pressure rose sharply. the physician attempted to push and pull the shaft to see if ar had improved; however, it did not, and the pump was pulled out. ucg also showed that the left leaflet was damaged, and the patient underwent an emergency aortic valve replacement(avr). the patient¿s hemodynamics were maintained by administering dob (dobutamine) to nad, which was continuously administered even after impella was removed. the patient is currently stable. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37311 | 1220648-2020-00530 | 10386446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10386446&pc=OZD | 07/10/2020 | 08/10/2020 |  | 08/10/2020 |  |  | 07/16/2020 | 08/10/2020 |  |  | 03/04/2022 |  | 08/10/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1465351 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 41 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/16/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that after the patient was moved to the intensive care unit (icu), red urine was observed. blood in the urine improved by adjusting the position of impella. additionally, haptoglobin was administered. the physician suspected the issue was associated with the use of impella. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37312 | 1220648-2020-00526 | 10378307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10378307&pc=OZD | 07/13/2020 | 09/23/2020 |  | 08/07/2020 |  |  | 07/14/2020 | 08/07/2020 |  |  | 10/31/2021 |  | 08/07/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1450376 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 43 YR | [5 empty] | Injury | No | Yes | Pumping Stopped | [6 empty] | No Consequences Or Impact To Patient | [41 empty] | Initial submission | Followup |  |  |  | 07/14/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 5.0 pump inserted in the right axillary for the purpose of left ventricle unload after mitra clip procedure. it was reported that after a few hours on support the impella device stopped. from the first stop in the evening until the following morning there were a total of five pump stops. the device stopped completely when attempting to switch the patient to impella cp. impella 5.0 was successfully explanted and the physical was able to insert impella cp from the femoral artery. when the pump catheter was removed several kinks, marks were observed on the shaft near the motor. it was mentioned that the patient had a subclavian artery diameter of only 5mm, and the physician proceeded to insert impella 5.0. the physician had a right-angle anastomosis without a bevel during the graft anastomosis. although it managed to pass, considerable tension was applied to the cannula and shaft during the insertion procedure of impella 5.0. the physician stated that resistance was felt as impella 5.0 was being inserted and believes this may have been the cause of the kink. no further information was provided. | [5 empty] | The product was returned for investigation. lt log review shows purge system blocked and high purge pressure alarms throughout the case, followed by rising motor current and multiple pump stops. the pump was examined in the lab. the pump would not start and the high purge pressure was reproduced. blood ingress was found in the purge filter. a catheter kink was confirmed just distal to the motor housing. thus, the root cause of the pump stop was verified to be blood ingress due to a kinked purge line. dhr was reviewed and found no manufacturing related issues with this lot. there is one other complaint related this failure mode in this lot. abiomed will continue to monitor these types of events. no further information was provided. the manufacturer is closing the file on this event. | The impella 5.0 has been returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37313 | 1220648-2020-00527 | 10382320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10382320&pc=OZD | 02/13/2018 | 08/07/2020 |  | 08/07/2020 |  |  | 02/13/2018 | 08/07/2020 |  |  | 11/30/2019 |  | 08/07/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | CP C-7 14FR PUMP SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1321070 |  | 0048-0020 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 18 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/13/2018 | * | Invalid/Incomplete |  |  | We received a publication from artificial organs, ¿biventricular impella use in pediatric patients with severe graft dysfunction from acute rejection after heart transplantation¿, published 20 august 2019. an (b)(6) years old caucasian male had impella cp pump inserted in the left femoral. the patient developed significant bleeding from the impella cp site requiring multiple transfusions. | [5 empty] | Device discarded by customer. the cause of the access site bleeding was unable to be determined as both data and product were not returned for evaluation.
37314 | 1220648-2020-00524 | 10367452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10367452&pc=OZD | 06/20/2020 | 06/29/2021 |  | 08/05/2020 |  |  | 07/07/2020 | 08/05/2020 |  |  | 12/31/2021 |  | 08/05/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1465639 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 40 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 07/07/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old japanese male patient had impella cp pump inserted in the right femoral for cardiomyopathy. it was reported the patient developed impella access site bleeding. the patient was administered blood products to treat the injury, specifics were not provided. the patient was also treated with surgical extubation and hemostasis. it was noted that the patient was also on extracorporeal membrane oxygenation (ecmo) at the time of the event. no further information was provided. | Additional information was received on (b)(6) 2021, indicating that during impella support, the patient developed hemolysis. as a result, infusion was administered. the event is reportable. no further information was provided. |  |  |  |  | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella cp was discarded by the customer. and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37315 | 1220648-2020-00525 | 10367648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10367648&pc=OZD | 07/08/2020 | 08/05/2020 |  | 08/05/2020 |  |  | 07/08/2020 | 08/05/2020 |  |  | 02/29/2020 |  | 08/05/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1331385 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/08/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old japanese male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that despite there being no progression of aortic regurgitation (ar) during the impella support. a transthoracic echocardiogram (tte) revealed moderate ar with augmentation of a regurgitant jet to the mitral anterior leaflet after impella was removed. the patient seemingly developed moderate ar due to the removal of the impella device. the left ventricular ejection fraction (lvef) improved to 38% with moderate impaired systolic function, and moderate ar remained. the patient was discharged from the hospital without symptoms after being in the hospital for 40 days. four weeks later, the patient required re-admitted due to progressive dyspnea on exertion. tte showed lvdd/ds of 61/52 mm, lvef of 35% with severe hypokinesis and severe. following improvement in heart failure, aortic valve replacement (avr) was performed to treat the recurrent heart failure due to ar. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37316 | 1220648-2020-00523 | 10364790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10364790&pc=OZD | 07/06/2020 | 08/04/2020 |  | 08/04/2020 |  |  | 07/07/2020 | 08/04/2020 |  |  | 03/31/2022 |  | 08/04/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1477975 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/07/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male patient presenting for high risk percutaneous coronary intervention, using the impella cp for hemodynamic support. the device was inserted and initiated without any reported issue, however, the patient ultimately exhibited lower limb ischemia in the impella inserted leg. as treatment, a bypass sheath was administered to feed blood from the opposing limb to the affected limb. support continued successfully until the patient was weaned. | [5 empty] | The impella cp was not returned and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37317 | 1220648-2020-00522 | 10356787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10356787&pc=OZD | 07/04/2020 | 08/24/2020 |  | 08/03/2020 |  |  | 07/04/2020 | 08/03/2020 |  |  | 11/30/2021 |  | 08/03/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1452909 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 69 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 07/04/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient had impella 5.5 pump inserted for additional support and longevity after having impella 5.0 pump. the impella 5.5 pump was inserted without any issue. the pump stopped over the duration of a couple hours. the patient's family withdrew from care to make patient comfort measures only (cmo) and the patient expired. | [5 empty] | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The pump was returned for evaluation. there was biomaterial found in the outlet when visually inspected. the pump stop was likely caused by ingested biomaterial.
37318 | 1220648-2020-00518 | 10339031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10339031&pc=OZD | 07/07/2020 | 07/29/2020 |  | 07/29/2020 |  |  | 07/07/2020 | 07/29/2020 |  |  | 08/31/2020 |  | 07/29/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1369160 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | LITERA |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 07/07/2020 | * | Invalid/Incomplete |  |  | A social media post was made via (b)(6) in which an acute myocardial infarction patient was supported by an impella cp for hemodynamic support. the pump was placed in the left ventricle (lv) and was supporting the patient when suction alarms were noted on day 6 of support. upon review by echo thrombus was seen within the lv. the patient suffered signs of a stroke. the ct showed m1 distribution stroke. the patient had his impella explanted, and intra-aortic balloon pump was placed, and a sentinel cerebral protection device. the twitter post cited the case publication in jacc. the jacc piece did note that publications on lv clot are not many and that the use of cerebral protection devices may be considered. when reviewing the abiomed database, the case support was discovered, but no allegation of impella use causing or contributing the stroke was noted by the abiomed field team member. | [5 empty] | There was no investigation possible at this time. the product was discarded in 2018 and there was no imaging analysis possible. the root cause of the stroke can not be tied to the use of impella. there was no evidence to suggest the stroke was caused by the use of the impella pump and the thrombus was noted to be within the left atrium prior to the use of the impella. the failure mode will be trended and monitored.
37319 | 1220648-2020-00519 | 10339032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10339032&pc=OZD | 07/16/2020 | 07/29/2020 |  | 07/29/2020 |  |  | 07/16/2020 | 07/29/2020 |  |  | 05/31/2021 |  | 07/29/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED EUROPE |  |  | GM |  | 1424953 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Positioning Problem | Device Handling Problem | [5 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | [28 empty] | Initial submission |  |  |  |  | 07/16/2020 | * | Invalid/Incomplete |  |  | The complainant in (b)(6) admitted a male patient with ami/cgs who had infarcted 3 weeks prior. the medical team escalated his care from a iabp and made the choice to place an impella 5.0 via the axillary artery, but due to delivery issues at the subclavian, they had to place the pump instead via the femoral artery. the angioplasty with a 8mm balloon failed to dilate the anatomy for the placement via the axillary. due to the access attempts at his axillary artery, a hematoma developed which prompted intervention. a drain was placed to the hematoma. | [5 empty] | The medical team in (b)(6) discarded the product. no data logs were forwarded for analysis. the further clinical details of whether or not blood products were infused or the vessel size and morphology were not shared. the only intervention shared, to treat the hematoma, was that the team in (b)(6) placed a drain. should further details be shared, a follow-up will be submitted.
37320 | 1220648-2020-00520 | 10339040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10339040&pc=OZD | 07/10/2020 | 07/29/2020 |  | 07/29/2020 |  |  | 07/10/2020 | 07/29/2020 |  |  | 03/31/2022 |  | 07/29/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED EUROPE |  |  | GM |  | 1479957 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [22 empty] | Initial submission |  |  |  |  | 07/10/2020 | * | Invalid/Incomplete |  |  | The complainant decided to place an impella cp as care escalation from an intra-aortic balloon pump (iabp) in a cardiomyopathy patient. the team removed the iabp and exchanged sheaths for the cp to be placed. the previous iabp insertion was done at a steep angulation and there was concern for bleeding with the larger cp pump and it's sheath. when in the icu, for hemodynamic monitoring, the cp site was seen to be oozing. to treat the blood loss, the team held pressure, and later gave 2 units of blood replacement. the cp pump successfully supported the patient for 5 days, in conjunction with ecmo, til wean and explant. | [5 empty] | The investigation into the access site bleeding was possible, as the product was returned for analysis. when the product was returned no defects were found that would have caused or contributed to the bleeding. the root cause was unable to be determined. the failure mode will be monitored and trended.
37321 | 1220648-2020-00521 | 10341453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10341453&pc=OZD | 07/15/2020 | 08/13/2020 |  | 07/29/2020 |  |  | 07/15/2020 | 07/29/2020 |  |  | 12/31/2021 |  | 07/29/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED EUROPE |  |  | GM |  | 1477340 |  | 005042 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Difficult to Insert | Malposition of Device | [5 empty] | Perforation of Vessels | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 07/15/2020 | * | Invalid/Incomplete |  |  | A female patient with multiple known coronary risk factors was admitted for a high risk angioplasty with atherectomy. the team made the choice to place an impella 2.5 for hemodynamic support during the angioplasty. while attempting to access the femoral artery and place the 13fr introducer there was an observation made of vessel damage. the introducer had exited and perforated the artery. the introducer was replaced by a 14fr introducer and later the perforation was treated by placement of a viabahn stent, when ballooning alone did not remedy the perforation. after stenting the iliac was intact and all bleeding had ceased. the patient survived with pulses intact and flow to the foot. | [5 empty] | Since the initial report was filed, the investigation has completed. it was determined that the physician's failure to use fluoroscopy imaging to place the introducer sheath caused the vascular damage. the failure mode will be monitored and trended. | The investigation is on-going at this time. upon completion of the investigation, a final report will be submitted.
37322 | 1035166-2020-00043 | 10331773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10331773&pc=DYB | 07/03/2020 | 10/05/2020 |  | 07/28/2020 |  |  | 07/07/2020 | 07/28/2020 |  |  | 06/01/2021 |  | 07/28/2020 |  | INTRODUCER, CATHETER, PRODUCT CODE: DYB |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-14301 |  | 0052-3021 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 81 YR | [5 empty] | Malfunction | No | Yes | Crack | [6 empty] | Death | [41 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that on (b)(6) 2020, there was a 23 fr introducer sheath that cracked during a post cardiac-cardiogenic shock (pccs) right ventricle (rv) failure procedure. the user was placing the sheath when the rp sheath valve bleeding with insertion over stiff wire occurred. there was a procedure delay around 5 minutes. the patient expired days later, however, not related to the sheath issue. | It was reported that the patient lost about 20 cc of blood. |  |  |  |  | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. | The following sections were updated in follow-up. b4, b5, g3, g6, h1, h2, h6, h10. the device was used for treatment. the device was explanted and discarded, therefore the clinical observation could not be confirmed. the investigation will focus on a review of product documentation. as per procedure (abiomed introducer sheath in-process and final inspection), the devices in this lot were inspected as follows: 9.2.2 visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. per section 12.2, visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. perform leak test per procedure. the leak test is performed by qa on 100% of the sheaths. per ifu: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. infusion through the sideport can be done only after all air is removed from the unit. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
37323 | 1220648-2020-00514 | 10326718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10326718&pc=OZD | 06/29/2020 | 09/23/2020 |  | 07/27/2020 |  |  | 06/29/2020 | 07/27/2020 |  |  | 03/31/2022 |  | 07/27/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1473835 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 39 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Not Applicable | [41 empty] | Initial submission | Followup |  |  |  | 06/29/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old black/african american male presenting with cardiomyopathy. the impella 5.5 pump was selected for hemodynamic support and inserted without issue. approximately 143 hours into support, the automated impella controller console emitted a "purge flow lumen blocked" alarm, with an elevated purge pressure reading and rising motor current. thereafter, the pump stopped and patient was taken to the operating room where the device was removed and replaced with a new impella 5.5. upon removal, a thrombus was found entrained within the pump. the patient did not endure any harm. | [5 empty] | The device and data logs were received and an investigation was completed. data logs confirmed a rise in motor current prior to the pump stop. physical inspection found blood ingress, which was determined to have caused the pump stop. | The impella 5.5 was returned by the customer and an investigation of device is underway. a supplemental mdr will be filed upon completion.
37324 | 1220648-2020-00516 | 10326751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10326751&pc=OZD | 07/02/2020 | 08/13/2020 |  | 07/27/2020 |  |  | 07/02/2020 | 07/27/2020 |  |  | 03/31/2022 |  | 07/27/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED INC. |  |  | US |  | 1477966 |  | 0550-0008 |  | Yes | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 07/02/2020 | * | Invalid/Incomplete |  |  | A cardiac patient was admitted for a 5 coronary artery bypass surgery (cabg) and an aortic valve replacement. to assist with getting off the bypass the team placed an impella 5.5 for hemodynamic support. the 5.5 supported the patient with some suction alarms for 11 days when the team chose to explant the pump due to concerns for hemolysis. the labs were hemolyzing and despite repositioning the hemolysis signs were not resolving. | [5 empty] | Since the original report was filed, the investigation has come to conclusion. the medical center returned the pump's data logs but, not the actual impella pump. the root cause of the hemolysis could not be determined without the pump. no benchtop hemolysis testing was possible. the failure mode will be monitored and trended. | The investigation into the presumed hemolysis caused by the impella 5.5 is on-going at this time. no product or data logs have yet been returned. upon completion of the investigation, a final report will be submitted.
37325 | 1220648-2020-00515 | 10326753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10326753&pc=OZD | 06/30/2020 | 07/27/2020 |  | 07/27/2020 |  |  | 07/01/2020 | 07/27/2020 |  |  | 02/28/2022 |  | 07/27/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1477174 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/01/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male patient presenting for high risk percutaneous coronary intervention. the impella cp optical was selected for hemodynamic support, prepped, and inserted without any reported issue. during support, the patient developed an access site bleed and hematoma. as treatment, a surgical cutdown and hematoma evacuation was performed, a "wound vac" was placed, and 3 units of packed red blood cells were delivered. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37326 | 1220648-2020-00517 | 10327000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10327000&pc=OZD | 06/29/2020 | 07/27/2020 |  | 07/27/2020 |  |  | 06/29/2020 | 07/27/2020 |  |  |  |  | 07/27/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1476547 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/29/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp. the device was prepped and inserted into the patient's left femoral artery without any reported issue. during support, the patient developed ischemia that prompted premature removal of the device. it should be known that this patient had peripheral artery disease which caused narrowing of the vessel around the impella, leading to ischemia. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37327 | 10307207 | 10307207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10307207&pc=OZD | 06/12/2020 | 07/13/2020 | 07/13/2020 | 07/22/2020 | 07/13/2020 | 07/22/2020 |  | 07/22/2020 |  |  |  |  | 07/22/2020 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | IMPELLA |  |  | OZD |  |  | 0042-0000-US |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 24090 DA | [5 empty] | Malfunction |  | Yes | Computer Operating System Problem | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella alarming 'motor current high' at p0, mean arterial pressure in the 40's, right ventricular assist device (rvad) flows <3. attempted to reset impella to p8, immediately returned to p0, attempted to unplug and restart impella per rep, unable to restart pump. impella removed at the bedside. instructed by impella rep to save device and pump. bedside echo performed.
37328 | 1220648-2020-00512 | 10301108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10301108&pc=OZD | 06/18/2020 | 09/15/2020 |  | 07/21/2020 |  |  | 06/25/2020 | 07/21/2020 |  |  | 10/31/2021 |  | 07/21/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1446023 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 18 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 06/25/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) old white male patient presenting with cardiomyopathy. the impella cp was selected for hemodynamic support and inserted without any noted issues. during support, the patient exhibited signs of hemolysis. intervention was taken to prematurely remove the device. a new pump was placed and the hemolysis resolved. | [5 empty] | The device and data logs were reviewed for investigation. data logs confirm low flow and suction alarms, however, physical investigation and simulated use testing of the device found no issues. we can not definitively determine the root cause of the hemolysis reported. | The impella cp was has not been received from the customer and therefore, investigation of device was not possible. should we be successful in obtaining the device, an investigation will be completed and a supplemental mdr will be filed.
37329 | 1220648-2020-00513 | 10302412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10302412&pc=OZD | 11/27/2019 | 07/21/2020 |  | 07/21/2020 |  |  | 06/22/2020 | 07/21/2020 |  |  | 06/30/2021 |  | 07/21/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1432560 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Unspecified Infection | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/22/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) male patient had impella 2.5 pump inserted in the right femoral for high risk percutaneous coronary intervention (hrpci). it was reported that after the pump was removed an infection at the insertion site was noted. no treatment was administered, and the patient was discharged. a few weeks later the patient had to be re-admitted as signs and symptoms of fever and swelling and pain at the access site developed. the patient underwent vascular repair with the diagnosis of pseudo-aneurysm. intravenous antibiotics were administered to treat the access site infection. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37330 | 1220648-2020-00509 | 10296248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10296248&pc=OZD | 06/21/2020 | 07/20/2020 |  | 07/20/2020 |  |  | 06/22/2020 | 07/20/2020 |  |  | 03/31/2022 |  | 07/20/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1477153 |  | 0550-0008 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 32 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/22/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient had impella 5.5 pump inserted. an epical bleed was noted due to the catheter pushing against the wall of the left ventricle (lv) during attempts to insert the outlet out of the graft. lv apical perforation occurred and as a result two (2) units of packed red blood cells (prbcs), one (1) unit of fresh frozen plasma (ffp), and one (1) unit of platelets was administered. | [5 empty] | Device discarded by customer. the impella 5.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37331 | 1220648-2020-00510 | 10296270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10296270&pc=OZD | 06/23/2020 | 07/20/2020 |  | 07/20/2020 |  |  | 06/23/2020 | 07/20/2020 |  |  | 02/28/2022 |  | 07/20/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED EUROPE |  |  | GM |  | 1473964 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Positioning Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 06/23/2020 | * | Invalid/Incomplete |  |  | A medical center had a patient admitted for a high risk angioplasty of the right coronary artery. the team chose to use the 14 fr abiomed sheath for the placement of the impella 2.5 pump and the angioplasty rotational atherectomy equipment. the patient experienced limb ischemia, and the medical team chose to place a bypass via the opposite limb. when placing the bypass sheath there was an access site bleed that prompted blood replacement products to be infused, both red cells and ffp. the limb ischemia needed no further intervention. the limb ischemia experienced was noted to secondary to the known peripheral arterial disease. | [5 empty] | The medical center did not return the impella pump or introducer product for analysis. the team did return the data logs, however they do not lead to a root cause for the limb ischemia. the medical team was aware of the patient's peripheral arterial disease which contributed to the limb ischemia and need for external bypass. the failure mode will be monitored and trended.
37332 | 1220648-2020-00511 | 10296284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10296284&pc=OZD | 08/01/2019 | 07/20/2020 |  | 07/20/2020 |  |  | 06/26/2020 | 07/20/2020 |  |  | 12/31/2020 |  | 07/20/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1391206 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/26/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted in the right femoral for myocarditis. it was reported the patient developed hemolysis during pump support. patient's lactate dehydrogenase decreased, and urine was red/brown. to treat the patient haptoglobin was administered (amount unknown) and impella's position and level were changed. no further information was required. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37333 | 1220648-2020-00508 | 10290345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10290345&pc=OZD | 06/22/2020 | 07/17/2020 |  | 07/17/2020 |  |  | 06/22/2020 | 07/17/2020 |  |  | 02/28/2022 |  | 07/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1474678 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/22/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) old white male patient presenting for high risk percutaneous coronary intervention. the impella cp optical was selected for hemodynamic support and inserted without issue. after the patient was transferred to the intensive care unit, blood was found leaking between the repositioning sheath and blue butterfly suture pad. the bleed was resolved with manual pressure, a femostop, and 2 units of blood products. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37334 | 1220648-2020-00503 | 10270463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10270463&pc=OZD | 06/23/2020 | 07/14/2020 |  | 07/14/2020 |  |  | 06/23/2020 | 07/14/2020 |  |  | 06/30/2020 |  | 07/14/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1353068 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | LITERATURE |  |  |  |  | 41 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/23/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 5.0 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed aortic insufficiency which was associated with the use of impella. the patient required surgical intervention, details were not provided. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37335 | 1220648-2020-00504 | 10270471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10270471&pc=OZD | 01/20/2020 | 07/14/2020 |  | 07/14/2020 |  |  | 06/18/2020 | 07/14/2020 |  |  | 09/30/2021 |  | 07/14/2020 |  | : TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1446739 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/18/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) female patient had impella cp pump inserted in the left femoral for high risk percutaneous coronary intervention (hrpci). it was reported during pump support the patient experienced cerebrovascular accident (cva). as a result, the patient was given drug therapy. the patient ultimately expired due to patient status and another adverse event. per the pms report, the death was not related to impella. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37336 | 1220648-2020-00505 | 10270497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10270497&pc=OZD | 06/18/2020 | 07/14/2020 |  | 07/14/2020 |  |  | 06/18/2020 | 07/14/2020 |  |  | 02/28/2021 |  | 07/14/2020 |  | : TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1412004 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/18/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that prior to impella placement a rota was performed without passing the wire. the patient was shocked during rota and impella was implanted independently. the following day pseudoaneurysm was noted at the access site. as a result, an embolization was performed in the angio chamber. per the radiologist, the insertion procedure and heparin contributed to the bleeding. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37337 | 1220648-2020-00506 | 10270524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10270524&pc=OZD | 06/17/2020 | 07/14/2020 |  | 07/14/2020 |  |  | 06/18/2020 | 07/14/2020 |  |  | 12/31/2021 |  | 07/14/2020 |  | : TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1461513 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/18/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that upon explant of the implella, a blood vessel dissociation at the right cfa was noted. as a result, a stent was placed as treatment. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37338 | 1220648-2020-00507 | 10270555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10270555&pc=OZD | 08/22/2019 | 07/14/2020 |  | 07/14/2020 |  |  | 06/17/2020 | 07/14/2020 |  |  | 04/30/2021 |  | 07/14/2020 |  | : TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1417620 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | HEALTH PROFESSIONAL | U |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/17/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral for high risk percutaneous coronary intervention (hrpci). it was reported that the patient experienced hemorrhagic shock. the patient was sedated for intubation management. a ct imaging revealed acute cerebral infarction suspected of emboli in the bilateral cerebellum, bilateral occipital lobe, left temporal lobe, and left frontal-parietal lobe. per the physician prolonged emboli or post-operative hemorrhagic shock suggest possible impella-related embolism. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37339 | 1035166-2020-00040 | 10259632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10259632&pc=DYB | 06/10/2020 | 11/30/2020 |  | 07/10/2020 |  |  | 06/15/2020 | 07/10/2020 |  |  | 07/01/2021 |  | 07/10/2020 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3001 |  |  | OSCOR INC. |  |  | US |  | C1-14497 |  | 0052-3001 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN | [10 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Myocardial Infarction | Tachycardia | Cardiogenic Shock | [38 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported patient came from intensive care unit (icu) in a rather critical state. the 23f sheath was used for acute myocardial infraction (ami) and a cardiogenic shock (cgs) procedure. issue with the 23f sheath was bleeding through the valve, it was identified when sheath was inserted into the femoral vein, physician tried to stop bleeding with 0.027" wire (different vendors product), then 12f sheath, then rp pump (in that order) unsuccessfully. only brief delay of the procedure due to bleeding, no impact on patient condition. patient outcome is death ((b)(6) 2020), however, death is not related to device failure. physician reported cause of patient death as ventricular tachycardia (attempts to achieve sinusrhythm were unsuccessful) followed by no output from left ventricle (lv). additionally, the physician said that the blood loss was not contributing to the death of the patient since it was not therapeutically relevant. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. | The device was used in treatment. the device was not returned for analysis, as it was discarded. the investigation will focus on a review of product documentation. the device history records were reviewed to confirm that the device passed all applicable in process and final inspections. per qa procedure abiomed introducer sheath in-process and final inspections: 12.3 vacuum leak test: perform pressure and vacuum leak test per procedure, latest revision. this procedure is performed by qa 100%. per instructions for use (ifu): when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
37340 | 1220648-2020-00498 | 10244596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10244596&pc=OZD | 06/11/2020 | 09/08/2020 |  | 07/08/2020 |  |  | 06/11/2020 | 07/08/2020 |  |  | 08/31/2021 |  | 07/08/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1438295 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 37 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Insufficiency | Mitral Regurgitation | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 06/11/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock for support with the impella cp. during support, the patient's urine was noted as red/brown in color. an echocardiogram was completed and found the device was "flipped" toward the mitral valve. upon explant and escalation to an impella 5.0, the patient's posterior mitral valve leaflet was found torn and resulted in mitral regurgitation. impella 5.0 was inserted and patient was noted as stable. patient continued support on impella 5.0 for approximately 15 days with potential plans to repair the valve, however, the patient expired. there was no allegation of either device's involvement in the patient's outcome as care was withdrawn. | [5 empty] | The impella cp was received and a failure investigation was completed. upon examination of the pump, no damage was found. data logs were reviewed and confirm multiple suction issues and positioning alarms took place early in the case. a review of the device history record found no manufacturing issues, reworks, or complaints are associated with this failure mode from this pump lot. we are unable to definitively determine the root cause of the heart valve damage, however, data logs analysis suggests improper positioning resulted in the reported hemolysis. | The impella cp was returned and an investigation is underway. a supplemental mdr will be filed upon completion.
37341 | 1220648-2020-00499 | 10244597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10244597&pc=OZD | 06/09/2020 | 07/08/2020 |  | 07/08/2020 |  |  | 06/12/2020 | 07/08/2020 |  |  | 01/31/2022 |  | 07/08/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1470953 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/12/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed bleeding at the impella access site, there was side leakage from the insertion portion of the indwelling sheath. as a result, surgical hemostasis was performed, and the patient was administered 10 units of blood products. patient's act's were around 160. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37342 | 1220648-2020-00500 | 10244598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10244598&pc=OZD | 06/08/2020 | 07/08/2020 |  | 07/08/2020 |  |  | 06/12/2020 | 07/08/2020 |  |  | 12/31/2021 |  | 07/08/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1465639 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/12/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed bleeding at the impella access site, there was side leakage from the insertion portion of the indwelling sheath. as a result, surgical hemostasis was performed, and the patient was administered 4 units of blood products. patient's act's were 160-180. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37343 | 1220648-2020-00501 | 10244603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10244603&pc=OZD | 06/16/2020 | 07/08/2020 |  | 07/08/2020 |  |  | 06/16/2020 | 07/08/2020 |  |  | 03/31/2022 |  | 07/08/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1476932 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 45 YR | [5 empty] | Malfunction | No | Yes | No Apparent Adverse Event | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 06/16/2020 | * | Invalid/Incomplete |  |  | The us complainant had placed an impella 5.5 for hemodynamic support for a patient with cardiomyopathy. the patient was suffering from nonischemic cardiomyopathy with shortness of breath and pleural effusions. the 5.5 was placed and was to support the patient while team meetings were ongoing for a possible heart transplant. unfortunately, the pump stopped after just 1 hour of support. this prompted the team to explant the pump. | [5 empty] | The complainant returned the impella 5.5 for a product analysis. the data logs from the support were returned as well. the pump was found to have ingested biomaterial that caused the stop. the biomaterial was not created by the pump but, rather thought to be pre-existing. the failure mode will be monitored and trended.
37344 | 1220648-2020-00502 | 10247705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10247705&pc=OZD | 05/29/2020 | 07/08/2020 |  | 07/08/2020 |  |  | 05/29/2020 | 07/08/2020 |  |  | 09/30/2021 |  | 07/08/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1445031 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR | [5 empty] | Malfunction | No | Yes | Detachment of Device or Device Component | Infusion or Flow Problem | No Apparent Adverse Event |  |  |  |  | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission |  |  |  |  | 05/29/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) old white femail patient presenting for high risk percutaneous coronary intervention. the impella cp was selected for hemodynamic support and inserted without any noted issue. approximately 5 minutes into use, the device flows dropped from approximately 3.5 l/min to approximately 1 l/min. the issue could not be resolved with normal troubleshooting, therefore, the pump was removed and replaced without any adverse impact to the patient. the device was returned and broken impeller blades were identified. | [5 empty] | The impella cp was received and a failure investigation was completed. a simulated use test was performed and successfully reproduced the reported low flows. examination of the pump found a broken impeller blade. a review of the device history record was conducted and found no manufacturing issues, reworks, or complaints associated with this failure mode from this pump lot. since the console's data logs were not returned and no imagery was provided, we can not definitively determine the cause of the low flow or broken blade.
37345 | 1220648-2020-00493 | 10233738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10233738&pc=OZD | 10/03/2019 | 07/06/2020 |  | 07/06/2020 |  |  | 06/08/2020 | 07/06/2020 |  |  | 05/31/2021 |  | 07/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1421773 |  |  |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 12 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/08/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella 2.5 pump inserted in the right femoral for cardiomyopathy. it was reported that patient experienced a hematoma at the impella access site during support. as a result, the patient received an unknown amount of blood products. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37346 | 1220648-2020-00494 | 10233793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10233793&pc=OZD | 06/11/2020 | 07/06/2020 |  | 07/06/2020 |  |  | 06/11/2020 | 07/06/2020 |  |  | 09/30/2021 |  | 07/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1446739 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 35 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/11/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed hemolysis. impella support was lowered and the patient was administered haptoglobin. it was noted the patient was also on ecmo support. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37347 | 1220648-2020-00495 | 10233802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10233802&pc=OZD | 06/05/2020 | 07/06/2020 |  | 07/06/2020 |  |  | 06/10/2020 | 07/06/2020 |  |  | 06/30/2021 |  | 07/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1426675 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/10/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) male patient had impella 2.5 pump inserted in the right femoral for myocarditis. it was reported the patient experienced hemolysis. as a result, the impella device was explanted. pfhg measurements were not provided; however, the site stated that due to decreasing of some blood cell component by blood sampling, hemolysis was confirmed. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37348 | 1220648-2020-00496 | 10233813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10233813&pc=OZD | 06/10/2020 | 07/06/2020 |  | 07/06/2020 |  |  | 06/12/2020 | 07/06/2020 |  |  | 12/31/2021 |  | 07/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1461512 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 47 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/12/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) male patient had impella cp pump for acute heart failure. insertion site was not provided. it was reported that patient whose physique was large and whose femoral artery was deep experience oozing from the puncture site. as a result, the access site was surgically opened, and hemostasis was completed. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37349 | 1220648-2020-00497 | 10233833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10233833&pc=OZD | 06/11/2020 | 07/06/2020 |  | 07/06/2020 |  |  | 06/12/2020 | 07/06/2020 |  |  | 01/31/2022 |  | 07/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1470941 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/12/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed bleeding at the impella access site, there was a flank leak from the insertion site of the indwelling sheath. as a result, surgical hemostasis was performed, and the patient was administered 8 units of blood products. patient's act's were 160-170. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37350 | 1220648-2020-00492 | 10226387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10226387&pc=OZD | 06/06/2020 | 07/02/2020 |  | 07/02/2020 |  |  | 06/06/2020 | 07/02/2020 |  |  | 09/30/2021 |  | 07/02/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1452120 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Malposition of Device | [6 empty] | Ventricle, Abnormality Of | Blood Loss | Pericardial Effusion | [10 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 06/06/2020 | * | Invalid/Incomplete |  |  | The medical center had a patient admitted in cardiogenic shock suffering from an acute myocardial infarction. she had an impella cp placed for hemodynamic support for an angioplasty and thrombectomy of the left anterior descending coronary artery. the first impella pump placed had low flows and was replaced. after placement of the second impella cp the patient the patient condition deteriorated and an echo revealed a effusion. the team attempted to perform a pericardiocentesis to remove the fluid. during this time she was receiving multiple rounds of defibrillation and cpr. the patient expired. the medical team attributed the effusion to the impella guidewire causing a perforation. | [5 empty] | The complainant has discarded the impella pump and guidewire product. analysis is not possible. there was a request made for return of echo imaging and the impella console data logs. they have not yet been returned for analysis. should more information be shared into the cause of the ventricular perforation, and a root cause be determined, another report will be filed.
37351 | 1220648-2020-00490 | 10214702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10214702&pc=OZD | 06/10/2020 | 06/30/2020 |  | 06/30/2020 |  |  | 06/10/2020 | 06/30/2020 |  |  | 09/30/2021 |  | 06/30/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED EUROPE |  |  | GM |  | 1448721 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Malposition of Device | [6 empty] | Ventricle, Abnormality Of | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 06/10/2020 | * | Invalid/Incomplete |  |  | A (b)(6) old female was admitted and evaluated for her coronary artery disease. the cardiothoracic surgery team turned her down and so the team chose to perform a high risk angioplasty for her multivessel disease. an impella cp was placed for hemodynamic support during the left main coronary and circumflex angioplasty. the right femoral artery was accessed for both the impella and angioplasty procedure. there was difficulty felt by the physician in placement of the impella and an extra "push" was needed for placement. as the angioplasty was begun the patient condition deteriorated. a pericardial effusion was observed. the pericardiocentesis failed and so a window was performed. the patient survived the window and multiple rounds of cpr. later the condition again deteriorated and her chest was opened at which time clot was removed and the impella catheter was seen to have perforated through the ventricular wall. the patient's perforation was repaired, blood products infused, and impella removed. | [5 empty] | The complainant returned the impella pump. no product defects were discovered. the patient condition and imaging were not shared for analysis, but the underlying disease and patient condition was the most likely cause of the ventricular wall perforation. the failure mode will be monitored and trended.
37352 | 1220648-2020-00491 | 10214720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10214720&pc=OZD | 06/01/2020 | 06/30/2020 |  | 06/30/2020 |  |  | 06/01/2020 | 06/30/2020 |  |  | 03/31/2022 |  | 06/30/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1471612 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/01/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) old white male patient had impella 5.5 pump inserted in the right axillary. the pump was removed due to the unresolved elevation in lactate dehydrogenase (ldh), creatinine, and suspicion of hemolysis. care was withdrawn with unknown results of plasma-free hemoglobin (pfhgb). | [5 empty] | The data logs showed the pump started on p-9, flow was 4.8 l/min for about 40 minutes then p level was adjusted to p-6, flow dropped to 3 l/min and suction alarms were triggered. the pump was returned for evaluation but was missing pump plug as the catheter was cut off when explanted. there were some biomaterial left in the returned package. visual inspection found no obstruction in if cage, cannula & pump housing, no scratch found on the ph wall, and the impeller was concentric. there was no issue was found on the returned pump. hemolysis testing was unable to be conducted as the pump was damaged. the cause of the suspected hemolysis was likely biomaterial as reported by the clinical team despite the lack of biomaterial presence on the returned pump since no other issues were observed.
37353 | 1220648-2020-00486 | 10206033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10206033&pc=OZD | 10/15/2019 | 06/29/2020 |  | 06/29/2020 |  |  | 03/10/2020 | 06/29/2020 |  |  | 06/30/2021 |  | 06/29/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1428735 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/10/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year-old (b)(6) female patient had impella 2.5 pump inserted in the right femoral for myocarditis. it was reported that the patient developed hemolysis during pump support. as a result, the patient was administered drug treatment, specifics were not provided. explanting the impella device also helped resolve hemolysis markers. hemolyis was confirmed via pfhg measurements. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37354 | 1220648-2020-00487 | 10206050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10206050&pc=OZD | 06/18/2020 | 06/29/2020 |  | 06/29/2020 |  |  | 06/18/2020 | 06/29/2020 |  |  | 06/04/2014 |  | 06/29/2020 |  | TEMPORARY NON ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | 5.0 CATHETER |  |  | ABIOMED EUROPE |  |  | GM |  | 0066006 |  | 004683 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | LITERATURE |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Mitral Regurgitation | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 06/18/2020 | * | Invalid/Incomplete |  |  | The abiomed (b)(6) team conducted a literature review and found a canadian publication from a 2013 abiomed impella case. the team had admitted a patient in cardiogenic shock and placed an impella 5.0 via the axillary artery for hemodynamic support. while on support a pericardial effusion was noted as was impella pump malposition. the patient suffered from an acute mitral valve regurgitation (mr) due to chordal rupture. the thought expressed in the publication was that the device coming out of position allowed for mitral valve leaflet damage and mr. the patient expired with multiorgan failure, when deemed too critical for intervention. | [5 empty] | The pump product was discarded in 2013. no data logs can be retrieved. the literature was reviewed. the patient injury was most likely due to the pump being out of position and interacting with the mitral apparatus. the cause of the improper positioning was unknown.
37355 | 1220648-2020-00489 | 10206143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10206143&pc=OZD | 06/02/2020 | 06/29/2020 |  | 06/29/2020 |  |  | 06/02/2020 | 06/29/2020 |  |  | 01/31/2022 |  | 06/29/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1471210 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/02/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white female patient presenting with post cardiotomy cardiogenic shock. after mitral valve repair (post coronary artery bypass surgery), the patient was unable to sustain hemodynamics. as such, the impella cp optical was chosen for support and inserted into the right femoeral artery without any reported issues. during support, the patient developed poor distal circulation in the impella leg. as treatment, the device was removed and surgical repair was performed to rectify damage of the artery that occurred at the access site. it was unclear as to the specific damage that had occurred. after addressing the circulation issues and access site damage, a new impella was inserted into the patient's right axillary artery without any further issues. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37356 | 1220648-2020-00488 | 10206144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10206144&pc=OZD | 06/04/2020 | 07/29/2020 |  | 06/29/2020 |  |  | 06/05/2020 | 06/29/2020 |  |  | 05/31/2022 |  | 06/29/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1476801 |  | 005060 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 56 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 06/05/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient had impella 5.0 pump inserted in the right axillary for acute myocardial infarction (ami)/cardiogenic shock (cgs). the pump stopped, three (3) unsuccessful attempt to turn the pump on before changing to another pump. | [5 empty] | Device not returned by customer. the impella 5.0 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed. | Correction on d10. the product was returned and evaluated. there was no biomaterial found in the outflow cage, cannula, or entraining the impeller. the electrical resistance of the motor phases were measured and found to be within specification. the pump was able to start in the lab and ran within specification. therefore, the root cause of the pump being unable to start was unable to be determined as it could not be reproduced.
37357 | 1220648-2020-00484 | 10190875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10190875&pc=OZD | 05/26/2020 | 07/09/2020 |  | 06/24/2020 |  |  | 05/26/2020 | 06/24/2020 |  |  | 10/31/2021 |  | 06/24/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED EUROPE |  |  | GM |  | 1454788 |  | 0048-0032 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage, Cerebral | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 05/26/2020 | * | Invalid/Incomplete |  |  | The hospital admitted a patient for a coronary artery bypass surgery (cabg). the team chose to place an impella cp prior to the surgery for hemodynamic support. after 5 days of support the patient was seen to have suffered from a neurologic event (cva). the etiology was unknown. there was speculation of a thrombotic event due to the prior rapid atrial fibrillation and supraventricular tachycardia that was treated with multiple shocks. though the use of impella could not be ruled out as a contributing factor/cause. after another 2 days the patient was diagnosed with heparin induced thrombocytopenia (hit+) and the impella pump stopped. the pump was able to be restarted, though only to stop again in the subsequent days. the pump was eventually explanted after a total of 10 days of hemodynamic support. she has survived the cva. | [5 empty] | Since the initial report was filed. the investigation has concluded. the product was not returned for analysis. imaging was not shared for review. there was no evidence of the impella pump causing the stroke event. the pump was used beyond indicated use. the failure mode will be monitored and trended. | To date the investigation into the event is on-going. upon completion of the analysis a follow-up report will be filed.
37358 | 1220648-2020-00485 | 10190903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10190903&pc=OZD | 05/28/2020 | 07/10/2020 |  | 06/24/2020 |  |  | 05/28/2020 | 06/24/2020 |  |  | 03/31/2022 |  | 06/24/2020 |  | TEMPORARY NON ROLLER TYPE LEFT HEART SUPPOR BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1475274 |  | 0550-0008 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO. | [22 empty] | Initial submission | Followup |  |  |  | 05/28/2020 | * | Invalid/Incomplete |  |  | A patient was transferred into the tertiary medical center suffering from a myocardial infarction. she had a coronary angioplasty, dc shocks in the catheterization lab, and an intra-aortic balloon pump (iabp) placed at the first hospital. upon admission to the tertiary center the iabp was removed and an impella 5.5 was placed. due to her poor condition, she was being shocked at 5.5 insertion, she also had ecmo placed. when her condition was not improving she was taken to the operating department for further evaluation and found to have a left ventricle (lv) perforation. the tissues was noted to be "like butter". the impella was seen thought to have perforated the already fragile necrotic tissue of the lv. the laceration was repaired and the patient remained on support with the impella 5.5. the pump supported for 6 days until care was withdrawn and she expired. | [5 empty] | The investigation into the lv perforation is on-going at this time. the product was discarded at the time of her death. upon completed analysis of the event, a final report will be filed. | Since the initial report was filed, the investigation has come to a conclusion. the impella pump product was not returned for analysis, nor were the data logs from the console. as such only the clinical report was available for consideration. the cause of the lv laceration was determined to be patient condition, as the medical team shared with the abiomed team. the patient condition of necrotic heart tissue allowed for the lv perforation. the failure mode will be monitored and trended.
37359 | 1220648-2020-00481 | 10160197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10160197&pc=OZD | 05/21/2020 | 06/16/2020 |  | 06/16/2020 |  |  | 05/23/2020 | 06/16/2020 |  |  | 12/31/2021 |  | 06/16/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1461512 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/23/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the right femoral for cardiomyopathy. it was reported that hemolysis occurred immediately after the introduction of impella. repositioning and flow rate adjustment were performed but was unsuccessful. ultimately, the patient was administered haptoglobin and continuous hemofiltration/hemodiafiltration (chdf) was performed. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37360 | 1220648-2020-00482 | 10160209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10160209&pc=OZD | 05/23/2020 | 06/16/2020 |  | 06/16/2020 |  |  | 05/23/2020 | 06/16/2020 |  |  | 01/31/2022 |  | 06/16/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/23/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for myocarditis. it was reported the patient¿s hematuria and ldh levels increase, hb decrease and free hemoglobin increase after three days of being on impella support. hematuria was confirmed visually and with pfhg reading. the patient was administered mannitol, adenine, and phosphate (map) and haptoglobin and the position of impella was adjusted. by adjusting the position and lowering the auxiliary level the issue was resolved. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37361 | 1220648-2020-00477 | 10153360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10153360&pc=OZD | 05/18/2020 | 06/30/2020 |  | 06/15/2020 |  |  | 05/18/2020 | 06/15/2020 |  |  | 02/28/2022 |  | 06/15/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1472474 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Mechanical Problem | [6 empty] | Perforation of Vessels | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 05/18/2020 | * | Invalid/Incomplete |  |  | The medical team placed an impella cp via the femoral artery. the cp was to support the patient for a high risk coronary angioplasty. shortly after insertion the pump stopped, troubleshooting measures were taken, though eventually the team chose to replace the pump. after the exchange the team observed a common femoral artery and iliac artery dissection with distal embolization. they treated the dissection by angioplasty and explanted the cp pump. | [5 empty] | Since the initial report was filed, the investigation has come to a conclusion. the medical center did not return the product for analysis nor, did they return imaging as requested. the root cause of the vascular damage was most likely the patient condition. the patient was reported to have peripheral arterial disease and was noted to have calcified anatomy. the native anatomy was either damaged by the introducer sheath or by the dry closure method utilized. the failure mode will be monitored and trended. | The investigation is on-going at this time. upon completion a final report will be filed.
37362 | 1220648-2020-00478 | 10153367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10153367&pc=OZD | 05/01/2020 | 06/15/2020 |  | 06/15/2020 |  |  | 05/19/2020 | 06/15/2020 |  |  | 08/31/2021 |  | 06/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1441282 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 35 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/19/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian female patient had impella cp pump inserted in the right femoral for cardiogenic shock and severe heart failure due to fulminant myocarditis without any problem. the patient had bleeding at the insertion site. as treatment 14 units of mannitol, adenine, and phosphate (map), 10 units of fresh frozen plasma (ffp) and 80 units of platelets (pc) was given in one week and surgical repair was used. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37363 | 1220648-2020-00479 | 10153372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10153372&pc=OZD | 05/17/2020 | 06/15/2020 |  | 06/15/2020 |  |  | 05/17/2020 | 06/15/2020 |  |  | 11/30/2021 |  | 06/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1466419 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/17/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male patient had impella cp optical pump inserted in the right femoral for acute myocardial infarction (ami) with shock. the patient¿s right leg pulses were faint then gone, patient was brought to the cath lab and a distal perfusion catheter was placed, issue did not resolve and the patient needed further support so cp optical pump was removed and 5.0 pump was inserted. at the time of the vascular closure a fogerty balloon was used to remove any thrombus at the access site then the following day the patient received a below the knee amputation (bka) to the right lower extremity (rle). | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37364 | 1220648-2020-00480 | 10155137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10155137&pc=OZD | 05/22/2020 | 06/15/2020 |  | 06/15/2020 |  |  | 05/22/2020 | 06/15/2020 |  |  | 01/31/2022 |  | 06/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1469839 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemostasis | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/22/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient had impella cp optical pump inserted in the left femoral for acute myocardial infarction (ami) with shock without issue. patient had hematoma and bleeding around the repositioning sheath and as a result a total of ten (10) plus units of red blood cells (rbcs) was administered. the family decided to put the patient on do not resuscitate (dnr) prior to withdrawing from care. | [5 empty] | Device discarded by customer. the impella cp optical was discarded by the customer therefore a failure analysis investigation cannot be completed.
37365 | 1220648-2020-00476 | 10140331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10140331&pc=OZD | 04/24/2020 | 06/10/2020 |  | 06/10/2020 |  |  | 05/12/2020 | 06/10/2020 |  |  | 04/30/2021 |  | 06/10/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1415058 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/12/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that after percutaneous coronary intervention (pci) was performed under the support of impella the patient experienced bleeding from access site. as a result, surgical hemostasis was performed in the catheterization room. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37366 | 10133119 | 10133119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10133119&pc=PBL | 02/10/2020 | 05/26/2020 | 05/26/2020 | 06/09/2020 | 05/26/2020 | 06/09/2020 |  | 06/09/2020 |  |  |  |  | 06/09/2020 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PBL |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  | 227389 |  |  |  | Yes | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 25915 DA | [5 empty] | Death | Yes | Yes | Break | [6 empty] | Foreign Body In Patient | [12 empty] | Death | [28 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella device broke in the left ventricular (lv) after it was positioned, physician tried to reposition the device and exerted pressure that made the device loop in the lv. then he tried to pull the device back to un-loop it, device broke. manufacturer response for impella cp 3.5, impella cp 3.5 (per site reporter), unknown.
37367 | 1220648-2020-00474 | 10133664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10133664&pc=OZD | 05/28/2020 | 06/09/2020 |  | 06/09/2020 |  |  | 05/28/2020 | 06/09/2020 |  |  | 02/28/2022 |  | 06/09/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE |  |  | GM |  | 1467018 |  | 0048-0002 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 83 YR | [5 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Great Vessel Perforation | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 05/28/2020 | * | Invalid/Incomplete |  |  | The medical team in (b)(6) placed an impella cp for a patient admitted in cardiogenic shock. the implant of the impella was done independent of an abiomed representative. the pump was placed via the left femoral artery for hemodynamic support during a left main and right coronary arteries angioplasty. the placement of the pump was difficult due to a calcified aortic valve. shortly after implant the team noted contrast spillage and a false lumen at the ascending aorta. the team transferred the patient to another medical center for a vascular repair to the aorta. the pump was explanted and discarded at the second medical center. the cause of the damage to the ascending aorta was thought to be perhaps due to the impella pump placement. | [5 empty] | The complainant in the eu discarded the product and data log retrieval is not possible with covid19 regulations. a root cause of the ascending aorta perforation is not determined. the failure mode will be monitored and trended.
37368 | 1220648-2020-00475 | 10135456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10135456&pc=OZD | 05/15/2020 | 07/08/2020 |  | 06/09/2020 |  |  | 05/15/2020 | 06/09/2020 |  |  | 10/31/2021 |  | 06/09/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1453055 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 05/15/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral. the pump was used for seven (7) days, when the pump was weaned and removed the computerized tomography (ct) revealed a thrombus and as a result both femoral had ischemia. an emergency thrombectomy was performed, endovascular treatment, and bilateral thrombus aspiration was performed on both femoral. | [5 empty] | After the pump was weaned and removed, lower limb ischemia reportedly occurred. a blood clot was confirmed in the lower limb via ct scan and the vessel was aspirated. along with a thrombus, a foreign substance about 5 cm long was also removed. images returned confirm the clot and substance. blood and biomaterial looked to be adhered to a thin piece of plastic on the sample that was received in a separate vial. there was also what looked like threads or fibers adhered to the sample returned. it was unknown where the foreign material or fibers came from. the returned pump did not show damage or evidence that the foreign material came from impella. the introducer was not returned for review but the foreign material did not resemble any components of the introducer. data logs were not returned. the pump supported the patient successfully for 5+ days which is beyond design intent. the root cause of the thrombus was not definitively determined, although the coagulation conditions could be a factor. the root cause of the foreign substance was unknown. | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37369 | 1220648-2020-00469 | 10130983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10130983&pc=OZD | 05/12/2020 | 06/08/2020 |  | 06/08/2020 |  |  | 05/12/2020 | 06/08/2020 |  |  | 01/31/2022 |  | 06/08/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1465351 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 61 YR | [5 empty] | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/12/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported patient experienced a decreased in hemoglobin levels after returning to the icu. a ct scan showed bleeding from the impella access site to the retroperitoneum. the patient underwent an emergency operation to stop the bleeding. when the blood vessel was exposed and the access site was visually confirmed, a side leakage of the sheath and the blood vessel was noted. due to cabg, impella was switched to ecmo. no further information was provided. | [5 empty] | The impella cp has not been returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37370 | 1220648-2020-00470 | 10130989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10130989&pc=OZD | 12/11/2019 | 06/08/2020 |  | 06/08/2020 |  |  | 05/12/2020 | 06/08/2020 |  |  | 05/31/2021 |  | 06/08/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1424965 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/12/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 2.5 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed a hemorrhage at the impella access site. the reported information states the event was cause by the impella device, patient status, and concomitant medication. the patient received 10 units of red concentrated cells, 10 units of fresh frozen plasma and 10 units of platelet count. additionally, surgical hemostasis was applied, and treatment was escalated to impella 5.0. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37371 | 1220648-2020-00471 | 10130995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10130995&pc=OZD | 05/13/2020 | 07/08/2020 |  | 06/08/2020 |  |  | 05/14/2020 | 06/08/2020 |  |  | 12/31/2021 |  | 06/08/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 PUMP 550 SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 1465546 |  | 0550-0001 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 71 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 05/14/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) old female patient had impella 5.5 pump inserted in the right axillary. the impella stopped while weaning out of the patient. | [5 empty] | The impella 5.5 pump experienced a purge filter break, high purge pressure, and low purge flow during the case. the team performed a purge bypass which initially solved the problem. the pump eventually stopped during the weaning process and was not replaced as the patient was stable. blood had pushed its way back into the purge line creating purge system blocked alarms, which were confirmed by the data logs that was returned and as shown by field photos. the returned pump confirmed blood in the purge gap and in the purge line as described in the clinical account. the filter had a noticeable crack and leaked when pressurized with water. the root cause of the pump stop was a blood ingress into the motor caused by a purge system break. it was unknown how the break occurred. | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37372 | 1220648-2020-00472 | 10131008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10131008&pc=OZD | 05/14/2020 | 06/08/2020 |  | 06/08/2020 |  |  | 05/14/2020 | 06/08/2020 |  |  | 11/30/2021 |  | 06/08/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  |  |  | 1466689 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 46 YR | [5 empty] | Malfunction | No | Yes | No Apparent Adverse Event | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Hospitalization | [28 empty] | Initial submission |  |  |  |  | 05/14/2020 | * | Invalid/Incomplete |  |  | A patient who was covid19+ was found to be suffering from myocarditis while hospitalized. the medical team chose to place an impella 5.5 via the axillary artery as she as in cardiogenic shock as well. the pump ran for just over 12 hours when the pump stopped. the pump had been running and supporting the patient successfully til stoppage occurred. the team noted that the patient did not have heparin anticoagulation running for many hours of the support. the pump was explanted and not replaced. she was stable enough without additional hemodynamic support. | [5 empty] | The 5.5 pump was returned for analysis, as were the pump data logs. the analysis revealed the pump to have stopped due to blood ingress, most likely from ingestion of biomaterial during the support. biomaterial was found around the impeller of the pump and the logs revealed a sudden event that caused the purge to be blocked. the failure mode will be monitored and trended.
37373 | 1220648-2020-00473 | 10131015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10131015&pc=OZD | 03/26/2018 | 06/08/2020 |  | 06/08/2020 |  |  | 03/26/2018 | 06/08/2020 |  |  | 11/30/2019 |  | 06/08/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE |  |  | GM |  | 1323506 |  | 0048-0020 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | LITERA |  |  |  |  | 19 YR | [5 empty] | Injury | Yes | No | Difficult to Insert | [6 empty] | Perforation | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 03/26/2018 | * | Invalid/Incomplete |  |  | Discovered in a literature review back in 2018 a (b)(6) old female patient, admitted with a rejection of her heart transplant, had an access site injury at the impella cp access site. she was known to have small vessels and be in vasospasm. the team observed the dissection and placed a contralateral balloon via the right femoral artery and achieved hemostasis after a 20 minute ballooning. the patient was supported by the cp after this repair. | [5 empty] | The publication was discovered years after the case in a literature review conducted by abiomed personnel. the field team was contacted to try to learn more clinical details of this young patient's support in 2018. no product was returned for review. the root cause of the vascular damage was most likely due to patient condition, as the clinical account reported she had small vasculature and the vessels were in vasospasm. the failure mode will be monitored and trended.
37374 | 1220648-2020-00468 | 10126679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10126679&pc=OZD | 05/08/2020 | 06/07/2020 |  | 06/07/2020 |  |  | 05/08/2020 | 06/07/2020 |  |  | 05/31/2021 |  | 06/07/2020 |  | : TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1424928 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report | [8 empty] | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/08/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) old japanese female patient had impella 2.5 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that hemolysis was confirmed via pfhg levels. the physician attempted to adjust the pumps position via echo; however, was unsuccessful. volume, but the patient had polycythemia, water escaped to the third space and lungs, and respiratory status deteriorated as pulmonary edema. during the evening, the impella device was explanted and the patient was converted to aibp. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37375 | 1220648-2020-00467 | 10113876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10113876&pc=OZD | 05/07/2020 | 06/03/2020 |  | 06/03/2020 |  |  | 05/08/2020 | 06/03/2020 |  |  | 01/31/2022 |  | 06/03/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1465351 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/08/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male patient had impella cp pump inserted in the left femoral for high-risk percutaneous coronary intervention (hrpci). it was reported that a subcutaneous hematoma developed at the insertion site after the impella cp was implanted. the patient received 4 units of mannitol, adenine, and phosphate. per the physician, there were no major adverse event. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37376 | 1220648-2020-00456 | 10098653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10098653&pc=OZD | 04/28/2020 | 06/29/2020 |  | 05/28/2020 |  |  | 04/28/2020 | 05/28/2020 |  |  | 10/31/2021 |  | 05/28/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1450139 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 04/28/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white female patient had impella cp pump inserted in the right femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). overnight the patient developed a retroperitoneal bleeding as a result a covered stent was placed in the iliac along with one (1) unit of fresh frozen plasma (ffp) and three (3) units of packed red blood cells (prbcs) was administered. | [5 empty] | The pump was returned for evaluation and no damage was found to the guide wire repositioning unit (gwru). the introducer was not returned for evaluation. the root cause of the access site bleeding was unable to be determined without additional clinical details. | Device was not returned by customer. the impella cp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37377 | 1220648-2020-00457 | 10098659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10098659&pc=OZD | 04/29/2020 | 05/28/2020 |  | 05/28/2020 |  |  | 04/29/2020 | 05/28/2020 |  |  | 01/31/2022 |  | 05/28/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1471133 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/29/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) old white male patient had impella cp optical pump inserted in the right femoral for acute myocardial infarction (ami) with shock. the physician had to put a fem by fem bypass technique for distal perfusion to treat the oozing from the repositioning sheath. | [5 empty] | Device was not returned by customer. the impella cp optical pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37378 | 1220648-2020-00458 | 10098691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10098691&pc=OZD | 04/29/2020 | 05/28/2020 |  | 05/28/2020 |  |  | 04/29/2020 | 05/28/2020 |  |  | 01/31/2022 |  | 05/28/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1465640 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/29/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old asian male patient had impella 5.0 pump inserted in the right axillary for acute myocardial infarction (ami) with shock. the patient had hemolysis, pfhgb spiked up to 200 then down to 170s-150. as treatment two (2) units of packed red blood cells (prbcs) was administered. | [5 empty] | Device was not returned by the customer. the impella 5.0 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37379 | 1220648-2020-00459 | 10098696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10098696&pc=OZD | 04/30/2020 | 05/28/2020 |  | 05/28/2020 |  |  | 04/30/2020 | 05/28/2020 |  |  | 12/31/2021 |  | 05/28/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1467549 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/30/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old black or african american male patient had impella cp optical pump inserted in the left femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had suspected hemolysis, the pump was repositioned and then removed as hemolysis did not resolve. | [5 empty] | Device not returned by the customer. the impella cp optical pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37380 | 1220648-2020-00460 | 10098708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10098708&pc=OZD | 05/02/2020 | 05/28/2020 |  | 05/28/2020 |  |  | 05/02/2020 | 05/28/2020 |  |  | 01/31/2022 |  | 05/28/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1464387 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/02/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white female patient had impella cp optical pump inserted in the right femoral. the oscor 14frx25cm introducer kinked upon insertion and displacing a previous placed stent. intervention was required to retrieve the displaced stent, however, intervention resulted in inadvertent vascular dissection and occlusion of blood flow to the lower leg. further intervention was needed to repair the dissection and restore partial flow to the leg with a 16fr gore sheath and two (2) units of packed red blood cells (prbcs). the patient received multiple angioplasty attempts along with viabahn stenting to both extremities on the prior day. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37381 | 1220648-2020-00462 | 10098723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10098723&pc=OZD | 04/30/2020 | 05/28/2020 |  | 05/28/2020 |  |  | 04/30/2020 | 05/28/2020 |  |  | 09/30/2021 |  | 05/28/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | DANVERS, MA |  |  | US |  | 1444716 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 59 YR | [5 empty] | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/30/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) male patient had impella cp pump inserted in the left femoral for cardiomyopathy. it was reported the patient developed a hemorrhage at the impella access site during pump support. as a result, the patient was administered 18 units of red blood cells and 8 units of frozen plasma. the impella was successfully weaned and explanted. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37382 | 1220648-2020-00463 | 10098737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10098737&pc=OZD | 11/07/2019 | 05/28/2020 |  | 05/28/2020 |  |  | 05/01/2020 | 05/28/2020 |  |  | 05/31/2021 |  | 05/28/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1424928 |  | 005048 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 57 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/01/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed anemia during pump support. as a result, the patient received 46 units of red concentrated cells and 20 units of fresh frozen plasma. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37383 | 1220648-2020-00461 | 10098757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10098757&pc=OZD | 09/16/2019 | 05/28/2020 |  | 05/28/2020 |  |  | 04/30/2020 | 05/28/2020 |  |  | 10/31/2020 |  | 05/28/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1376920 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/30/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) male patient had impella 2.5 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient went into hemorrhagic shock. the patient underwent cardiac surgery, specific details were not provided. the patient was also administered 22 units of red blood cells and 20 units of platelets. the physician stated that the impella attributed to the hemorrhagic shock and was explanted due to the adverse event. the patient was also on pcps, it is not known which device cause the adverse event. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37384 | 1220648-2020-00464 | 10098779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10098779&pc=OZD | 03/16/2020 | 05/28/2020 |  | 05/28/2020 |  |  | 04/10/2020 | 05/28/2020 |  |  | 09/30/2021 |  | 05/28/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1449886 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 47 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/10/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old black or african american male patient had impella 5.0 pump inserted in the right axillary artery without issue. the patient¿s mitral valve (mv) was torn that lead to mitral regurgitation (mr), the mv was surgically repaired by a cardiothoracic surgeon. | [5 empty] | Device not returned by customer. the impella 5.0 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37385 | 1220648-2020-00465 | 10098800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10098800&pc=OZD | 05/06/2020 | 06/29/2020 |  | 05/28/2020 |  |  | 05/06/2020 | 05/28/2020 |  |  | 02/28/2022 |  | 05/28/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED EUROPE |  |  | GM |  | 1470101 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | IMPELLA RP | [22 empty] | Initial submission | Followup |  |  |  | 05/06/2020 | * | Invalid/Incomplete |  |  | A male patient was admitted in cardiogenic shock suffering from an acute myocardial infarction due to right coronary artery disease. an impella cp was placed after the patient deteriorated with no reflow to the coronary artery and active cpr being performed. the patient's leg began to suffer limb ischemia. the team chose to place another catheter for distal perfusion to the limb via the contralateral limb. the leg color began to improve and the impella remained on for support. in addition to the cp the patient had a right sided rp impella placed. his prognosis was not a good one. he had many medications on board that were causing vasoconstriction and was transfused with blood products . | [5 empty] | Since the initial report was filed, the investigation has concluded into the reported limb ischemia. the product was not returned for investigation. the cause of the ischemia was patient condition since both legs experienced the documented and treated ischemia. the patient was on reported vessel restricting medication which caused the limb ischemia. the failure mode will be monitored and trended. | The investigation into the limb ischemia is on-going at this time. upon completion of the investigation a final report will be submitted.
37386 | 1220648-2020-00466 | 10098805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10098805&pc=OZD | 04/25/2020 | 06/29/2020 |  | 05/28/2020 |  |  | 04/25/2020 | 05/28/2020 |  |  | 01/31/2022 |  | 05/28/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1469901 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 04/25/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male patient had impella cp optical pump inserted in the right femoral for acute myocardial infarction (ami) with shock. there was blood in the patient¿s urine and as a result the pump was removed due to suspected hemolysis. | [5 empty] | Hemolysis was not mentioned in the clinical account, but a conversation with the abiomed representative mentions that the pump was removed due to suspected hemolysis, although labs were not drawn and hemolysis not confirmed. the cause of the hemolysis was not determined. | Device was not returned by customer. the impella cp optical pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37387 | 1220648-2020-00455 | 10091579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10091579&pc=OZD | 03/10/2020 | 06/08/2020 |  | 05/27/2020 |  |  | 03/10/2020 | 05/27/2020 |  |  | 10/31/2021 |  | 05/27/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1453055 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 03/10/2020 | * | Invalid/Incomplete |  |  | The medical team in (b)(6) admitted a patient for a planned bypass surgery, but instead the team performed a high risk angioplasty of multiple coronary arteries. the plan was for rotational atherectomy of the left main coronary and left anterior descending branch, and additionally the right coronary artery. due to the high risk, an impella cp was placed via the right femoral artery. the angioplasty was completed and then the team removed the impella 14fr introducer and advanced the pump's repositioning sheath. upon that removal and placement there was pulsatile blood loss. the medical team consulted and chose to remove the impella after surgical treatment of the bleed site. the exact method of treatment is unknown at this time. | [5 empty] | The root cause of the access site bleeding was unable to be determined. no data logs or product was returned from japan. the clinical details of the support were not extensive enough to come to a root cause. the team in japan was able to share that 300ml of blood was lost. the failure mode will be monitored and trended. | The complainant in (b)(6) discarded the product in (b)(6) 2020 when the event occurred. clinical questions have been asked but, to date no response has occurred. when the investigation is completed a final report will be filed.
37388 | 1220648-2020-00454 | 10089351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10089351&pc=OZD | 05/17/2020 | 07/06/2020 |  | 05/26/2020 |  |  | 05/17/2020 | 05/26/2020 |  |  | 11/30/2021 |  | 05/26/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1451735 |  | 0048-0002 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Hospitalization | Required Intervention |  |  |  |  | CRRT HEMOFILTRATION. | [22 empty] | Initial submission | Followup |  |  |  | 05/17/2020 | * | Invalid/Incomplete |  |  | The complainant in (b)(6) admitted a patient with myocarditis. to give the patient hemodynamic support an impella cp was placed via the right femoral artery. when on support the limb with the impella experienced limb ischemia that was treated with a contralateral leg external bypass. within 24 hours the pump was explanted and replaced due to concern of hemolysis. a second impella was successfully placed for the continuation of the support. | [5 empty] | Since the original report was filed the investigation has completed. the impella pump was returned for analysis but, not the introducer sheath. when the pump was examined biomaterial was round entraining the impeller. there was no damage or rough surfaces to the pump's repositioning sheath. the data logs were retuned and analyzed. the logs did reveal that the pump was not in an optimal position, and did reveal an alarm of the pump on valve. the hemolysis was due to the pump being improperly positioned. the cause of the ischemia was most likely due to the patient's known peripheral arterial disease. the failure modes of hemolysis and limb ischemia will be monitored and trended. | The complainant in (b)(6) has not yet returned the pump product nor the imaging as requested. the data logs have been returned for analysis. the analysis is on-going at this time. upon completion, a final report will be submitted.
37389 | MW5094552 | 10074428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10074428&pc=OZD | 05/07/2020 | 05/15/2020 |  | 05/19/2020 |  |  |  | 05/19/2020 |  |  |  |  | 05/19/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED INC. |  |  | US | [5 empty] | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 71 YR | [5 empty] | Malfunction | No | Yes | Crack | No Device Output | [5 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Purge rate suddenly increased, placement signal indicated impella no longer in correct place, micron filter was cracked and unable to be changed without changing entire pump, pump then suddenly stopped. fda safety report id# (b)(4).
37390 | 1220648-2020-00445 | 10063411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10063411&pc=OZD | 04/18/2020 | 05/15/2020 |  | 05/15/2020 |  |  | 04/18/2020 | 05/15/2020 |  |  | 08/31/2021 |  | 05/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1448311 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/18/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella cp optical pump inserted in the left femoral. the pump was removed due to suspected hemolysis. the patient expired two (2) days after the pump was removed due to multi-organ failure (mof). | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37391 | 1220648-2020-00446 | 10063470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10063470&pc=OZD | 01/18/2020 | 05/15/2020 |  | 05/15/2020 |  |  | 04/22/2020 | 05/15/2020 |  |  | 09/30/2021 |  | 05/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC |  |  | US |  | 1446739 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/22/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for myocarditis. it was reported the patient developed limb ischemia during pump support. per the pms report impella was not the only direct cause of lower limb ischemia, circulatory failure due to shock condition and right sma being cto was also a cause. the original vascular lesion is unknown. to treat the injury the patient received blood transfer via side branch of pcps. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37392 | 1220648-2020-00447 | 10063476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10063476&pc=OZD | 01/01/2020 | 05/15/2020 |  | 05/15/2020 |  |  | 04/22/2020 | 05/15/2020 |  |  | 09/30/2021 |  | 05/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1446479 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/22/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed limb ischemia at the left and right groin where impella and pcps were inserted. to treat the injury the patient received blood transfer via side branch of pcps. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37393 | 1220648-2020-00448 | 10063508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10063508&pc=OZD | 04/24/2020 | 06/15/2020 |  | 05/15/2020 |  |  | 04/24/2020 | 05/15/2020 |  |  | 11/30/2021 |  | 05/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT SIDED HEART PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1466927 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Blood Loss | [11 empty] | Hospitalization | Required Intervention |  |  |  |  | TAVR COREVALVE | [22 empty] | Initial submission | Followup |  |  |  | 04/24/2020 | * | Invalid/Incomplete |  |  | A patient was admitted in cardiogenic shock and suffering from an acute myocardial infarction. earlier in the month the patient had been admitted and had a tavr corevalve placed. the valve replacement was done with the tavr corevalve product, manufactured by medtronic. the patient was supported by an intra-aortic balloon pump (iabp) upon this second admission. when it did not support the patient effectively, it was explanted and an impella cp was placed via the left femoral artery. during his support the access site of the impella was oozing blood and he was also experiencing signs and symptoms of hemolysis. the patient received an infusion of fluids and blood products. after 2 days the care had to be escalated to the impella 5.0. | [5 empty] | Since the initial report was filed the investigation has completed. the pump product was returned and tested on the bench top. the root cause of the bleeding and presumed hemolysis was unable to be determined. there was a thought shared by the medical team that the use of tpa could have caused the excessive blood loss. the blood products given for the blood loss also benefitted the patient for the presumed hemolysis. the bench top testing did reveal however that the tavr had interacted with the impella. the impeller was found to be damaged. this damage lead to the low flows experienced during the case support. the failure mode will be monitored and trended. | The complainant has returned the product for analysis. the analysis is on-going at this time. upon completion of the investigation, a final report will be filed.
37394 | 1220648-2020-00449 | 10063532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10063532&pc=OZD | 09/05/2019 | 05/15/2020 |  | 05/15/2020 |  |  | 04/24/2020 | 05/15/2020 |  |  | 02/28/2021 |  | 05/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1406291 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 57 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/24/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient experienced a hemorrhage the device's access site. additionally, the patient also developed hemolysis, pfhg levels were not provided. the patient was administered 4 units of concentrated red cells. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37395 | 1220648-2020-00450 | 10063534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10063534&pc=OZD | 09/18/2019 | 05/15/2020 |  | 05/15/2020 |  |  | 04/24/2020 | 05/15/2020 |  |  | 02/28/2021 |  | 05/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1406291 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/24/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed a hemorrhage at the device's access site. as a result, the patient received 8 units of concentrated red blood cells. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37396 | 1220648-2020-00451 | 10063556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10063556&pc=OZD | 12/10/2019 | 05/15/2020 |  | 05/15/2020 |  |  | 04/24/2020 | 05/15/2020 |  |  | 08/31/2021 |  | 05/15/2020 |  | : TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC |  |  | US |  | 1434848 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 37 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | No Code Available | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/24/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the right femoral to unload the left ventricle. it was reported the patient developed a hemorrhage at the device's access site. additionally, the patient developed thrombocytopenia during pump support. as a result, the patient received 50 units of concentrated platelets, 12 units of concentrated red blood cells and 6 units of fresh frozen plasma, wound sutured were also placed. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37397 | 1220648-2020-00452 | 10063561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10063561&pc=OZD | 04/26/2020 | 05/15/2020 |  | 05/15/2020 |  |  | 04/27/2020 | 05/15/2020 |  |  | 10/31/2021 |  | 05/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1453164 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/27/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the left femoral for myocarditis. it was reported the patient was on ecmo + impella cp when the patient lower limb color became poorer than before impella insertion. the patient underwent anterograde puncture on both lower limbs on the day of impella insertion. the blood flow seemed to improve temporarily; however, although it may be due to the bending of the punctured sheath, ischemia (ata closure) of the left lower extremity (impella side) was observed, and there was a suspicion of compartment syndrome. on (b)(6) 2020 a decompression incision on the left lower limb was performed. per the physician the cause of the event is unknown. no further information was provided. | [5 empty] | Disposition of the device is unknown. should any new information be returned, a supplemental mdr will be filed.
37398 | 1220648-2020-00453 | 10063562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10063562&pc=OZD | 01/19/2020 | 05/15/2020 |  | 05/15/2020 |  |  | 04/23/2020 | 05/15/2020 |  |  | 10/31/2021 |  | 05/15/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1450243 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 87 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/23/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) female patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed a hemorrhage at the right groin. as a result, the patient received 20 units of concentrated platelets, 10 units of fresh frozen plasma and 12 units of mannitol, adenine, and phosphate (map). no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37399 | 1220648-2020-00443 | 10051635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10051635&pc=OZD | 01/28/2020 | 05/13/2020 |  | 05/13/2020 |  |  | 04/16/2020 | 05/13/2020 |  |  | 04/30/2021 |  | 05/13/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1441946 |  |  |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/16/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year-old (b)(6) female patient had impella 2.5 pump inserted in the right femoral, indication was not provided. it was reported after the impella device was explanted the patient experienced cerebrovascular accident (cva). as a result, the patient received drug treatment, specifics were not provided. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37400 | 1220648-2020-00444 | 10053158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10053158&pc=OZD | 05/04/2020 | 07/27/2020 |  | 05/13/2020 |  |  | 05/04/2020 | 05/13/2020 |  |  | 01/31/2022 |  | 05/13/2020 |  | TEMPORARY NONROLLER TYPE LEFT SIDE HEART SUPPORT PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1464339 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage, Cerebral | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | ECMO | [21 empty] | Initial submission | Followup |  |  |  | 05/04/2020 | * | Invalid/Incomplete |  |  | A patient with cardiomyopathy was admitted and placed upon impella support. the team chose to escalate care from the impella cp to 5.0 pump after 4 days of support. the impella 5.0 was placed via the right axillary artery. after 2 days of support on the impella 5.0 there was a neurological event. the patient was diagnosed via ct imaging and found to have had a cerebral event with a right sided deficit. the etiology of the event is not known, though impella use could have caused or contributed. | [5 empty] | The device was received and an investigation was conducted. although biomaterial was identified during visual inspection of the device, data log analysis found no elevation in motor current that would suggest the pump had operated out of specification. there is no evidence to indicate the stroke was related to the impella. a review of the device history record found no manufacturing issues or reworks associated with this pump lot. there are no other complaints related to this failure mode associated with pumps from this lot. given that no issues were found during review of the pump, data logs, nor device history record, we can not definitively determine a root cause for the reported neurological event. | The investigation is on-going at this time. upon completion a final report will be submitted.
37401 | MW5094489 | 10055705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10055705&pc=PBL | 04/24/2020 | 05/12/2020 |  | 05/13/2020 |  |  |  | 05/13/2020 |  |  |  |  | 05/13/2020 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA CP |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  |  |  | 230995 |  |  |  | No | [5 empty] | No | No Information |  |  | Voluntary report | [14 empty] | Injury | No | Yes | Device Alarm System | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Pt came to the cath lab for a complex pci, dr. (b)(6) placed an impella catheter for temporary hemodynamic support. upon insertion 'purge cassette pressure alarm" was alerting. the team after unsuccessful troubleshooting reached out to the rep ((b)(6)) who informed the team that it will still work for the short term duration of the procedure. near the end of the procedure after searching for kinks in the line and resetting multiple times with no success, the physician opted to shut it off and proceed. the case was completed and the pt. went back to the unit after a successful pci. the impella device was sequestered for manager review. fda safety report id# (b)(4).
37402 | 1220648-2020-00438 | 10049156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10049156&pc=OZD | 01/28/2020 | 05/12/2020 |  | 05/12/2020 |  |  | 04/13/2020 | 05/12/2020 |  |  | 05/31/2021 |  | 05/12/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1420671 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/13/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed anemia during pump support. as a result, the patient received 2 units of red blood cells. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37403 | 1220648-2020-00439 | 10049177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10049177&pc=OZD | 11/01/2019 | 05/12/2020 |  | 05/12/2020 |  |  | 04/14/2020 | 05/12/2020 |  |  | 11/30/2020 |  | 05/12/2020 |  | : TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1383954 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/14/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella 2.5 pump inserted in the left femoral for myocarditis. it was reported the patient developed renal failure during pump support. it was noted the event was cause by the impella device. the patient received dialysis treatment. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37404 | 1220648-2020-00440 | 10049315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10049315&pc=OZD | 11/29/2019 | 05/12/2020 |  | 05/12/2020 |  |  | 04/15/2020 | 05/12/2020 |  |  | 07/31/2021 |  | 05/12/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1430738 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/15/2020 | * | Invalid/Incomplete |  |  | "the complainant reported a (b)(6) male patient had impella 5.0 pump inserted in the right axillary for circulation manage and to perform radical cure (vsp repair) in two phases. it was reported that on (b)(6) 2019 pci was performed; however, patient¿s recovery was not very good. impella was removed on (b)(6) 2020 and the patient experienced thrombotic complication (non-central venous system). on (b)(6) 2020, patient¿s blood flow was insufficiency and upper right limb was noted. thrombus was removed using a forgaty catheter. per the physician, thrombus flew when the impella was explanted. the event was deemed an impella device-related complication. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37405 | 1220648-2020-00441 | 10049327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10049327&pc=OZD | 01/21/2020 | 05/12/2020 |  | 05/12/2020 |  |  | 04/15/2020 | 05/12/2020 |  |  | 05/31/2021 |  | 05/12/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1424928 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/15/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(4) female patient had impella 2.5 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed hemolysis during pump support. the pumps p levels were lowered from p4 to p2, and the patient received 6 units of mannitol, adenine, and phosphate. the patient concurrently supported by ecmo, dates were not provided. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37406 | 1220648-2020-00442 | 10049346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10049346&pc=OZD | 11/16/2019 | 05/12/2020 |  | 05/12/2020 |  |  | 04/17/2020 | 05/12/2020 |  |  | 05/31/2020 |  | 05/12/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1349348 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/17/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that after the impella device was explanted and a clot was noted at the apex. as a result, the patient received drug treatment. per the physician it is unknown whether it was caused by the impella device. the patient ultimately recovered and mechanical support was no longer needed. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37407 | 1220648-2020-00437 | 10032242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10032242&pc=OZD | 04/08/2019 | 05/07/2020 |  | 05/07/2020 |  |  | 04/08/2020 | 05/07/2020 |  |  | 05/31/2020 |  | 05/07/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1349348 |  |  |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 87 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/08/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male patient had impella 2.5 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported during pump support the patient experienced cerebrovascular accident (cva). as a result, the patient was received non cardiac surgery, specifics were not provided. the patient ultimately recovered and mechanical support was no longer needed. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37408 | 1220648-2020-00432 | 10023896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10023896&pc=OZD | 02/03/2020 | 05/05/2020 |  | 05/05/2020 |  |  | 04/06/2020 | 05/05/2020 |  |  | 08/31/2021 |  | 05/05/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1444039 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/06/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 2.5 pump inserted in the left femoral for myocarditis. it was reported the patient experienced a hemorrhage at the left groin. the patient received 30 units of mannitol, adenine, and phosphate and 20 units if platelet count. impella was later successfully weaned and explanted. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37409 | 1220648-2020-00433 | 10023911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10023911&pc=OZD | 01/16/2020 | 05/05/2020 |  | 05/05/2020 |  |  | 04/07/2020 | 05/05/2020 |  |  | 08/31/2021 |  | 05/05/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1430528 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/07/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for cardiogenic shock due to acute exacerbation of chronic heart failure. it was reported the patient developed hemolysis during pump support. pfhg levels were not provided. the patient received 6 units of red blood cells. impella was later successfully weaned and explanted. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37410 | 1220648-2020-00434 | 10026136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10026136&pc=OZD | 04/14/2020 | 05/05/2020 |  | 05/05/2020 |  |  | 04/14/2020 | 05/05/2020 |  |  | 12/31/2021 |  | 05/05/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1463225 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 85 YR | [5 empty] | Injury | Yes | Yes | Material Separation | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/14/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) white male patient presenting for pre-operative hemodynamic support with the impella cp optical. against best suggested best practice, the physician chose to insert the impella without 0.018" guide wire. the catheter became kinked and attempts to straighten resulted in a separation between the cannula and inlet/pigtail. the fragmented piece remained in the patient. access was gained through the patient's opposing femoral artery and the remainder of the device extracted using a surgical snare and forceps. no vascular damage was endured. a second impella was opened and inserted without issue; using the 0.018" guide wire. | [5 empty] | The impella cp optical was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37411 | 1220648-2020-00435 | 10026138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10026138&pc=OZD | 01/16/2009 | 10/09/2020 |  | 05/05/2020 |  |  | 06/06/2019 | 05/05/2020 |  |  |  |  | 05/05/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 005042 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | LITERATURE | [5 empty] | 32 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 06/06/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella 2.5 pump. the pump was inserted without any reported issue and the patient was transferred to another hospital. when arriving at the new hospital, the patient was "severely vasoplegic requiring vasopressors and hydrocortisone, eventually responding to methylene blue." inspection of the heart, under transesophageal echocardiogram found a perforation in the inferolateral wall. the device was pulled back into the aorta and suture were applied to repair to perforation. ec-vad was placed thereafter. ultimately, care was withdrawn and patient expired. this event was identified during a review of a case reported published by wiley. | [5 empty] | Correction: g(4): upon review of this complaint, it was determined the incorrect aware date was provided in the initial report. the aware date should be 2020-04-13. | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37412 | 1220648-2020-00436 | 10026143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10026143&pc=OZD | 04/23/2020 | 05/05/2020 |  | 05/05/2020 |  |  | 04/23/2020 | 05/05/2020 |  |  | 12/31/2021 |  | 05/05/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1467502 |  | 0048-0024 |  | No | [5 empty] | Y | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombus | Atrial Perforation | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/23/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) hispanic male patient presenting with cardiomyopathy for hemodynamic support with the impella cp optical device. the patient was inserted with this device at (b)(6) and then transferred to (b)(6) for further support. approximately 8 to 12 hours after arrival at transfer hospital, the patient was exhibiting signs of hemolysis and a low partial thromboplastin time (ptt). the pump was removed and a damaged profunda was identified. within the profunda, a clot was found. as treatment, the vascular repair was administered and the clot was removed. it should be noted that the physician attributed the damaged profunda was caused during placement of the device and the hemolysis was caused by the patient's sub-therapeutic ptt. the patient was ultimately escalated to an impella 5.5 without any further issues reported. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37413 | 1220648-2020-00430 | 10016089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10016089&pc=OZD | 11/10/2019 | 04/30/2020 |  | 04/30/2020 |  |  | 04/02/2020 | 04/30/2020 |  |  | 02/28/2021 |  | 04/30/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1418858 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 85 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemostasis | Blood Loss | [10 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/02/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6)-year-old (b)(6) female patient had impella 5.0 pump inserted in the right femora for assist circulation due to vsp. it was reported the patient developed a hemorrhage at the abdominal wall. as a result, the patient received 8 units of red blood cells. hemostasis was applied by catheter embolization and the hematoma was removed. per the physician, the impella and guidewire did not contact the patient's abdominal wall; however, there was multiple bleeding seen at the peripheral branch of the iliac artery where the impella was insertion. this was confirmed by angiography. the event is unlikely to occurred if impella did not consume platelets or von willebrand factor. patient's anticoagulation regimen was not provided. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37414 | 1220648-2020-00431 | 10016255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10016255&pc=OZD | 03/28/2020 | 06/15/2020 |  | 04/30/2020 |  |  | 04/03/2020 | 04/30/2020 |  |  | 12/31/2020 |  | 04/30/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1400793 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 22 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Septic Shock | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 04/03/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 5.0 pump inserted in the right axillary for cardiomyopathy. it was reported the patient's cardiac functioned improved once impella 5.0 was inserted; however, the patient experienced a fever which is believed to have been due to long-term detention of impella. the patient developed septic shock. the physician stated the infection subsided shortly after the removal of the device. patient was administered (b)(6) for drug administration which were replaced when the impella device was removed. the physician felt the causal relationship with the impella was unknown. the patient was safely weaned and there were no troubles when removing the device. after the removal, the heart failure treatment was continued by intravenous drip, but the hemodynamics were stable. | [5 empty] | The impella 5.0 and data logs were returned for review. there were no issues with the product or data logs relating to the failure mode. the cause of the septic shock on the patient was unable to be determined. there was no evidence indicate the relation between the impella and the septic shock. all impella products were sterilized before sending to the customer. dhr was reviewed and found no manufacturing related issues with this lot. there are no other complaints related this failure mode in this lot. infection is listed in the tempubunsho as a potential adverse event that may be associated with the use of the impella device. abiomed will continue to monitor these types of events. the manufacturer is closing the file on this event. | The impella 5.0 was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37415 | 1220648-2020-00427 | 10008906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10008906&pc=OZD | 11/11/2019 | 10/09/2020 |  | 04/29/2020 |  |  | 11/11/2019 | 04/29/2020 |  |  |  |  | 04/29/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, US | [7 empty] | UNKNOWN |  | 0048-0003 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 15 YR |  |  | 69 |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 11/11/2019 | * | Invalid/Incomplete |  |  | An article from ¿utilization of an abiomed impella device as a rescue therapy for acute ventricular failure in a fontan patient¿ reported a (b)(6) years old female patient, (b)(6) kg, with hypoplastic left heart syndrome had impella cp pump inserted in the left axillary for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient developed worsening suspected hemolysis and suspected thrombocytopenia. the patient developed anuric renal failure requiring continuous renal replacement therapy (crrt) and hepatic injury that precluded durable ventricular assist device (vad) or cardiac candidacy. the patient died 89 days after admission. | [5 empty] | Correction: g(4): upon review of this complaint, it was determined the incorrect aware date was provided in the initial report. the aware date should be (b)(6) 2020. | Manufacturing name, city, and state; expiration date and unique identifier (udi)#; these fields were left empty as we do not know where the pump was manufacture. we could not identify the pump's serial number in the account's history. device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37416 | 1220648-2020-00425 | 10008909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10008909&pc=OZD | 04/06/2020 | 05/28/2020 |  | 04/29/2020 |  |  | 04/06/2020 | 04/29/2020 |  |  | 12/31/2021 |  | 04/29/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1463387 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 60 YR | [5 empty] | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | Myocardial Infarction | Electric Shock | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 04/06/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient had impella cp optical pump inserted in the right femoral artery (rfa) without difficulty for acute myocardial infarction (ami) with shock. upon advancing the repo device and advancing completely into the left ventricle (lv) the cannula kinked, the physician tried to remove the kink by pulling into the descending aorta, but the kink remained. the physician then had to use a snare that was inserted via right radial access. the pigtail snared and cannula was straightened out. support resumed without issue. | [5 empty] | The impella cp optical was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The pump and data logs was returned for evaluation. the data log review demonstrated a ventricular placement signal (ps) and a loss of motor current pulsatility, indicating that the pump was ventricularized shortly after insertion. the pump was evaluated confirming a cannula kink two centimeters from the inflow cage. the root cause of the cannula kink was the physician accidentally advancing the pump completely into the left ventricle upon advancing the repositioning sheath.
37417 | 1220648-2020-00426 | 10008922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10008922&pc=OZD | 11/11/2019 | 10/09/2020 |  | 04/29/2020 |  |  | 11/11/2019 | 04/29/2020 |  |  | 08/31/2018 |  | 04/29/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1259924 |  | 005042 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 14 YR |  |  | 69 |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 11/11/2019 | * | Invalid/Incomplete |  |  | An article from ¿utilization of an abiomed impella device as a rescue therapy for acute ventricular failure in a fontan patient¿ reported a (b)(6) female patient, (b)(6) kg, with hypoplastic left heart syndrome had impella 2.5 pump inserted in the right femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). after nine (9) days of support the patient began to develop signs of right lower extremity ischemia and microthrombi within the impella apparatus, so the transfemoral device was removed. | [5 empty] | Correction: g(4): upon review of this complaint, it was determined the incorrect aware date was provided in the initial report. the aware date should be (b)(6)2020. | Device discarded by customer. the impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37418 | 1220648-2020-00428 | 10008929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10008929&pc=OZD | 04/10/2020 | 06/01/2020 |  | 04/29/2020 |  |  | 04/10/2020 | 04/29/2020 |  |  | 11/30/2021 |  | 04/29/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1457308 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | Cardiac Perforation | [11 empty] | Hospitalization | Required Intervention |  |  |  |  | TEMPORARY PACER, IVUS | [22 empty] | Initial submission | Followup |  |  |  | 04/10/2020 | * | Invalid/Incomplete |  |  | The medical facility was treating a patient admitted for a high risk angioplasty of the left main and right coronary arteries. prior to the angioplasty the team placed an impella cp for hemodynamic support. the team also placed a temporary pacer. during the angioplasty the patient condition deteriorated and the team found she was experiencing a pericardial effusion with left ventricle wall perforation. the team pulled off 2000ml of blood and sent her to the operating department for repair. | [5 empty] | Since the initial report was filed, the investigation into the left ventricular (lv) perforation has concluded. the complainant did not return the pump product and ancillary equipment for investigation and analysis. the data logs were returned and analysis of them was completed. in addition to data log analysis a conversation with the operating physician was held. the data logs revealed no malposition of the pump in the lv nor any abrupt movement. the conversation with the physician revealed that the damage to the lv was more of a slice than a perforation and the thought was that the guidewire had caused the damage, not the impella pump. the guidewire was not returned for investigation. the failure mode will be monitored and trended. | To date the medical facility has not yet returned the product for investigation. also, further clinical details and imaging have been requested. upon completion of the investigation a final report will be filed. until said time, the failure mode of left ventricle perforation will continue to be trended.
37419 | 1220648-2020-00429 | 10011660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=10011660&pc=OZD | 11/01/2019 | 04/29/2020 |  | 04/29/2020 |  |  | 04/01/2020 | 04/29/2020 |  |  | 04/30/2021 |  | 04/29/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1417620 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/01/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported during pump support the patient experienced cerebrovascular accident (cva). as a result, the patient was given drug therapy. the patient ultimately recovered and mechanical support was no longer needed. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37420 | 1220648-2020-00424 | 9996877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9996877&pc=OZD | 03/30/2020 | 04/24/2020 |  | 04/24/2020 |  |  | 03/30/2020 | 04/24/2020 |  |  | 05/31/2021 |  | 04/24/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1426699 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 62 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | Sepsis | [11 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 03/30/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) female patient had impella 5.0 pump inserted in the left femoral for post-cardiotomy cardiogenic shock (pccs). it was reported the patient developed a fever and was soon diagnosed with and infection. the patient ultimately expired from septic shock. the source of the infection is unknown. per the physician, causal relationship with impella is unknown; however, other comorbidities such as vsp operation and swan-ganz could have attributed to the event. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37421 | 1220648-2020-00336 | 9993224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9993224&pc=OZD | 01/07/2020 | 04/23/2020 |  | 04/23/2020 |  |  | 03/09/2020 | 04/23/2020 |  |  | 09/30/2021 |  | 04/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1450509 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old hispanic female patient presenting with a proximal left anterior descending lesion. percutaneous coronary intervention (pci), in the form of a drug eluding stent placement, was performed. after pci, a decision was made to insert an impella cp optical device for ventricle unloading purposes. during support, despite optimal positioning and adequate volume, the patient exhibited signs of hemolysis. as treatment the device was prematurely removed, blood products were delivered, and dialysis was administered. | [5 empty] | The pump was returned and an investigation was completed. the data logs identified positioning issues occurring during the case, however, simulated use testing for hemolysis was performed and the failure was unable to be reproduced. a review of the dhr was conducted and found no manufacturing issues, reworks, or complaints associated with this failure mode from this pump lot. as such, we are unable to determine a root cause for the reported failure. | The impella cp optical has been received and an investigation is underway. upon completion of the investigation, a supplemental mdr will be submitted. please note, this mdr was initially submitted on 02/02/2020 with all 3 acknowledgements received. when attempting to submit a supplemental mdr on 04/09/2020, acknowledgement 3 failed due to "section: report identifier error message:error: initial report / prior supplement has not been received. the initial report is missing. " we reached out to the esg helpdesk and was advised the original submission should have failed due to selecting "both hispanic and not hispanic options for ethnicity." this has been revised and we are submitting this initial mdr to you, as it was originally submitted. follow-up mdr to follow today.
37422 | 1220648-2020-00420 | 9993798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9993798&pc=OZD | 12/10/2019 | 04/23/2020 |  | 04/23/2020 |  |  | 03/28/2020 | 04/23/2020 |  |  | 12/31/2020 |  | 04/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1391294 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/28/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient's right leg developed limb ischemia during pump support. per the physician, the ischemia was cause by the impella device; however, there was no increase in blood potassium observed. the ischemia was treated by adding a sheath in the antegrade direction. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37423 | 1220648-2020-00421 | 9993802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9993802&pc=OZD | 03/28/2020 | 04/23/2020 |  | 04/23/2020 |  |  | 03/28/2020 | 04/23/2020 |  |  | 12/31/2021 |  | 04/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1459326 |  | 0550-0008 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 57 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission |  |  |  |  | 03/28/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) african american male presenting with cardiomyopathy for hemodynamic support using the impella 5.5 device. approximately 5 days into support, the pump stopped and could not be restarted. the pump was removed and a permanent left ventricular assist device (lvad) was implanted. the permanent lvad was planned, however, implanted sooner than expected due to impella pump stop. | [5 empty] | The impella 5.5 was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37424 | 1220648-2020-00422 | 9993804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9993804&pc=OZD | 04/02/2020 | 05/15/2020 |  | 04/23/2020 |  |  | 04/02/2020 | 04/23/2020 |  |  | 02/28/2022 |  | 04/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 1466939 |  | 0550-0008 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 74 YR | [5 empty] | Malfunction | No | Yes | Mechanical Problem | [6 empty] | No Code Available | [12 empty] | Hospitalization | [28 empty] | Initial submission | Followup |  |  |  | 04/02/2020 | * | Invalid/Incomplete |  |  | The complainant had an impella 5.5 pump placed for a patient with cardiomyopathy. in addition the patient was covid 19 positive. the 5.5 was placed as an escalation of care from the intra-aortic balloon pump (iabp). the iabp was not enough hemodynamic support. the impella 5.5 was placed via the right axillary artery to the left ventricle. the 5.5 supported the patient for 7 days and then it stopped. the team was unable to restart the pump despite troubleshooting measures. the patient later expired. the physician did state the death was due to clotting secondary to the covid 19+. | [5 empty] | To date the product has not been returned for analysis. the data logs have been returned and are being evaluated for the cause of the pump stop. upon completion of the investigation, a final report will be filed. | Since the initial report was filed, the investigation into the pump stop has concluded. the team returned the data logs for analysis but not the physical pump. the logs revealed that the team had coagulation difficulties during support. during this time the team opened the purge system for an unknown reason. when the pump ran without purge fluid the pump began to fail. the pump without proper purge fluid is then allowing for blood ingress to the motor which can cause the pump to stop. the physician is aware of the issue caused by staff failing to have proper purge fluid for the pump motor and proper operation. the failure mode will be monitored.
37425 | 1220648-2020-00423 | 9993814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9993814&pc=OZD | 11/25/2019 | 04/23/2020 |  | 04/23/2020 |  |  | 03/27/2020 | 04/23/2020 |  |  | 08/31/2021 |  | 04/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1430528 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/27/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for post-cardiotomy cardiogenic shock (pccs). it was reported the patient's impella's extraction site wound had to be dissected after device removal. the patient received vac therapy and debridement. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37426 | 1220648-2020-00415 | 9978459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9978459&pc=OZD | 01/31/2020 | 04/20/2020 |  | 04/20/2020 |  |  | 03/23/2020 | 04/20/2020 |  |  | 09/30/2021 |  | 04/20/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1446739 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/23/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed hemolysis during pump support. pfhg levels were not provided. the pumps position was switched, and the patient received 18 units if red blood cells, 8 units of fresh frozen plasma and 1000ml of albumin. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37427 | 1220648-2020-00414 | 9978464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9978464&pc=OZD | 11/27/2019 | 04/20/2020 |  | 04/20/2020 |  |  | 03/23/2020 | 04/20/2020 |  |  | 07/31/2021 |  | 04/20/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1430942 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 57 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/23/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed a hemorrhage at the impella access site. as a result, the patient received 22 units of red blood cells, 10 units of platelet concentrate, 16 units of fresh frozen plasma and 27.5g of albumin. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37428 | 1220648-2020-00416 | 9978465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9978465&pc=OZD | 03/25/2020 | 05/15/2020 |  | 04/20/2020 |  |  | 03/25/2020 | 04/20/2020 |  |  | 12/31/2021 |  | 04/20/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5. PUMP |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1453797 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 65 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 03/25/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male patient had impella 5.5 pump inserted in the right axillary. after six (6) days of using the pump, it stopped, therefore this 5.5 pump was removed and another 5.5 pump was inserted. | [5 empty] | The impella 5.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The clinical describes that the pump was out of position and off for an extended period of time which can exacerbate clot formation. there was no high purge pressure during the case except for when the pump was unable to start. the pump was returned for evaluation and biomaterial was found on the impeller and in the purge gap, there was no blood was found in the purge lumen directly proximal to the motor. the pump was unable to be tested due to covid mitigations. the root cause of the pump stop was ingested biomaterial from the pump off for an extended period of time.
37429 | 1220648-2020-00417 | 9978476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9978476&pc=OZD | 01/05/2020 | 04/20/2020 |  | 04/20/2020 |  |  | 03/26/2020 | 04/20/2020 |  |  | 04/30/2021 |  | 04/20/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1415058 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 89 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/26/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed a hemorrhage at the impella access site during pump support. as a result, the patient received platelets, exact amount was not provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37430 | 1220648-2020-00418 | 9978486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9978486&pc=OZD | 08/01/2019 | 04/20/2020 |  | 04/20/2020 |  |  | 03/26/2020 | 04/20/2020 |  |  | 05/31/2021 |  | 04/20/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1420671 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 84 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/26/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6)-year-old (b)(6) male patient had impella 2.5 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that during pump support the patient developed hemolysis. it was noted that the adverse event was cause by the impella device and by the insertion of percutaneous cardiopulmonary support (pcps). as a result, the patient received 10 units of red blood cells, 8 units of fresh frozen plasma and 20 units of pc. plasma free hemoglobin (pfhg) levels were not provided. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37431 | 1220648-2020-00419 | 9978906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9978906&pc=OZD | 03/26/2020 | 04/20/2020 |  | 04/20/2020 |  |  | 03/26/2020 | 04/20/2020 |  |  | 07/31/2021 |  | 04/20/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1430942 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/26/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 2.5 pump inserted in the right femoral for unloading purposes. it was reported the physician observed hematuria three hours after the impella pump was turned on. patient showed an increased in both ld and bilirubin levels and signs of renal failure. the patient was placed on continuous hemofiltration/hemodiafiltration which resolved the hemolysis. plasma free hemoglobin (pfhg) levels were not provided. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37432 | 1220648-2020-00409 | 9967180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9967180&pc=OZD | 01/23/2020 | 04/16/2020 |  | 04/16/2020 |  |  | 03/17/2020 | 04/16/2020 |  |  | 08/31/2020 |  | 04/16/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | AACHEN, GERMANY |  |  | US |  | 1367668 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/17/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) female patient had impella 5.0 pump inserted in the left femoral for post-cardiotomy cardiogenic shock (pccs). it was reported the patient developed a hemorrhage at the femoral artery and internal carotid artery during pump support. as a result, the patient received 22 units of red blood cells, 25 units of platelets and 33 units fresh frozen plasma. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37433 | 1220648-2020-00410 | 9967225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9967225&pc=OZD | 01/14/2020 | 04/16/2020 |  | 04/16/2020 |  |  | 03/19/2020 | 04/16/2020 |  |  | 08/31/2021 |  | 04/16/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | DANVERS, MA |  |  | US |  | 1441282 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 50 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Pneumothorax | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/19/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left axillary for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient experienced a pneumothorax at the left lung, as a result, the patient required a chest drain to be inserted. additionally, the patient developed an infection at the insertion site which required drug treatment. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37434 | 1220648-2020-00411 | 9969866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9969866&pc=OZD | 03/19/2020 | 04/16/2020 |  | 04/16/2020 |  |  | 03/19/2020 | 04/16/2020 |  |  | 10/31/2021 |  | 04/16/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1462073 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/19/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp optical pump. during support, the physician found a lack of distal pulses in the impella inserted leg. as treatment, the device was prematurely removed. the patient was deemed stable, therefore, no new left ventricular assist device was required. | [5 empty] | The impella cp optical was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37435 | 1220648-2020-00412 | 9969872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9969872&pc=OZD | 07/11/2019 | 04/16/2020 |  | 04/16/2020 |  |  | 03/19/2020 | 04/16/2020 |  |  | 12/31/2020 |  | 04/16/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1391180 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Unspecified Infection | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/19/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 2.5 pump inserted in the right femoral for myocarditis. it was reported the patient developed an infection at the insertion site for which drug treatment was administered. per the physician, the causal relationship between the reported adverse events and the impella device is unknown. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37436 | 1220648-2020-00413 | 9969877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9969877&pc=OZD | 03/22/2020 | 04/16/2020 |  | 04/16/2020 |  |  | 03/22/2020 | 04/16/2020 |  |  | 10/31/2020 |  | 04/16/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1384654 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 13 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/22/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old african american patient presenting with cardiomyopathy. the impella cp pump was inserted for hemodynamic support without any reported issue. during support, that patient's access site developed a bleed around the suture holding the blue butterfly in place. as treatment, the suture was removed, thrombin powder was applied, and 2 liters of packed red blood cells were administered. | [5 empty] | The impella cp was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.