37437 | 1220648-2020-00406 | 9956781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9956781&pc=OZD | 03/18/2020 | 04/14/2020 |  | 04/14/2020 |  |  | 03/18/2020 | 04/14/2020 |  |  | 05/31/2021 |  | 04/14/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1421375 |  | 005040 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Regurgitation | Death | [11 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 03/18/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella 2.5. during support, the patient had experienced severe aortic regurgitation due to the impella restricting leaflet motion of the patient's aortic bioprosthesis. the pump was removed due to patient's "respiratory issues" and septic shock. respiratory issues and septic shock were not alleged to have been caused or contributed to by the impella device. after removal of the impella device, the aortic regurgitation did not resolve. the patient's status was not deemed stable enough for surgical intervention to correct the aortic insufficiency and the patient expired. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37438 | 1220648-2020-00407 | 9956784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9956784&pc=OZD | 07/02/2019 | 04/14/2020 |  | 04/14/2020 |  |  | 03/18/2020 | 04/14/2020 |  |  | 04/30/2020 |  | 04/14/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1340134 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 37 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Injury | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/18/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 5.0 pump inserted in the right axillary for cardiomyopathy. it was reported the patient developed aortic valve injury. as a result, the patient received park's stitch at the aortic valve. the physician is unsure as to when the injury took place and if it was related to impella. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37439 | 1220648-2020-00408 | 9956804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9956804&pc=OZD | 01/08/2020 | 04/14/2020 |  | 04/14/2020 |  |  | 03/18/2020 | 04/14/2020 |  |  | 08/31/2021 |  | 04/14/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1436348 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/18/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed a hematoma at the impella access site. as a result, the patient received 12 units of red blood cells. 20 units of pc and 10 units of fresh frozen plasma. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37440 | 1220648-2020-00404 | 9933074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9933074&pc=OZD | 12/21/2019 | 04/07/2020 |  | 04/07/2020 |  |  | 03/13/2020 | 04/07/2020 |  |  | 03/31/2021 |  | 04/07/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1418897 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 03/13/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella 5.0 pump inserted in the left axillary for cardiomyopathy. it was reported the patient developed hemolysis during pump support which was confirmed via pfhg measurements. to improve the situation impella¿s level was reduced from p8 to p4. additionally, the patient was administered four units of red blood cells. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37441 | 1220648-2020-00405 | 9933091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9933091&pc=OZD | 12/19/2019 | 04/07/2020 |  | 04/07/2020 |  |  | 03/16/2020 | 04/07/2020 |  |  | 08/31/2021 |  | 04/07/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1432131 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 51 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/16/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed hemolysis during pump support which was confirmed via pfhg measurements. the patient was administered 2000 units of hemoglobin/100mlv 2 bottles. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37442 | 1220648-2020-00399 | 9912894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9912894&pc=OZD | 01/01/2020 | 04/01/2020 |  | 04/01/2020 |  |  | 03/17/2020 | 04/01/2020 |  |  | 04/30/2021 |  | 04/01/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1415058 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/17/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient required continuous hemofiltration/hemodiafiltration (chdf) to treat acute kidney injury (aki). per the physician the aki was cause by hemolysis which the patient developed during pump support. the patient did not receive treatment for hemolysis. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37443 | 1220648-2020-00400 | 9913131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9913131&pc=OZD | 03/06/2020 | 04/01/2020 |  | 04/01/2020 |  |  | 03/06/2020 | 04/01/2020 |  |  | 11/30/2021 |  | 04/01/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1465233 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/06/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting for high risk percutaneous coronary intervention using the impella cp optical for hemodynamic support. the pump was inserted using the single access method with an abiomed 14fr x 13cm introducer. during support, the access site developed a bleed from the introducer. as treatment, the device was removed, a 2 unit transfusion of packed red blood cells was administered, and hemostasis was achieved by deploying perclose and manual compression. | [5 empty] | Neither the impella cp optical nor the 14fr x 13cm introducer were received for investigation. should any new information be returned, a supplemental mdr will be filed.
37444 | 1220648-2020-00401 | 9913153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9913153&pc=OZD | 03/09/2020 | 04/27/2020 |  | 04/01/2020 |  |  | 03/09/2020 | 04/01/2020 |  |  | 09/30/2021 |  | 04/01/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1452465 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Blood Loss | Vascular Dissection | [11 empty] | Hospitalization | [28 empty] | Initial submission | Followup |  |  |  | 03/09/2020 | * | Invalid/Incomplete |  |  | A patient admitted with a myocardial infarction (st elevation mi) was treated in the cardiac catheterization lab and had 5 coronary stents placed. the patient had an impella 5.0 placed percutaneously via the right axillary for hemodynamic support. the placement of the impella was a difficult one that required aggressive manipulation for the steep angulation at the insertion site. it was noted that force was needed to transition the pump from the introducer to the artery. the axillary artery need a balloon to tamponade the artery as result of the multiple and aggressive placement attempts. in addition blood product were transfused back to the patient. the impella pump remained on for support for 2 days despite this access site complication and provided hemodynamic support to the patient. | [5 empty] | The investigation into the vessel damage and blood loss in on-going at this time. product and data log return from the medical facility have been affected by the covid 19 virus and facility access. upon completion of the investigation, a final report will be drafted and submitted. | Since the initial report was filed the investigation has completed. data logs were not returned for review. product was returned, but due to covid-19 a complete analysis is not possible. based upon the clinical account and patient history, the root cause of the vascular damage is attributed to the excessive force used on insertion as a result of the steep insertion angulation and patient anatomy. the failure mode will be monitored and trended.
37445 | 1220648-2020-00402 | 9913180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9913180&pc=OZD | 03/12/2020 | 04/29/2020 |  | 04/01/2020 |  |  | 03/12/2020 | 04/01/2020 |  |  | 08/31/2021 |  | 04/01/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  |  |  | 1441282 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Nerve Damage | Renal Failure | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 03/12/2020 | * | Invalid/Incomplete |  |  | The medical team in (b)(6) admitted a patient suffering from an acute myocardial infarction and cardiogenic shock and was found to have left main coronary artery occlusion. the team placed an impella cp for hemodynamic support and did a coronary intervention. the impella successfully supported the patient for 9 days and was explanted. months later the medical team notified abiomed (b)(6) of adverse events that took place during the impella support. the patient had an underlying renal history and the team alleged that the renal function further deteriorated on impella support. also, the team had limb palsy issues that were attributed to positional and compression issues the nurses had with the impella pump. there was an allegation made of peroneal nerve palsy, which was treated with rehabilitation and the wearing of orthotic. | [5 empty] | The team in (b)(6) has discarded the impella pump in december of 2019 when this patient was treated. there can be no investigation into the nerve palsy or renal failure being caused by the device. further clinical details have been asked of the team in (b)(6), but due to the covid 19 virus the information gathering is hampered. upon the completion of the investigation a final report will be filed. | Since the initial report was filed the investigation has concluded. the medical team in (b)(6) did not return the pump product for analysis and hemolysis testing. the data logs were returned and analysis was done. logs revealed no evidence of pump positioning or suction alarms that would be inclusive of hemolysis events and eventual renal failure. the root cause of the renal failure could not be determined without product return. in addition, the history of chronic kidney disease prior to insertion is of note as a probable cause of the renal failure. the failure mode will be monitored and trended.
37446 | 1220648-2020-00403 | 9913197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9913197&pc=OZD | 12/27/2019 | 04/01/2020 |  | 04/01/2020 |  |  | 12/27/2019 | 04/01/2020 |  |  | 03/31/2021 |  | 04/01/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1422808 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/27/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient presenting with non-ischemic cardiomyopathy for hemodynamic support with the impella 5.0 device. the pump was delivered successfully, however, the patient's condition deteriorated over the next week. the physician noted signs of mitral regurgitation, however, an echo was found no evidence of free leaflets or torn chordae. the decision was made to withdraw care due to patient's bi-ventricular dysfunction and the pump was removed. the physician alleged the impella may have played a role in the patient's mitral regurgitation. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37447 | 1220648-2020-00397 | 9897907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9897907&pc=OZD | 01/10/2019 | 03/30/2020 |  | 03/30/2020 |  |  | 06/25/2019 | 03/30/2020 |  |  | 10/31/2020 |  | 03/30/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1375181 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/25/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 2.5 pump inserted in the right axillary for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient experienced a hemorrhage at the implella access site, as a result the impella device was removed. no other treatment was provided. | [5 empty] | The root cause was unable to be determined because the product was not returned for review. data logs do not pertain to the reported failure mode. dhr was reviewed and found no manufacturing related issues with this lot. there are no other complaints related this failure mode in this lot.
37448 | 1220648-2020-00398 | 9897929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9897929&pc=OZD | 01/15/2020 | 03/30/2020 |  | 03/30/2020 |  |  | 03/09/2020 | 03/30/2020 |  |  | 09/30/2021 |  | 03/30/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1447303 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/09/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) japanese male patient had impella cp pump inserted for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the physician suspected hemolysis during pump support. patients plasma free hemoglobin data were not provided. the patient was administered bilirubin adsorption. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37449 | 1220648-2020-00392 | 9886990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9886990&pc=OZD | 11/15/2019 | 06/03/2020 |  | 03/26/2020 |  |  | 03/03/2020 | 03/26/2020 |  |  | 07/31/2020 |  | 03/26/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1359038 |  | 05066-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 03/03/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male presenting for hemodynamic support with the impella 5.0 device. during support, the patient exhibited clinical signs of hemolysis via tests for hemoglobin, total bilirubin, and ldh. plasma-free hemoglobin was not tested, thus, we cannot definitively confirm the patient's hemolysis. as treatment, several units of blood products were delivered. | [5 empty] | The impella 5.0 was received and an investigation was completed. no abnormalities were found with the device. simulated use testing was performed and pump flows were measured within specification. the root cause of the hemolysis was attributed to improper positioning with the pump outflow on the aortic valve. | The impella 5.0 was received (b)(6) 2019 under complaint case-(b)(4) under a different symptom that was assessed as not reportable. the original complaint was closed, however, the customer has reported new information pertaining to a new symptom (hemolysis). thus, a new complaint record was opened. upon evaluation of this device for the newly reported symptom, a supplemental mdr will be filed.
37450 | 1220648-2020-00393 | 9887025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9887025&pc=OZD | 03/02/2020 | 04/20/2020 |  | 03/26/2020 |  |  | 03/02/2020 | 03/26/2020 |  |  | 08/31/2021 |  | 03/26/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1441282 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report | [8 empty] | 75 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [22 empty] | Initial submission | Followup |  |  |  | 03/02/2020 | * | Invalid/Incomplete |  |  | A patient was admitted to the medical facility in (b)(6) after having an acute myocardial infarction at an outlying community hospital. the patient was supported by pcps/ecmo and an intraaortic balloon pump (iabp). upon admission to the tertiary center the iabp was removed and an impella cp placed for escalated cardiac support. the patient's condition and prognosis were not good, as he had poor liver function and developed multiorgan failure. after 23 days of impella support the patient expired after the family made a decision to not extend measures, and he became a dnr. months later the medical facility informed the abiomed (b)(6) representatives of a hemorrhage that occurred during the impella support. the team had given the patient extensive blood product infusions. the (b)(6) medical team noted units of 261 erythrocytes, 46 ffp, 160 of platelets, and 55 units of blood. | [5 empty] | Since the initial report was filed the investigation has concluded. the investigation of the product was not possible. no product was returned for analysis from japan. the data logs were returned and found that the pump operated to specification. the cause of the bleeding that prompted the team to infuse a total of 261 units of blood product to the patient is unknown. the failure mode will be monitored and trended. | The medical team in (b)(6) has discarded the impella pump in (b)(6) of 2019 when the patient expired. clinical details and the pump data logs are being analyzed. upon completion of the investigation, a final report will be filed.
37451 | 1220648-2020-00394 | 9887032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9887032&pc=OZD | 01/29/2020 | 03/26/2020 |  | 03/26/2020 |  |  | 03/04/2020 | 03/26/2020 |  |  | 08/31/2021 |  | 03/26/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1436348 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/04/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp device. during support, an access site bleed formed and was treated with a blood transfusion. approximately 10 days later, the patient developed leg ischemia and was treated with unspecified drug therapy. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37452 | 1220648-2020-00396 | 9887048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9887048&pc=OZD | 03/03/2020 | 05/05/2020 |  | 03/26/2020 |  |  | 03/03/2020 | 03/26/2020 |  |  | 10/31/2021 |  | 03/26/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1444788 |  | 0550-0008 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | Followup |  |  |  | 03/03/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella 5.5 device. during support, the pump exhibited lower than expected purge flow that led to a purge flow blocked alarm after 10 days of use. on day 11 of support, the pump stopped and was unable to be restarted. the device was removed and a second impella was placed. | [5 empty] | The impella 5.5 was returned by the customer and an investigation of device was completed. a visual inspection identified blood ingress in the purge lumen within the catheter. the root cause of the reported pump stop was due to blood ingress within the pump's motor. we could not definitively determine the cause of the low purge flow, however. data logs were not returned for analysis. a review of the dhr was conducted and found no manufacturing issues, reworks, or complaints associated with this failure mode from this pump lot. | The impella 5.5 was not received by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37453 | 1220648-2020-00395 | 9887049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9887049&pc=OZD | 03/06/2020 | 03/26/2020 |  | 03/26/2020 |  |  | 03/06/2020 | 03/26/2020 |  |  | 11/30/2021 |  | 03/26/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1460151 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/06/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella cp optical pump inserted in the left femoral artery (lfa) for acute myocardial infarction (ami)/ cardiogenic shock (cgs). upon removal of the peel-away sheath post procedure, it was noted that the lfa site was bleeding and despite attempts to angle match the catheter and "snug" the perclose sutures around the sheath the site continued to bleed. the physician decided to remove the pump and obtain hemostasis on the lfa site with perclose. a second pump was placed due to hemodynamic instability. a central line was placed for blood/fluids to be administered. the patient was initially given two (2) units of whole blood and three (3) more units of whole blood was giving during the case. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37454 | 1220648-2020-00390 | 9871039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9871039&pc=PYX | 03/13/2020 | 04/29/2020 |  | 03/24/2020 |  |  | 03/13/2020 | 03/24/2020 |  |  | 05/31/2021 |  | 03/24/2020 |  | TEMPORARY NONROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED EUROPE, GMBH |  |  | GM |  | 1434427 |  | 0046-0011 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 38 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [22 empty] | Initial submission | Followup |  |  |  | 03/13/2020 | * | Invalid/Incomplete |  |  | A female patient of (b)(6) was admitted with cardiomyopathy and placed on ecmo. when she was unable to be weaned from the ecmo, the team chose to place bipella support with both an impella cp and impella rp. after 4 days of support there was note of massive hemolysis. the team chose to explant the rp. after the removal the hemolysis began to be rectified over the subsequent 3 days. the team also infuse blood products, in the form of 23 transfusions over 4 days. | [5 empty] | Since the initial report was filed, the investigation into the hemolysis allegation has concluded. the complainant in the eu did not return the product for hemolysis testing and product analysis. the team returned partial data logs from the console. during the support no positioning or suction events were found in the data logs that would indicate impella was the cause of the hemolysis. the cause of the hemolysis can not be determined. insufficient clinical details were shared and no product was returned. the failure mode will be monitored and trended. | The complainant in the eu has not returned product for investigation. the analysis into the clinical report and data logs is on-going at this time. a final report will be submitted upon completion.
37455 | 1220648-2020-00391 | 9871244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9871244&pc=OZD | 02/22/2020 | 07/08/2020 |  | 03/24/2020 |  |  | 02/24/2020 | 03/24/2020 |  |  | 10/31/2021 |  | 03/24/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1453002 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 22 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup | Followup |  |  | 02/24/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella cp pump inserted in the right femoral for cardiomyopathy. it was reported the physician suspected hemolysis during pump support. pfhg levels were not provided. pumps flows were lowered; however, there was no improvement. as a result, treatment was escalated to 5.0. no further information was provided. | [5 empty] | The returned pump was visually inspected and no delaminated epo-tek, impeller damage, or biomaterial were found. an orange material was observed on the edge of the cannula delo joint and within the cannula to outlet cage bond. the cannula was inspected via borescope and a rough transition edge was observed between the cannula and the outlet cage. the pump passed electrical testing. the returned pump was cleaned and prepped for hemolysis testing. the pump failed hemolysis testing. intermittent suction was observed throughout the hemolysis test. the orange material was likely blood or denatured proteins that had entered from the cannula id to within the cannula to outlet cage bond and out the delo joint. the cause of the hemolysis was most likely a partial cannula detachment from the outflow cage resulting in a rough transition edge. dhr was reviewed and found no manufacturing related issues with this lot. there are no other complaints related this failure mode in this lot. hemolysis is listed in the tempubunsho as a potential adverse event that may be associated with the use of the impella device. abiomed will continue to monitor these types of events. the manufacturer is closing the file on this event. | The impella cp was returned; however, due to covid-19 work from home policies, the returned product was unable to be visually inspected. the root cause of the hemolysis cannot be determined. data logs were not provided. dhr was reviewed and found no manufacturing related issues with this lot. there are no other complaints related to this failure mode in this lot. | The impella cp is expected to be return. a supplemental mdr will be filed with the results of the investigation, upon completion.
37456 | 1035166-2020-00022 | 9862240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9862240&pc=DYB | 02/07/2020 | 06/05/2020 |  | 03/20/2020 |  |  | 02/25/2020 | 03/20/2020 |  |  | 02/01/2022 |  | 03/20/2020 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-15664 |  | 0052-3021 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 79 YR | [5 empty] | Injury | Yes | Yes | Break | [6 empty] | Myocardial Infarction | Perforation of Vessels | Cardiogenic Shock | [10 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported during acute myocardial infraction (ami) and cardiogenic shock (cgs) procedure while inserting an oscor 23fr introducer into the access site, the distal end of the device split. the introducer was removed and replaced with a new 23 fr introducer. no further issues were experienced. decision made to insert impella rp for right heart support. additional information received on march 4, 2020, when the sheath split it tore the vein. they were able to stent, no surgery required. no significant delay in therapy. addition information has been requested. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. | The device was used for treatment.the device was not returned for analysis therefore, the clinical observation could not be confirmed. the investigation will focus on a review of product documentation. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. as per procedure abiomed introducer sheath in-process and final inspection the devices in this lot were inspected as follows:visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. slight discoloration from tipping process is acceptable. verify proper tip shape according to drawing. per section visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. the instructions for use (ifu) informs the user: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
37457 | 1220648-2020-00388 | 9863037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9863037&pc=OZD | 02/24/2020 | 05/12/2020 |  | 03/20/2020 |  |  | 02/24/2020 | 03/20/2020 |  |  | 05/31/2021 |  | 03/20/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1433793 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 02/24/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella cp optical pump inserted in the right femoral for post cardiac-cardiogenic shock (pccs). hemolysis was first noticed 20-30 minutes after impella placed and continued to get worse until impella was removed. hemolysis began to resolve shortly after impella removal. | [5 empty] | The returned pump was returned, during evaluation a strand of biomaterial was found wrapped around the impeller. high motor current and pump stops were recreated during a test run with the biomaterial laden pump. after the biomaterial was removed, the pump ran within specification. biomaterial was ingested since the alarms occurred early in the case. the biomaterial was of a single strand thus creating a whipping motion as it rotated with the impeller. the cause of hemolysis was ingested biomaterial in the pump. a hemolysis test was not conducted since it was determined to be a non-essential activity while covid-19 mitigations were in place. | The impella cp optical was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37458 | 1220648-2020-00389 | 9863038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9863038&pc=OZD | 11/22/2019 | 03/20/2020 |  | 03/20/2020 |  |  | 02/27/2020 | 03/20/2020 |  |  | 08/31/2021 |  | 03/20/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1435812 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/27/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male patient had impella cp pump inserted in the right femoral for acute myocardial infarction (ami) with shock. the patient had a hemorrhage at the impella access site during pump support. as a result 14 units of red blood cells (rbcs) was administered for treatment. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37459 | 1220648-2020-00387 | 9857145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9857145&pc=OZD | 01/19/2020 | 10/09/2020 |  | 03/19/2020 |  |  | 01/19/2020 | 03/19/2020 |  |  | 07/31/2021 |  | 03/19/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1445555 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/19/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp optical device. during support, the patient exhibited signs of hemolysis via hemolyzed labs, however, no plasma-free hemoglobin tests were performed to confirm. approximately 6 hours into support, the impella was rapidly weaned in response to the suspected hemolysis. | [5 empty] | Correction: g(4): upon review of this complaint, it was determined the incorrect aware date was provided in the initial report. the aware date should be 2020-02-28. | The impella cp was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37460 | 1220648-2020-00385 | 9848467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9848467&pc=OZD | 02/17/2020 | 03/18/2020 |  | 03/18/2020 |  |  | 02/17/2020 | 03/18/2020 |  |  | 11/30/2021 |  | 03/18/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1454868 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 68 YR |  |  | 141 |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/17/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) old white male patient had impella cp optical pump inserted in the left femoral for acute myocardial infarction (ami)/cardiogenic shock (cgs). the patient had limb ischemia despite fem and the pump was explanted. another pump was inserted in the right axillary. | [5 empty] | Device has not been returned. the impella cp optical pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if the device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37461 | 1220648-2020-00386 | 9849464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9849464&pc=OZD | 02/20/2020 | 03/18/2020 |  | 03/18/2020 |  |  | 02/20/2020 | 03/18/2020 |  |  | 07/31/2021 |  | 03/18/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1436531 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 50 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/20/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp pump. the physician reported pump not properly positioning within the ventricle, therefore, an attempt was made to remove and re-insert the device. during attempted removal, the device failed to exit the introducer due to looping of the guidewire. the introducer, pump, and guidewire were removed as one unit and a new pump was inserted, however, the patient developed a hematoma at the pump's insertion site. after patient's hemodynamics stabilized with the second pump, surgical removal was performed. surgical removal was required due to hematoma that afflicted the insertion site from the pump in question. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37462 | 1035166-2020-00021 | 9850021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9850021&pc=DYB | 02/22/2020 | 03/19/2020 |  | 03/18/2020 |  |  | 02/28/2020 | 03/18/2020 |  |  | 10/01/2022 |  | 03/18/2020 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-17964 |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 28 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that during cardiogenic shock procedure, there was bleeding at the insertion site. user reported that there was no issue with introducer. they noticed excessive bleeding after giving heparin for anticoagulation once the sheath was in. the first cp was placed had excessive bleeding in the right groin. then the second cp was placed in the left groin once the first cp was removed from the sheath. patient loss 800ml blood. two (2) units of red blood cells (rbcs) were given and gore sheath was placed for hemostasis in right groin and new cp in left groin. the sheath was sutured to the skin. the bleeding caused a delay in putting the rp in, which we never got to because of the condition of the patient. patient was very sick. patient expired on 22nd february from cardiac arrest. the patient was in cardiac standstill. there was no product malfunction reported. device was discarded. no other information is available. | [5 empty] | The device was used in the treatment. the device was discarded and not returned for analysis. there was no specific performance related failure reported by the user. the device history record was reviewed to confirm that the device passed all applicable inprocess and final inspections. no further investigation is required. based on the investigation, a capa is not required as as findings did not identify a design, labeling or manufacturing nonconformity. the event will be re-evaluated if additional information becomes available. no further follow-up is required. | The following sections were updated in follow-up. b4,d4,g4,g7,h2,h10. correction : d4 : expiration date corrected oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
37463 | 1220648-2020-00376 | 9845460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9845460&pc=OZD | 02/19/2020 | 03/17/2020 |  | 03/17/2020 |  |  | 02/19/2020 | 03/17/2020 |  |  | 10/31/2021 |  | 03/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1453002 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/19/2020 | * | Invalid/Incomplete |  |  | The complainant reported an asian male patient had impella cp pump inserted in the left femoral for high risk percutaneous coronary intervention (hrpci). the patient had limb ischemia, pump was removed and endovascular treatment (evt) was performed. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37464 | 1220648-2020-00377 | 9845465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9845465&pc=OZD | 09/01/2019 | 03/17/2020 |  | 03/17/2020 |  |  | 02/17/2020 | 03/17/2020 |  |  | 05/31/2021 |  | 03/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1420671 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/17/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient experienced a hemorrhage at the access site. to treat the adverse event the patient was administered 10 units of red blood cells. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37465 | 1220648-2020-00378 | 9845467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9845467&pc=OZD | 02/23/2020 | 04/09/2020 |  | 03/17/2020 |  |  | 02/23/2020 | 03/17/2020 |  |  | 10/31/2021 |  | 03/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUM |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED INC. |  |  | US |  | 1453045 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention |  |  |  |  | VA ECMO | [22 empty] | Initial submission | Followup |  |  |  | 02/23/2020 | * | Invalid/Incomplete |  |  | A patient was transferred in to the tertiary medical center for cardiology care. the patient was suffering from an acute myocardial infarction. she had decreased ejection fraction and coronary artery disease in need of intervention/angioplasty. for hemodynamic support an impella cp was placed via the left femoral artery. an external bypass was also placed due to diminished pulses in the lower limbs. she was anticoagulated with aspirin, heparin, aggrastat, and brilanta and supported with the impella for approximately 1 day when there was an observation made of a hematoma at the impella access site. the groin injury and associated bleeding prompted a vascular surgery repair and 2 units of blood product replacement. the patient remained on the impella support with the impella cp for 5 days. | [5 empty] | Since the initial report was filed, the investigation has concluded. the clinical details shared by the medical facility noted multiple anticoagulants were used for the patient care, and as such the blood was thin. the hematoma and access site bleeding were due to the thin blood. the device was returned and no damage to it was observed that could have caused the bleeding and hematoma injury. the failure mode will be monitored and trended. | The complainant has returned the impella pump product for analysis. the analysis of the product and clinical details in on-going at this time. upon the completion of the investigation, a final report will be submitted.
37466 | 1220648-2020-00379 | 9845468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9845468&pc=OZD | 02/22/2020 | 03/17/2020 |  | 03/17/2020 |  |  | 02/22/2020 | 03/17/2020 |  |  | 09/30/2021 |  | 03/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1452705 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 28 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Death | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/22/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) old white male patient had impella cp optical pump inserted in the right femoral for cardiogenic shock (cgs). the patient had a massive groin bleeding, as treatment two (2) units of red blood cells (rbcs) and 16f gore sheath placed for hemostasis and gave protamine. the physician tried everything, but the patient expired from cardiac arrest as he was in cardiac standstill. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37467 | 1220648-2020-00380 | 9845469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9845469&pc=OZD | 11/20/2019 | 03/17/2020 |  | 03/17/2020 |  |  | 02/20/2020 | 03/17/2020 |  |  | 08/31/2021 |  | 03/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1435812 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 65 YR | [5 empty] | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Sepsis | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/20/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed sepsis while on impella support. the physician reported many catheters (other than impella) were inserted and therefore were unable to determined which catheter cause the sepsis. however, there is a possibility that the catheter-related bloodstream infection (crbsi) occurred due to pneumonia. the patient received drug treatment to treat sepsis. additionally, the patient developed limb ischemia on the left foot¿s first toe. the patient received non cardiac surgery. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37468 | 1220648-2020-00382 | 9845478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9845478&pc=OZD | 12/02/2019 | 03/17/2020 |  | 03/17/2020 |  |  | 02/20/2020 | 03/17/2020 |  |  | 12/31/2020 |  | 03/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  |  |  | 1391512 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 87 YR | [5 empty] | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/20/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) female patient had impella 5.0 pump inserted in the right axillary for standby operation purpose. it was reported the patient developed hemolysis during pump support. the account attempted to troubleshoot the issue by changing impella's placement. the patient was also administered four units of red cell count. no further information was provided. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37469 | 1220648-2020-00381 | 9845480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9845480&pc=OZD | 06/06/2019 | 03/17/2020 |  | 03/17/2020 |  |  | 02/20/2020 | 03/17/2020 |  |  | 10/31/2020 |  | 03/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1377030 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 34 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/20/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 2.5 pump inserted in the right femoral for cardiomyopathy. it was reported the patient developed an infection at the insertion site. to treat the infection the patient was administered antibacterial drugs. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37470 | 1220648-2020-00383 | 9845484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9845484&pc=OZD | 03/02/2020 | 05/12/2020 |  | 03/17/2020 |  |  | 03/02/2020 | 03/17/2020 |  |  | 12/31/2021 |  | 03/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1455282 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 66 YR | [5 empty] | Malfunction | No | Yes | Separation Problem | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 03/02/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) old white male patient had impella cp pump inserted in the left femoral for acute myocardial infarction (ami) with shock takotsubo <10%. the cannula broke inside the patient¿s left ventricle (lv) and as a result a special lasso wire was used to remove the broken cannula. the patient expired due to takotsubo <10% which is unrelated to impella. | [5 empty] | The returned data logs confirmed that the pump was initially in a good position but went deep into the ventricle during activation, as evidenced by a ventricular waveform placement signal. the slack was not reported to have been removed from the catheter and the tuohy-borst was not reported to have been locked, this could result in the pump going deep into the ventricle. the returned pump was inspected; the cannula detachment was confirmed with a snare still latched on the pigtail. the detachment was not a delo bond failure nor a nitinol failure. the cannula polyurethane composite was found to be sheared and torn off the inflow indicating that the cannula caught on something, possibly the mitral apparatus or another device since the pump was deep in the ventricle. there were scratches observed on the inflow and pigtail indicating an interaction. improper use prior to insertion or during the insertion process can cause the cannula to kink next to the inflow. excessive force was applied during pump removal as reported by the clinical. the patient had a pre-existing condition of takotsubo cardiomyopathy which can cause the patient's left ventricle to resemble a lobster trap with a narrow neck extending into a round ventricle. the improper pump positioning deep in the left ventricle (lv) likely caused the pump cannula to kink and excessive force was applied without imaging causing the cannula to detach. | The impella cp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37471 | 1220648-2020-00384 | 9845511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9845511&pc=OZD | 09/26/2019 | 03/30/2020 |  | 03/17/2020 |  |  | 09/26/2019 | 03/17/2020 |  |  | 04/30/2021 |  | 03/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1415058 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemostasis | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 03/02/2020 | * | Invalid/Incomplete |  |  | The complainant reported a 70 years old asian female patient had impella cp pump inserted in the right femoral for preoperative myocardial recovery. during the case, the patient received a total of 36 units of rbcs, 20 units of fresh frozen plasma (ffp), and 75 grams of albumin for access site bleeding. | The complainant reported a (b)(6) asian female patient had impella cp pump inserted in the right femoral for preoperative myocardial recovery. the patient received two (2) units of red blood cells (rbcs) for suspected hemolysis. during the case, the patient received a total of 36 units of rbcs, 20 units of fresh frozen plasma (ffp), and 75 grams of albumin for access site bleeding. |  |  |  |  | Corrected date recv'd by mfr from 26-sept-2019 to 02-mar-2020. corrected patient from 1886 and 1895 to 1895. | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37472 | 1220648-2020-00373 | 9826681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9826681&pc=OZD | 09/21/2019 | 03/12/2020 |  | 03/12/2020 |  |  | 02/14/2020 | 03/12/2020 |  |  | 04/30/2021 |  | 03/12/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1417620 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/14/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese female patient had impella cp pump inserted in the left femoral for heart failure. the complainant reported the developed sepsis while on impella support. the patient received 310 units of platelets, 28 units of red blood cells, 30 units of plasma, 400 (25%) ml of albumin, 1000 (5%) of albumin and medical observation. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37473 | 1220648-2020-00374 | 9826683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9826683&pc=OZD | 02/12/2020 | 03/12/2020 |  | 03/12/2020 |  |  | 03/02/2020 | 03/12/2020 |  |  | 05/31/2021 |  | 03/12/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1423572 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | DISTRI |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/02/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) old female patient presenting post coronary artery bypass procedure with acute myocardial infarction and cardiogenic shock. the impella cp was selected for hemodynamic support and inserted without any reported issue. during the first night of support, the patient developed an access site bleed and went into hemorrhagic shock. the site was surgically repaired with a tobacco purse suture and controlled. on day 7 of support, however, bleeding resumed and surgical intervention was again performed to control the bleeding. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37474 | 1220648-2020-00375 | 9826694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9826694&pc=OZD | 02/12/2020 | 03/12/2020 |  | 03/12/2020 |  |  | 02/12/2020 | 03/12/2020 |  |  | 10/31/2021 |  | 03/12/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1453032 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Death | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | Cardiac Tamponade | Vascular Dissection | [10 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 02/12/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the right femoral for unloading left ventricle. it was reported the patient was rushed to the emergency room for unexplained severe heart failure or cardiogenic shock. the physician stated insertion was difficult due to a strong resistance and during the insertion dissection of the abdominal aorta occurred. the physician stated the cause of the abdominal dissection is unknown as it could have been cause by percutaneous cardiopulmonary support (pcps) or the impella device. the patient ultimately expired due to aortic dissection and cardiac tamponade. the physician stated the relationship between the device and death is unknown as he was unable to determine if the cause of the cardiac tamponade was due to aortic dissection or cardiac rupture due to thin ventricular walls. it is also unknown whether the patients ruptured heart was cause by the impella or by the patient's blood pressure jumping to 200 when the adrenaline was administered. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37475 | 9811348 | 9811348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9811348&pc=PBL | 11/06/2019 | 01/14/2020 | 01/17/2020 | 03/10/2020 | 01/17/2020 | 03/10/2020 |  | 03/10/2020 |  |  |  |  | 03/10/2020 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PBL |  |  | IMPELLA CP PUMP SET |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003.0814 |  | No | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 28470 DA |  |  | 41 |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Occlusion | [12 empty] | Other | [28 empty] | Initial submission | [5 empty] | Y | Invalid/Incomplete |  |  | Occluded left common femoral artery at site of impella insertion during pci.
37476 | 1035166-2020-00019 | 9813687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9813687&pc=DYB | 02/03/2020 | 03/10/2020 |  | 03/10/2020 |  |  | 02/13/2020 | 03/10/2020 |  |  | 03/01/2021 |  | 03/10/2020 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-14183 |  | 0052-3021 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 84 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that patient was presented to physician for right ventricle (rv) failure. acute myocardial infarction (ami)/cardiogenic shock (cgs) procedure there was bleeding at the sheath insertion site while using a 23 fr introducer. femstop was applied after 30 plus minutes of manual pressure. femstop was removed by cardiothoracic surgery (ct) surgeon. ct surgeon added additional mattress suture, then re-applied femstop. patient loose 250cc of blood loss. initially the sutures placed by the ic were probably shallow to the skin. the surgical pa and ct surgeon went deeper with their sutures. they did an ultrasound to confirm that there was not a perforation. femstop was in place for 6 hours and hemostasis achieved. the patient was on support for four (4) days . there was no product performance issue reported. as per physician, size discrepancy between the rp pump and the repo sheath was the cause for the bleed. no other intervention was needed. no additional information is available. | [5 empty] | The device was used for treatment. the device was discarded and there was no specific device performance related failure reported by the user. cause of the bleeding is related to non oscor device used in procedure. the investigation will focus on a review of product documentation. the device history record was reviewed to confirm that the device passed all applicable inprocess and final inspections.based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
37477 | 1220648-2020-00370 | 9814409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9814409&pc=OZD | 11/27/2019 | 03/10/2020 |  | 03/10/2020 |  |  | 02/10/2020 | 03/10/2020 |  |  | 05/31/2021 |  | 03/10/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1424928 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/10/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the physician suspected hemolysis during impella support. as a result, the patient was placed on continuous hemofiltration/hemodiafiltration (chdf). no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37478 | 1220648-2020-00371 | 9814422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9814422&pc=OZD | 02/11/2020 | 04/24/2020 |  | 03/10/2020 |  |  | 02/11/2020 | 03/10/2020 |  |  | 11/30/2020 |  | 03/10/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1397086 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 02/11/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 5.0 pump inserted in the right axillary for post-cardiotomy cardiogenic shock (pccs). it was reported the physician suspected hemolysis. the initial position of the pump was verified by echo. the physician attempted to re-position the impella; however, this did not resolve the issue. continuous hemofiltration/hemodiafiltration (chdf) was started to treat the patient and the impellas position was re-adjusted. the physician stated the patient underwent aortic valve replacement (avr) and mitral valve repair (mvr) surgery, due to this the position of the aortic valve was different from that of a normal patient; therefore, hemolysis was likely to occur. | [5 empty] | The impella 5.0 was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion. | The pump was returned for review; however, only the distal portion was returned. the reported hemolysis was confirmed. the root cause was most likely improper positioning. data logs were not returned for review. dhr was reviewed and found no manufacturing related issues with this lot. there are no other complaints related this failure mode in this lot.
37479 | 1220648-2020-00372 | 9814430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9814430&pc=OZD | 02/28/2020 | 03/10/2020 |  | 03/10/2020 |  |  | 02/28/2020 | 03/10/2020 |  |  | 12/31/2021 |  | 03/10/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1455279 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 0 YR | [5 empty] | Malfunction | No | Yes | Detachment of Device or Device Component | Infusion or Flow Problem | No Apparent Adverse Event |  |  |  |  | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission |  |  |  |  | 02/28/2020 | * | Invalid/Incomplete |  |  | The complainant reported a patient of unknown age, gender, or race presenting for hemodynamic support using the impella cp device. the complainant reported pump stopping immediately after initiation and could not be restarted. the pump was removed and replaced without any adverse impact to the patient. the device was returned and broken impeller blades were identified. | [5 empty] | The impella cp and data logs were returned for evaluation. data logs found the device did not encounter ta pump stop, however, low flow and suction was present during use. the device was visually inspected and noted to have a fractured impeller blade. simulated use testing in a basin of water successfully reproduced the low pump flows as indicated in the data logs. a review of the dhr was conducted and found no manufacturing issues or reworks associated with this pump lot. there are no other complaints related to this failure mode associated with pumps in this lot. due to lack of clinical data, we can not definitively determine the cause of the damaged impeller blades.
37480 | 9792497 | 9792497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9792497&pc=PBL | 11/25/2019 | 01/16/2020 | 01/16/2020 | 03/05/2020 | 01/16/2020 | 03/05/2020 |  | 03/05/2020 |  |  |  |  | 03/05/2020 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP SET |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003 |  | Yes | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 32120 DA | [5 empty] | Malfunction | No | Yes | Fracture | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | 0.018 impella wire unwound and fractured in the patient. the entire wire was retrieved and there was no evidence of any retained wire in the patient. a new wire was obtained from another kit and the case continued. impella wire advanced through pigtail catheter. pigtail removed. impella catheter advanced over the wire to lv, wire removed. wire fragment noticed on x-ray in impella. impella catheter and wire removed. 6fr. pigtail advanced over the wire to lv, wire removed. impella wire advanced through pigtail catheter. pigtail removed. impella catheter advanced over the wire to lv, wire removed.
37481 | 1220648-2020-00368 | 9795380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9795380&pc=PYX | 06/11/2018 | 03/05/2020 |  | 03/05/2020 |  |  | 02/18/2019 | 03/05/2020 |  |  | 01/31/2020 |  | 03/05/2020 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1322181 |  | 004334 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/11/2018 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) patient presenting with post cardiotomy cardiogenic shock requiring right ventricular support. the impella rp was inserted without any reported issues, however, the pump migrated and hemolysis was suspected. as treatment, the device was removed and a second rp was inserted. | [5 empty] | The impella rp was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37482 | 1220648-2020-00367 | 9795382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9795382&pc=PYX | 02/18/2019 | 03/05/2020 |  | 03/05/2020 |  |  | 02/18/2019 | 03/05/2020 |  |  | 07/31/2019 |  | 03/05/2020 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1301336 |  | 004334 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/18/2019 | * | Invalid/Incomplete |  |  | The complainant reported a patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella rp. during use, the patient's access site developed a bleed. as treatment, the patient received 2 units of packed red blood cells. | [5 empty] | The impella cp was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37483 | 1220648-2020-00369 | 9795412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9795412&pc=OZD | 09/26/2019 | 03/05/2020 |  | 03/05/2020 |  |  | 02/06/2020 | 03/05/2020 |  |  | 06/30/2021 |  | 03/05/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1428735 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Sepsis | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/06/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that patient developed hemolysis during impella support which was confirmed via pfhg measurements. impella was re-positioned and p-level's were adjusted. the patient also received 20 units of red blood cell fluid, 480ml of fresh frozen plasma and 500ml albuminar. it was also reported the patient developed sepsis. per the physician, heart failure after acute myocardial infarction combined with sepsis resulted in long-term impella placement and difficult removal. the patient was upgraded to impella 5.0 for infection and long-term rehabilitation management. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37484 | 1220648-2020-00360 | 9787675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9787675&pc=OZD | 02/05/2020 | 03/04/2020 |  | 03/04/2020 |  |  | 02/05/2020 | 03/04/2020 |  |  | 08/31/2021 |  | 03/04/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1442376 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 76 YR | [5 empty] | Malfunction | No | Yes | Detachment of Device or Device Component | Infusion or Flow Problem | No Apparent Adverse Event |  |  |  |  | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission |  |  |  |  | 02/05/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) old white male patient presenting for high risk percutaneous coronary intervention using the impella cp pump for hemodynamic support. the pump was inserted without any reported issue, however, flows dropped from 3.6 l/min to 0.5 l/min after initiation. normal troubleshooting attempts did not resolve the issue, therefore, the device was removed and replaced with a second impella pump. hemodynamic support was successfully achieved, thereafter. on 2020-02-05, upon investigation of the returned device, we identified broken impeller blades. | [5 empty] | The impella cp and data logs were returned for evaluation. the device was visually inspected and noted to have fractured impeller blades. simulated use testing in a basin of water successfully reproduced low pump flows reported. follow up with the clinical account indicated no other devices were in use within the patients heart. we could not definitively determine the cause of the damaged impeller blades. a review of the dhr was conducted and found no manufacturing issues or reworks associated with this pump lot. there are no other complaints related to this failure mode associated with pumps in this lot.
37485 | 1220648-2020-00361 | 9787729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9787729&pc=OZD | 02/13/2020 | 03/04/2020 |  | 03/04/2020 |  |  | 02/13/2020 | 03/04/2020 |  |  | 07/31/2021 |  | 03/04/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1446852 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/13/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) white female patient had impella cp optical pump inserted in the right femoral artery (rfa) for high risk percutaneous coronary intervention (hrpci). the patient¿s left ventricle was perforated and as a result the physician put in an intra-aortic balloon pump (iabp), wean and explanted the pump, and they reversed the heparin on board, the effusion slowed considerably. | [5 empty] | Device discarded by customer. the impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37486 | 1220648-2020-00362 | 9787739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9787739&pc=OZD | 02/10/2020 | 04/07/2020 |  | 03/04/2020 |  |  | 02/10/2020 | 03/04/2020 |  |  | 12/31/2021 |  | 03/04/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1458877 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 02/10/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella cp pump inserted in the right femoral for cardiogenic shock (cgs). the patient¿s right leg looked mottled and the decision was made to remove the pump. this patient looked to have peripheral arterial disease (pad). thrombus was noted in the iliac system and as a result angiovac and balloon angioplasty was done to restore flow. the patient was brought to the operating room (or) for thrombectomy and fasciotomy was needed post procedure. | [5 empty] | Imaging prior to impella insertion showed that the patient had peripheral artery disease (pad). the returned pump was visually inspected and no abnormalities related to the ischemia were observed. there was some light scratching down the length of the repositioning unit sheath and some notching next to the blue suture pad. the data log was reviewed showing the pump operated to specification. the recommended 5% dextrose and 50 iu/ml heparin concentration were added to the purge 6 hours into the run. the cause of the leg ischemia was a combination of peripheral artery disease and thrombus found in the iliac artery after pump removal as reported in the clinical. after the pump was removed, the ischemia persisted until the thrombus was removed from the iliac artery to resolve the ischemia. the patient's acts were not reported to indicate the source of the thrombus. | The impella cp was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37487 | 1220648-2020-00363 | 9787746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9787746&pc=PYX | 02/07/2020 | 03/04/2020 |  | 03/04/2020 |  |  | 02/07/2020 | 03/04/2020 |  |  | 07/31/2021 |  | 03/04/2020 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP, OZD |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1446362 |  | 004334 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Venipuncture | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/07/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) old white male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella rp. upon insertion of the 23 fr introducer, the tip split and caused a venous tear. as treatment, the introducer was removed, the vein was stented, and a new 23 fr introducer was placed without issue. | [5 empty] | The 23 fr introducer was retained by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37488 | 1220648-2020-00364 | 9787769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9787769&pc=OZD | 02/07/2020 | 04/24/2020 |  | 03/04/2020 |  |  | 02/07/2020 | 03/04/2020 |  |  | 05/31/2021 |  | 03/04/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1422971 |  | 005040 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 85 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 02/07/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male patient presenting for high risk percutaneous coronary intervention. the impella 2.5 was selected for support. the device was inserted and used without issue. upon successful hemodynamic support, the device was removed and the site was closed with a closure device. soon after closure, the patient became hemodynamically unstable and required cardiopulmonary resuscitation. two perforations in the apex of the ventricle were found, resulting in pericardial effusion. surgical intervention was performed to resolve the injury and the patient has since recovered. although no impella device malfunctions or deficiencies had occurred, the physician feels the impella still may have caused or contributed to this injury. | [5 empty] | The root cause of the cardiac apex perforations could not be determined as review of the returned pump found no defects that could have caused or contributed to the reported injury. a device history record review was performed found no manufacturing issues or reworks associated with this pump lot. there are no other complaints related to this failure mode associated with pumps in this lot. | The impella 2.5 was returned and an investigation is underway. upon completion, a supplemental mdr will be filed.
37489 | 1220648-2020-00365 | 9787793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9787793&pc=OZD | 04/12/2019 | 03/04/2020 |  | 03/04/2020 |  |  | 02/06/2020 | 03/04/2020 |  |  | 04/30/2020 |  | 03/04/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1347814 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 47 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/06/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 5.0 pump inserted in the right axillary post-cardiotomy cardiogenic shock (pccs). it was reported the physician suspected the patient had developed hemolysis during impella support. impella's position was verified via echo. the patient was administered concentrated red blood cells as treatment. no further information was provided | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37490 | 1220648-2020-00366 | 9787796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9787796&pc=OZD | 01/06/2020 | 04/07/2020 |  | 03/04/2020 |  |  | 01/06/2020 | 03/04/2020 |  |  | 01/31/2021 |  | 03/04/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1404634 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemolysis | [12 empty] | Hospitalization | [5 empty] | ECMO | ECMO | [21 empty] | Initial submission | Followup |  |  |  | 01/06/2020 | * | Invalid/Incomplete |  |  | In (b)(6) a patient was admitted in cardiogenic shock and had an extracorporeal membrane oxygenation (ecmo) and impella 2.5 placed for hemodynamic support. the impella supported the patient for over 4 days when it was explanted. at the explant date the patient had his angioplasty of the coronary arteries. at explant the patient was noted to be stable and no longer in need of mechanical support. months after pump explant the medical team informed abiomed of presumed hemolysis during the impella support. the team had treated the hemolysis, presumed to be caused by impella, by infusion of blood products. the japan team infused the japanese infusion of 36 units of blood product. the unit in japan is not equal to that of the usa unit of measure. | [5 empty] | Since the initial report was filed the investigation into the presumed hemolysis experienced by the user in (b)(6) has completed. the team in (b)(6) was unable to return for investigation either the pump product or the data logs. without the product or data logs the root cause of the hemolysis is unable to be determined. either the impella or the pcps/ecmo could have caused the hemolysis. the failure mode will be monitored and trended. | No product has been returned, and is not expected to be. the event occurred in (b)(6) of 2019 and abiomed was notified in 2020. questions have been asked of the team in (b)(6), but knowledge sharing has been hampered by the coronavirus/covid-19 in (b)(6). upon completion of the investigation, a final report will be filed.
37491 | 1035166-2020-00015 | 9781630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9781630&pc=DYB | 01/29/2020 | 06/01/2020 |  | 03/03/2020 |  |  | 02/07/2020 | 03/03/2020 |  |  | 09/24/2019 |  | 03/03/2020 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-17775 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 73 YR | [5 empty] | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | No Consequences Or Impact To Patient | [12 empty] | Other | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | During primary percutaneous coronary intervention 14 fr x 23 cm introducer that kinked. the physician placed 2 covered stents in the iliacs after he saw how bent the introducer was just to make sure there was no internal bleeding. there was no internal damage, covered stents were placed as precaution. there was 30- 45 min delay reported to place covered stents and changing devices. procedure was not completed and patient outcome is stable. no other information is available. | On may 29, 2020 customer reported correct sheath size used in procedure is 14fr x 25cm. |  |  |  |  | Our investigation is still in progress, follow up report will be submitted if we find any further additional information. | Corrected data: b5 correct size used in procedure is 14fr x 25cm. the device was used in treatment. one 14f, 25 cm long abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. blood was found on and inside the sheath and sideport tubing. three kinks were observed on the sheath tube. upon evaluation of the returned sheath under 10x magnification, there were three kinks observed in the sheath tube at approximately 7, 9 and 10cm from the distal tip. the distal tip of the sheath looked normal along its circumference edges. the hemostatic valve looked normal with its helical cuts intact. the sheath was measured with a pin gauge and was within manufacturing specifications. the impella device used in the procedure was not returned to oscor so a fit check could not be performed. the sheath (including the inner diameter) is within manufacturing specifications. potential contributing factors to the kinks could be if the sheath was advanced through an area of resistance or if the sheath was subjected to unusual stresses. no manufacturing defects were found. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure abiomed introducer sheath in-process and final inspection: these steps are performed by qa as per sampling plan ansi z 1.4, gen level i, aql 0.65. measure dimensions: b) measure tip i.d. of sheath using appropriate pin gauges. visual inspection: this step is performed by qa for sample size of 100% inspection a) with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. the instructions for use (ifu) informs the user: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
37492 | 1035166-2020-00014 | 9776856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9776856&pc=DYB | 01/31/2020 | 05/07/2020 |  | 03/02/2020 |  |  | 02/03/2020 | 03/02/2020 |  |  | 07/01/2021 |  | 03/02/2020 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3001 |  |  | OSCOR INC. |  |  | US |  | C1-14497 |  | 0052-3001 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | Yes | Break | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that patient was suffering from cardiomyopathy and during access at the femoral vein for placement of an impella rp, sheath spilt and placement of sheath was unsuccessful. there was delay in procedure reported as no back up sheath was available. perforation of vein observed which was fixed with vessel stenting. patient outcome is stable. no additional information is available. | [5 empty] | Our investigation is still in progress, follow up report will be submitted if we find any further additional information. | The following sections were updated in follow-up. b4,d10,g4,g7,h2,h3,h6,h10 the device used in treatment. one 23f abiomed introducer sheath was returned from the customer with the dilator. there were no other accessories. blood was found on and inside the sheath and dilator. upon evaluation of the returned sheath, it was found that there was a split in the sheath tube along the scoreline at approximately 10.5cm from the distal tip. the distal tip of the dilator has slight damage. there is also damage around the circumference of the distal tip edge of the sheath and the hemostatic valve was slightly torn. the dilator was inserted into the sheath for a fit check and passed through the sheath fully and smoothly. when measured od of the dilator was within specification. due to the condition of the tip, the tipped id could not be measured with a pin gauge. instead, the sheath was cut in half in order to measure the non-tipped portion of the sheath. the id of the non-tipped portion of the sheath was within specification. when measured reveal length of the dilator within specification. returned device analysis revealed the sheath and dilator are within manufacturing specifications. potential contributing factors to the damage (split) could be if the sheath was advanced through an area of resistance or if the sheath was subjected to unusual stresses. since the hemostatic valve was torn and the distal tip of the dilator was damaged, probably the sheath tip was damaged during use. no manufacturing defects were found. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure abiomed introducer sheath in-process and final inspection, the devices in this lot were inspected as follows: visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. this is performed by qa on the first 5 consecutive properly tipped parts. with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. this is performed by qa on 100% of the sheaths. assembly fit check: the dilator reveal length, when the dilator assembly is completely engaged with the sheath assembly, shall be 50.8 ± 6.4 mm for 23f. per procedure abiomed introducer dilator in-process and final inspection: first 5 properly printed parts, inspect 100% 9.1.1 verify correct french size (number 0.035" of guidewire) is printed on dilator by verifying outer diameter (od) using laser micrometer. per ifu never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
37493 | 1220648-2020-00357 | 9772703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9772703&pc=OZD | 11/17/2019 | 03/01/2020 |  | 03/01/2020 |  |  | 01/31/2020 | 03/01/2020 |  |  | 08/31/2021 |  | 03/01/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1433531 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 42 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/31/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella cp pump inserted in the left femoral for unknown indication. it was reported the patient developed lower limb ischemia on the left limb. the patient received retrograde perfusion to treat the event. the physician stated the event was not related to impella because the lower limb ischemia was caused by a decrease in blood pressure due to circulatory failure. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37494 | 1220648-2020-00359 | 9772704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9772704&pc=OZD | 08/25/2019 | 03/01/2020 |  | 03/01/2020 |  |  | 02/04/2020 | 03/01/2020 |  |  | 04/30/2021 |  | 03/01/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1417620 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 02/04/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed limb ischemia when the impella device was being explanted. the patient received non-cardiac surgery, specifics were not provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37495 | 1220648-2020-00358 | 9772705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9772705&pc=OZD | 02/10/2020 | 03/30/2020 |  | 03/01/2020 |  |  | 02/10/2020 | 03/01/2020 |  |  | 10/31/2021 |  | 03/01/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1454567 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | Yes | Fracture | Device Dislodged or Dislocated | [5 empty] | Aortic Insufficiency | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 02/10/2020 | * | Invalid/Incomplete |  |  | A patient admitted with coronary artery disease and mitral regurgitation was decompensating. due to his cardiogenic shock he was taken to the cardiac catheterization lab for evaluation of his coronary arteries and was to have an impella cp placed for hemodynamic support. when the cp was placed in the left ventricle (lv), the pump placement was not optimal and so the physician attempted to adjust position. when adjusting the position the cp became kinked. as the physician was trying to remedy the kink and straighten the catheter the pump separated. to two pieces. a portion of the pump remained in the lv and a portion was across the valve. the patient condition began to decline and so the team chose to treat the underlying coronary artery, the left main. cpr was administered and he portion of the impella not within the lv was removed and an intra-aortic balloon pump was placed. despite these efforts the patient expired. | [5 empty] | The investigation into the separation of the impella cp, which allowed part of the pump to remain within the patient's lv, is on-going. product has not been returned and more clinical details are expected. as more clinical details are shared, the cause of the reported aortic insufficiency and death will be analyzed. when the investigation concludes a final report will be filed. | The investigation has concluded since the initial report was filed. the medical team returned for analysis the data logs, pump images, and fluoroscopy imaging. the data logs revealed the pump was ventricularized and had a motor current spike from a possible excessive force. the pump became kinked which caused the inflow cage to become damaged. the force caused the inflow cage to detach and become stuck across the aortic valve. the failure mode will be monitored and trended.
37496 | 1035166-2020-00012 | 9765348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9765348&pc=DYB | 01/24/2020 | 04/28/2020 |  | 02/27/2020 |  |  | 01/31/2020 | 02/27/2020 |  |  | 11/01/2022 |  | 02/27/2020 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-17774 |  | 0052-3000 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN | [10 empty] | Injury | Yes | Yes | Break | [6 empty] | Cardiogenic Shock | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A female patient (age unknown) was presented to the physician under resuscitation with cardiogenic shock (cgs) signs. patient went in the intensive care unit (icu) with va-extracorporeal membrane oxygenation (ecmo) being in cardiogenic shock (cgs). the devices were used for cgs due to acute myocardial infarction (ami). two oscor (2) 14fr introducers sheaths failed, the sheath's split at the tip and rolled up, and they were discarded; they also used a competitor's 13f introducer (trial for dilatation). after successful implantation of va-ecmo an angiologist had to intervene, because there were inner bleedings in the groin. patient is obese, difficult to locate arteria under cpr, puncturing with needle. physician reported the sheaths probably failed because they were too short and too soft/ fragile, as a lot of tissue needed to be crossed, due to obese patient. obviously they injured a side branch from artery femoralis during access attempts (puncturing the arteria). it was reported the representative saw the groin angiology/ fluoroscopy - there was obviously no peripheral arterial disease (pad) or kinking of the arteria femoralis. additional information has been requested. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete. | The device was used for treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. as per procedure abiomed introducer sheath in-process and final inspection, the devices in this lot were inspected as follows: 9.2.2 visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. slight discoloration from tipping process is acceptable. verify proper tip shape according to drawing. per section 12.2, visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. the instructions for use (ifu) informs the user: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
37497 | 1220648-2020-00356 | 9765841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9765841&pc=OZD | 01/29/2020 | 02/27/2020 |  | 02/27/2020 |  |  | 01/29/2020 | 02/27/2020 |  |  | 02/28/2021 |  | 02/27/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1406291 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/29/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese female patient had impella 2.5 pump inserted in the left femoral for high risk percutaneous coronary interventions (hrpci). it was reported the physician suspected hemolysis one hour after impella was implanted. impella¿s position was adjusted in order to resolve the issue; however, this was unsuccessful. continuous hemofiltration/hemodiafiltration (chdf) was performed and the device was explanted due to the hemolysis. no further information was provided. | [5 empty] | Product has not been returned; investigation has not been completed.
37498 | 1220648-2020-00353 | 9758808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9758808&pc=PYX | 01/31/2020 | 03/30/2020 |  | 02/26/2020 |  |  | 01/31/2020 | 02/26/2020 |  |  | 10/31/2020 |  | 02/26/2020 |  | TEMPORAY NONROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1378760 |  | 0046-0011 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | Yes | Insufficient Information | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [5 empty] | IMPELLA CP | [22 empty] | Initial submission | Followup |  |  |  | 01/31/2020 | * | Invalid/Incomplete |  |  | The medical team in (b)(6) had a male patient in need of both right and left ventricular support. the patient was suffering from biventricular heart failure with influenza, myocarditis, and cardiomyopathy. the team had an access site complication when placing the impella rp for the right ventricular support. there was a femoral vein perforation that was treated with a covered stent of approximately 7cm, and blood products were infused. once stented, the patient was placed on support with the impella successfully. both the rp and impella cp were placed. | [5 empty] | The investigation into the vessel dissection that necessitated blood products and stenting is on-going at this time. upon completion of the investigation, a final report will be filed. | The investigation has completed since the initial report was filed. the sheath product was returned for analysis and investigation. the 23fr x 30cm introducer sheath was returned and found to have a sheath tip split. the product malfunction caused the patient injury. the sheath damage caused the medical team to infuse blood product and stenting. the failure mode will be monitored and trended.
37499 | 1220648-2020-00355 | 9758821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9758821&pc=OZD | 01/31/2020 | 04/20/2020 |  | 02/26/2020 |  |  | 01/31/2020 | 02/26/2020 |  |  | 01/31/2021 |  | 02/26/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1400754 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/31/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 2.5 pump inserted in the right femoral for vsp. it was reported the patient developed hemolysis during impella support. patient's plasma free hemoglobin measurements confirmed the event. the patient received two units of fresh frozen plasma-leucocyte reduced (rcc-lr) for anemia progression which was due to hemolysis. no further information was provided. | [5 empty] | The impella 2.5 was returned under another case. the product performed within specification during the case; therefore, the cause of hemolysis cannot be determined. dhr was reviewed and found no manufacturing related issues with this lot. there were two other complaints in this device lot related to the failure mode. | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37500 | 1220648-2020-00354 | 9758831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9758831&pc=OZD | 01/31/2020 | 03/30/2020 |  | 02/26/2020 |  |  | 01/31/2020 | 02/26/2020 |  |  | 11/30/2021 |  | 02/26/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1454868 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 01/31/2020 | * | Invalid/Incomplete |  |  | A patient admitted with a complex medical history was admitted to the medical facility with difficulty breathing. after a right heart catheterization was performed, the team chose to place an impella cp for hemodynamic support. after insertion a hematoma developed at the access site. the hematoma was treated by femostop compression. the patient was bleeding from the access site for the swan as well. due to the blood loss the team infused blood product. after 54 hours of support the team chose to explant the impella. shortly after explant the patient experienced signs of limb ischemia. there was discussion of possible amputation, but this was not done. | [5 empty] | Since the initial report was filed, the investigation into the leg ischemia experienced during impella use has completed. the pump was not returned for analysis. the data logs and clinical report were shared for analysis and review. the clinical report revealed that the patient was on high doses of pressors and was known to have venous congestion. the data logs were shown to reveal that the patient was not always appropriately anticoagulated with heparin throughout the support. the data logs showed that the pump ran to specification throughout the support. the root cause was not found to be due to the use of the impella but, rather to the patient underlying condition of venous congestion and pressor medication. | To date the medical facility has not yet returned any impella product for analysis. when the investigation is completed a final report will be filed.
37501 | 1220648-2020-00350 | 9745111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9745111&pc=OZD | 01/26/2020 | 04/27/2020 |  | 02/24/2020 |  |  | 01/26/2020 | 02/24/2020 |  |  | 10/31/2021 |  | 02/24/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  |  |  | 1453339 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 26 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/26/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. the device was inserted without issue, however, a large hematoma developed on the access site during use. although the hematoma appeared as not actively growing, blood counts were monitored and "a few units" of blood products were delivered. | [5 empty] | On 2020-03-16, clarification was received from the abiomed clinical field representative regarding the initially reported access site hematoma. the clinical representative, who was on site during support, clarified that the patient's hematoma originated from a malpositioned reprofusion sheath and not the impella. there was no malfunction or deficiency with the impella device, thus, this is not deemed a complaint or an abiomed reportable event. we respectfully request the initial mdr to be retracted. | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37502 | 1220648-2020-00351 | 9746876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9746876&pc=OZD | 01/31/2020 | 02/24/2020 |  | 02/24/2020 |  |  | 01/31/2020 | 02/24/2020 |  |  | 10/31/2021 |  | 02/24/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1450243 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 87 YR | [5 empty] | Death | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | Perforation | Cardiac Tamponade | [10 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 01/31/2020 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6)-year-old (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that findings from a transthoracic echocardiogram (tte) showed a rupture within the heart. the physician suspected the patient had developed a heart tamponade. an echocardiogram (echo) showed blood was stored in the pericardium. the inlet part of impella was in the pericardium and the outlet was in the left ventricle. the physician was unsure whether the impella had penetrated the left ventricular wall or whether it had moved forward due to a rupture of the left ventricle. soon after the patient was moved to the incentive care unit and expired. per the physician the cause of death was due to heart rupture. it is unclear whether it was related to impella. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37503 | 1220648-2020-00352 | 9746881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9746881&pc=OZD | 01/28/2020 | 04/16/2020 |  | 02/24/2020 |  |  | 01/28/2020 | 02/24/2020 |  |  | 05/31/2021 |  | 02/24/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1424965 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/28/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that the patient¿s urine was noted to be red while the patient was in the intensive care unit (icu). hemolysis was suspected. the physician adjusted to pumps level and position; however, this did not resolve the issue. due to suspected hemolysis and decreased in blood pressure the physician made the decision to explant the device. | [5 empty] | The impella 2.5 was returned for review and no damage or defects were found. no biomaterial was found on the returned device. the pump was hemolysis tested and passed. the root cause of hemolysis was unable to be determined. data logs were not returned. dhr was reviewed and found no manufacturing related issues with this lot. there is one other complaint related this failure mode in this lot. | The impella 2.5 was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37504 | 1220648-2020-00347 | 9742057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9742057&pc=OZD | 01/24/2020 | 03/10/2020 |  | 02/21/2020 |  |  | 01/24/2020 | 02/21/2020 |  |  | 11/30/2021 |  | 02/21/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, OZD |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1449532 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [5 empty] | 14FR INTRODUCER. | 14FR INTRODUCER. | [21 empty] | Initial submission | Followup |  |  |  | 01/24/2020 | * | Invalid/Incomplete |  |  | The complainant reported a female patient with cardiogenic shock (cgs). the physician tried to access with two (2) different 14fr sheaths and one (1) 13fr sheath but failed to be introduced into the vessel. the sheath ¿splitted¿ at the tip and rolled up so they were broken. this patient is obese and the arterial femoral was ¿deep¿ under the tissue, there was groin bleedings after several punctures. the bleedings was intervened by an angiologist with a covered stent. | [5 empty] | Corrected data included: d11 concomitant medical products are two (2) 14fr introducer that was used during this case. | The device was not returned by the customer. the 13fr sheath was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37505 | 1220648-2020-00348 | 9742061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9742061&pc=OZD | 01/27/2020 | 03/30/2020 |  | 02/21/2020 |  |  | 01/27/2020 | 02/21/2020 |  |  | 08/31/2021 |  | 02/21/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1439840 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 01/27/2020 | * | Invalid/Incomplete |  |  | A (b)(6) old patient was admitted in cardiogenic shock. the team placed an impella cp for support via the right femoral artery. the team consulted cardiothoracic surgery for possible bypass. while in the icu, awaiting bypass surgery, the team observed the impella flow to drop. troubleshooting was attempted. the team observed a thrombus present within the left ventricle (lv). the presence of clot within the lv is a noted concern within the ifu. the team chose to explant the impella. at removal the medical team also had to care for the patient as the right leg became ischemic. a thrombectomy was done. | [5 empty] | Since the initial report was filed the investigation has completed. the pump was returned for analysis, though the data logs from the console were not. no imaging of the peripheral vasculature was returned for review. the pump was run successfully in the benchtop testing and not damage was found. the pump dimensions were found to be within specification. the root cause of the limb ischemia was unable to be determined. the failure mode will be monitored and trended. | The investigation into the low flows and limb ischemia are on-going at this time. upon investigation completion, a final report will be submitted.
37506 | 1220648-2020-00349 | 9742073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9742073&pc=OZD | 01/23/2020 | 02/21/2020 |  | 02/21/2020 |  |  | 01/23/2020 | 02/21/2020 |  |  | 10/31/2020 |  | 02/21/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1380163 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 48 YR | [5 empty] | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Infarction, Cerebral | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/23/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old (b)(6) female patient had impella 5.0 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that after the impella device was successfully weaned and explanted the patient experienced a cerebral infarction. the patient was administered glyceol to treat the adverse event. the physician stated the impella, percutaneous cardiopulmonary support (pcps) and intra-aortic balloon pump (iabp)caused the adverse event. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37507 | 1220648-2020-00342 | 9726516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9726516&pc=OZD | 01/21/2020 | 04/09/2020 |  | 02/19/2020 |  |  | 01/21/2020 | 02/19/2020 |  |  | 02/29/2020 |  | 02/19/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1332857 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/21/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp pump. the device was inserted without issue, however, the patient's access site developed a large hematoma during support. as treatment, the device was removed, blood products were delivered, and a second impella was placed in the opposing femoral artery. | [5 empty] | The device was returned in 2 pieces as the catheter was cut intentionally cut during removal of the pump, therefore, bench testing could not be performed. visual inspection found no issues with the pigtail, inflow cage, cannula, motor housing, or endcap of the catheter. no damage was found on the repositioning sheath. a review of the dhr was conducted and found no manufacturing issues or reworks associated with this failure mode from this pump lot. as such, we cannot definitively determine a root cause for the reported issue. | The impella cp was returned and an investigation is underway. upon completion, a supplemental mdr will be filed.
37508 | 1220648-2020-00343 | 9726529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9726529&pc=OZD | 01/21/2020 | 03/24/2020 |  | 02/19/2020 |  |  | 01/21/2020 | 02/19/2020 |  |  | 01/31/2020 |  | 02/19/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP 550 SET, EU |  |  | ABIOMED INC. |  |  | US |  | 1441071 |  | 0550-0001 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 19 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 01/21/2020 | * | Invalid/Incomplete |  |  | The (b)(6) cardiomyopathy patient was admitted to a medical facility in (b)(6) and placed on the support of an impella 5.5 pump. during the support the patient exhibited signs of hemolysis, in the form of elevated plasma free hemoglobin levels. due to the rising lab values the team chose to explant and replace the impella pump after approximately 197 hours. the new 5.5 was successfully placed. the pump was explanted and upon observation of it, there was biomaterial observed despite assertions that the young patient was therapeutically anticoagulated. | [5 empty] | The investigation is on-going at this time. upon completion a final report will be filed. | Since the initial report was filed the investigation has concluded. the medical facility did return a portion of the impella pump, but not its entire length, and they returned the data logs for analysis. only the distal portion of the pump was returned. some biomaterial was observed near the pump's outlet cage. the pump had another/new handle soldered onto it and the pump was run for hemolysis testing. the pump when reconstructed in this way did pass hemolysis testing. the data logs revealed the pump was in proper position with some periods of suction and low patient pulsatility. the root cause of the reported hemolysis was unable to be determined. the failure mode will be monitored and trended.
37509 | 1220648-2020-00344 | 9726538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9726538&pc=OZD | 01/24/2020 | 03/24/2020 |  | 02/19/2020 |  |  | 01/24/2020 | 02/19/2020 |  |  | 04/30/2021 |  | 02/19/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1429219 |  | 05062 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | CONSUMER | HEALTH PROFESSIONAL |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 01/24/2020 | * | Invalid/Incomplete |  |  | A cardiomyopathy patient was admitted for a workup for a possible left ventricular assist device (lvad) and ablation to treat underlying ventricular tachycardia (vt). he had an impella 5.0 placed via the right axillary artery, despite the observation made of a thrombus in the left ventricle, which is listed as a contraindication in the ifu for the impella 5.0. the patient had his vt ablation procedure while on the support of the 5.0. post ablation he was sent to the icu for continued monitoring. while in the icu he exhibited signs of hemolysis. the team decided to explant the 5.0 and not replace the pump. the patient had been supported by the impella for approximately 2 days. | [5 empty] | The investigation into the presumed hemolysis is on-going at this time. the medical facility has not yet released the pump for analysis and hemolysis testing. upon completion of the investigation a final report will be filed. | Since the initial report was filed the investigation into the hemolysis presumed to be caused by the use of impella has completed. the product was not returned for any investigation of benchtop hemolysis testing. at explant the hospital team stated biomaterial was observed, but this could not be confirmed without pump return. the data logs were returned for analysis. the logs do reveal many suction alarms during the support but it is unknown if this correlated to when the hemolysis was reported. the root cause of the hemolysis could not be determined. the failure mode will be monitored and trended.
37510 | 1220648-2020-00345 | 9726541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9726541&pc=OZD | 01/23/2020 | 02/19/2020 |  | 02/19/2020 |  |  | 01/23/2020 | 02/19/2020 |  |  | 11/30/2020 |  | 02/19/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC |  |  | US |  | 1385057 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Code Available | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/23/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 2.5 pump inserted in the right femoral for cardiomyopathy. it was reported the patient developed thrombocytopenia during pump support. the patient was administered 20 units of irradiated concentrated platelets. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37511 | 1220648-2020-00346 | 9726552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9726552&pc=OZD | 10/17/2019 | 03/26/2020 |  | 02/19/2020 |  |  | 01/23/2020 | 02/19/2020 |  |  | 02/28/2021 |  | 02/19/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1412004 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/23/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 2.5 pump inserted in the right femoral for exacerbation of mitral regurgitation. the physician reported hemolysis was suspected during pump support. the patient was administered 2 units of irradiated red blood cells. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella 2.5 was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37512 | 1220648-2020-00340 | 9707361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9707361&pc=OZD | 01/20/2020 | 03/30/2020 |  | 02/13/2020 |  |  | 01/20/2020 | 02/13/2020 |  |  | 02/28/2021 |  | 02/13/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1406393 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/20/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) old male patient with acute myocardial infarction (ami)/cardiogenic shock (cgs) had impella cp pump inserted in the right femoral. after three hours the patient¿s urine was dark red and labs hemolyzed. the decision was made to remove the device early as the patient¿s hematocrit was at 55. | [5 empty] | The product was returned for evaluation. the pump was visually inspected and found free of damage, residue of unknown origin was found on the outflow cage. the pump could not be hemolysis tested due to an unrelated issue with the returned product and the root cause of the hemolysis was unable to be determined. | The impella cp was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37513 | 1220648-2020-00341 | 9709033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9709033&pc=OZD | 01/14/2020 | 03/30/2020 |  | 02/13/2020 |  |  | 01/14/2020 | 02/13/2020 |  |  | 07/31/2021 |  | 02/13/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1430738 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/14/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) old japanese female patient had impella 5.0 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that on the 8th day of impella support the physician suspected hemolysis. as a result, the device was explanted. | [5 empty] | The impella 5.0 was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion. | The pump was returned. no biomaterial was found. the device operated to specification when tested. the root cause of the hemolysis was unable to be determined because blood lab results were not reported, and the returned pump passed hemolysis testing. data logs showed that on day 1, outflow blocked alarms triggered possibly from the presence of biomaterial. dhr was reviewed and found no manufacturing related issues with this lot. there are no other complaints related this failure mode in this lot.
37514 | 1220648-2020-00339 | 9703427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9703427&pc=OZD | 01/13/2020 | 03/30/2020 |  | 02/12/2020 |  |  | 01/14/2020 | 02/12/2020 |  |  | 09/30/2021 |  | 02/12/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1447303 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 37 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/14/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese female patient had impella cp pump inserted in the left axillary for cardiomyopathy. it was reported the patient experienced hematuria which became slightly darker. the physician suspected hemolysis which was believed was due to impella. impellas position was confirmed via echocardiogram (echo). due to severity of the suspected hemolysis the impella device was explanted. after explant the patient remained stable. | [5 empty] | This information is based entirely on journal literature. this event occurred outside the us. patient information is limited due to confidentiality concerns. of note, multiple patients were noted in the article; however, a one to one correlation could not be made with unique product lot/serial numbers. since no device id was provided, it is unknown if this event has been previously reported. a request for additional information will be made and upon receipt a supplemental report will be submitted accordingly. referenced article: ¿improved electrical performance/stability of a novel active fixation coronary sinus lead compared to passive fixation leads: a multi-centre study.¿ european heart journal, volume 40, issue supplement_1, october 2019, ehz746.0623, https://doi.org/10.1093/eurheartj/ehz746.0623 poster session 6: cardiac resynchronisation therapy: p5681. esc congress 2019 together with world congress of cardiology. 31 august ¿ 4 september 2019, paris - france. if information is provided in the future, a supplemental report will be issued. | The pump was returned for review and no damage was found to any of the components in the flow of blood. the reported biomaterial was not found indicating the biomaterial evidence was likely removed during the hospital's pathology dept. investigation. the pump passed hemolysis testing. the root cause was unable to be determined because the failure could not be reproduced. data logs were not returned for review. dhr was reviewed and found no manufacturing related issues with this lot. there are no other complaints related this failure mode in this lot.
37515 | 9686451 | 9686451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9686451&pc=DYB | 12/18/2019 | 01/14/2020 | 01/17/2020 | 02/10/2020 | 01/17/2020 | 02/10/2020 |  | 02/10/2020 |  |  |  |  | 02/10/2020 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | DYB |  |  | IMPELLA CP PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | C1-17020 |  | 0052-3025 |  | No | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 29200 DA |  |  | 60 |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [28 empty] | Initial submission | Followup |  |  |  |  | Y | Invalid/Incomplete |  |  | Impella inserted right femoral artery. the next day it was removed, patient developed ischemic limb after removal and had to have a thrombectomy. | Impella inserted right femoral artery. the next day it was removed, patient developed ischemic limb after removal and had to have a thrombectomy.
37516 | 1220648-2020-00337 | 9682647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9682647&pc=OZD | 01/13/2020 | 04/01/2020 |  | 02/07/2020 |  |  | 01/13/2020 | 02/07/2020 |  |  | 09/30/2021 |  | 02/07/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1452465 |  | 005062-RO |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/13/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient presenting with cardiomyopathy for hemodynamic support with the impella 5.0 pump. during support, the patient exhibited signs of hemolysis via hemolyzed labs. the issue could not be resolved with normal troubleshooting attempts, therefore, the device required premature removal. | [5 empty] | The impella 5.0 was received and an investigation is underway. a supplemental report will be filed upon completion of the investigation. | The device was returned and an investigation was performed. simulated use testing for hemolysis was conducted and the device passed successfully. no data logs were returned from the customer. a review of the dhr was conducted and found no manufacturing issues, reworks, or complaints associated with this failure mode from this pump lot. since the reported device failure could not be replicated, and data logs were unavailable for analysis, we cannot determine the cause of the hemolysis.
37517 | 1220648-2020-00338 | 9682704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9682704&pc=OZD | 01/09/2020 | 02/07/2020 |  | 02/07/2020 |  |  | 01/10/2020 | 02/07/2020 |  |  | 05/31/2021 |  | 02/07/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  |  |  | 1424928 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/10/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed lower limb ischemia at the impella insertion site. the impella device was removed and the patient was surgically treated for the event. the patient remained stable. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37518 | 1220648-2020-00334 | 9675778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9675778&pc=OZD | 01/12/2020 | 03/24/2020 |  | 02/06/2020 |  |  | 01/12/2020 | 02/06/2020 |  |  | 05/31/2021 |  | 02/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA RP PUMP |  |  | OZD |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1437014 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 80 YR | [5 empty] | Malfunction | No | Yes | Separation Problem | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/12/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella rp pump inserted. the physician was removing the rp pump and upon removal the catheter had fractured from the motor housing. only the catheter was removed. x-ray confirmed that the rp motor housing was right at the left iliac. the pump remnants were removed through multiple means including snaring and forceps followed by cut down. plastic coating was noted to be chipped off by forceps, but it was able to be extracted. | [5 empty] | Corrected g4 pma/510(k) number from p140003 to p170011. the wire used to insert the replacement rp in the lfv was slightly larger at .032" vs the normal .027" diameter wire used for rp. the team was said to have difficulty placing the wire, but not the introducer. only data from the attempted restarts after the pump stop was returned. the returned pump confirmed the break at the motor housing, and the outflow cage was shown to have been bent in one direction. a returned image from the field shows the severed pump with the bent cage inside the patient, indicating possible placement difficulty. without further clinical information and complete logs, the root cause of the pump damage was unable to be definitively determined. | The impella rp was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37519 | 1220648-2020-00335 | 9675785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9675785&pc=OZD | 01/08/2020 | 02/06/2020 |  | 02/06/2020 |  |  | 01/08/2020 | 02/06/2020 |  |  | 07/31/2020 |  | 02/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | CP C-7 NO ACCESS PUMP SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1360428 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR | [5 empty] | Malfunction | No | Yes | Detachment of Device or Device Component | Infusion or Flow Problem | [5 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission |  |  |  |  | 01/08/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male tavr corevalve patient presenting for hemodynamic support with the impella cp pump. during support, the device experienced lower than expected flows. attempts to reposition the device did not resolve the issue, therefore, the pump was removed and replaced with a second impella pump. upon receipt and inspection of the device, damage to the device's impeller was identified. | [5 empty] | The impella cp was received and a failure investigation was completed. data log review and simulated use testing confirmed the clinical report of lower than expected flows during use of the device. upon physical examination of the pump, a chip was identified in on of the impeller blades. evidence suggests the damage was caused by an interaction between the impella and the patients mechanical aortic valve, however, this could not be confirmed. a review of the device history record found no manufacturing issues or reworks associated with this pump lot. there are no other complaints associated with pumps from this lot.
37520 | 1220648-2020-00333 | 9664208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9664208&pc=PYX | 01/06/2020 | 04/01/2020 |  | 02/04/2020 |  |  | 01/06/2020 | 02/04/2020 |  |  | 03/31/2021 |  | 02/04/2020 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1429348 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | Yes | Break | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/06/2020 | * | Invalid/Incomplete |  |  | A (b)(6) male presented post cardiotomy cardiogenic shock while on left ventricle hemodynamic support with an impella cp. an attempt to wean and remove the cp was unsuccessful as patient's condition would intolerant. a transthoracic echocardiogram was completed and the patient's right ventricle was found dilated. therefore, the decision was made to insert an impella rp to help improve the patient's condition enough to remove the cp. the physician chose to insert the rp in the left femoral vein due to the placement of the impella cp and due to the right common femoral vein being assessed as inaccessible with ultrasound guidance. despite recommendations in the impella instructions for use, and advisement of clinical field representative, the physician inserted the impella rp into the patients left femoral vein. during insertion, the device endured difficulty passing the pulmonary artery. the physician felt adequate placement was achieved, despite fluoroscopic imaging showing device "far from the bifurcation." upon removal of the guide wire, the device separated at the distal ball which connects the device's outlet cage to the pigtail. the rp was removed and a snare was utilized to remove the broken component. | [5 empty] | The impella rp was returned and an investigation was performed. physical observation confirmed a separation of the ball-pigtail from the inflow cage struts. multiple kinks in the catheter was also found, aligning with the user's report of difficult delivery. as stated in the initial mdr, the physician went against ifu recommendations by attempting delivery of the device into the opposing femoral vein. the right femoral vein is recommended due to ease of insertion as the pump has a straighter path to the ivc from the right side, applying less stress on the pump during delivery. a review of the dhr was conducted and found no manufacturing issues, reworks, or complaints associated with this failure mode from this pump lot. the root cause of the reported pump damage was excess stress on the device that led to a deformed outflow cage. | The impella rp was returned and an investigation is currently underway. upon completion, a supplemental mdr will be filed.
37521 | 1220648-2020-00331 | 9662111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9662111&pc=OZD | 01/06/2020 | 02/03/2020 |  | 02/03/2020 |  |  | 01/06/2020 | 02/03/2020 |  |  | 08/31/2021 |  | 02/03/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1434959 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/06/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the physician suspected the patient developed hemolysis. to treat the suspected hemolysis the impellas position was adjusted and the patient was administered an unknown amount of haptoglobin. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37522 | 1220648-2020-00332 | 9662170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9662170&pc=OZD | 01/06/2020 | 03/24/2020 |  | 02/03/2020 |  |  | 01/06/2020 | 02/03/2020 |  |  | 08/31/2021 |  | 02/03/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1434959 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 37 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/06/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the right axillary for cardiomyopathy. the physician reported the pump's p-level had to be lowered to p-3 due to a suction alarm. the suction alarm persisted with a decreased in pump flow. on (b)(6) 2020 the impella device was replaced with another impella device. upon removal, thrombus-like substance was observed at the inflow part and the tip of the pigtail. | [5 empty] | The impella cp was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion. | The impella cp was returned. the root cause of the low pump flow was most likely ingested biomaterial occluding pump flow. the returned device confirmed evidence of biomaterial on the pigtail. no biomaterial was found on the inlet/outlet or in the cannula. the pump ran to specification with good pump flows in the lab. the data logs confirmed good pump flow, which was within specification for five minutes prior to a high motor current event that was followed by low pump flow and suction alarms, which was evident of biomaterial ingestion. dhr was reviewed and found no manufacturing related issues with this lot. there are no other complaints related this failure mode in this lot.
37523 | 1220648-2020-00329 | 9654401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9654401&pc=PYX | 01/06/2020 | 04/01/2020 |  | 01/31/2020 |  |  | 01/06/2020 | 01/31/2020 |  |  | 03/31/2021 |  | 01/31/2020 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1429348 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Break | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/06/2020 | * | Invalid/Incomplete |  |  | A (b)(6) white male presented post cardiotomy cardiogenic shock while on left ventricle hemodynamic support with an impella cp. an attempt to wean and remove the cp was unsuccessful as patient's condition would intolerant. a transthoracic echocardiogram was completed and the patient's right ventricle was found dilated. therefore, the decision was made to insert an impella rp to help improve the patient's condition enough to remove the cp. the physician chose to insert the rp in the left femoral vein due to insertion of cp and right common femoral vein inaccessible under ultrasound guidance. despite recommendations in the impella instructions for use, and advisement of clinical field representative, the physician inserted the impella rp into the patients left femoral vein. during insertion, the device failed to access the pulmonary artery. the physician attempted to remove the device by pulling the guide wire, prior to pulling the pump. thereafter, the pump separated in the patient's inferior vena cava, at the distal ball which connects the device's outlet cage to the pigtail. the rp was removed and a snare was utilized to remove the detached component. | [5 empty] | The impella rp was returned and an investigation was performed. physical observation confirmed a separation of the ball-pigtail from the inflow cage struts. a kink in the catheter was also found, aligning with the user's report of difficult delivery. as stated in the initial mdr, the physician went against ifu recommendations by attempting delivery of the device into the opposing femoral vein. the right femoral vein is recommended due to ease of insertion as the pump has a straighter path to the ivc from the right side, applying less stress on the pump during delivery. a review of the dhr was conducted and found no manufacturing issues, reworks, or complaints associated with this failure mode from this pump lot. the root cause of the reported pump damage was excess stress on the device that led to a deformed outflow cage. | The impella rp was returned and an investigation is currently underway. upon completion, a supplemental mdr will be filed.
37524 | 1220648-2020-00330 | 9655890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9655890&pc=OZD | 01/03/2020 | 01/31/2020 |  | 01/31/2020 |  |  | 01/03/2020 | 01/31/2020 |  |  | 10/31/2021 |  | 01/31/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1446324 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | Perforation | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 01/03/2020 | * | Invalid/Incomplete |  |  | The complainant reported a white male patient had impella cp pump inserted for acute myocardial infarction (ami) with shock. an emergency echocardiography was done and found a pericardial effusion/pericardial tamponade, so an emergency pericardial drainage was performed. the physician believes that the impella could have possibly perforated the left ventricle as pericardial effusion/tamponade occurred relatively immediately after insertion of the impella pump. the physician does not know the cause of the patient's death. | [5 empty] | The patient's date of death is unknown. the impella cp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37525 | MW5092675 | 9659183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9659183&pc=PBL | 01/22/2020 | 01/29/2020 |  | 01/31/2020 |  |  |  | 01/31/2020 |  |  |  |  | 01/31/2020 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA CP |  |  | PBL |  |  | 42413 |  |  | ABIOMED, INC. | [8 empty] | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 54 YR |  |  | 85 |  |  | Injury | No | Yes | Positioning Problem | [6 empty] | Death | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | On (b)(6) 2020, impella implanted; (b)(6) 2020, the device required repositioned. after the catheter was repositioned and placement was confirmed correct, the device started alarming with the following message. placement signal not reliable. placement signal and suction monitoring and suspended. monitor pt hemodynamics and impella position with imaging; check the pt cable for kinks. call placed to the impella rep (b)(6) for guidance. the pt had an unusual wave form. on (b)(6) 2020 pt expired and device was removed and released to the impella rep. fda safety report id# (b)(4).
37526 | 1220648-2020-00327 | 9650113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9650113&pc=OZD | 01/05/2020 | 03/30/2020 |  | 01/30/2020 |  |  | 01/05/2020 | 01/30/2020 |  |  | 02/28/2021 |  | 01/30/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1405009 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/05/2020 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted in the right axillary artery for post-cardiotomy cardiogenic shock (pccs). it was reported the physician suspected hemolysis during impella support. to treat the suspected hemolysis the patient was transfused, specifics of the transfusion were not provided. the physical stated that due to impella being used in combination with extracorporeal membrane oxygenation (ecmo), it is unknown whether impella was the cause of the suspected hemolysis. no further information was provided. | [5 empty] | The impella 2.5 was returned. the root cause of the hemolysis was unable to be determined. there were no defects noted on the returned pump with the exception of biomaterial in the purge gap. data logs did not indicate any issues that may have contributed to the hemolysis. dhr was reviewed and found no manufacturing related issues with this lot. there are no other complaints related this failure mode in this lot. | The impella 2.5 was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37527 | 1220648-2020-00328 | 9650469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9650469&pc=OZD | 01/03/2020 | 03/03/2020 |  | 01/30/2020 |  |  | 01/20/2020 | 01/30/2020 |  |  | 08/31/2021 |  | 01/30/2020 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1435405 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG | [10 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | ECMO | [22 empty] | Initial submission | Followup |  |  |  | 01/03/2020 | * | Invalid/Incomplete |  |  | The complainant in (b)(6) had a male patient of unknown age admitted in cardiogenic shock and acute myocardial infarction. the team chose to place an impella cp for support, in combination with extra corporeal membrane oxygenation (ecmo). after two days of support there were signs of hemolysis. the medical team chose to assess the pump position and give volume. the impella was stated to be in good position. the following day the team drew labs and found the plasma free hemoglobin (pfhg) level to have risen above 130mg/dl. due to these signs of hemolysis, the impella pump was explanted and the team infused blood products. | [5 empty] | The complainant in (b)(6) has answered all clinical questions and returned the product for analysis. the analysis is on-going at this time. upon completion of the investigation, a final report will be submitted. | Since the initial report was filed the product and data investigation has completed. the complainant returned from the eu the pump and hemolysis testing was done. when the pump was returned, biomaterial was noted to be within the pump impeller. upon removal of the biomaterial the pump passed hemolysis testing. the data logs revealed the pump to be in appropriate position and without any suction or motor current alarms. the etiology of the biomaterial is not known, and as such the root cause of the hemolysis was unable to be determined. the failure mode will be monitored and trended.
37528 | MW5092537 | 9636482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9636482&pc=PBL | 01/19/2020 | 01/23/2020 |  | 01/27/2020 |  |  |  | 01/27/2020 |  |  |  |  | 01/27/2020 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | PBL | [5 empty] | ABIOMED EUROPE GMBH |  |  | US | [5 empty] | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 58 YR | [5 empty] | Death | Yes | Yes | Device Alarm System | Gas/Air Leak | [5 empty] | Air Embolism | Death | Hemolysis | Hypoxia | Renal Failure | [8 empty] | Death | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Impella cp placed for cardiogenic shock, purge solution line leak went undetected; no alarms came on, air embolism occurred, echo showed bubbles in the lv likely originating from the device; air embolism, severe hemolysis, acute renal failure and severe hypoxemia, lactic acidosis and pump malfunction ensued resulting in rapid death. fda safety report id# (b)(4).
37529 | 1220648-2020-00326 | 9624433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9624433&pc=OZD | 12/17/2019 | 01/23/2020 |  | 01/23/2020 |  |  | 12/25/2019 | 01/23/2020 |  |  | 08/31/2021 |  | 01/23/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | DANVERS, MA |  |  | US |  | 1441282 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/25/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)-year-old japanese male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient developed hemolysis for which blood products were administered. amount of blood product were not provided. no further information was provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37530 | 1220648-2020-00323 | 9616322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9616322&pc=OZD | 12/27/2019 | 02/21/2020 |  | 01/22/2020 |  |  | 12/27/2019 | 01/22/2020 |  |  | 08/31/2021 |  | 01/22/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED EUROPE, GMBH |  |  | GM |  | 1446358 |  | 05062 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 50 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Insufficiency | Mitral Regurgitation | [11 empty] | Hospitalization | Required Intervention |  |  |  |  | VA-ECMO | [22 empty] | Initial submission | Followup |  |  |  | 12/27/2019 | * | Invalid/Incomplete |  |  | A (b)(6) male patient admitted with cardiomyopathy was awaiting a heart transplant and being hemodynamically supported by the impella 5.0. the patient had been supported for 14 days when his condition began to deteriorate. the team placed va-ecmo to supplement the impella. there was concern of possible mitral valve damage from the impella. the impella remained on for support despite this until he received his transplant. in total the patient had been supported by impella for 18 days. | [5 empty] | Since the initial report was filed, the investigation has completed. the physician had a concern for mitral valve damage due to the use of the impella 5.0 pump. the pump was returned and found to have damage to it. there was kinking to the pigtail, a deformed inflow strut, a cut cannula, and hemostat-like markings on the catheter. all of these damages could have been due to the removal process, but this is not known definitively. clinical details of the event were not shared for analysis and investigation. the data logs and imaging were not returned for analysis. without further clinical details, data logs, and imaging of the patient, the mitral valve damage could not be confirmed, nor investigated. the failure mode will be monitored and trended. | The investigation into the presumed mitral valve damage is on-going at this time. the product has been returned, along with data logs, but the investigation has not yet completed. there is more clinical report still pending. upon completion a final report will be filed.
37531 | 1220648-2020-00324 | 9616402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9616402&pc=OZD | 12/27/2019 | 03/30/2020 |  | 01/22/2020 |  |  | 12/27/2019 | 01/22/2020 |  |  | 08/31/2021 |  | 01/22/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1436358 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup | Followup |  |  | 12/27/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted for acute myocardial infarction (ami)/cardiogenic shock (cgs) via the right femoral. the impella cp was removed 30 minutes after hemolysis was suspected by the clinical staff. hemolysis resolved 10 hours after removal of the impella cp. patient's urine was reported to be the color of red wine. | [5 empty] | We received information that hemolysis resolved with repositioning and the patient had pre-existing condition of mitral regurgitation (mr) and chordal cord rupture before insertion of impella. these events were unrelated to impella. | The cp pump was suspected of hemolysis during the case. it was unknown if hemolysis was confirmed through lab data. it was described that there was positioning issues that vary from the pump being too deep, with the outflow possibly interacting with the aortic valve, to the pump being pulled back while the patient was being moved, leaving the pump too shallow. the returned data logs confirmed that the positioning issues described occurred and the returned pump did not appear to have sharp surfaces near the shear zone, although the impeller was eccentric to the housing. the pump passed hemolysis testing per dir id (b)(4) in document (b)(4). the root cause of the suspected hemolysis was unable to be definitively determined because there were multiple compounding variables that could've contributed to the reported failures. | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37532 | 1220648-2020-00325 | 9616419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9616419&pc=OZD | 12/26/2019 | 01/22/2020 |  | 01/22/2020 |  |  | 12/26/2019 | 01/22/2020 |  |  | 09/30/2021 |  | 01/22/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1450138 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 56 YR | [5 empty] | Malfunction | No | Yes | Failure to Unfold or Unwrap | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/26/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient had impella cp optical pump inserted in the right femoral. impella folded on itself and a snare was used to unfold the pump. | [5 empty] | The impella cp optical pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37533 | 1220648-2020-00322 | 9610419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9610419&pc=OZD | 01/09/2020 | 03/12/2020 |  | 01/21/2020 |  |  | 01/09/2020 | 01/21/2020 |  |  | 08/31/2021 |  | 01/21/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | AACHEN, GERMANY |  |  | GM |  | 1436620 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 57 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | Cardiac Perforation | [11 empty] | Required Intervention | [5 empty] | VA-ECMO | [22 empty] | Initial submission | Followup |  |  |  | 01/09/2020 | * | Invalid/Incomplete |  |  | The complainant in belgium was utilizing an impella cp for hemodynamic support of a patient admitted in cardiogenic shock of unknown etiology. during the support with impella the patient also had va-ecmo for gas exchange. the patient was hemodynamically unstable despite these interventions. the team observed a right atrium tamponade that was treated with a surgical patch. the following day a left ventricle perforation and tamponade was treated by a surgical patch. the impella was not explanted at this time, rather left within the left ventricle. on the following morning a third perforation was observed. this third perforation was again surgically treated by a thoracotomy. when this was done the impella was seen to have perforated the ventricle. the impella was explanted after 66 hours of support. the patient remains on va-ecmo support. | [5 empty] | Since the initial report was filed, the investigation into the left ventricle (lv) perforation has concluded. the team in the eu did return the pump product and the data logs for investigation. the pump was returned with a slight bend and defect, this defect was not of such a degree that it would cause the pump to be more predictive of a puncture/perforation to the lv. the data logs revealed that there were some erratic placement signals during the patient's cpr. the signal's erratic period was then followed by a flattening of the placement signal, which could be indicative of the time of the lv perforation. the clinical details were also analyzed. the noted 3 perforations, 2 to the right atrium and 1 to the lv, are indicative of weak heart tissue. all of the investigation does not point to a definitive root cause of the lv perforation, though the assumption of the heart tissue being weak and cpr during impella support, are possible contributing factors. the failure mode will be monitored and trended. | To date the investigation into the left ventricular perforation has not concluded. the product has not yet been returned for investigation from belgium. further clinical details and clarification is pending. upon the completion of the investigation the final report will be submitted.
37534 | 1220648-2020-00318 | 9602754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9602754&pc=OZD | 12/19/2019 | 03/12/2020 |  | 01/17/2020 |  |  | 12/19/2019 | 01/17/2020 |  |  | 08/31/2021 |  | 01/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1448718 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 34 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Other | [28 empty] | Initial submission | Followup |  |  |  | 12/19/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient presenting with an ejection fraction of 5- 10% with acute myocardial infarction and cardiogenic shock. an impella cp optical pump was inserted for hemodynamic support and initiated without any reported issues. the device was repositioned, at which time air bubbles were found in the patient's left ventricle under echo imaging. despite this issue, no adverse symptoms exhibited by the patient. | [5 empty] | The impella cp was received and a failure investigation was completed. data log review found no instances of low purge pressure that could indicate air entering the purge line. examination of the purge system found no leaks or damage to the air air filter. simulated use testing was performed, however, was unsuccessful in reproducing the reported air bubbles. a review of the device history record found no manufacturing issues or reworks associated with this pump lot. there are no other complaints related to this failure mode associated with pumps from this lot. we are unable to determine a root cause for the reported air bubbles in the patient's ventricle. | The impella cp optical was returned and an investigation is underway. upon completion, a supplemental mdr will be filed.
37535 | 1220648-2020-00319 | 9602800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9602800&pc=OZD | 12/20/2019 | 03/12/2020 |  | 01/17/2020 |  |  | 12/20/2019 | 01/17/2020 |  |  | 11/30/2021 |  | 01/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1451969 |  | 0550-0008 |  | Yes | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/20/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient presenting with post cardiotomy cardiogenic shock for hemodynamic support using the impella 5.5 device. after successfully placing the device in the patient's ventricle, it was started without issue. during support, the anesthesiologist inaccurately interpreted the placement of the device as shallow, despite warnings from clinical representative, and advanced the device 2 cm. as a result, the device punctured the ventricular wall. as treatment, patient was placed on cardiopulmonary bypass, 2 sutures were placed, and blood products were given. after repair, support continued successfully. | [5 empty] | Correction to d(10): the initial mdr erroneously had "device available for evaluation?" reported as no and should have been reported as yes. this has been corrected. h(10): the impella 5.5 was received and a failure investigation was completed. upon examination of the pump, evidence of excessive force was identified via kink in the catheter near the yoke. a review of the device history record found no manufacturing issues or reworks associated with this pump lot. there are no other complaints related to this failure mode associated with pumps from this lot evidence identified during the investigation suggests the root cause of the ventricular perforation was due to a combination of patient condition and human error. | The impella cp was returned and an investigation is underway. upon completion, a supplemental mdr will be filed.
37536 | 1220648-2020-00320 | 9602809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9602809&pc=OZD | 12/20/2019 | 01/17/2020 |  | 01/17/2020 |  |  | 12/20/2019 | 01/17/2020 |  |  | 06/30/2021 |  | 01/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1440565 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/20/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp optical device. the device was inserted without any reported issues, however, the patient developed intermittent loss of pedal pulses during the case. approximately 12 hours after device removal, the patient's leg required amputation. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37537 | 1220648-2020-00321 | 9603765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9603765&pc=OZD | 12/15/2019 | 01/17/2020 |  | 01/17/2020 |  |  | 12/24/2019 | 01/17/2020 |  |  | 02/28/2021 |  | 01/17/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1406291 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/24/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted in the left femoral for severe heart failure. it was reported that patient developed aortic regurgitation after impella was successfully weaned and explanted. aortic regurgitation was confirmed via echocardiogram. the patient received unspecified medical treatment and was placed under observation. no further information was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37538 | 1035166-2020-00004 | 9586194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9586194&pc=DYB | 12/13/2019 | 04/07/2020 |  | 01/14/2020 |  |  | 12/17/2019 | 01/14/2020 |  |  | 11/01/2022 |  | 01/14/2020 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-17774 |  | 0052-3000 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN | USER FACI |  |  |  | 83 YR |  |  | 120 |  |  | Malfunction | No | Yes | Break | [6 empty] | Death | Cardiogenic Shock | [40 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported the patient was presented to the physician with hemodynamic support for cardiogenic shock (cgs) and the clinic decided to implant impella cp. before using the abiomed 14f introducer, they used terumo radufocus introducer ii with 8f. the obturator part had been introduced without problems, everything runs in normal way. as they used the 14f introducer, they dilate with 10f / 12f. then they took the 14f introducer peel away system. before launching in the vessel, the top of the peel away system split in two pieces. so they took a second peel away system and everything runs normal. customer reported no medical or surgical intervention was reported, however there was a delay for several minutes. on (b)(6), 2020, the customer reported the therapy stopped the following night and the patient expired on friday, (b)(6), 2019. as reported by physician the patient was in poor condition, cause of death cpr in icu after very bad prognoses in cath lab. additional information has been requested. | [5 empty] | The following sections were updated in follow-up 1: b4, g4, g7, h2, h6 and h10: the device was used for treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. multiple attempts were made to obtain the information regarding device location however, no information is available. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. as per abiomed introducer sheath in-process and final inspection procedure, the devices in this lot were inspected as follows: 9.2.2 visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. slight discoloration from tipping process is acceptable. verify proper tip shape according to drawing. per section 12.2, visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. the instructions for use (ifu) informs the user: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, our investigation is still in progress. a follow-up report will be submitted as soon as the investigation is complete.
37539 | 1220648-2020-00317 | 9580552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9580552&pc=OZD | 11/12/2019 | 01/31/2020 |  | 01/13/2020 |  |  | 12/18/2019 | 01/13/2020 |  |  | 12/31/2020 |  | 01/13/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1391206 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [41 empty] | Initial submission | Followup |  |  |  | 12/18/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese male patient had impella 2.5 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed a hemorrhage at the impella insertion site. as a result, the patient received 14 units of red blood cells and 10 units of plasma. no further information was provided. | The complainant reported a 74 year old japanese female patient had impella 2.5 pump inserted in the right femoral. |  |  |  |  | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | A2, a3, b5: correction.
37540 | 1220648-2020-00316 | 9571383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9571383&pc=OZD | 12/17/2019 | 02/13/2020 |  | 01/09/2020 |  |  | 12/17/2019 | 01/09/2020 |  |  | 11/30/2020 |  | 01/09/2020 |  | TEMPORARY NON ROLLER LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP, APAC |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1399579 |  | 0048-0035 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [22 empty] | Initial submission | Followup |  |  |  | 12/17/2019 | * | Invalid/Incomplete |  |  | Upon admission to the cardiac cath lab an impella cp was placed for hemodynamic support. the patient suffered a hematoma at the impella access site. the team called for vascular surgery to perform a direct mattress suture and infused blood products. the team in hong kong noted 4-5 bags of blood product, but exact amount in cc's is unknown. the patient had his impella pump weaned and was supported afterwards by ecmo (extracorporeal membrane oxygenation). | [5 empty] | To date the product has not yet been returned from the user facility in (b)(6). without the product the investigation has not yet completed. upon return of the product and analysis a final report will be drafted. until that time the hematoma event will continue to monitored and trended. | Since the initial report was filed, the complainant in hong kong returned the product. the physician had noted excessive bleeding upon placement of the impella repositioning sheath. the product was returned with significant damage but, the cause of the damage was not determined. the team had clamped off the product during removal, and as such the etiology of the product damage could not be determined. the root cause of the bleeding at the impella access site is unknown. the further clinical details were not provided. the anticoagulation regimen was not shared. the failure mode will be monitored and trended.
37541 | 1220648-2020-00311 | 9563731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9563731&pc=OZD | 12/13/2019 | 03/26/2020 |  | 01/08/2020 |  |  | 12/13/2019 | 01/08/2020 |  |  | 11/30/2020 |  | 01/08/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1385057 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/13/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the physician suspected the patient developed hemolysis. the patient was administered six units of red blood cells and 8000 units of haptoglobin. | Additional information was provided indicating patient¿s pre-existing condition of renal failure worsened because of the hemolysis. as a result the patient received dialysis. no further information was provided. |  |  |  |  | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37542 | 1220648-2020-00312 | 9564752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9564752&pc=OZD | 12/13/2019 | 01/08/2020 |  | 01/08/2020 |  |  | 12/13/2019 | 01/08/2020 |  |  | 08/31/2019 |  | 01/08/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1435812 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 23 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/13/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese female patient had impella cp pump inserted in the left femoral for myocarditis. it was reported that when explanting impella, bleeding was noted, location was not provided. balloon technique was used to stop the bleeding; however, it was not successful; therefore, protamine was administered. lower limb angiography confirmed the impella insertion site looked abnormal. the patient did not complain of any pain; however, lower limb ischemia and thrombus formation were noted. lower limb embolization and thrombus aspiration was performed for treatment. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37543 | 1220648-2020-00313 | 9564796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9564796&pc=OZD | 10/29/2019 | 01/08/2020 |  | 01/08/2020 |  |  | 12/12/2019 | 01/08/2020 |  |  | 12/31/2020 |  | 01/08/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1391180 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 51 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/12/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese male patient had impella 2.5 pump inserted in the right femoral for cardiomyopathy. it was reported the patient developed a hemorrhage at the impella access site. as a result, the patient was administered 42 units of total blood transfusion. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37544 | 1220648-2020-00314 | 9566148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9566148&pc=OZD | 10/28/2019 | 01/08/2020 |  | 01/08/2020 |  |  | 10/28/2019 | 01/08/2020 |  |  | 01/31/2021 |  | 01/08/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1417542 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/28/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient had impella cp pump inserted for cardiogenic shock (cgs). the patient¿s leg did not have pulses distal to right leg, so pump was removed for concerns of leg ischemia. as treatment a balloon was also inserted, and the patient¿s leg recovered. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37545 | 1220648-2020-00315 | 9566187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9566187&pc=OZD | 05/29/2019 | 01/08/2020 |  | 01/08/2020 |  |  | 12/13/2019 | 01/08/2020 |  |  | 02/28/2021 |  | 01/08/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1405009 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/13/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted in the right femoral for cardiomyopathy. the physician reported impella interfered with the mitral valve leaflets or chordae when the inlet sucked in either the leaflets or chordae. the patient experienced intermittent exacerbation of the mitral complex. the reported issue was confirmed via echocardiogram. as a result, the impella device was explanted. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37546 | 1220648-2020-00309 | 9557978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9557978&pc=OZD | 12/13/2019 | 02/07/2020 |  | 01/07/2020 |  |  | 12/13/2019 | 01/07/2020 |  |  | 05/31/2021 |  | 01/07/2020 |  | TEMPORARY NON ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1436304 |  | 005062 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 53 YR | [5 empty] | Malfunction | No | Yes | Contamination | [6 empty] | Unspecified Infection | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION) | [22 empty] | Initial submission | Followup |  |  |  | 12/13/2019 | * | Invalid/Incomplete |  |  | The complainant was utilizing an impella 5.0 for hemodynamic support, in conjunction with ecmo, for a patient in biventricular failure. during the following four days of support, the patient was noted to have diarrhea while on support in the icu. the diarrhea was found to have contaminated the impella pump's purge reservoir. the team returned to the operating suite to replace the pump. in addition to the pump replacement the patient received antibiotics. | [5 empty] | Since the initial report was filed, the investigation into the matter has concluded. the team did return for analysis the data logs, but not the impella 5.0 pump that was contaminated with fecal matter. the data logs did not reveal any large purge system leak that could have allowed for fecal matter ingress. the root cause of the ingress of the fecal matter could not be determined. the medical team did inform abiomed of the culture results. there was no infection present however, the team had given preventative antibiotics to the patient and monitored for a few weeks post impella explant. the failure mode will be monitored and trended. | The investigation into the purge contamination with the fecal matter/diarrhea is on-going at this time. upon the completion of the investigation a final report will be submitted.
37547 | 1220648-2020-00310 | 9559193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9559193&pc=OZD | 12/12/2019 | 01/07/2020 |  | 01/07/2020 |  |  | 12/12/2019 | 01/07/2020 |  |  | 12/31/2020 |  | 01/07/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1382261 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Insufficiency, Valvular | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/12/2019 | * | Invalid/Incomplete |  |  | The complainant reported a male patient, of unknown age or race, presenting with a ventricular septal defect (vsd) with acute myocardial infarction. the impella cp was selected for hemodynamic support and inserted without issue. the patient exhibited signs of "severe" hemolysis during support that was suspected as a result of the device being placed too deeply. the procedure to repair the vsd was successful and the device was removed, however, aortic insufficiency was found and aortic valve replacement was conducted. during surgery, abrasions on the valve leaflets were identified and suspected as the result of the inflow cage being too close. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37548 | 1220648-2020-00305 | 9554524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9554524&pc=OZD | 12/10/2019 | 03/12/2020 |  | 01/06/2020 |  |  | 12/10/2019 | 01/06/2020 |  |  | 09/30/2021 |  | 01/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1442444 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/10/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp pump. approximately 2 days after placement, the patient exhibited diminished pulses in the foot of the impella inserted leg. as a result, the impella was removed. no new pump insertion was required, as the patient's hemodynamics had improved. | [5 empty] | The impella cp was returned and an investigation is underway. a supplemental mdr will be filed upon completion of the investigation. | The impella cp was received and a failure investigation was completed. upon examination of the pump's repositioning sheath, the outer diameter od was measured to be within the specification. a review of the device history record confirmed completion of incoming inspection for the guidewire repositioning sheath lots associated with this pump lot. no manufacturing issues, reworks, or complaints are associated with this failure mode from this pump lot. due to lack of clinical data, a root cause for the patient's leg ischemia could not be determined.
37549 | 1220648-2020-00306 | 9554530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9554530&pc=OZD | 12/11/2019 | 02/12/2020 |  | 01/06/2020 |  |  | 12/11/2019 | 01/06/2020 |  |  | 05/31/2021 |  | 01/06/2020 |  | TEMPORARY NON ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1433338 |  | 05062 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 12/11/2019 | * | Invalid/Incomplete |  |  | The complainant had a patient with multiple cardiac comorbidities admitted and placed on impella hemodynamic support. during the support there were signs of hemolysis. the medical team noted elevated plasma free hemoglobin labs. due to the patient's deteriorating condition and no chance for care to be escalated at another medical facility, the patient's family chose to withdraw care after the third day of impella support. the patient expired and the impella pump was not removed due to anatomic limitations. | [5 empty] | The investigation into the hemolysis reported during impella use is on-going at this time. upon completion of the investigation, and if the root cause is able to be determined, a final report will be filed. | Since the initial report was filed, the investigation into the presumed hemolysis has concluded. the pump product was not returned for analysis and abiomed's protocol of hemolysis testing on the benchtop. the data logs were returned and revealed numerous pump positioning alarms. the pump being out of position can cause hemolysis, though without pump return, the root cause and a confirmation of hemolysis can not be verified. the failure mode will be monitored and trended.
37550 | 1220648-2020-00307 | 9554550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9554550&pc=OZD | 11/08/2019 | 01/06/2020 |  | 01/06/2020 |  |  | 12/13/2019 | 01/06/2020 |  |  | 08/31/2021 |  | 01/06/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1430528 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/13/2019 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient developed hemolysis during impella support. patient's lactate dehydrogenase (ldh) levels were confirmed to be 3010. as a results impella device was explanted. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37551 | 1220648-2020-00308 | 9554607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9554607&pc=OZD | 12/12/2019 | 02/07/2020 |  | 01/06/2020 |  |  | 12/12/2019 | 01/06/2020 |  |  | 07/31/2021 |  | 01/06/2020 |  | TEMPORARY NON ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1443457 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 12/12/2019 | * | Invalid/Incomplete |  |  | The complainant had an obese patient admitted for a coronary intervention and angioplasty. the team chose to place an impella cp for hemodynamic support. the placement was complicated due to the patient body habitus. the pump position was not optimal and was therefore repositioned in the cardiac catheterization lab. when the pump was repositioned the team observed a large hematoma that necessitated multiple units of blood replacement. a second, new pump was placed via another introducer and the support proceeded til wean and explant. | [5 empty] | The investigation into the cause of the access site hematoma is on-going at this time. upon completion, a final report will be filed. | The investigation has completed. the pump was returned and no defects were observed. the team had noted that the patient had a large panus. this anatomy made for a long tissue track. this caused difficulty in hemostasis being achieved by the length of the impella's re-positioning sheath. when the longer 14fr x30cm sheath was placed, hemostasis was achieved. patient anatomy contributed to the blood loss and access site hematoma. the ifu does caution for obesity. the failure mode will be monitored and trended.
37552 | 1220648-2020-00303 | 9544145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9544145&pc=OZD | 12/09/2019 | 01/03/2020 |  | 01/03/2020 |  |  | 12/09/2019 | 01/03/2020 |  |  | 01/31/2021 |  | 01/03/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1404634 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | HEALTH PROFESSIONAL | U |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/09/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the physician suspected hemolysis during impella support. to treat the suspected hemolysis patient¿s support levels were decreased, anticoagulant therapy was enhanced, and intravascular volume was increased. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37553 | 1220648-2020-00304 | 9548324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9548324&pc=OZD | 12/09/2019 | 03/04/2020 |  | 01/03/2020 |  |  | 12/09/2019 | 01/03/2020 |  |  | 10/31/2020 |  | 01/03/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1376920 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/09/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese male patient had impella 2.5 pump inserted in the right femoral for cardiomyopathy. it was reported that after peel away and pushing away the repo sheath there was a blood leak observed from the side. the angle and depth of the sheath were adjusted and the issue was resolved. it was noted the leak was coming from a gap between the repo sheath and the catheter shaft. the gap was closed with bone wax. the patient received an unknown amount of blood products. | [5 empty] | The impella 2.5 was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion. | The impella 2.5 as returned for investigation. product analysis revealed no pump damage. the outer diameter of the repo sheath was within specification. the root cause of the access site bleeding cannot be determined as no issues were found with the pump and no clinical abnormalities were reported during the case. data logs do not pertain to the failure mode. dhr was reviewed and found no manufacturing related issues with this lot. there one other complaint related this failure mode in this lot.
37554 | 1220648-2020-00302 | 9540527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9540527&pc=OZD | 12/07/2019 | 03/12/2020 |  | 01/02/2020 |  |  | 12/07/2019 | 01/02/2020 |  |  | 08/31/2021 |  | 01/02/2020 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1436348 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Infarction, Cerebral | Ischemia | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/07/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the physician stated the patient showed signs of lower limb ischemia. to deliver blood flow to lower limbs a sheath was inserted into the left femoral artery (fa) and right femoral vein (fv) with volume being inserted due to volume loss. the right fa was bypassed from the left fa into the lower extremities. the patient¿s limbs exacerbated, and amputation of both limbs was performed 15cm from the knee. the physician indicated there was no issue with the puncture site. the distal end of the sheath was located at the third branch from the superficial femoral artery. the patient's blood vessels were also initially weak and became further blocked when the sheath was inserted. | [5 empty] | The impella cp was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion. | The impella cp was returned. no issues were found with the gwru body and its od was measured to be within specification. the root cause of the leg ischemia was most likely the patient's pre-existing condition of narrow vasculature, but was exacerbated by improper access technique. furthermore, the root cause of the cerebral infarction was unable to be determined. data logs do not pertain to the reported event. dhr was reviewed and found no manufacturing related issues with this lot. there are no other complaints related this failure mode in this lot.
37555 | 1220648-2019-00301 | 9535623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9535623&pc=OZD | 12/04/2019 | 02/03/2020 |  | 12/31/2019 |  |  | 12/04/2019 | 12/31/2019 |  |  | 08/31/2021 |  | 12/31/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1438723 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 46 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/04/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male patient had impella cp pump inserted in the right femoral for acute myocardial infarction (ami) with shock. after impella was removed, clot was noticed so the patient was taken to the operating room (or) for vascular surgery. | [5 empty] | The returned pump was seen to have a clot on the outflow cage, after removal the pump ran within specification in a basin of water. the data logs showed low flows and suction alarms occurred on (b)(6) 2019 after 4.5 days of successful support. heparin level in the purge was reduced prior to low flows. the root cause of the low flows was built up biomaterial which wound around the impeller over time and caused low flow. there was not a notable motor current spike indicating biomaterial ingestion. the root cause of the clot was due to built up biomaterial. | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37556 | 1035166-2019-00125 | 9536406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9536406&pc=DYB | 12/10/2019 | 03/19/2020 |  | 12/31/2019 |  |  | 12/11/2019 | 12/31/2019 |  |  | 01/01/2022 |  | 12/31/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | CR-02078 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 86 YR | [5 empty] | Malfunction | No | Yes | Break | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported a 14 fr x 23 cm that the dilator tip frayed (looks to be split on the distal tip). the issue was identified during procedure, not inserted into the body, never entered into the vessel. procedure was completed with an alternative dilator from a sheath kit, no patient complication or adverse event reported. additional information requested. | [5 empty] | The following sections were updated in follow-up 1: b4, d10, g4, g7, h2, h3, h6 and h10: the device was used in treatment. one dilator from an abiomed 14f introducer kit was returned from the customer. the introducer sheath was not returned and there were no other accessories. blood was found on and inside the dilator. upon evaluation of the returned product, it was found the distal tip of the dilator was damaged. the damage to the tip looks like a sharp object may have carved into the tip and almost sheared off a piece of the tip. the dilator is within manufacturing specifications. no manufacturing defects were found and the reason for return cannot be confirmed. per abiomed introducer dilator in-process and final inspection procedure: tipping - first 5 properly tipped parts, inspect 100%. visual inspection: using 10x microscope magnification, verify tip of dilator is round. verify tipped dilator is free of flash, splitting, cracks, or any other damages. verify proper tip shape according to drawing. insert a 0.037" mandrel or 0.038" guidewire into dilator hub through to tip and check if mandrel/guidewire inserts without excessive force or obstruction. also per abiomed 14f introducer kit packaging inspection for impella, 13cm and 25cm procedure: ensure dilator tips are perfectly round and well opened. this inspection is performed 100% by qa. the instructions for use (ifu) informs the user: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete.
37557 | 1220648-2019-00300 | 9531445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9531445&pc=OZD | 12/03/2019 | 12/30/2019 |  | 12/30/2019 |  |  | 12/03/2019 | 12/30/2019 |  |  | 09/30/2021 |  | 12/30/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1445812 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/03/2019 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old white male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp pump. after pump placement and initiation, the repositioning sheath was advanced and a hematoma formed. as treatment, vascular surgery of the vessel was performed and 2 units of blood products were delivered. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37558 | 1220648-2019-00299 | 9531488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9531488&pc=OZD | 12/04/2019 | 01/31/2020 |  | 12/30/2019 |  |  | 12/04/2019 | 12/30/2019 |  |  | 06/30/2021 |  | 12/30/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1442567 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Malposition of Device | Device Dislodged or Dislocated | [5 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 12/04/2019 | * | Invalid/Incomplete |  |  | The complainant had a patient admitted for a high risk coronary angioplasty of the left main and left anterior descending coronary arteries. the team placed an impella cp for hemodynamic support during the angioplasty. after the successful angioplasty, imaging was performed and a pericardial effusion and left ventricular perforation by the impella was diagnosed. the patient was treated surgically and the impella was positioned back into the appropriate location within the left ventricle. the patient received blood products. after two days of support, the impella was weaned and explanted successfully. | [5 empty] | Since the initial report was filed, the investigation into the left ventricular wall perforation has completed. the pump was returned and no major deformities were fond. the data logs revealed alarms of the pump being out of proper position following patient movement. the imaging which would confirm the pump being in the pericardial space was not shared for analysis. the root cause of the perforation could not be determined. however, the physician had noted that during patient movement the perforation could have taken place. the failure mode will be monitored and trended. | The investigation is on-going at this time. upon completion of the investigation, a final report will be filed regarding the left ventricle perforation.
37559 | 1220648-2019-00296 | 9519639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9519639&pc=OZD | 12/03/2019 | 12/26/2019 |  | 12/26/2019 |  |  | 12/03/2019 | 12/26/2019 |  |  | 11/30/2020 |  | 12/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1397086 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | No Code Available | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/03/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese male patient had impella 5.0 pump inserted in the right axillary for cardiomyopathy. it was reported that during impella support hemolysis was suspected. the patient was administered 36 units of red cell count to treat the suspected hemolysis. subsequently, the patient developed thrombocytopenia for which 40 units of platelets were administered. ultimately impella device was explanted due to suspected hemolysis and thrombocytopenia. | [5 empty] | (b)(4). the impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37560 | 1220648-2019-00295 | 9519640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9519640&pc=OZD | 11/29/2019 | 12/26/2019 |  | 12/26/2019 |  |  | 11/29/2019 | 12/26/2019 |  |  | 11/30/2020 |  | 12/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1385057 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 29 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | No Code Available | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/29/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese female patient had impella 2.5 pump inserted in the right femoral for myocarditis. it was reported the patient developed thrombocytopenia during pump support. for treatment the patient received 14 units of fresh frozen plasma and 20 units of concentrated irradiated platelets. subsequently, cyanosis in the right lower limb worsened. to alleviate cyanosis, bypassed from the right femoral artery to the percutaneous cardiopulmonary support (pcps) blood transfer route to secure blood flow to the right lower limb was performed. unfortunately, the patient developed lower limb ischemia at the right lower leg. the patient¿s right leg was ultimately amputated. | [5 empty] | (b)(4). the impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37561 | 1220648-2019-00297 | 9519647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9519647&pc=OZD | 12/03/2019 | 12/26/2019 |  | 12/26/2019 |  |  | 12/03/2019 | 12/26/2019 |  |  | 10/31/2020 |  | 12/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1378257 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/03/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese male patient had impella 5.0 pump inserted in the right axillary for ventricular tachycardia storm and unstable hemodynamics. it was reported the patient developed a hematoma at the impella access site. the wound was opened, and vacuum-assisted closure therapy was applied. the patient was also administered 2 units of mannitol, adenine, and phosphate (map). after impella device was explanted the patient experienced vascular complication, specifics were not provided; however, the patient required right subclavian artery reconstruction. the patient successfully recovered from hematoma and vascular complication. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37562 | 1220648-2019-00298 | 9519655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9519655&pc=OZD | 12/02/2019 | 12/26/2019 |  | 12/26/2019 |  |  | 12/02/2019 | 12/26/2019 |  |  | 08/31/2021 |  | 12/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1435812 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/02/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient had impella cp pump inserted in the right femoral for pre-op coronary artery bypass grafting (cabg). it was reported that while performing a thoracotomy due to left ventricle perforation, the impella was noted to be at the apex of the heart. the cause of the left ventricle perforation is unknown as it may have been cause when the physician deepened the position of the impella. the device was successfully explanted. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37563 | MW5091915 | 9525978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9525978&pc=OZD | 12/19/2019 | 12/23/2019 |  | 12/26/2019 |  |  |  | 12/26/2019 |  |  |  |  | 12/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED INC. |  |  | US | [5 empty] | No answer provided | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 66 YR |  |  | 98 |  |  | Injury | No | Yes | Fluid/Blood Leak | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Impella 5.5 purge reservoir found leaking, impella rep and ccu fellow at bedside and bypass it by using filtered needle and secured it with tape and splint. no leaking of purge fluid noted afterwards. a similar incident report was entered on (b)(6) 2019 concerning the same device on a different pt verge report (b)(4). fda safety report id# (b)(4).
37564 | 1220648-2019-00292 | 9516566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9516566&pc=OZD | 11/30/2019 | 12/24/2019 |  | 12/24/2019 |  |  | 11/30/2019 | 12/24/2019 |  |  | 08/31/2021 |  | 12/24/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1444442 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 38 YR | [5 empty] | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/30/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male patient had impella cp pump inserted in the right femoral for myocarditis. it was reported the patient developed hemolysis during impella support. the patient was administered haptoglobin to treat hemolysis, amount was not provided. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37565 | 1220648-2019-00293 | 9516572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9516572&pc=OZD | 11/26/2019 | 12/24/2019 |  | 12/24/2019 |  |  | 11/26/2019 | 12/24/2019 |  |  | 09/30/2020 |  | 12/24/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1367729 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 71 YR | [5 empty] | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/26/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient experienced access site bleeding during pump support. the patient was administered 4 units of irradiated red blood cells and leukocytes were reduced. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37566 | 1220648-2019-00294 | 9516599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9516599&pc=OZD | 12/02/2019 | 12/24/2019 |  | 12/24/2019 |  |  | 12/02/2019 | 12/24/2019 |  |  | 01/31/2021 |  | 12/24/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1396842 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/02/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese female patient had impella 2.5 pump inserted in the right femoral for cardiogenic shock caused by vsp. it was reported the physician suspected hemolysis during impella support. to treat the suspected hemolysis the patient was administered haptoglobin. support levels and diuretics were also reduced. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37567 | 1220648-2019-00291 | 9512779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9512779&pc=OZD | 11/29/2019 | 02/27/2020 |  | 12/23/2019 |  |  | 11/29/2019 | 12/23/2019 |  |  | 08/31/2021 |  | 12/23/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1434848 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 84 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 11/29/2019 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that at the time of impella explant there was a thrombus of about 1mm noted. the thrombus was removed using a fogarty embolectomy catheter. there was no blood flow obstruction after removal. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed. | The impella cp was returned for investigation. the root cause of the thrombus cannot be determined, as no information is known regarding the patient's anticoagulation regimen throughout the case. product analysis revealed a small amount of biomaterial in the outflow cage but no large thrombus was confirmed. dhr was reviewed and found no manufacturing related issues with this lot. there is one other complaint related this failure mode in this lot.
37568 | 1220648-2019-00290 | 9512792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9512792&pc=OZD | 11/26/2019 | 12/23/2019 |  | 12/23/2019 |  |  | 11/26/2019 | 12/23/2019 |  |  |  |  | 12/23/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH | [10 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/26/2019 | * | Invalid/Incomplete |  |  | The complainant reported a patient of unknown age or gender presenting for high rick percutaneous coronary intervention. in preparation to insert an impella cp pump, a 14 fr oscor peel-away sheath was placed. subsequently, an arterial perforation was discovered. a balloon tamponade was placed and the patient's artery was repaired. the patient's procedure was subsequently carried out without impella support. | [5 empty] | The 14 fr introducer was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37569 | 1220648-2019-00288 | 9502817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9502817&pc=OZD | 11/29/2019 | 12/20/2019 |  | 12/20/2019 |  |  | 11/29/2019 | 12/20/2019 |  |  | 07/31/2021 |  | 12/20/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1445555 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Improper or Incorrect Procedure or Method | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/29/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp optical pump. the pump was inserted without any reported issues, however, the physician opted to leave the 14 fr peel-away introducer in the access site. during support, the patient's foot became pale and dopplerable pulses could not be obtained. the peel-away introducer was then removed, however, the pump fell out of the ventricle and was unable to be properly re-advanced. the pump was removed, a new pump was placed, and a fem-fem bypass was administered. the patient's ischemia resolved. | [5 empty] | The impella cp optical and 14 fr introducer were discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37570 | 1220648-2019-00289 | 9503324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9503324&pc=OZD | 11/25/2019 | 02/21/2020 |  | 12/20/2019 |  |  | 11/25/2019 | 12/20/2019 |  |  | 07/31/2021 |  | 12/20/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1428650 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 71 YR | [5 empty] | Malfunction | No | Yes | Separation Problem | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 11/25/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella cp pump inserted in the right femoral for acute myocardial infarction (ami) with shock. the pump was bent and twisted in the ventricle, in the ascending aorta the pump separated into two (2) parts. the first piece was removed by the rest of the shaft and the second piece (inlet area and pigtail) was successfully removed after 15 minutes with andra snare. | [5 empty] | Per ifu 0042-9028 re, impella cp is not intended to be inserted into the lv without the use of a guidewire. the root cause of the cannula being detached was wireless re-access resulting in the pump being bent and twisted in the ventricle leading to the cannula detaching in the ascending aorta during pump removal. the root cause of the hematoma was unable to be determined due to insufficient clinical details. | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37571 | 1035166-2019-00124 | 9504739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9504739&pc=DYB | 11/26/2019 | 03/30/2020 |  | 12/20/2019 |  |  | 12/02/2019 | 12/20/2019 |  |  |  |  | 12/20/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | UNKNOWN |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [11 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Perforation | Cardiac Tamponade | [10 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | Additional information md felt that it was just oozing from the small 6fr sheath. angiomax was started, 14fr impella sheath was placed. angiomax stopped, balloon placed and artery was repaired. patient had pci without impella support. field rep believe's the perforation was caused by the initial stick by the physician. | It was reported that after placing an oscor 14fr introducer (material # 0052-3025, lot # unknown) a perforation in the femoral artery resulted and a hematoma formed. medical intervention - balloon inserted to spot of tamponade. percutaneous coronary intervention performed without impella pump use. additional information has been requested. |  |  |  |  | The following sections were updated in follow-up 1: b4, b5, g4, g7, h2, h6, and h10: the device was used in treatment. the device was discarded and not returned for evaluation. there was no specific performance issue related to failure reported by the user. no further investigation is required. the lot number of this device was not provided, therefore neither a review of the device history recored nor complaint history could be performed. inspection procedures require any oscor product to pass all in-process and qa final inspections before shipping to the customer. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. no further follow-up is required. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete.
37572 | 1220648-2019-00287 | 9497051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9497051&pc=OZD | 11/22/2019 | 12/19/2019 |  | 12/19/2019 |  |  | 11/22/2019 | 12/19/2019 |  |  | 09/30/2020 |  | 12/19/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1367729 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/22/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male patient had impella 2.5 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). during pump support hemolysis was suspected. the physician believes the suspected hemolysis was caused by the impella device. one unit of haptoglobin was administered as treatment. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37573 | MW5091750 | 9495056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9495056&pc=OZD | 07/01/2016 | 12/16/2019 |  | 12/18/2019 |  |  |  | 12/18/2019 |  |  |  |  | 12/18/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. | [8 empty] | No | [5 empty] | No | No Information |  |  | Voluntary report | [8 empty] | 16 YR |  |  | 59 |  |  | Injury |  | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Scar Tissue | No Code Available | [11 empty] | Hospitalization | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Unnecessary and harmful implanted defibrillator (icd) with ongoing contractions to the chest and heads from the wires and pulse generator constricting the chest. increasing scar tissue prevents removing without harm. implant by (b)(6) hospital to the side effects of midazolam by paramedics and er combined with intubation and arctic sun cooling pads. five questions about enforcing a device. first, is abiomd¿s (abmd) impella heart pump unnecessarily being used, and not for heart failure or cardiac disease as fda approved in 2016? second, are abmd and (b)(6) hospital promoting themselves with doctor photos and personal-interest story in the (b)(6) about impella and patient (b)(6) instead of reporting about medical practice to health journals? third, did the patient (b)(6) have syncope or fainting, and not a lapse of consciousness (loc) from cardiogenics? fourth, was (b)(6) medically-induced by (b)(6) hospital with sedation, intubation and cooling pads, and not his ventricles? the reports indicated the excessive and unnecessary treatment by (b)(6). five, did the (b)(6) year old patient have the short-term enlarged heart of a growing body and sports activity, and not long-term cardiac enlargement and cardiomyopathy for implanting a defibrillator? and was the defibrillator implanted without the fda-required 60 day waiting period for cardiac recovery such as the side effects of abiomd¿s impella heart pump on the ventricles? fda safety report id # (b)(4).
37574 | 1220648-2019-00285 | 9482660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9482660&pc=OZD | 11/19/2019 | 12/17/2019 |  | 12/17/2019 |  |  | 11/19/2019 | 12/17/2019 |  |  | 02/28/2021 |  | 12/17/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1412004 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/19/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that when explanting impella due to patient's do-not-resuscitate (dnr) order, a blood clot flew into superficial femoral artery causing lower limb ischemia. as the patient was dnr, no treatment was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37575 | 1220648-2019-00286 | 9482711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9482711&pc=OZD | 11/20/2019 | 12/17/2019 |  | 12/17/2019 |  |  | 11/20/2019 | 12/17/2019 |  |  | 02/28/2021 |  | 12/17/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1408367 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/20/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that during impella support the patient developed lower limb ischemia at impella access site. vascular suturing and blood removal were performed to treat limb ischemia. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37576 | 1220648-2019-00284 | 9475422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9475422&pc=OZD | 11/23/2019 | 01/24/2020 |  | 12/16/2019 |  |  | 11/23/2019 | 12/16/2019 |  |  | 07/31/2021 |  | 12/16/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  |  |  | 1443723 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 43 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Hospitalization | [28 empty] | Initial submission | Followup |  |  |  | 11/23/2019 | * | Invalid/Incomplete |  |  | The complainant was utilizing an impella cp for hemodynamic support of a patient admitted in cardiogenic shock. during the support of 46 hours the patient exhibited signs of hemolysis in urine color change and suction alarms of the pump. the team chose to explant the impella and replace it with another pump. | [5 empty] | Since the initial report was filed, the investigation into the presumed hemolysis has concluded. the pump and data logs were returned for analysis. in benchtop testing the pump passed abiomed's hemolysis testing. there was observed biomaterial ingested into the pump upon the pump's review. the data logs also revealed signatures indicative of biomaterial ingestion and expulsion during the support. the hemolysis, and lower that expected pump flows, was most likely caused by ingested biomaterial. the failure mode will be monitored and trended. | The investigation is on-going at this time. not all of the impella product has been returned. further clarification of the treatment for the presumed hemolysis is pending. to date the treatment beyond pump replacement is unknown. there will be a final report drafted upon completion of the investigation.
37577 | 1220648-2019-00280 | 9466721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9466721&pc=OZD | 11/15/2019 | 12/13/2019 |  | 12/13/2019 |  |  | 11/15/2019 | 12/13/2019 |  |  | 01/31/2021 |  | 12/13/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1400754 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 35 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/15/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male patient had impella 2.5 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported hemolysis was suspected due to urine and bilirubin values. the patient ultimately developed renal failure due to hemolysis. continuous renal replacement therapy (crrt) was used to treat renal failure. relation between the device and hemolysis is unknown. per the physician the patient was also on percutaneous cardiopulmonary support. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37578 | 1220648-2019-00282 | 9468113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9468113&pc=OZD | 11/20/2019 | 12/13/2019 |  | 12/13/2019 |  |  | 11/20/2019 | 12/13/2019 |  |  | 06/30/2021 |  | 12/13/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1442255 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/20/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old asian female patient had impella cp optical pump inserted for acute myocardial infarction (ami) with shock. the patient had an estimated blood loss of 1,000ml from the groin where impella was inserted. as a result, the physician had to repair the artery around the insertion site and four (4) units of red blood cells (rbcs) was given. | [5 empty] | The impella cp optical pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37579 | 1220648-2019-00281 | 9468126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9468126&pc=OZD | 11/18/2019 | 02/04/2020 |  | 12/13/2019 |  |  | 11/18/2019 | 12/13/2019 |  |  | 08/31/2021 |  | 12/13/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1444342 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 40 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 11/18/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) hispanic/latino male patient had impella 5.0 pump inserted. the patient had a clot and the clot was removed from the axillary graft with a fogarty catheter and forceps during the surgery. | [5 empty] | A large clot was lodged in the graft when the introducer was removed requiring extraction with a fogarty arterial embolectomy catheter. act at that time was checked and noted to be 157sec. the cause of the clot may have been related to the patient's anticoagulation status from the start of the case but was unable to be determined since act at the start of the case was unknown and insufficient clinical data was provided. | The impella 5.0 was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37580 | 1220648-2019-00283 | 9469155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9469155&pc=OZD | 11/15/2019 | 12/13/2019 |  | 12/13/2019 |  |  | 11/15/2019 | 12/13/2019 |  |  | 05/31/2021 |  | 12/13/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1421773 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/15/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) female patient had impella 2.5 pump inserted in the right femoral for high risk percutaneous coronary interventions (hrpci). it was reported that during pump support the patient gradually became symptomatic. an mri confirmed the patient had developed cerebral infarction. the physician stated that there was calcification, thrombus, and plaque in the aortic arch due to high risk in triple vessel disease. the patient is currently being treated for cerebral infarction; however, specifics of treatment have not been provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37581 | 1035166-2019-00122 | 9453401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9453401&pc=DYB | 10/11/2019 | 12/11/2019 |  | 12/11/2019 |  |  | 11/18/2019 | 12/11/2019 |  |  | 10/02/2020 |  | 12/11/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | C1-13804 |  | 0052-3006 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 49 YR | [5 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Blood Loss | [41 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that during cardiomyopathy procedure introducer was placed through a graft on the patient's left axillary artery. while advancing the impella 5.0 catheter through the introducer, blood was seen leaking around the shaft of the catheter. the impella's outlet was beyond the distal tip of sheath when this was noticed. amount of blood loss is unknown. there was no additional surgery or blood products required to control the blood loss. the issue was resolved immediately by using a tourniquet & advancing the impella's repositioning sheath. procedure completed with same device. the physician believes there was a hole in the hemostatic valve of the peel away sheath that caused the blood loss. however, hole in hemostatic valve was never confirm. patient outcome is stable. product was discarded. no additional information is available. no serious injury reported. | [5 empty] | The device was used for treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. the investigation will focus on a review of product documentation. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. per qa procedure abiomed introducer sheath in-process and final inspections:vacuum leak test: perform pressure and vacuum leak test per procedure (latest revision). this procedure is performed by qa 100%. per instructions for use (ifu): when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
37582 | 1035166-2019-00117 | 9443664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9443664&pc=DYB | 11/05/2019 | 03/13/2020 |  | 12/10/2019 |  |  | 11/13/2019 | 12/10/2019 |  |  | 09/01/2022 |  | 12/10/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-17687 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 51 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Consequences Or Impact To Patient | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | As reported information, the bleeding (oozing) occurred around competitor's sheath when micro puncture technique was used. | It was reported the patient was presented to physician for a primary percutaneous coronary intervention procedure (ppci). during the ppci procedure it was reported the physician had difficulty with competitor's device. physician used a micro puncture technique so that one single access site could be used to place an impella device and use a competitor's sheath for a high risk percutaneous coronary intervention (hrpci). bleeding (oozing) occurred around competitor's sheath, there was no reported malfunction with oscor's device. after hrpci was completed both the competitor and oscor's sheaths were removed and manual pressure was held for the recommended 45 minutes and adequate hemostasis was achieved. the oozing did cause an estimated blood loss of around 500 cc's, maybe less, but no harm was done to the patient and the procedure was able to be completed. patient did not receive any blood products of any kind and was discharged home the following day. |  |  |  |  | The following sections were up-dated in follow-up 1: b4, b5, d10, g4, g7, h2, h3, h6 and h10. the device was used in treatment. as reported information, the bleeding (oozing) occurred around competitor's sheath when micro puncture technique was used. returned device analysis revealed the 14f abiomed introducer sheath is within manufacturing specifications. the device passed leak testing that was performed per applicable iso standard as well as manual leak testing. the inner diameter of the device was also within manufacturing specifications. no manufacturing defects were found. per procedure adelante s2s introducer sheath in-process and final inspection: vacuum leak test - perform pressure and vacuum leak test per procedure latest revision. this test is performed by qa on 100% of the product. measure dimensions: b) measure tip i.d. of sheath using appropriate pin gauges. visual inspection: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. this test is performed by qa 100%. the instruction for use (ifu) in forms the user: precautions: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. directions for use: slowly retract the guidewire and dilator, leaving the sheath in position. the hemostasis valve will reduce the loss of blood and the inadvertent aspiration of air through the sheath. introduce the catheter through the hemostasis valve/sheath and advance it into position. no further investigation is required, as there was no issue reported for the device performance or quality. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | The device was used for treatment. there was no product performance issue reported. the device was not returned for evaluation, therefore, the clinical observation could not be confirmed. the device history records were reviewed to confirm that the device passed all applicable in process and final inspections before shipping to the customer. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
37583 | 1035166-2019-00119 | 9444607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9444607&pc=DYB | 11/13/2019 | 03/05/2020 |  | 12/10/2019 |  |  | 11/22/2019 | 12/10/2019 |  |  | 08/01/2022 |  | 12/10/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-17732 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 70 YR | [5 empty] | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | No Consequences Or Impact To Patient | [12 empty] | Other | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that during high risk percutaneous coronary intervention (hrpci) procedure impella was not able to advance because kinked peel away sheath. kink was visualized under fluoroscopy. percutaneous axillary technique used for insertion and no additional device used during procedure. product issue resolved by implanting another impella femorally. there was no patient side effects or intervention reported. there was 2 hour delay reported in surgery as md decided to try another access site. peel away sheath exchanged by 14fr cook sheath because perclose failed. user had to hold manual pressure which delayed the procedure. no other additional information is available. no other serious injury reported. | [5 empty] | The following sections were updated in follow-up. the device used in treatment. one 14f 25 cm long abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. blood was found on and inside the sheath and sideport tubing. four kinks were observed on the sheath tube. upon evaluation of the returned sheath under 10x magnification, there were four kinks observed in the sheath tube at 8cm, 10cm, 16cm and 17.5cm. the sheath distal tip along its circumference edges and the hemostatic valve looked normal with its helical cuts intact. according to the event description summary, when they tried to advance the impella cp, it would not advance through the sheath. when id of the sheath measured was within manufacturing specifications. the impella device used in the procedure was not returned to oscor so a fit check could be performed. potential causes for the damage (deformation/kinks) on the sheath tube could be due to the forceful or improper insertion/orientation of an auxiliary device. this might have subjected the sheath to stress/resistance during the procedure. no manufacturing defects were found. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure, abiomed introducer sheath in-process and final inspections: measure dimensions: measure tip i.d. of sheath using appropriate pin gauges. visual inspection: this step is performed by qa for sample size of 100% inspection. with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. per instructions for use, never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress ,follow up report will be submitted if we find any further additional information.
37584 | 1220648-2019-00278 | 9446891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9446891&pc=OZD | 11/12/2019 | 12/10/2019 |  | 12/10/2019 |  |  | 11/12/2019 | 12/10/2019 |  |  | 05/31/2021 |  | 12/10/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1420671 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Dissection | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/12/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) female patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that during the impella insertion procedure a blood vessel dissected to the ascending aorta when the peel away was inserted. the area was party thrombosed, however, the impella was successfully inserted. no treatment was provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37585 | 1220648-2019-00279 | 9446905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9446905&pc=OZD | 11/13/2019 | 12/10/2019 |  | 12/10/2019 |  |  | 11/13/2019 | 12/10/2019 |  |  | 05/31/2021 |  | 12/10/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1424942 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | [5 empty] | 70 YR | [5 empty] | Malfunction | No | Yes | No Apparent Adverse Event | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission |  |  |  |  | 11/13/2019 | * | Invalid/Incomplete |  |  | The impella cp was placed for hemodynamic support as patient¿s ejection fraction (ef) was 15%. there was an observation made of suction and pump flows decreased. the team attempted to troubleshoot the pump by adjusting the pump level and checking that the pump was in correct position within the left ventricle. the pump was weaned and removed after two (2) days of support. when the pump was returned for a complaint investigation there was an observation made of broken impeller blades. the pump damage was not observed by the customer, rather in bench top testing. | [5 empty] | The impella cp pump was returned for evaluation. there was damage to the impeller due to interaction with a malpositioned tavr.
37586 | 1035166-2019-00116 | 9437674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9437674&pc=DYB | 10/28/2019 | 03/13/2020 |  | 12/09/2019 |  |  | 11/12/2019 | 12/09/2019 |  |  | 08/01/2022 |  | 12/09/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-17244 |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 64 YR | [5 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Consequences Or Impact To Patient | [41 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported the patient was presented to physician in stable condition for elective procedure of primary percutaneous coronary intervention procedure (ppci). the ppci procedure was completed without any complication, the patient was discharged the following day. next day discharge is normal, as the patient was inpatient prior to the percutaneous coronary intervention (pci). during the procedure, when the dilator was removed from the sheath, prior to placing the impella the physician experienced oozing with the 13cm sheath at the valve. the procedure was delayed for 5 minutes, the issue was resolved by swapping out for another 14f oscor sheath. minimal blood was lost, no blood product or surgical technique was required to control the oozing. requesting additional information. | [5 empty] | The following sections were updated in follow-up. a3,b4,g4,g7,h2,h6,h10. the device was used for treatment.location of device is unknown and the device was not returned for analysis, therefore, the clinical observation could not be confirmed.the investigation will focus on a review of product documentation. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections per procedure adelante s2s introducer sheath in-process and final inspection, perform pressure and vacuum leak test per procedure,latest revision. this test is performed by qa on 100% of the product. per ifu.precuations: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. directions for use: slowly retract the guidewire and dilator, leaving the sheath in position. the hemostasis valve will reduce the loss of blood and the inadvertent aspiration of air through the sheath. introduce the catheter through the hemostasis valve/sheath and advance it into position. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete.
37587 | 1220648-2019-00276 | 9437795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9437795&pc=OZD | 11/14/2019 | 12/09/2019 |  | 12/09/2019 |  |  | 11/14/2019 | 12/09/2019 |  |  | 10/31/2020 |  | 12/09/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1376920 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 30 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/14/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese male patient had impella 2.5 pump inserted in the left femoral for myocarditis. it was reported that during pump support the patient experienced bleeding at the groin site. patient was administered 24 units of concentrated red blood cells and 10 units of platelets, along with bandage changed. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37588 | 1220648-2019-00277 | 9437819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9437819&pc=OZD | 11/13/2019 | 12/09/2019 |  | 12/09/2019 |  |  | 11/13/2019 | 12/09/2019 |  |  | 04/30/2021 |  | 12/09/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1415058 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/13/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese male patient had impella cp pump inserted in the right femoral for pre-op coronary artery bypass surgery. it was reported the patient experienced bleeding at the impella access site. the patient received two units of red blood cells. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37589 | 1220648-2019-00273 | 9430426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9430426&pc=OZD | 11/12/2019 | 12/06/2019 |  | 12/06/2019 |  |  | 11/12/2019 | 12/06/2019 |  |  | 08/31/2021 |  | 12/06/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1433531 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/12/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient was found to have lower limb ischemia requiring impella to be removed. a thrombectomy was performed using the fogarty embolectomy catheter. thrombus was found and removed. blood flow was improved after thrombus removal. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37590 | 1220648-2019-00272 | 9431694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9431694&pc=OZD | 11/11/2019 | 12/06/2019 |  | 12/06/2019 |  |  | 11/11/2019 | 12/06/2019 |  |  | 02/28/2021 |  | 12/06/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1408367 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Dissection | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/11/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese female patient had impella 2.5 pump inserted in the left axillary for post-cardiotomy cardiogenic shock (pccs). it was reported that when impella was being inserted there was a dissociation from the distal arch to the ascending aorta. additionally, the patient experienced bleeding at the insertion site for which hemostasis was applied and blood products were administered, exact amount is unknown. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37591 | 1220648-2019-00274 | 9432149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9432149&pc=OZD | 11/15/2019 | 02/19/2020 |  | 12/06/2019 |  |  | 11/15/2019 | 12/06/2019 |  |  | 03/31/2021 |  | 12/06/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1416670 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 11/15/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with an impella cp optical. during support, the patient developed hemolysis as a result of the pump's catheter being lodged within the patient's papillary muscle. there was no indication of a plasma-free hemoglobin draw to confirm hemolysis. as treatment, the device was prematurely removed, as optimal positioning could not be obtained to resolve the hemolysis. | [5 empty] | The complainant returned the device and an thorough investigation was completed. a simulated use test was performed, however, we could not reproduce the reported failure mode (hemolysis). evidence suggests improper positioning was to blame, however, we cannot definitively confirm this without further objective evidence. a device history review was performed and found no manufacturing issues or reworks associated with this pump lot. there are no other complaints related to this failure mode associated with pumps in this lot. | The impella cp optical was not returned by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37592 | 1220648-2019-00275 | 9432559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9432559&pc=OZD | 11/19/2019 | 12/06/2019 |  | 12/06/2019 |  |  | 11/19/2019 | 12/06/2019 |  |  | 06/30/2021 |  | 12/06/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1442255 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/19/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient presenting, post coronary artery bypass surgery, with acute myocardial infarction and cardiogenic shock. patient was escalated from an intra-aortic balloon pump to an impella cp optical for hemodynamic support. the following day, the patient endured diminished pulses in foot of the leg that which the impella was inserted. the device was rapidly weaned and removed on the 3rd day of support, as the patient's condition did not improve. distal pulses were not present, therefore, a below knee amputation was required. the patient remained in stable condition, thereafter. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37593 | 1220648-2019-00270 | 9417653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9417653&pc=OZD | 11/08/2019 | 12/05/2019 |  | 12/05/2019 |  |  | 11/08/2019 | 12/05/2019 |  |  | 09/30/2019 |  | 12/05/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1313154 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/08/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 5.0 pump inserted in the right axillary for cardiomyopathy. it was reported the patient experienced bleeding at the impella access site. the patient was administered 4 units of red blood cells. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37594 | 1035166-2019-00115 | 9417740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9417740&pc=DYB | 10/28/2019 | 02/06/2020 |  | 12/05/2019 |  |  | 11/12/2019 | 12/05/2019 |  |  | 05/01/2022 |  | 12/05/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3025-JP |  |  | OSCOR INC. |  |  | US |  | C1-16769 |  | 0052-3025-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN |  |  |  |  | 85 YR | [5 empty] | Injury | Yes | No | Break | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that during insertion of impella cp,14 fr peel away introducer peeled away before it was completely removed from the blood vessel. user did not reported any product malfunction. access site was leaking blood and leakage did not stop. amount of blood is unknown and the patient received 18 units of red blood cells. patient outcome is stable. impella was removed on 10/30. no other serious injury reported. no additional information is available. | [5 empty] | The device was used in treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. the investigation will focus on a review of product documentation. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections per procedure abiomed introducer sheath in-process and final inspection, the devices in this lot were inspected as follows: visual inspection:verify hub to be smooth, no cracks, short shots, sinking or unfilled areas, and check for fm. verify hub shape as per drawing. verify french size is properly embossed on hub. check for excessive flash (maximum allowable limit is 0.005") and tube pinch. using 10x microscope, look down into hub for irregularities. verify no splitting at break lines and a smooth transition between sheath and hub on inside of hub. this is performed by qa on the first 5 consecutive properly tipped parts. with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm. this is performed by qa on 100% of the sheaths. per ifu : never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress ,follow up report will be submitted if we find any further additional information.
37595 | 1220648-2019-00271 | 9419312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9419312&pc=OZD | 11/06/2019 | 12/05/2019 |  | 12/05/2019 |  |  | 11/06/2019 | 12/05/2019 |  |  | 07/31/2021 |  | 12/05/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1429067 |  | 0048-0034-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/06/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella cp pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the physician suspected infection. as a result, the impella was explanted. impella¿s pig tail testing confirmed a very small amount of staphylococcus epidermidis. patient was administered antibiotics. | [5 empty] | The impella cp pump was returned. no biomaterial was found in the pump. the cause of bacteria on the pump was unable to be determined. data logs were reviewed and did not indicate any issue that could be related to infection. sterilization records were reviewed, and the pump passed all incoming inspection. impella cp was properly sterilized, the bacteria was likely not present on the pump prior to use. dhr was reviewed and found no manufacturing related issues with this lot. there are no other complaints related this failure mode in this lot.
37596 | 1220648-2019-00269 | 9411080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9411080&pc=OZD | 11/05/2019 | 01/13/2020 |  | 12/04/2019 |  |  | 11/05/2019 | 12/04/2019 |  |  | 06/30/2020 |  | 12/04/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP SET |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1353056 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | Sepsis | [11 empty] | Death | [28 empty] | Initial submission | Followup |  |  |  | 11/05/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male patient had impella 5.0 pump inserted in the right axillary for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported the patient who had a do-not-resuscitate (dnr) order expired. the cause of death was due to sepsis. additionally, the physician stated the long use of impella could have indirectly caused sepsis. the impella worked as intended throughout the entire procedure. | [5 empty] | The impella 5.0 pump was returned. visual inspection found biomaterial in the if cage, cannula and the pump housing. no other issues were found on the rest of the pump. the cause of sepsis was unable to be determined. data logs were reviewed and showed that the pump operated within normal range since activation. sterilization records were reviewed, and the pump passed all incoming inspection. impella 5.0 was properly sterilized. dhr was reviewed and found no manufacturing related issues with this lot. there are no other complaints related this failure mode in this lot. | The impella cp was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37597 | 1220648-2019-00267 | 9401809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9401809&pc=OZD | 11/07/2019 | 12/03/2019 |  | 12/03/2019 |  |  | 11/07/2019 | 12/03/2019 |  |  | 08/31/2021 |  | 12/03/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1439072 |  | 0048-0032 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 86 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 11/07/2019 | * | Invalid/Incomplete |  |  | The hospital had a patient admitted in cardiogenic shock and had an impella cp placed via the left femoral artery for hemodynamic support. after almost 18 hours of support the team chose to explant the impella due to the signs of limb ischemia. the limb had become mottled and lost pulses. the team chose to place an iabp via the contralateral femoral. | [5 empty] | The complainant has discarded the pump product. there are questions and requests for more clinical details that are open at this time. the investigation is ongoing. should more information be shared that leads to a root cause of the ischemia, a final report will be filed.
37598 | 1220648-2019-00268 | 9401811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9401811&pc=OZD | 11/08/2019 | 12/03/2019 |  | 12/03/2019 |  |  | 11/08/2019 | 12/03/2019 |  |  | 04/30/2021 |  | 12/03/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  |  |  | 1432334 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 89 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 11/08/2019 | * | Invalid/Incomplete |  |  | The (b)(6) female was admitted to the er in atrial fibrillation and suffering from abdominal pain. she had an echo done which revealed a greatly depressed ejection fraction (ef of 10%) and was sent to the cardiac catheterization lab for diagnosis. she had an impella cp placed and a diagnosis of cardiomyopathy was made. the impella cp had to be explanted after just under 6 hours due to suspected impella caused hemolysis. | [5 empty] | The investigation is on-going at this time. upon completion of pump hemolysis testing a final report will be submitted.
37599 | 1220648-2019-00263 | 9394290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9394290&pc=OZD | 11/04/2019 | 01/13/2020 |  | 12/02/2019 |  |  | 11/04/2019 | 12/02/2019 |  |  | 03/31/2021 |  | 12/02/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1426738 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 11/04/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old asian male patient had impella cp optical pump inserted. the patient had suspected hemolysis; the physician tried to reposition but felt more resistance then normal so he decided to remove the pump. the patient¿s hemolysis resolved with pump removal and patient¿s kidney failed so dialysis was needed. | [5 empty] | The pump was found to be about 5cm deep and the inlet was not in the apex of the heart. the patient was in very poor condition and had rv failure that led to filling issues. the pump position or patient condition could have contributed to the hemolysis which was not diagnosed with plasma free hemoglobin. no damages was found on the returned pump, no scratches or delaminating epo-tek after washing, and no biomaterial was found in the pump. the pump was cleaned with tergazyme solution prior to hemolysis testing. the pump passed hemolysis testing. there are multiple clinical factors that could've contributed to hemolysis, the pump passed hemolysis testing when investigated at abiomed danvers. the root cause of hemolysis was unable to be definitively determined due to lack of additional clinical details and the data logs. | The impella cp optical was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37600 | 1220648-2019-00264 | 9394807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9394807&pc=OZD | 11/05/2019 | 02/04/2020 |  | 12/02/2019 |  |  | 11/05/2019 | 12/02/2019 |  |  | 05/31/2021 |  | 12/02/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1440207 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 11/05/2019 | * | Invalid/Incomplete |  |  | The complainant reported a 59 years old white male patient had impella cp optical pump inserted independently in the left femoral. the patient had hemolysis, lab results still showed hemolysis, and urine was still bloody, repositioning was unsuccessful even with fluoro and echocardiogram (echo) guidance so the decision was made to explant the device. | [5 empty] | The data logs were returned and showed multiple positioning alarms as well as suction alarms during the case. the pump was returned for evaluation and a visual inspection showed traces of biomaterial on the tip of the impellers and no visual deformities were found. the pump passed hemolysis testing. the cause of the hemolysis is unable to be determined but could be due to the improper positioning that was both reported clinically and observed via the returned data logs. | The impella cp optical was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37601 | 1220648-2019-00265 | 9394814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9394814&pc=OZD | 11/04/2019 | 01/02/2020 |  | 12/02/2019 |  |  | 11/04/2019 | 12/02/2019 |  |  | 06/30/2021 |  | 12/02/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED INC. |  |  | US |  | 1441368 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hematoma | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 11/04/2019 | * | Invalid/Incomplete |  |  | A cardiac patient was admitted and had a coronary intervention. while in the icu his condition deteriorated and so they placed an impella cp in a return visit to the cardiac catheterization lab. the coronary stents were reopened and the patient was sent back to the icu on the impella for hemodynamic support. while in the icu the patient was observed to have low impella flows due to incorrect positioning, as observed by echo. the patient was again taken to the catheterization lab. while transferring to the exam table, a hematoma was observed at the impella access site. the hematoma was treated with 2 units of blood and albumin infusion. in addition, the impella was explanted and replaced by a new pump. | [5 empty] | Since the initial medwatch report was filed, the investigation has concluded. the pump and introducer sheath were returned for analysis. scratches were found on the introducer and the pump's repositioning unit was found to have scratches and a kink. these damages could have been caused by excessive force or improper handling and technique. these assumptions and the root cause for the hematoma can not be determined however. the medical facility did not release all clinical details, and as such, a root cause can not be definitively determined. the failure mode will be monitored and trended. | The product has been returned for analysis. the analysis is on-going at this time. upon completion of the investigation a final medwatch will be filed.
37602 | 1220648-2019-00266 | 9394847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9394847&pc=OZD | 11/08/2019 | 12/02/2019 |  | 12/02/2019 |  |  | 11/08/2019 | 12/02/2019 |  |  | 12/31/2019 |  | 12/02/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.5 PUP |  |  | OZD |  |  | IMPELLA 5.5 PUMP 550 SET, EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1422474 |  | 0550-0001 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Insufficiency, Valvular | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/08/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male patient had impella 5.5 pump inserted in the right axillary with difficulty due to vascular anatomy and poor vision on fluoroscopy in the operating room (or). the physician decided to stop repositioning attempts in the or and sent the patient to the intensive care unit (icu) to optimize position under transesophageal echocardiography (tee). the team failed to reposition the pump; under echocardiogram (echo) inlet cage was visually entangled with the papillary muscle. as a result, venoarterial extracorporeal membrane oxygenation (va-ecmo) was added, the next day a left ventricular assist device (lvad) was implanted, and both the impella 5.5 and va-ecmo was explanted. | [5 empty] | Device discarded by customer. the impella 5.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37603 | 1220648-2019-00256 | 9375050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9375050&pc=OZD | 10/26/2019 | 11/26/2019 |  | 11/26/2019 |  |  | 10/31/2019 | 11/26/2019 |  |  | 02/29/2020 |  | 11/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1329826 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 31 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Rupture | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/31/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient had impella 5.0 pump inserted in the operating room (or) with no issue. physician noted that during explant there was a papillary muscle rupture. the patient was bridge to durable ventricular assist device (vad) and is now awaiting open heart transplant (oht). the physician stated impella may not have been the sole cause of the rupture but did provide mechanism for the rupture. | [5 empty] | Device discarded by customer. the impella 5.0 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37604 | 1220648-2019-00257 | 9375064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9375064&pc=OZD | 10/30/2019 | 01/17/2020 |  | 11/26/2019 |  |  | 10/30/2019 | 11/26/2019 |  |  | 04/30/2021 |  | 11/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1420727 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 57 YR | [5 empty] | Malfunction | No | Yes | No Apparent Adverse Event | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 10/30/2019 | * | Invalid/Incomplete |  |  | A patient admitted in cardiogenic shock had an impella cp placed for hemodynamic support. while placing the impella pump, and removal of the .018 guidewire, there was a malfunction in which the wire was unable to be removed. the pump became kinked. to remove the kink, the team had to snare the pump's pigtail from the contralateral leg access. there was no lasting harm and the patient was successfully supported by an impella from the contralateral leg access. | [5 empty] | Since the initial report was filed the investigation into the kinking of the impella, and it's associated snaring for removal has completed. the product was returned for investigation. the impella pump and wire showed extensive kinking. the team additionally returned imaging of the native anatomy and location of the pump and wire. the cause of the kinking and associated removal by snare was determined to be the user's wireless insertion attempts. the wireless insertion was attempted after the wire had backed out of the pump during initial insertion attempts. the failure mode will be monitored and trended. | The investigation into the kinked pump, and need for a snare to remove the product, is currently ongoing. upon completion of the investigation, a final report will be submitted.
37605 | 1220648-2019-00258 | 9375077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9375077&pc=OZD | 10/29/2019 | 11/26/2019 |  | 11/26/2019 |  |  | 10/29/2019 | 11/26/2019 |  |  | 11/30/2020 |  | 11/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1385057 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Disseminated Intravascular Coagulation (DIC) | Pseudoaneurysm | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/29/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male patient had impella 2.5 pump inserted in the left femoral for cardiomyopathy. the patient endured pseudoaneurysm at the impella removal site and disseminated intravascular coagulation (dic) during pump support. the patient received non cardiac surgery, 34 units of fresh frozen plasma, 42 units of irradiated red blood cell, and 40 units of irradiated concentrated platelets. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37606 | 1220648-2019-00259 | 9375124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9375124&pc=OZD | 11/04/2019 | 11/26/2019 |  | 11/26/2019 |  |  | 11/04/2019 | 11/26/2019 |  |  | 05/31/2021 |  | 11/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1440464 |  | 005042 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 85 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | Perforation | [11 empty] | Death | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 11/04/2019 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old white male patient presenting for an elective high risk percutaneous coronary intervention for hemodynamic support with an impella 2.5 device. during insertion, the physician advanced the catheter into place and initiated the device. after starting the pump, the device's motor current appeared flat and troubleshooting suggested the device was actually inserted into the descending aorta, rather than the left ventricle. the patient's blood pressure began to fall and dye insertion revealed a large perforation in the patient's aortic arch at the bend of the ascending aorta. prior to insertion of the device, a large "chunk" of calcium was noted at the site of the perforation. intervention in the form of two stent placements, blood products, and intubation were performed. the patient ultimately endured severe bradycardia, pulseless electrical activity, and expired. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37607 | 1220648-2019-00260 | 9375145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9375145&pc=OZD | 11/01/2019 | 11/26/2019 |  | 11/26/2019 |  |  | 11/01/2019 | 11/26/2019 |  |  | 03/31/2021 |  | 11/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1418897 |  | 005066-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 11/01/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male patient had impella 5.0 pump inserted in the right axillary for heart failure. during pump support the patient experienced a hemorrhage at the right subclavian artery placement site. the patient was administered 20 units of red blood cells and 6 units of fresh frozen plasma. additionally, hemostasis was applied, and the hematoma was surgically removed. | [5 empty] | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37608 | 1220648-2019-00261 | 9375149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9375149&pc=OZD | 10/31/2019 | 11/26/2019 |  | 11/26/2019 |  |  | 10/31/2019 | 11/26/2019 |  |  | 08/31/2021 |  | 11/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1434848 |  | 0048-0034-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 85 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/31/2019 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) japanese female patient had impella cp pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that while attempting to peel-away the introducer completely from the artery a large amount of blood was observed from the patients access site. the angle and depth of the sheath were adjusted; however, the bleeding continued. the patient received 18 units of red blood cells and the impella was explanted. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37609 | 1220648-2019-00262 | 9375188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9375188&pc=OZD | 10/31/2019 | 11/26/2019 |  | 11/26/2019 |  |  | 10/31/2019 | 11/26/2019 |  |  | 11/30/2020 |  | 11/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1385057 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/31/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male patient had impella 2.5 pump inserted in the left femoral for cardiomyopathy. during pump support hemolysis was suspected. adjusting impella flow rate did not resolve the issue and it was decided to explant the device. the patient required additional mechanical support. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37610 | 1035166-2019-00109 | 9371688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9371688&pc=DYB | 11/01/2019 | 11/25/2019 |  | 11/25/2019 |  |  | 11/04/2019 | 11/25/2019 |  |  | 12/01/2021 |  | 11/25/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC |  |  | US |  | C1-15477 |  | 0052-3006 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 69 YR | [5 empty] | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that patient was presented to physician for cardiomyopathy. during the placement of impella 5. 0, excessive leaking noted on the 23 fr sheath around the catheter shaft after advancing the pump/motor into the graft. hemostatic valve was leaking while advancing the product through valve. patient loss 250cc blood and required 1 unit of cells infused. md held the finger and gauze over the leaky area. procedure completed successfully. product was discarded. no other additional information is available. no other serious injury reported. | [5 empty] | The device was used for treatment. the device was discarded and not returned for analysis, therefore, the clinical observation could not be confirmed. the investigation will focus on a review of product documentation. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure adelante s2s introducer sheath in-process and final inspection: vacuum leak test - perform pressure and vacuum leak test per procedure ( latest revision) .this test is performed by qa on 100% of the product. per ifu: precuations: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane directions for use: slowly retract the guidewire and dilator, leaving the sheath in position. the hemostasis valve will reduce the loss of blood and the inadvertent aspiration of air through the sheath. introduce the catheter through the hemostasis valve/sheath and advance it into position. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing nonconformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
37611 | 1035166-2019-00108 | 9371920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9371920&pc=DYB | 10/11/2019 | 11/25/2019 |  | 11/25/2019 |  |  | 10/31/2019 | 11/25/2019 |  |  | 03/01/2021 |  | 11/25/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC |  |  | US |  | C1-14074 |  | 0052-3021 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Other | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that patient was presented to physician for cardiogenic shock. during access at the right femoral vein for an insertion of the impella rp patient loss the blood while physician removing the impella rp to re-position it. since the patient still needed the rp a 23fr peel away was taken from the other rp kit at the account. md was then able to advance the new sheath followed by the rp catheter. procedure completed successfully with another oscor 23fr sheath. amount of blood loss is unknown and there was no transfusion required. patient died on 10/11/2019. cause of the death is unknown. the product is noted to be discarded no other additional information is available. | [5 empty] | The device was used for treatment. the device was discarded and not returned for analysis, therefore, the clinical observation could not be confirmed. the investigation will focus on a review of product documentation. there was no specific performance related failure reported by the user. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. per ifu when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. infusion through the sideport can be done only after all air is removed from the unit. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the side port only. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
37612 | 1220648-2019-00253 | 9362130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9362130&pc=OZD | 10/25/2019 | 11/22/2019 |  | 11/22/2019 |  |  | 10/25/2019 | 11/22/2019 |  |  | 07/31/2021 |  | 11/22/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1430942 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/25/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 2.5 inserted in the left femoral for unloading purposes. the physician reported that advanced aortic regurgitation occurred after impella insertion. the patient was treated with direct cardioversion and drugs; however, a more aggressive treatment was not desired by the patient's family. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37613 | 1220648-2019-00254 | 9362158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9362158&pc=OZD | 10/29/2019 | 11/22/2019 |  | 11/22/2019 |  |  | 10/29/2019 | 11/22/2019 |  |  | 07/31/2021 |  | 11/22/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1430942 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/29/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male patient had impella 2.5 pump inserted in the left femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the physician reported the patient experienced retroperitoneal hemorrhage at the end of the deep femoral artery. catheter embolization was performed to treat the bleed. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37614 | 1035166-2019-00106 | 9364263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9364263&pc=DYB | 11/01/2019 | 04/17/2020 |  | 11/22/2019 |  |  | 11/30/2019 | 11/22/2019 |  |  | 08/01/2022 |  | 11/22/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-17244 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 72 YR | [5 empty] | Injury | Yes | Yes | Deformation Due to Compressive Stress | Adverse Event Without Identified Device or Use Problem | [5 empty] | Aortic Valve Stenosis | Peripheral Vascular Disease | No Known Impact Or Consequence To Patient | [10 empty] | Hospitalization | [28 empty] | Initial submission | Followup | Followup |  |  |  | * | Invalid/Incomplete |  |  | Addition information received that the patient had pvd (peripheral disease) and also had as (aortic stenosis. the delay and inability to place the pump was due to the heart valve allowing the pump to traverse it. this is not related to the oscor sheath. | It was reported a (b)(6) male, was presented to physician with coronary artery disease and high risk coronary angioplasty after being turned down by cardiothoracic surgery. during procedure, a 14f x 25cm introducer kinked during access to the right femoral artery for placement of the impella cp. device kinked in the middle of right femoral artery, did not try smaller sheath as vessel was calcified and tortuous. stick done under ultrasound, normal angle of approach. several dilations completed over abiomed 0.35. no damage to vessel. the team attempted for 4 hours to place the pump and aborted the case. it was report devices are to be returned for evaluation. |  |  |  |  | The following sections were updated in follow-up 1: b4, d10, g4, g7, h2, h3, h6 and h10: the device was used in treatment. one 14f 25 cm long abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. blood was found on and inside the sheath and sideport tubing. a kink was observed on the sheath tube. upon evaluation of the returned sheath, it was found that there was a kink in the sheath tube at approximately 12.5cm from the distal tip at 0.68x magnification. the rest of the sheath including the distal tip and the hemostatic valve looked normal at 10x mag. according to the event description summary, the sheath kinked in middle of right femoral artery (rfa), when they tried to advance the impella cp. it was determined that the patient's vessel was calcified and tortuous. per drawing the id of the sheath should be .198" ± .002". the sheath was measured with a pin gauge and was within manufacturing specifications at .197". the impella device used in the procedure was not returned to oscor so a fit check could not be performed. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. returned device analysis revealed sheath (including the inner diameter) is within manufacturing specifications. potential contributing factors to the kink could be if the sheath was advanced through an area of resistance or if the sheath was subjected to unusual stresses. according to the event description summary, the patient's vessel was calcified and tortuous which likely caused the sheath to kink during the procedure. no manufacturing defects were found. per procedure abiomed introducer sheath in-process and final inspections: 10.2.2 measure dimensions: b) measure tip i.d. of sheath using appropriate pin gauges. 12.2 visual inspection: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. this is performed by qa 100%. the instructions for use (ifu) informs the user: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | The following sections were updated in follow-up 2: b1, b4, b5, g4, g7, h2, and h10: h2. correction: b1 - no product problem h6 - device codes h2: additional information: b5: the patient had pvd (peripheral disease) and also had as (aortic stenosis). the delay and inability to place the pump was due to the heart valve allowing the pump to traverse it. this is not related to the oscor sheath. h6: patient codes oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete.
37615 | 1220648-2019-00255 | 9364767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9364767&pc=OZD | 10/28/2019 | 01/17/2020 |  | 11/22/2019 |  |  | 10/28/2019 | 11/22/2019 |  |  | 02/21/2021 |  | 11/22/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP SET |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1408138 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 10/28/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). hematuria was seen in the patient¿s urinary catheter and hemolysis was suspected. the physician stated that due to the patient status and the severity of the suspected hemolysis the impella was explanted. | [5 empty] | The impella 2.5 was returned and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion. | The impella 2.5 was returned. there was no damage seen on the returned pump and it passed hemolysis testing. the root cause of the suspected hemolysis was unable to be determined as the issue was not able to be reproduced. blood lab results were also not provided to confirm the hemolysis. dhr was reviewed and found no manufacturing related issues with this lot. there are no other complaints related this failure mode in this lot.
37616 | 1035166-2019-00105 | 9353321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9353321&pc=DYB | 10/14/2019 | 11/21/2019 |  | 11/21/2019 |  |  | 10/24/2019 | 11/21/2019 |  |  | 07/01/2022 |  | 11/21/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | C1-17210 |  | 0052-3006 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 65 YR | [5 empty] | Injury | Yes | No | Fluid/Blood Leak | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | After attempting to wean off of coronary artery bypass procedure patient arrested and endured acute respiratory failure. during post cardiotomy cardiogenic shock, 23fr introducers valve damage after failed insertion attempts due to patient's anatomical conditions and unable to pass through the artery to the aortic valve. 23fr could no longer seal to the graft and resulted in oozing. 23fr was removed as bleeding would have occurred if the soft-jaw clamp was removed from the graft. the physician believed the 14fr x 25cm sheath would provide conduit beyond the graft/into the artery and might make it easier to pass the smaller cp pump. the physician successfully placed the 14fr and cp, thereafter, and bleeding subsided. the introducer was removed and replaced with a 14 fr x 25cm and an impella cp was delivered successfully. patient was successfully inserted with impella cp via 14 fr introducer. care was ultimately withdrawn on 2019-10-15 and patient expired while on support. there was no causal relationship with either the impella or oscor device. as per additional information, .035 wire was placed before the sheath. therefore the sheath was actually back loaded over the wire in retrograde fashion. user able to pass the 5.0 impella through the sheath but unable to pass it into the distal vessel. after taking the impella out, the valve did not seal. decision was made to peel away and remove the 23 fr sheath and utilize the 14 fr sheath for cp placement. amount of blood loss is unknown. blood loss was minimal. blood loss mitigated by physician thumb over valve and soft-jaw clamp at graft anastomosis. valve was no longer hemostatic. the sheath did not crack. valve sustained the tear. no other additional information is available. | [5 empty] | The device was used in treatment. the device was not returned for analysis. the investigation will focus on a review of product documentation. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure a biomed introducer sheath in-process and final inspections : visual inspection: verify hub to be smooth, no cracks, short shots, sinking or unfilled areas, and check for fm using 10x microscope, look down into hub for irregularities. verify no splitting at break lines and a smooth transition between sheath and hub on inside of hub. vacuum leak test perform pressure and vacuum leak test per procedure. leak test perform 1005 of the sheath. per instructions for use (ifu) ab-17-001z-x: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. no further follow-up required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
37617 | 1035166-2019-00107 | 9355647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9355647&pc=DYB | 10/19/2019 | 04/03/2020 |  | 11/21/2019 |  |  | 10/23/2019 | 11/21/2019 |  |  | 04/01/2021 |  | 11/21/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-14191 |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 69 YR | [5 empty] | Malfunction | No | Yes | Device Dislodged or Dislocated | [6 empty] | Blood Loss | [41 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that patient was admitted in cardiogenic shock and having acute myocardial infarction. during insertion at femoral artery while advancing 0.035" wire the sheath valve popped out. as per additional information received on, october 23,201 "the valve popped out of the orange retaining ring that holds it in place. the angiographic pigtail catheter that was backloaded on an 0.035" guide wire and was being advanced through the oscor sheath valve. while pushing it through the valve into the body, the valve popped out of its position in the sheath. we had no hemostasis with back bleeding from a straight 14f cannula. the valve failed completely. it was detached from the sheath and up on the pigtail catheter away from the sheath itself about 2-3cm. we removed the pigtail and wire with the valve still on it and placed the dilator in the sheath lumen to control hemostasis. the orange ring was cracked but not at the notch where it breaks when we peel away. instead, the orange retaining ring failure was over the white sheath wing on the same side as the side arm. patient loss 75c of blood, no blood products given. no intervention required. no other additional information is available. | [5 empty] | The following sections were updated in follow-up. b4,g4,g7,h2,h6,h10 the device was used for treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. the investigation will focus on a review of product documentation. per procedure adelante s2s introducer sheath in-process and final inspections: use a 10x microscope. along the entire length of the hub, verify that there is no break in the hub's molded plastic where the extrusion is visible. vacuum leak test - perform pressure and vacuum leak test per procedure latest revision. this testing performed 100% by manufacturing or quality assurance. per ifu:when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress ,follow up report will be submitted if we find any further additional information.
37618 | 1220648-2019-00252 | 9356114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9356114&pc=OZD | 10/29/2019 | 11/21/2019 |  | 11/21/2019 |  |  | 10/29/2019 | 11/21/2019 |  |  | 11/30/2020 |  | 11/21/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1385057 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/29/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that the patient developed mitral regurgitation (mr) during impella pump support. it is believed that mr was caused by impella and was removed. there were no impella positional issues reported. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37619 | 1035166-2019-00104 | 9349071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9349071&pc=DYB | 10/11/2019 | 02/13/2020 |  | 11/20/2019 |  |  | 10/23/2019 | 11/20/2019 |  |  | 04/01/2022 |  | 11/20/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-16469 |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Consequences Or Impact To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | As per additional information, there was no damage to introducer reported. the interventional cardiologist discovered the perforation prior to inserting the cp into the 14 fr sheath. patient was stable from the previous day, due to the urgent necessity of treating the patient the decision was made to attempt the ppci again the following day. introducer is not available as it was discarded by the customer. | We have been notified of a complaint involving an oscor 14fr x 13cm introducer (material #0052-3025, lot # c1-16469). the physician accessed the patient's right common femoral artery without issue and proceeded to dilate the insertion site with a 8fr, 10fr, 12fr dilators, respectively. the introducer was placed without any reported issues, however, the physician noticed a perforation in the artery. nothing stood out on introducer sheath that indicated there was any damage with the introducer that caused or contributed to this event. the procedure was aborted and successfully completed the next day with a new introducer (14fr x 13 cm oscor) and new impella. no intervention/surgical repair due to not being able to access contralaterally because of previous hip replacement wound healing. the perforation was just proximal to the insertion sight, manual compression applied until hemostasis was achieved. there was no adverse impact to patient. additional information has been requested. |  |  |  |  | The device was used for treatment. there was no product performance issue reported. the device was not returned for evaluation, therefore, the clinical observation could not be confirmed. the device history records were reviewed to confirm that the device passed all applicable in process and final inspections before shipping to the customer. based on the investigation, a capa is not required as there was no issue reported for the device performance or quality. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete.
37620 | 1220648-2019-00251 | 9349686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9349686&pc=OZD | 10/19/2019 | 11/20/2019 |  | 11/20/2019 |  |  | 10/24/2019 | 11/20/2019 |  |  |  |  | 11/20/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US | [7 empty] | UNKNOWN |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemostasis | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/24/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient had impella cp pump inserted for acute myocardial infarction (ami) with shock-impella pre-(pci) percutaneous coronary intervention. heparin was not added to the purge for 2-3 hours post implant due to a great deal of groin bleeding at site. within 24 hours a pfhgb was drawn which resulted in the 300s, the pump was then explanted due to hemolysis. | [5 empty] | Device discarded by customer. manufacturer name, city and state, serial number, lot number, and unique identifier (udi) # , and manufacturing site name and address could not be provided as product was discarded before abiomed employee was made aware of the incident and could gather the information. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37621 | 1220648-2019-00249 | 9340748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9340748&pc=OZD | 10/24/2019 | 11/19/2019 |  | 11/19/2019 |  |  | 10/24/2019 | 11/19/2019 |  |  | 05/31/2021 |  | 11/19/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1420671 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/24/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 2.5 inserted in the right femoral for cardiogenic shock from frequent ventricular tachycardia. the patient developed lower limb ischemia for which a sheath femorodemoral (ff) bypass was placed for both blood removal and delivery. subsequently, the patient developed a hematoma. manual pressure was placed, and two units of red blood cells were administered. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37622 | 1220648-2019-00250 | 9343115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9343115&pc=OZD | 10/25/2019 | 12/20/2019 |  | 11/19/2019 |  |  | 10/25/2019 | 11/19/2019 |  |  | 03/31/2021 |  | 11/19/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART ASSIST BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED INC. |  |  | US |  | 1417939 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 10/25/2019 | * | Invalid/Incomplete |  |  | The complainant had a patient admitted for a high risk coronary angioplasty. the patient had a history of coronary artery disease with a previous angioplasty and stenting. in addition he was suffering from the comorbidities of cancer, leukemia, kidney disease and diabetes. these comorbidities had caused the cardiothoracic surgeons to refuse bypass surgery. the patient had the angioplasty done and his condition deteriorated. this deterioration caused the team to remove the iabp and place an impella. when the pump products were exchanged there was a bleed at the access site. when transferred to the icu the patient re-bled at the access site. a femostop was placed and 5 units of blood products were infused. | [5 empty] | Since the initial medwatch was filed the investigation has come to a close. the medical facility did not return the product for analysis. without the pump the root cause of the access site bleeding could not be determined. the clinical report did discuss patient movement in transfer between hospitals and possible excessive anticoagulation could have contributed to the blood loss. the failure mode will be monitored and trended. | The investigation is on-going at this time. the root cause of the access site bleed and need for intervention in the form of blood replacement is not known. the product return is expected and upon completion of the investigation, a final medwatch report will be filed.
37623 | 1220648-2019-00248 | 9334985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9334985&pc=OZD | 10/19/2019 | 03/04/2020 |  | 11/18/2019 |  |  | 10/19/2019 | 11/18/2019 |  |  | 04/30/2021 |  | 11/18/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1432279 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 10/19/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting for hemodynamic support using the impella cp optical for ami/cgs. the device was inserted into the right axillary artery without any reported issues, however, ischemia developed. as treatment, the physician performed a fasciotomy. note: despite multiple good faith efforts, we were unable to determine an exact date of this event. therefore, we have used the same date the device was inserted. should we succeed in gathering this information, a supplemental mdr will be filed. | [5 empty] | The pump was received and a thorough investigation was performed. the guide wire repositioning unit diameter was measured within specification. we cannot conclusively determine the root cause of the reported ischemia without further objective evidence. a dhr review was performed and found no manufacturing issues or reworks associated with this pump lot or guide wire repositioning unit lot. there are no other complaints related to this failure mode associated with pumps in this lot. | A good faith efforts have been made to retrieve the impella cp optical, however, we have not yet received the device. should this be received, a supplemental mdr will be filed upon completion of the investigation.
37624 | 1035166-2019-00099 | 9335544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9335544&pc=DYB | 09/25/2019 | 11/18/2019 |  | 11/18/2019 |  |  | 10/23/2019 | 11/18/2019 |  |  | 04/01/2022 |  | 11/18/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-16469 |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient endured a retroperitoneal bleed caused by the introducer. as treatment, heparin was reduced and the bleed was contained. after bleeding was contained, heparin was resumed. on (b)(6) 2019, the patient exhibited an irregular heart rhythm, coded and expired. it was reported there was no allegation that the oscor device caused or contributed to the death. however it was reported the oscor device was alleged to have caused the bleed, the customer reported the death was not reported as related to the bleed. no specific product problem was reported. the device was discarded. additional information has been requested. | [5 empty] | Device used in treatment. the device was discarded and will not be returned for evaluation. there was no specific performance related failure reported by the user. no further investigation is required. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections before shipping to the customer. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. no further follow-up is required.
37625 | 1220648-2019-00247 | 9328555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9328555&pc=OZD | 10/19/2019 | 11/15/2019 |  | 11/15/2019 |  |  | 10/19/2019 | 11/15/2019 |  |  | 11/30/2020 |  | 11/15/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1393569 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/19/2019 | * | Invalid/Incomplete |  |  | The complainant reported a patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp pump. the device was inserted and initiated without any issue, however, the decision was made to withhold from administering heparin (both systemic and in purge fluid) in anticipation of patient receiving a coronary artery bypass procedure. not adding heparin to the purge solution is against the impella's instructions for use recommendations. approximately 3 days after implant, the patient developed a peripheral thrombosis in impella leg that required removal of device and a fasciotomy. patient recovered and later received a successful coronary artery bypass procedure. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37626 | 1220648-2019-00246 | 9321296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9321296&pc=OZD | 10/16/2019 | 12/13/2019 |  | 11/14/2019 |  |  | 10/16/2019 | 11/14/2019 |  |  | 05/31/2021 |  | 11/14/2019 |  | TEMPRORAY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED INC. |  |  | US |  | 1437717 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 10/16/2019 | * | Invalid/Incomplete |  |  | A patient admitted in cardiogenic shock had an impella cp placed for hemodynamic support via the left femoral artery. the patient was decompensating during the angioplasty. after just after five hours of support the team chose to explant the impella due to hemolysis concerns. the labs were hemolyzing and the urine output was tinged in color. the team placed an intra-aortic balloon pump for support. | [5 empty] | Since the original medwatch was filed, the investigation into the reported hemolysis with the use of impella has concluded. the pump was returned, but not the pump data logs. the pump was evaluated on the bench top and run through hemolysis testing. once on the bench top, blue fibers were found inside the pump and attached to the impeller. once these fibers were removed the pump was run and passed hemolysis testing. the blue fibers were found to be the cause of the hemolysis. the fibers were most likely ingested when the pump was run outside the body. the fibers were introduced by the user in the cardiac catheterization lab. the timing of the ingestion can not be confirmed due the lack of pump data logs. the failure mode will be monitored and trended. | The product has been returned for analysis and hemolysis testing. the investigation is on-going at this time. upon completion a final medwatch will be filed.
37627 | MW5091093 | 9328291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9328291&pc=OZD | 11/08/2019 | 11/13/2019 |  | 11/14/2019 |  |  |  | 11/14/2019 |  |  |  |  | 11/14/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT & BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. | [8 empty] | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 77 YR |  |  | 64 |  |  | Death | Yes | No | Break | Device Contamination with Chemical or Other Material | [5 empty] | Death | [12 empty] | Death | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Pt was turned by primary rn and ccp during cleanup of bowel movement. impella alarmed retrograde flow and stopped. code called. troubleshooting done: catheter position checked with ultrasound, no kinks in the catheter. attempts x 3 to restart impella failed. impella rep at bedside for those attempts. backup console put into service but exhibited the same alarm. decision was made to pull the impella back to the descending aorta until it could be removed. purge solution continue to infuse throughout these events. upon further examination, the impella catheter was broken at the distal end; additionally, there was something red at the pigtail end and some tissue-like shreds that were expelled with the catheter when it was pulled were sent for pathology. the impella catheter was sequestered for examination by biomedical engineering. fda safety report id# (b)(4).
37628 | 1035166-2019-00097 | 9315868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9315868&pc=DYB | 10/15/2019 | 02/06/2020 |  | 11/13/2019 |  |  | 10/16/2019 | 11/13/2019 |  |  | 06/01/2022 |  | 11/13/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-17434 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN | [10 empty] | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | During setup of the system, while only 0.18" wire was in the 14f introducer, bleeding through the valve started. after impella placement the bleeding increased. intervention needed, procedure was restarted with utilization of a new introducer kit due to the heavy bleeding through the first one damaged. procedure was completed. no transfusion was performed on the patient. product was discarded. | [5 empty] | The following sections were updated in follow-up 1: b4, d10, g4, g7, h2, h3, h6 and h10: the device was used in treatment. upon initial inspection, the hemostatic valve was observed to be slightly damaged. the valve was then inspected in more detail under 10x magnification. the damage was found to be a tear that stemmed off of the valve split line. no leakage was found from the valve or hub when manually leak tested at the site where the sideport tubing and stopcock attaches to the hub. this test was then repeated the same way, but with a guidewire inserted into the introducer as stated in the event description summary. still there was no leakage observed. device analysis revealed the 14f abiomed introducer sheath is within manufacturing specifications. the device passed leak testing that was performed per applicable iso standard. the reason for return cannot be confirmed. no manufacturing defects were found. per adelante s2s introducer sheath in-process and final inspection vacuum leak test - perform pressure and vacuum leak test per procedure latest revision. this test is performed by qa on 100% of the product. the instructions for use (ifu) informs the user: precautions: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. directions for use: slowly retract the guidewire and dilator, leaving the sheath in position. the hemostasis valve will reduce the loss of blood and the inadvertent aspiration of air through the sheath. introduce the catheter through the hemostasis valve/sheath and advance it into position. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete.
37629 | 1220648-2019-00244 | 9311882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9311882&pc=OZD | 10/18/2019 | 11/12/2019 |  | 11/12/2019 |  |  | 10/18/2019 | 11/12/2019 |  |  | 01/31/2021 |  | 11/12/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1404634 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/18/2019 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old japanese male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient experienced a hemorrhage at the insertion site. as a result impella therapy was stopped and the patient received a total of 8 units of fresh frozen plasma (ffp), 20 units of mannitol, adenine, and phosphate (map20u), and 20 units if platelets (pc20u). | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37630 | 1220648-2019-00245 | 9312041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9312041&pc=OZD | 10/16/2019 | 11/12/2019 |  | 11/12/2019 |  |  | 10/16/2019 | 11/12/2019 |  |  | 06/30/2021 |  | 11/12/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1428735 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 82 YR | [5 empty] | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Regurgitation | Ventricular Fibrillation | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/16/2019 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old (b)(6) female patient had impella 2.5 pump inserted in the right femoral for myocarditis. when impella was inserted the patient was unable to maintain hemodynamics and therefore percutaneous cardiopulmonary support (pcps) were added. during impella support, the patient developed ventricular fibrillation and prompted direct cardioversion that resulted in the pump falling out of position. impellas pigtail entered between the papillary muscles making it difficult to change the position. in addition, the patient developed aortic insufficiency believed to have been caused by impella positioning issue. the physician stated the patient is older and had some degree of valve degeneration; however the leaflets could have been pushed by repositioning. impella was removed due to positional change and treatment was not provided. patient was stable when impella was explanted. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37631 | 1035166-2019-00096 | 9304813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9304813&pc=DYB | 07/31/2019 | 01/21/2020 |  | 11/11/2019 |  |  | 10/16/2019 | 11/11/2019 |  |  | 06/01/2022 |  | 11/11/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-16889 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [11 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Consequences Or Impact To Patient | [41 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that patient was presented to physician in cardiogenic shock. during insertion to femoral artery 14fr oscor introducer had leak at valve while wire was inside. seldinger technique was use for insertion. product was replaced with another oscor 14fr. there was no medical intervention or additional surgery required. patient outcome is stable. no additional information is available. | [5 empty] | The following sections were updated in follow-up. b4,g4,g7,h2,h6,h10 the device used in treatment. one 14f 13cm long abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. blood was found on and inside the sheath and inside the sideport tubing. the sheath was kinked directly under the hub. upon evaluation of the returned product, the hemostasis valve was found to be completely intact, looked normal and there were no splits found down the sheath tube. the sheath was kinked directly under the hub, but there was nothing split open. no leakage was found from the valve or hub when manually leak tested at the site where the sideport tubing and stopcock attaches to the hub. this test was then repeated the same way, but with a guidewire inserted into the introducer as stated in the event description summary. still there was no leakage observed. the introducer was then tested using the iso standard liquid leakage test. the device was connected at the distal tip and pressurized to 38kpa and then examined for liquid leakage. the part passed the leak test. the part was then pressurized to over 300kpa for 30 seconds and held pressure for the 30 second duration. the part passed per the iso standard.returned device analysis revealed the 14f abiomed introducer sheath is within manufacturing specifications. the reason for return cannot be confirmed. no manufacturing defects were found. the device history record was reviewed to confirm that the lead passed all applicable in-process and final inspections. per procedure adelante s2s introducer sheath in-process and final inspection:vacuum leak test - perform pressure and vacuum leak test per procedure latest revision this testing may be performed by manufacturing or quality assurance. occlusion test - perform occlusion test per procedure, latest revision. this testing may be performed by manufacturing or quality assurance. per : precautions: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane directions for use: slowly retract the guidewire and dilator, leaving the sheath in position. the hemostasis valve will reduce the loss of blood and the inadvertent aspiration of air through the sheath. introduce the catheter through the hemostasis valve/sheath and advance it into position. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress ,follow up report will be submitted if we find any further additional information.
37632 | 1220648-2019-00240 | 9295161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9295161&pc=OZD | 10/13/2019 | 11/08/2019 |  | 11/08/2019 |  |  | 10/13/2019 | 11/08/2019 |  |  | 03/31/2021 |  | 11/08/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1411481 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Swelling | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/13/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp device. during use of the device, bleeding and swelling in the impella leg developed. as treatment, a surgical cut down was performed and 2 units of blood products were delivered. the patient remained in stable condition, thereafter. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37633 | 1220648-2019-00241 | 9295230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9295230&pc=OZD | 10/14/2019 | 01/06/2020 |  | 11/08/2019 |  |  | 10/14/2019 | 11/08/2019 |  |  | 03/31/2021 |  | 11/08/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1420277 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 10/14/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. during support, leg ischemia was endured. as treatment, the pump was removed and surgical removal of the patient's leg was required. | [5 empty] | The product was returned and a thorough investigation was completed. simulated use testing was completed in a basin of water and flows were within specification. the pump and guidewire repositioning unit were inspected and found free of defects. the outer diameter of the repositioning unit was found within specification. the device history record was reviewed and found no manufacturing issues or reworks associated with this pump lot. no issues were identified with the device that would cause or contribute to the reported condition. due to lack clinical information regarding the patient's anticoagulation status, we cannot definitively determine a root cause for the reported ischemia. | The impella cp optical was returned and an investigation is underway. a supplemental mdr will be filed, upon completion.
37634 | 1220648-2019-00242 | 9295245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9295245&pc=OZD | 10/17/2019 | 11/08/2019 |  | 11/08/2019 |  |  | 10/17/2019 | 11/08/2019 |  |  | 02/28/2021 |  | 11/08/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1419262 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/17/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp optical device. the physician attempted re-crossing the aortic valve without a guide wire, after the pump had been sitting in the aorta during a coronary artery bypass procedure. in doing so, the left ventricle was perforated. the physician surgically repaired the perforation and the patient remained in stable condition. prior to attempted wireless re-crossing, the abiomed field rep advised the physician that crossing the aortic valve without a guide wire is not recommended. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37635 | 1220648-2019-00243 | 9295257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9295257&pc=OZD | 10/16/2019 | 01/06/2020 |  | 11/08/2019 |  |  | 10/16/2019 | 11/08/2019 |  |  | 06/30/2021 |  | 11/08/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1428058 |  | 005082 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | Yes | Break | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 10/16/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male patient presenting with cardiomyopathy for hemodynamic support with the impella ld. during device removal, the cannula became trapped within the graft and separated from the device. while no surgical intervention was required, the physician manually manipulated the component through the graft and out of the patient. the graft was trimmed and sutured closed without issue. the patient was stable, thereafter. | [5 empty] | The device was returned and a thorough investigation was completed. a visual inspection found the device's catheter completely separated from the pump motor housing; confirming the reported issue. a dhr was conducted and found no manufacturing issues or reworks associated with this pump lot. the root cause of cannula detachment is attributed to the pump being trapped in the graft while pulling back the device. | The impella ld was returned an investigation is underway. a supplemental mdr will be filed, upon completion.
37636 | 1220648-2019-00238 | 9283868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9283868&pc=OZD | 10/10/2019 | 11/06/2019 |  | 11/06/2019 |  |  | 10/10/2019 | 11/06/2019 |  |  | 05/31/2021 |  | 11/06/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1424928 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/10/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese female patient had impella 2.5 pump inserted in the right femoral for myocardial. during impella support the patient developed hemolysis and required 8 units of red blood cells over a period of two days. it is unclear whether the impella or extracorporeal membrane oxygenation (ecmo) caused the hemolysis. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37637 | 1220648-2019-00239 | 9283872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9283872&pc=OZD | 10/11/2019 | 11/06/2019 |  | 11/06/2019 |  |  | 10/11/2019 | 11/06/2019 |  |  | 02/28/2021 |  | 11/06/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1412004 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/11/2019 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) japanese male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction and cardiogenic shock (ami/cgs). the patient experience urinary bleeding sometime after impella insertion. the bleeding is believe to have been caused by hemolysis which is believed to have been cause by the impella. due to the suspected hemolysis the impella was explanted. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37638 | 9278441 | 9278441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9278441&pc=PBL | 10/10/2019 | 10/28/2019 | 10/28/2019 | 11/05/2019 | 10/28/2019 | 11/05/2019 |  | 11/05/2019 |  |  |  |  | 11/05/2019 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP | [5 empty] | US |  | 17877869 |  |  |  | Device was returned to manufacturer | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 24090 DA | [5 empty] | Malfunction | No | Yes | Device Displays Incorrect Message | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella device incurred an error during surgical procedure. device rep. notified, and several troubleshooting interventions were done. impella still would not work, so device was removed and replaced with new device.
37639 | 1220648-2019-00237 | 9272673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9272673&pc=OZD | 10/09/2019 | 11/04/2019 |  | 11/04/2019 |  |  | 10/09/2019 | 11/04/2019 |  |  | 07/31/2021 |  | 11/04/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1429067 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/09/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese patient had impella cp pump inserted for acute myocardial infarction and cardiogenic shock (ami/cgs). it was reported that when the 14fr sheath and repo sheath were replaced, the patient experienced bleeding due to unrest. a large hematoma of the puncture site formed. the patient required 2 units of fresh frozen plasma. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37640 | 1035166-2019-00092 | 9274843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9274843&pc=DYB | 09/24/2019 | 01/16/2020 |  | 11/04/2019 |  |  | 10/09/2019 | 11/04/2019 |  |  | 06/01/2022 |  | 11/04/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-17243 |  | 0052-3000 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that an (b)(6) white male patient was admitted for chronic total occlusion of the coronary artery and the products were used for a high risk percutaneous coronary intervention procedure. while attempting to insert the introducer into the patient's right femoral artery, the patient endured a "massive vasovagal reaction" and required cpr, intubation, and high dose catecholamine's. the insertion site was checked and no dissection was noted. the physician proceeded to insert the introducer into the left femoral artery without issue (no specification of time, change took place very quickly). the impella was inserted successfully, however, the site developed a bleed at the introducer. no further information available as to site of leak, there was blood leaking originating from the outside/surrounding the introducer insertion site. the patient lost approximately 500ml of blood and required a transfusion. both the introducer and impella were removed and the site was closed with manual compression and a femstop. no new device was used and the patient survived. after implantation of impella, the patient's hemodynamics improved and became more stable. it was reported approximately two units of blood were given due to patient's hemoglobin was lower (approximately 10 g/dl). no introducer problem was indicated. additional information has been requested. | [5 empty] | The device was used for treatment. there was no product performance issue reported. the device was not returned for evaluation, therefore, the clinical observation could not be confirmed. the investigation will focus on the documentation review. the device history records were reviewed to confirm that the devices passed all applicable in process and final inspections. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete.
37641 | 1220648-2019-00236 | 9268656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9268656&pc=OZD | 10/07/2019 | 11/01/2019 |  | 11/01/2019 |  |  | 10/07/2019 | 11/01/2019 |  |  | 04/30/2021 |  | 11/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1433342 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/07/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old male patient had impella cp optical pump inserted. a day after the pump was explanted, limb ischemia was discovered. the patient was taken to the operating room (or) to remove the pump and pulse could not be restore after surgical intervention. | [5 empty] | The impella cp optical pump was discarded by the customer therefore a analysis of the device could not be completed.
37642 | 1220648-2019-00235 | 9257249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9257249&pc=OZD | 10/02/2019 | 12/13/2019 |  | 10/30/2019 |  |  | 10/02/2019 | 10/30/2019 |  |  | 12/31/2020 |  | 10/30/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | AACHEN, GERMANY |  |  | GM |  | 1391180 |  | 005048-JP |  | Yes | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 15 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | ECMO | [22 empty] | Initial submission | Followup |  |  |  | 10/02/2019 | * | Invalid/Incomplete |  |  | The team in (b)(6) was utilizing an impella 2.5 to provide hemodynamic support to the patient admitted with myocarditis. the team chose to place the impella in addition to the already placed ecmo. during the 8 days of support the patient was showing signs of hemolysis. the medical team presumed there was hemolysis due to the rising ldh lab values and hematuria. the impella pump position was assessed and found to not be in what the team felt was good position, which could cause or contribute to hemolysis. the patient received an infusion of haptoglobin to treat the hemolysis and the impella and ecmo were explanted. | [5 empty] | Since the initial medwatch was filed no new information has been shared by the japanese medical team. the product was not returned for hemolysis testing nor were data logs or echo imaging shared. the root cause of the suspected hemolysis can not be determined. the medical team in japan had a supposition that the improper positioning of the pump and concurrent use of ecmo could have caused the hemolysis. the failure mode will be monitored and trended.
37643 | 1220648-2019-00231 | 9250838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9250838&pc=OZD | 10/02/2019 | 10/29/2019 |  | 10/29/2019 |  |  | 10/02/2019 | 10/29/2019 |  |  | 05/31/2021 |  | 10/29/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1424928 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/02/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese male patient with a large physique had impella 2.5 pump inserted in the left femoral during an electro physiology procedure due to patient condition worsening. the patient experienced bleeding during the procedure. patient's blood pressure decreased and required blood products. amount of blood given was not provided. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37644 | 1220648-2019-00232 | 9252038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9252038&pc=OZD | 10/04/2019 | 12/20/2019 |  | 10/29/2019 |  |  | 10/04/2019 | 10/29/2019 |  |  | 03/31/2021 |  | 10/29/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1424748 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 10/04/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient presenting for an aortic valvuloplasty due to critical aortic stenosis. the impella cp optical device was chosen for hemodynamic support. the physician elected to place the impella device deeper into the ventricle, contrary to the impella's instructions for use (ifu). the physician advised that a deeper placement of the device was necessary as the patient's anatomy made for a difficult insertion, and if the device fell out of the ventricle, re-crossing the aortic valve would not be possible. subsequently, the patient developed hemolysis that was confirmed with plasma-free hemoglobin measurements of 1080 mg/dl and 800 mg/dl, respectively. the device was prematurely removed, as treatment. a second device was not used as the patient's hemodynamic state improved. | [5 empty] | The device and data logs were returned an a thorough investigation was completed. data logs revealed "pump in ventricle" alarms in the beginning of the case which coincide with the issues positioning the pump initially. a visual inspection of the device was completed and no deformities were found. hemolysis testing was performed and the pump passed without issue. evidence suggests the root cause of the reported hemolysis was due to pump being seated intentionally deeper into the patient's ventricle than recommended. a review of the dhr was conducted and found no manufacturing issues or reworks associated with this failure mode or pump lot. there are no other complaints related to this failure mode associated with pumps in this lot. | The impella cp optical device was returned and an investigation is underway. once the evaluation is completed, a supplemental mdr will be filed.
37645 | 1220648-2019-00233 | 9252050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9252050&pc=OZD | 10/02/2019 | 10/29/2019 |  | 10/29/2019 |  |  | 10/02/2019 | 10/29/2019 |  |  | 03/31/2021 |  | 10/29/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1417939 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/02/2019 | * | Invalid/Incomplete |  |  | The complainant reported a patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp device. during support, a retroperitoneal bleed was discovered and attributed to the insertion of the 14 fr oscor introducer. as a result, the patient was delivered approximately 4 to 5 units of blood over the course of 2 days. | [5 empty] | The pump and introducer were discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37646 | 1220648-2019-00234 | 9252440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9252440&pc=OZD | 06/11/2019 | 03/26/2020 |  | 10/29/2019 |  |  | 10/01/2019 | 10/29/2019 |  |  | 09/30/2020 |  | 10/29/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1374528 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemolysis | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 10/01/2019 | * | Invalid/Incomplete |  |  | Additional information was received from the physician revising the injury from "exacerbation of anemia" to "hemolysis." additionally, the physician revised the date of this adverse event from (b)(6) 2019 to (b)(6) 2019. | The complainant reported a (b)(6) year old patient presenting with myocarditis for hemodynamic support using the impella 5.0 device. during treatment, the physician alleged the patient's pre-existing anemic condition had worsened, however, could not attribute this to a single cause. the physician felt the combination of simultaneous support from percutaneous cardiopulmonary support, hematoma, heparin induced thrombocytopenia, and impella support contributed to the exacerbation of the patient's pre existing condition. the relationship of the impella and this adverse event could not be conclusively determined. as treatment, 2 units of blood products were delivered. |  |  |  |  | Updated b3, b5, and h6 to include corrected information from customer. | The impella 5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37647 | 1220648-2019-00230 | 9244775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9244775&pc=OZD | 09/30/2019 | 10/28/2019 |  | 10/28/2019 |  |  | 09/30/2019 | 10/28/2019 |  |  | 05/31/2021 |  | 10/28/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EU |  |  | GM |  | 1420671 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 09/30/2019 | * | Invalid/Incomplete |  |  | A patient was admitted in cardiogenic shock to the medical facility in (b)(6). the patient had deteriorated on route to the facility and was cardioverted for ventricular fibrillation. upon admission he had coronary angioplasty while supported by an impella 2.5 pump. the pump remained on for hemodynamic support for 51 hours when the team chose to explant the pump due to hemolysis concerns. the team had attempted to remedy the hemolysis with pump repositioning and infusion of haptoglobin and red blood cells, to no effect. after pump explant the patient was stated to be recovered and stable without lasting harm from the hemolysis. | [5 empty] | The medical facility in (b)(6) did not return the impella product for analysis. the product was discarded upon explant. the medical facility did not release further clinical details or medical imaging to review impella pump positioning. should more details be shared that will lead to root cause, a final medwatch will be filed. the root cause of the hemolysis is not known. the failure mode will be monitored and trended.
37648 | 1220648-2019-00229 | 9226853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9226853&pc=OZD | 05/21/2019 | 10/23/2019 |  | 10/23/2019 |  |  | 10/01/2019 | 10/23/2019 |  |  | 09/30/2020 |  | 10/23/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1374516 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 10/01/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had impella 5.0 pump inserted in the ascending aorta. during pump support the patient had hemolysis. as treatment a total transfusion of 4 units of mannitol, adenine, phosphate (map) was performed and pump was removed. | [5 empty] | Device discarded by customer. the impella 5.0 pump was discarded by the customer therefore a failure analysis investigation of the device cannot be completed. should additional information is received a supplemental report will be submitted.
37649 | 1220648-2019-00227 | 9218739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9218739&pc=OZD | 09/27/2019 | 12/06/2019 |  | 10/22/2019 |  |  | 09/27/2019 | 10/22/2019 |  |  | 01/31/2021 |  | 10/22/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1401005 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | ECMO | [21 empty] | Initial submission | Followup |  |  |  | 09/27/2019 | * | Invalid/Incomplete |  |  | The medical team had a patient admitted in cardiogenic shock via the emergency department. the team placed an impella cp for hemodynamic support in the cardiac catheterization lab during coronary angioplasty procedure. the team then transferred the patient to the tertiary center for care escalation. the team placed ecmo for respiratory support. the patient began to experience signs of hemolysis. the plasma free hemoglobin level was elevated in a lab draw. the team changed the ecmo circuit, but found the hemolysis did not clear. the team explanted the impella as the presumed cause of the hemolysis. the pump had supported the patient for six days. | [5 empty] | The investigation is on-going at this time. upon completion of the investigation, the final medwatch report will be filed. | Since the original medwatch report was filed, the investigation has completed. the complainant returned the pump for hemolysis testing and analysis, but not the data logs. the pump was evaluated and run through internal benchtop hemolysis testing. the pump passed the testing. as the pump passed the test, the root cause of the reported hemolysis could not be determined. the request for imaging and further medical history, inclusive of renal history, was not provided. the failure mode will be monitored and trended.
37650 | 1220648-2019-00228 | 9219981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9219981&pc=OZD | 09/27/2019 | 12/06/2019 |  | 10/22/2019 |  |  | 09/27/2019 | 10/22/2019 |  |  | 05/31/2021 |  | 10/22/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1434603 |  | 0048-0024 |  | Yes | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 09/27/2019 | * | Invalid/Incomplete |  |  | The medical team placed a cardiomyopathy patient on impella hemodynamic support, and admitted to the patient to the icu for continued monitoring. the support lasted only 26 hours due to hemolysis concerns. as the team observed lab values changing, lft and creatinine specifically, and hemolyzed labs the team attempted to remedy the hemolysis by pump re-positioning. the change in pump position did not rectify the hemolysis and so the pump was explanted and the plan was to place a vad. | [5 empty] | Since the initial report was filed, the investigation has come to a close. the product and data logs were not returned for analysis. the case report stated that upon repositioning of the pump, the hemolysis did not clear. without further information or clinical imaging to show pump position and/or malposition, the root cause of the hemolysis can not be determined. the failure mode will be monitored and trended. | The complaint investigation is on-going. the impella pump is not yet returned for investigation. upon completion of the investigation, the final medwatch will be filed.
37651 | 1220648-2019-00226 | 9214385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9214385&pc=OZD | 09/13/2019 | 10/21/2019 |  | 10/21/2019 |  |  | 09/24/2019 | 10/21/2019 |  |  | 04/30/2021 |  | 10/21/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1417620 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/24/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) japanese male patient had an impella cp pump inserted in the left femoral. we received information that the physician commented the relation was unknown between sepsis and patient¿s death. the pump was not returned for investigation and the clinical details provided were not sufficient to determine relationship of the adverse event to the impella. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37652 | 1220648-2019-00224 | 9203104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9203104&pc=PYX | 09/20/2019 | 12/06/2019 |  | 10/17/2019 |  |  | 09/20/2019 | 10/17/2019 |  |  | 05/31/2021 |  | 10/17/2019 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1404623 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | Yes | Material Separation | Physical Resistance/Sticking | [5 empty] | Blood Loss | Device Embedded In Tissue or Plaque | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 09/20/2019 | * | Invalid/Incomplete |  |  | The complainant reported a 78 year old male patient presenting with right ventricle failure post implantation of a permanent left ventricular assist device. the impella rp device was selected for hemodynamic support. during attempted insertion, the introducer kinked and the impella's outlet cage failed to pass through the opening at the end of the 23 fr introducer. it should be noted that the physician chose to insert the pump in the patient's left femoral vein, which is against recommendations, as indicated in the device's instructions for use. device insertion was aborted, however, the impella's pigtail and tear drop separated from the catheter and remained inside the patient's anatomy. as treatment, a surgical snare was administered and the fragmented components were successfully removed. the patient remained in stable condition. | The complainant reported a (b)(6) male patient presenting with right ventricle failure post implantation of a permanent left ventricular assist device. the impella rp device was selected for hemodynamic support. during attempted insertion, the impella's outlet cage failed to pass through the opening at the end of the 23 fr introducer. device insertion was aborted, however, the impella's pigtail and tear drop separated from the catheter and remained inside the patient's anatomy. as treatment, a surgical snare was administered and the fragmented components were successfully removed. the patient remained in stable condition. |  |  |  |  | Correction: revised pma number, as it was erroneous in the initial mdr. revised event description as initial narrative was incomplete. the device was received for evaluation and the reported device separation was confirmed. the outer diameter of the pump was measured and found to be within specification. evidence suggests the force applied to push the pump through the reported kink and to the opening of the 23 fr introducer had resulted in the ultimate separation of the device. a review of the device history record was conducted and found no manufacturing issues or reworks associated with this pump lot. there are no other complaints related to this failure mode associated with pumps in this lot. | The impella rp is expected to be returned for analysis, but not yet received. should the device be received, a supplemental mdr will be filed after completion of an investigation.
37653 | 1220648-2019-00225 | 9204036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9204036&pc=OZD | 09/24/2019 | 10/17/2019 |  | 10/17/2019 |  |  | 09/24/2019 | 10/17/2019 |  |  | 05/31/2021 |  | 10/17/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1422673 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | Vaso-Vagal Response | [11 empty] | Life Threatening | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 09/24/2019 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old caucasian male patient presenting for high rick percutaneous coronary intervention using the impella cp for hemodynamic support. during attempted insertion of the 14fr introducer into the right femoral artery, the patient endured a vasovagal rection that required cardiopulmonary resuscitation, intubation, and high dose catecholamines. the artery was inspected and no dissection or otherwise reported damage was present. a second insertion attempt took place on the opposing femoral artery with the same introducer and was successful. the introducer was left in the patient's artery. a large bleed eventually developed around the outside of the introducer that prompted removal of both the introducer and pump. approximately 2 units of blood was delivered. no new impella insertion was required, thereafter. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37654 | 1220648-2019-00223 | 9191751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9191751&pc=OZD | 09/17/2019 | 10/15/2019 |  | 10/15/2019 |  |  | 09/18/2019 | 10/15/2019 |  |  | 11/30/2020 |  | 10/15/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1385393 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 60 YR | [5 empty] | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Insufficiency, Valvular | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/18/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted in the right femoral. the patient¿s papillary muscle ruptured and mitral valve replacement (mvr) was performed as a result. the pump was removed and the patient was put on veno arterial venous (vav)- extracorporeal membrane oxygenation (ecmo). | [5 empty] | Device discarded by customer. the impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed. should any additional information is received, a supplemental will be submitted.
37655 | 1220648-2019-00222 | 9186109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9186109&pc=OZD | 09/15/2019 | 10/14/2019 |  | 10/14/2019 |  |  | 09/15/2019 | 10/14/2019 |  |  | 12/31/2020 |  | 10/14/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1397071 |  | 0048-0032 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Aortic Insufficiency | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 09/15/2019 | * | Invalid/Incomplete |  |  | A patient was transferred to the tertiary medical center for continued care on an already placed impella cp. upon transfer and admission to the tertiary center the team conducted an echo and observed aortic insufficiency (ai). the team attempted to adjust the pump position to reduce the ai, but all repositioning failed to remedy the issue. the team explanted the pump. | [5 empty] | To date the clinical information shared does not yet lead to a conclusion as to the etiology of the aortic insufficiency. the team discarded the pump product and has not shared echo imaging for analysis. the root cause is not known. should more information be shared that will lead to a root cause, a supplemental report will be filed. the failure mode will be monitored and trended.
37656 | 1220648-2019-00219 | 9176289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9176289&pc=OZD | 07/03/2019 | 10/09/2019 |  | 10/10/2019 |  |  | 09/12/2019 | 10/10/2019 |  |  | 12/31/2020 |  | 10/10/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1391180 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/12/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old (b)(6) male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction (ami) with shock. the patient had leg ischemia on the fourth day after the pump was removed. surgical repair to treat the ischemia included a sheath inserted in the right common iliac artery to transfer blood. | [5 empty] | Device discarded by customer. the impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37657 | 1220648-2019-00218 | 9176290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9176290&pc=OZD | 09/12/2019 | 10/09/2019 |  | 10/10/2019 |  |  | 09/12/2019 | 10/10/2019 |  |  | 04/30/2021 |  | 10/10/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1421620 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Material Deformation | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/12/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella cp pump inserted in the right femoral. the patient was having some hematuria with labs hemolyzing, repositioning was done but catheter was kinked like an ¿s¿ and a vascular surgeon repaired the artery via cut down and suture was used. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore an evaluation of the device cannot be completed.
37658 | 1220648-2019-00220 | 9176377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9176377&pc=OZD | 07/22/2019 | 11/27/2019 |  | 10/10/2019 |  |  | 09/18/2019 | 10/10/2019 |  |  | 12/31/2020 |  | 10/10/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1387596 |  | 005060 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | DISTRI |  |  |  |  | 60 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | Hemolysis | [11 empty] | Death | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 09/18/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with myocardial infarction and requiring high risk percutaneous coronary intervention (hrpci). the hrpci was completed successfully, however, the patient developed signs of hemolysis via falling haptoglobin, rising ldh, rising bilirubin, and red colored urine. in addition to the reported hemolysis, the patient endured renal failure of an unknown cause. repositioning was successful, however, the hemolysis did not resolve and the patient expired of multiple organ failure. the physician believes the hemolysis had been a contributing factor in the patient's death. | [5 empty] | Correction: the date of occurrence was erroneous in the initial report, and has been corrected. the pump was returned for analysis and a through investigation was performed. visual inspection found no deficiencies, abnormalities, or malfunctions that could have caused or contributed to the reported hemolysis. simulated use testing was conducted and the pump operated to specification. we are unable to definitively determine a root cause as the reported failure could not be reproduced. | The impella 5.0 was returned and an investigation is underway. a supplemental mdr will be filed upon completion of the investigation.
37659 | 1220648-2019-00221 | 9176397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9176397&pc=OZD | 09/18/2019 | 12/02/2019 |  | 10/10/2019 |  |  | 09/18/2019 | 10/10/2019 |  |  | 06/30/2021 |  | 10/10/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1430231 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 37 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 09/18/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old african american male presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp. during support, the patent endured signs of hemolysis via red urine and hemolyzed labs. a test for plasma-free hemoglobin was not measured. normal troubleshooting was unable to resolve this problem, therefore, the device was removed and a new impella 5.0 was inserted. the patient was ultimately transferred to a new hospital, however, hemolysis had cleared before he left the hospital. | [5 empty] | The impella cp device and data logs were received and an investigation was performed. analysis of the console's data logs found several suction alarms throughout the case, however, these alarms were resolved upon repositioning of the device. a visual inspection found no defects or abnormalities with the pump and simulated use testing for hemolysis passed successfully. we are unable to definitively determine a root cause as the reported issue was unable to be reproduced, however, information received from the clinical account suggests improper positioning of the device contributed to the patient's hemolysis. | The impella cp was requested from the customer, but not received to date. should the device or any new information be returned, a supplemental mdr will be filed.
37660 | 1220648-2019-00215 | 9168216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9168216&pc=OZD | 09/13/2019 | 10/08/2019 |  | 10/08/2019 |  |  | 09/13/2019 | 10/08/2019 |  |  | 07/31/2021 |  | 10/08/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1431305 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 79 YR | [5 empty] | Death | Yes | No | Improper or Incorrect Procedure or Method | [6 empty] | Death | Cardiac Perforation | Blood Loss | [10 empty] | Death | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 09/13/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female caucasian presenting for high risk percutaneous coronary intervention (hrpci) while using the impella cp for hemodynamic support. during the procedure, the patient became unstable and required 2 rounds of cardiopulmonary resuscitation. a transthoracic echocardiogram revealed the impella pig tail had perforated the left ventricle, resulting in a cardiac tamponade. the patient was taken to the operating room and the perforation was repaired, however, the patient ultimately expired. the cause of death was attributed to blood loss during this event. during a post event interview between an abiomed clinical field representative and the physician, it was believed that the impella catheter was not held tightly enough during the insertion of the hrpci catheter, resulting in excess friction causing advancement of the impella catheter further into the ventricle than intended. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37661 | 1220648-2019-00216 | 9168377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9168377&pc=OZD | 08/10/2019 | 10/08/2019 |  | 10/08/2019 |  |  | 09/13/2019 | 10/08/2019 |  |  | 05/31/2021 |  | 10/08/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1421773 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 41 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/13/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient with a large physique had impella 2.5 pump inserted in the left femoral for myocarditis / cardiogenic shock (cgs). the patient had leg ischemia during pump support as there was coldness in the lower limb; tactile and doppler were poor. interventional treatment was needed including forward puncture, percutaneous bypass and as a result the patient recovered. | [5 empty] | Device discarded by customer. the impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37662 | 1220648-2019-00217 | 9168392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9168392&pc=OZD | 09/14/2019 | 11/22/2019 |  | 10/08/2019 |  |  | 09/14/2019 | 10/08/2019 |  |  | 02/28/2021 |  | 10/08/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1406058 |  | 0048-0032 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 09/14/2019 | * | Invalid/Incomplete |  |  | A post-operative patient had an intra aortic balloon pump (iabp) placed. the iabp was not sufficient support and so his care was escalated to the impella cp. the patient also had an pcps/ecmo (extracorporeal membrane oxygenation) for respiratory support. after insertion of the impella cp there was an observation made of bleeding at the access site. when the vascular surgeon performed a cut-down to the site there was damage noted to be at the impella pump's reposheath. the access site bleed was maintained with femostop and manual compression, in addition to infusion of blood products. the pump remained on for hemodynamic support for 13 days until explant. | [5 empty] | Since the initial report was filed, the investigation into the ischemia and bleeding from the access site has completed. the returned pump showed no signs of damage. the pump diameter was assessed to be to specification. the diameter of the pump was appropriated for the intended vascular access site and vessel diameter. the cause of the ischemia and bleeding could not be determined. the failure mode will be monitored and trended. | To date the investigation is ongoing. no product has yet been returned. should the medical team return the product or share more details that lead to a cause of the vessel access site bleed, a final medwatch will be filed with root cause analysis.
37663 | 1220648-2019-00209 | 9163377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9163377&pc=OZD | 09/11/2019 | 10/07/2019 |  | 10/07/2019 |  |  | 09/11/2019 | 10/07/2019 |  |  | 08/31/2020 |  | 10/07/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1363621 |  | 005048-JP |  | Yes | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/11/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted in the right femoral for acute myocardial infarction (ami) with shock. the patient had leg ischemia. surgical repair to treat the ischemia included a sheath inserted to transfer blood. | [5 empty] | Device has not been returned. the impella 2.5 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37664 | 1220648-2019-00210 | 9163425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9163425&pc=OZD | 09/10/2019 | 10/07/2019 |  | 10/07/2019 |  |  | 09/10/2019 | 10/07/2019 |  |  | 02/28/2021 |  | 10/07/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPOR BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1425175 |  | 005042 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 09/10/2019 | * | Invalid/Incomplete |  |  | The patient was admitted for a high risk coronary angioplasty (hrpci) and for hemodynamic support during the angioplasty procedure, had an impella 2.5 placed. the medical team knew that the patient's vasculature was not of a size to accommodate the 13fr introducer and pump, though it was placed despite this observation. the access site limb was noted to have signs of limb ischemia. the ischemia was observed while in the icu. the team chose to explant the impella pump. the pump had been supporting the patient for only 8 hours. the limb did recover pulses and the coolness dissipated. there was no further intervention done to the limb, other than impella explant. | [5 empty] | The complainant did not return the product, data logs, or patient imaging for analysis. without the product a root cause of the limb ischemia can not be determined. though, patient condition, specifically small diameter vessels and levophed (a vasoconstrictor) could have contributed to the ischemia. the failure mode will be monitored and trended.
37665 | 1220648-2019-00211 | 9163700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9163700&pc=PYX | 09/10/2019 | 12/02/2019 |  | 10/07/2019 |  |  | 09/10/2019 | 10/07/2019 |  |  | 01/31/2021 |  | 10/07/2019 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1403979 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | Yes | Material Separation | Physical Resistance/Sticking | [5 empty] | Blood Loss | Device Embedded In Tissue or Plaque | [11 empty] | Required Intervention | [5 empty] | IMPELLA CP | [22 empty] | Initial submission | Followup | Followup |  |  | 09/10/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female caucasian presenting with acute myocardial infarction and cardiogenic shock for right heart support with the impella rp device. during device insertion, some resistance was endured when passing the device through the introducer. after passing through introducer, prior to inserting into the pulmonic valve, the pump's teardrop had separated from the cannula and traveled into the right atrium. as treatment, a surgical snare was inserted and the device was removed. the patient endured blood loss which required 3 units of blood product delivery. the patient tolerated the procedure well and did not endure any additional adverse impact. the patient did not require a new rp device and was successfully supported with milrinone and continued impella cp support. | [5 empty] | Correction: in the initial report, the device' pma approval number was not included. we have added this information in this report. | A thorough investigation of the returned introducer and pump were conducted. both the reported introducer kink and detached outflow cage were confirmed. based on the clinical report, we believe the cause of the introducer kink was due to steep insertion angle required by the patient¿s anatomy. the kink in the introducer increased the forces required to deliver the device and these forces caused the outflow cage to deform and ultimately detach. a review of the device history record found no manufacturing issues, reworks, or other complaints associated with this failure mode for this pump lot. | The impella cp was returned and an investigation is underway. upon completion of the investigation, a supplemental mdr will be filed.
37666 | 1220648-2019-00212 | 9163723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9163723&pc=OZD | 09/06/2019 | 10/07/2019 |  | 10/07/2019 |  |  | 09/10/2019 | 10/07/2019 |  |  | 12/31/2020 |  | 10/07/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1397692 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/10/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient had impella 2.5 pump inserted for acute myocardial infarction (ami) with shock without any problems. the pump was removed due to hemolysis. | [5 empty] | Device discarded by customer. the impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37667 | 1220648-2019-00213 | 9164061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9164061&pc=OZD | 09/12/2019 | 10/07/2019 |  | 10/07/2019 |  |  | 09/12/2019 | 10/07/2019 |  |  | 12/31/2020 |  | 10/07/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1391206 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 84 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/12/2019 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella 2.5 pump. during support, the patient exhibited signs of leg ischemia (coldness) on the same leg that which the device was inserted. as treatment, an antegrade sheath was inserted. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37668 | 1220648-2019-00214 | 9164072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9164072&pc=OZD | 09/12/2019 | 10/07/2019 |  | 10/07/2019 |  |  | 09/12/2019 | 10/07/2019 |  |  | 09/30/2020 |  | 10/07/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1370748 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Insufficiency | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/12/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella 2.5 device. during support, the patient endured aortic insufficiency as a result of impella use. as treatment, the device was removed, the patient recovered, and no further treatment was performed. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37669 | 1220648-2019-00208 | 9144889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9144889&pc=OZD | 09/09/2019 | 11/26/2019 |  | 10/02/2019 |  |  | 09/09/2019 | 10/02/2019 |  |  | 04/30/2021 |  | 10/02/2019 |  | TEMPORARY NON ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1423917 |  | 0048-0024 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 89 YR | [5 empty] | Malfunction | No | Yes | Material Fragmentation | Difficult to Remove | [5 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 09/09/2019 | * | Invalid/Incomplete |  |  | When the impella was placed the team inserted through a long 14fr x25 cm sheath. the longer sheath was utilized due to peripheral arterial disease. the pump supported the patient hemodynamically for the lm and lad artery to be treated with rotational atherectomy and stenting. upon the completion of the stenting the pump was removed. as the team was removing the pump there was pump damage. the pump fractured into two pieces within the introducer. the motor housing and outlet cage remained within the introducer. all parts of the pump were removed without surgical intervention or the use of a snare. | [5 empty] | Since the initial medwatch was filed the investigation has concluded. the clinical report noted the pump separation occurred at removal through the introducer sheath. there was a nonabiomed provided .035 guidewire in the introducer at the time of pump removal. the guidewire was presumed to have caused additional resistance forces at removal of the pump. the higher than expected forces during removal allowed the cannula to separate. the failure mode will be monitored and trended. | To date the medical facility has not yet returned to abiomed the malfunctioning product. the investigation into the root cause of the pump fracture has not begun. upon the completion of the investigation, a final medwatch will be filed.
37670 | 1220648-2019-00205 | 9141787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9141787&pc=OZD | 09/06/2019 | 11/20/2019 |  | 10/01/2019 |  |  | 09/06/2019 | 10/01/2019 |  |  | 04/30/2021 |  | 10/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1427299 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 57 YR | [5 empty] | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 09/06/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) african american patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp. during support, the patient hemolysis that was confirmed with plasma-free hemoglobin levels exceeding 400mg/dl. as treatment, the device was removed and replaced with an intra-aortic balloon pump. the patient's hemolysis cleared, thereafter. | [5 empty] | The pump and data logs were returned and a thorough investigation was completed. data logs indicate the pump had operated normally during use. visual inspection found no abnormalities and simulated use testing was unable to reproduce the reported hemolysis. based on the evaluation, we are unable to definitively determine a root cause. | The impella cp was returned and an investigation is underway. a supplemental mdr will be filed upon the completion of the investigation.
37671 | 1220648-2019-00206 | 9141794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9141794&pc=OZD | 09/04/2019 | 10/01/2019 |  | 10/01/2019 |  |  | 09/04/2019 | 10/01/2019 |  |  | 09/30/2020 |  | 10/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1376425 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 84 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 09/04/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) caucasian male presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp. during support, a large hematoma formed around the access site. the next day, the decision was made to remove the impella due to site enduring "pulsatile bleeding." the site required surgical closure and 4 units of blood products, as treatment. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed.
37672 | 1220648-2019-00207 | 9141839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9141839&pc=OZD | 05/22/2019 | 10/09/2020 |  | 10/01/2019 |  |  | 05/22/2019 | 10/01/2019 |  |  |  |  | 10/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | FOREIGN | HEALTH PROFESSIONAL | L |  |  |  | 37 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemolytic Anemia | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 05/22/2019 | * | Invalid/Incomplete |  |  | A scientific publication reported an adverse event endured by an impella cp patient. the article reports a (b)(6) female patient presenting with cardiogenic shock for hemodynamic support with the impella cp device. the device was inserted without any reported problems. despite "careful repositioning" evidence of hemolysis and hemolytic anemia were observed approximately 24 hours later. weaning of the device was not successful, therefore, patient was considered for an lvad. approximately 5 days later, the device was surgically removed following a successful lvad implantation. post removal of the impella device, a partial thrombosis was identified on the impeller blades. as treatment, a thrombectomy was completed and surgical reconstruction of the artery was required. the hemolysis resolved and patient was ultimately discharged in "very good condition." | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation. | Correction: g(4): upon review of this complaint, it was determined the incorrect aware date was provided in the initial report. the aware date should be 2019-07-08.
37673 | 1220648-2019-00201 | 9124428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9124428&pc=OZD | 09/17/2019 | 12/05/2019 |  | 09/26/2019 |  |  | 09/17/2019 | 09/26/2019 |  |  | 02/09/2019 |  | 09/26/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE |  |  | GM |  | 1417457 |  | 0550-0001 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 76 YR | [5 empty] | Malfunction | No | Yes | Device Alarm System | Alarm Not Visible | [5 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 09/17/2019 | * | Invalid/Incomplete |  |  | The cardiomyopathy patient was admitted to a medical facility in (b)(6) and placed on an impella 5.5 for hemodynamic support. this was done pre-lvad insertion per physician protocol. the pump came out of the left ventricle and was repositioned with imaging guidance. sometime later the pump came out of position again but this was not noted by the automated controller (aic). the aic did not alarm for being out of position. this could have caused harm or injury to the patient. when the malposition was noted the team transferred the patient to the operating suite for the lvad placement, but instead deferred and placed an impella 5.0. the patient remains successfully supported by the impella 5.0 for 22 days now, as of the filing of this report. | [5 empty] | Since the initial report was filed, the investigation has concluded. originally, with the user report information supplied, the complaint was thought to be against the console. upon further investigation of the console logs, the abiomed team has discovered that the issue was pump related, rather than the console. the pump was out of position due to the pump's motor current pulsatility remaining higher than the threshold for triggering the out of position alarm on the console. | The investigation into the failure of the aic to alarm for a pump being malpositioned, is ongoing. upon the completion of the investigation, a final report will be filed.
37674 | 1220648-2019-00202 | 9124435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9124435&pc=OZD | 08/29/2019 | 09/26/2019 |  | 09/26/2019 |  |  | 08/29/2019 | 09/26/2019 |  |  | 04/30/2021 |  | 09/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1428376 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/29/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp. the device was placed without incident, however, patient developed hemolysis that prompted early removal. subsequently, the hemolysis cleared and the patient did not require any further intervention. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37675 | 1220648-2019-00203 | 9126039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9126039&pc=OZD | 09/06/2019 | 09/26/2019 |  | 09/26/2019 |  |  | 09/06/2019 | 09/26/2019 |  |  | 12/31/2020 |  | 09/26/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1393820 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 09/06/2019 | * | Invalid/Incomplete |  |  | In (b)(6) a patient who was admitted with a vf (ventricular fibrillation) storm during an acute myocardial infarction. the patient had both ecmo and an iabp placed for support due to shock. when the care needed to be escalated, the iabp was removed and an impella 2.5 was placed via the right axillary artery. femoral access was not possible due to an aaa. when the patient condition improved the ecmo was removed and the impella remained as the sole support device. on the fourth day of support the impella was also explanted due to improved patient condition. months after explant the team notified abiomed that there had been a bleed at the impella access site, on the second day of support, which was treated with blood product replacement. | [5 empty] | To date further clinical details have not been shared regarding cause of the access site bleed. the product was discarded and so no product analysis is possible. should more details be shared from the complainant in (b)(6), a final medwatch with a root cause analysis will be filed. this failure mode will be monitored and trended.
37676 | 1220648-2019-00204 | 9126045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9126045&pc=OZD | 08/31/2019 | 09/26/2019 |  | 09/26/2019 |  |  | 08/31/2019 | 09/26/2019 |  |  | 04/30/2021 |  | 09/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1417280 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/31/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) caucasian patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp. post implant, the patient was transferred to a new hospital. upon arrival, the pump was found sitting too deep within the ventricle and was attributed to the patient's ambulance transfer. as treatment, positioning was adjusted. despite repositioning, the patient exhibited signs of hemolysis, via red urine and hemolyzed labs, that took approximately 12 hours to clear. the decision was made to remove the device and replace with an intra-aortic balloon pump. although hemolysis was alleged, no tests for plasma-free hemoglobin were measured. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37677 | 1220648-2019-00198 | 9107187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9107187&pc=OZD | 04/09/2019 | 09/23/2019 |  | 09/23/2019 |  |  | 08/27/2019 | 09/23/2019 |  |  | 10/31/2020 |  | 09/23/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1376920 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/27/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old (b)(6) male had impella 2.5 pump inserted in the left femoral. the patient had lower left leg ischemia, an antegrade sheath was inserted into the left thigh so blood could be supplied from right lower limb sheath to the left leg in order to reduce ischemia. | [5 empty] | The impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37678 | 1220648-2019-00199 | 9107194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9107194&pc=OZD | 08/28/2019 | 09/23/2019 |  | 09/23/2019 |  |  | 08/28/2019 | 09/23/2019 |  |  | 04/30/2021 |  | 09/23/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1418299 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 51 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/28/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old black or (b)(6) male patient had impella cp pump inserted in the right femoral for acute myocardial infarction (ami). when the impella was removed, the patient did not have much flow from arteriotomy. distal pulse in the foot was still intact at the time. as treatment, fasciotomy was done due to vascular injury. | [5 empty] | The impella cp pump was discarded by the customer therefore an evaluation of the device cannot be completed.
37679 | 1220648-2019-00200 | 9107224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9107224&pc=OZD | 09/01/2019 | 11/22/2019 |  | 09/23/2019 |  |  | 09/01/2019 | 09/23/2019 |  |  | 06/30/2021 |  | 09/23/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1430461 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | DISTRI |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pseudoaneurysm | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 09/01/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female patient had impella cp pump inserted in the right femoral for acute myocardial infarction (ami). the pump was removed and upon follow up the nurse stated the patient ended up having a pseudo aneurysm in the groin. as treatment the groin was injected with thrombin. | [5 empty] | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The introducer was not returned for evaluation, however the returned repositioning unit was inspected and no abnormalities were observed. the root cause of the suspected injury of pseudo aneurysm in the groin was unable to be determined due to the lack of clinical information.
37680 | 1220648-2019-00191 | 9097279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9097279&pc=PYX | 08/21/2019 | 11/07/2019 |  | 09/20/2019 |  |  | 08/21/2019 | 09/20/2019 |  |  | 02/28/2021 |  | 09/20/2019 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1415968 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 94 YR | [5 empty] | Injury | Yes | Yes | Material Fragmentation | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [5 empty] | IMPELLA CP | IMPELLA CP | [21 empty] | Initial submission | Followup |  |  |  | 08/21/2019 | * | Invalid/Incomplete |  |  | The complainant a (b)(6) year old male presenting with acute myocardial infarction and cardiogenic shock and right ventricle failure for hemodynamic support with the impella rp. upon inserting the device, the distal end of the pump fractured and separated from the device and remained in the pulmonary artery, upon insertion. as treatment, a snare was used to remove the broken component and ecmo was placed due to patient requiring oxygenation. | [5 empty] | The impella rp was received from the customer and a device evaluation is underway. a supplemental mdr will be filed with the results of the investigation, upon completion. | The device was returned with cage and pigtail severed. visual inspection of the introducer revealed scratches and a kink. evidence suggests the cause of the reported failure was related to calcification within the vessels and/or difficult vasculature/delivery. a device history review was conducted and found no manufacturing issues or other complaints related to this failure mode from this pump lot. due to lack of clinical details, a root cause cannot be conclusively determined.
37681 | 1220648-2019-00197 | 9097998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9097998&pc=OZD | 08/23/2019 | 09/20/2019 |  | 09/20/2019 |  |  | 08/23/2019 | 09/20/2019 |  |  | 04/30/2021 |  | 09/20/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1417620 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/23/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old (b)(6) female had impella cp pump inserted in the left femoral. the physician stated he was using the 14fr long sheath which might have damaged the vessels that might have caused the patient to have retroperitoneal hemorrhage at the insertion site. as treatment six (6) units of red cells concentrates (rcc) was administered by noradrenaline (nad) drip. | [5 empty] | Device has not been returned. the device was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37682 | MW5089963 | 9104397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9104397&pc=OZD | 06/24/2019 | 09/16/2019 |  | 09/20/2019 |  |  |  | 09/20/2019 |  |  | 03/31/2021 |  | 09/20/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SET HOSPITAL ITEM # 83682 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  |  |  | 203747A |  |  |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Voluntary report | [8 empty] | 66 YR |  |  | 99 |  |  | Injury |  | Yes | Difficult to Remove | [6 empty] | No Consequences Or Impact To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Impella was placed in ventricle over 018 impella wire, upon trying to remove wire it was extremely difficult to remove. wire when separated and was uncoiled upon removal. physician opted to remove entire impella system and used a different device. no significant injury noted to pt, separate balloon device was used and cath was successful following events. mfr's rep (b)(6) collected product.
37683 | 1220648-2019-00194 | 9087249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9087249&pc=OZD | 08/21/2019 | 11/07/2019 |  | 09/19/2019 |  |  | 08/21/2019 | 09/19/2019 |  |  | 04/30/2021 |  | 09/19/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  |  |  | 1418732 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 32 YR | [5 empty] | Malfunction | No | Yes | Contamination | [6 empty] | No Consequences Or Impact To Patient | [41 empty] | Initial submission | Followup |  |  |  | 08/21/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp. aftert approximately 2 days on support, patient had a large bowel movement. fecal matter breached the purge system and traveled to the device's bacterial filter, despite reports of a secure connections within the system. as treatment, the device was removed and a second impella was placed. the patient did not endure any adverse impact. | [5 empty] | The device was returned and an investigation was performed. visual inspection found fecal matter on the exterior of the impella, not in the fluid path. the reported allegation was not confirmed as the fluid path was inspected and found clear and free of fecal matter. | The impella cp was obtained by an abiomed field representative and is currently being shipped back for device evaluation. a supplemental mdr will be filed with the results of the investigation, upon completion.
37684 | 1220648-2019-00195 | 9087493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9087493&pc=OZD | 08/23/2019 | 11/27/2019 |  | 09/19/2019 |  |  | 08/23/2019 | 09/19/2019 |  |  | 05/31/2021 |  | 09/19/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1425224 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 08/23/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian female patient presenting for high risk percutaneous coronary intervention with the impella cp. signs of hemolysis were exhibited during use of the device (red urine), however, this was not clinically confirmed via plasma-free hemoglobin testing. repositioning attempts were not successful, therefore, device removal was required and a new impella device was inserted and operated without issue. | [5 empty] | The impella cp was received and an investigation is underway. a supplemental mdr will be filed with the results of the investigation upon completion. | The device was returned for investigation and the console's data logs were returned for analysis. a visual inspection found no damage, deficiencies, or abnormalities with the pump. hemolysis testing was conducted and the pump operated within specification. despite failing to reproduce the reported failure mode, multiple pump positioning alarms were identified within the data logs, suggesting the pump had been out of position for the majority of support.
37685 | 1220648-2019-00196 | 9088066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9088066&pc=OZD | 08/28/2019 | 09/19/2019 |  | 09/19/2019 |  |  | 08/28/2019 | 09/19/2019 |  |  | 06/30/2021 |  | 09/19/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | AACHEN, GERMANY |  |  | GM |  | 1428570 |  | 004-0032 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | USER F |  |  |  |  | 31 YR | [5 empty] | Injury | No | No | Insufficient Information | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/28/2019 | * | Invalid/Incomplete |  |  | A (b)(6) was admitted after arresting in his home and being shocked out of a vf arrest by ems. the team had intubated him in the field and restored his heart rhythm. upon admission to the medical facility he was taken to the cardiac catheterization lab and found to have coronary artery occlusion that necessitated angioplasty. the angioplasty would be done with hemodynamic support of an impella pump. the pump position was not optimal and the team replace the pump with a second impella cp pump. at the sheath exchange and placement of the repositioning unit there was blood loss. this blood loss prompted blood replacement in the form of 2 units. | [5 empty] | The investigation into the root cause of the access site bleeding is on-going. the root cause will be relayed to the fda if the complainant returns any product or further clinical report. to date no product or extensive clinical report has been shared. no vessel imagery has been shared. the cause of the in use event may remain unknown.
37686 | MW5089922 | 9098109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9098109&pc=OZD | 09/05/2019 | 09/05/2019 |  | 09/19/2019 |  |  |  | 09/19/2019 |  |  |  |  | 09/19/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 005042 |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 53 YR |  |  | 90 |  |  | Injury | No | Yes | Defective Component | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Motor current high, controller failure on impella pump 2.5 - device removed and replaced with a new one that worked. fda safety report id# (b)(4).
37687 | 1220648-2019-00192 | 9081023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9081023&pc=OZD | 06/30/2019 | 09/18/2019 |  | 09/18/2019 |  |  | 06/30/2019 | 09/18/2019 |  |  | 12/31/2020 |  | 09/18/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1398263 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Malfunction | No | Yes | No Apparent Adverse Event | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 06/30/2019 | * | Invalid/Incomplete |  |  | The complainant placed an impella cp to hemodynamically support a cardiomyopathy patient admitted with possible myocarditis. upon placement of the impella the pump flows did not rise to expected levels, and so the team explanted the pump and replaced it with a second impella. the second pump supported the patient successfully. upon return of the pump the low flow were not initially reproduced. with continued running of the pump the impeller blade fractured. upon review of the pump and imagery of it, the pump had been returned with a blade crack that fully fractured with lab use. | [5 empty] | The product returned was found to have a damaged impeller blade. the blade fully fractured in the analysis and lab runs. the crack and eventual fracture was characteristic of a hard foreign material getting caught between the outlet and impeller blade. the medical facility did not report any calcification or foreign material, or device, that could have caused the blade fracture. this failure mode will be monitored and trended.
37688 | 1220648-2019-00193 | 9081029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9081029&pc=OZD | 08/22/2019 | 11/01/2019 |  | 09/18/2019 |  |  | 08/22/2019 | 09/18/2019 |  |  | 07/31/2021 |  | 09/18/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1429791 |  | 0046-0023 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 41 YR | [5 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 08/22/2019 | * | Invalid/Incomplete |  |  | A (b)(6) male with a history of cardiomyopathy was admitted to a tertiary medical center after being transferred from the community. his condition was deteriorating and was in need of hemodynamic support. the tertiary center placed an impella 5.0 via surgical access with a cutdown to the right axillary artery. the team placed the pump successfully through the graft and began support. the site did have a bleeding complication. the pump's valve closest to the suture pad was bleeding and so that team applied an internal suture with a silk tie. this remedied the issue and support proceeded with minimal blood lost. | [5 empty] | Since the initial medwatch was filed, the investigation has completed. the pump was returned for analysis. the request for lab values and anticoagulation regimen was not answered. the root cause of the access site bleeding was unable to be determined. the pump was returned without any damage to it that could have caused or contributed to the bleed. the failure mode will be monitored and trended. | The investigation is on-going at this time. a final medwatch will be filed at the completion of the analysis.
37689 | 9081130 | 9081130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9081130&pc=OZD | 08/20/2019 | 08/28/2019 | 08/28/2019 | 09/18/2019 | 08/28/2019 | 09/18/2019 |  | 09/18/2019 |  |  |  |  | 09/18/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No answer provided | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 23360 DA | [5 empty] | Malfunction | No | Yes | Output Problem | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella catheter error. when plugged in it stated that it was a console failure. a second console was brought in and it stated the same thing. therefore we replaced the catheter and no failure was stated when plugged into the console. catheter was never inserted into the patient.
37690 | 1220648-2019-00189 | 9069546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9069546&pc=PYX | 08/21/2019 | 10/25/2019 |  | 09/16/2019 |  |  | 08/21/2019 | 09/16/2019 |  |  | 04/30/2020 |  | 09/16/2019 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1345236 |  | 004334 |  | Yes | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Perforation | Blood Loss | [11 empty] | Hospitalization | Required Intervention |  |  |  |  | IMPELLA 2.5 | [22 empty] | Initial submission | Followup |  |  |  | 08/21/2019 | * | Invalid/Incomplete |  |  | The patient was transferred into the tertiary medical center for care escalation. the patient was in need of left and right hemodynamic support for a planned coronary angioplasty. the impella rp was placed over the .027 guide wire. upon placement there was an observation of blood in the patient's endotracheal tube from an apparent wire dissection of the pulmonary artery. the team adjusted the rp pump level to allow the pressure to close the perforation. the patient also received blood transfusions, in the amount of 4 units of blood. the impella 2.5 was then placed for the left sided support and the angioplasty. | [5 empty] | Since the initial medwatch was filed, the investigation has completed. the impella rp patient had a bleed observed in the endotracheal tube, which was indicative of internal bleeding. a full investigation into whether the impella pump or guidewire had caused a perforation and blood loss could not be done. no product nor data logs were returned for analysis. the physician believes that the guidewire likely punctured the left pulmonary artery due to deep placement. the failure mode will be monitored and trended. | The medical team has yet to release the product for investigation. the product and data logs are expected to be returned for investigation. upon the completion of the investigation, a final medwatch will be filed.
37691 | 1220648-2019-00190 | 9072318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9072318&pc=OZD | 08/22/2019 | 09/16/2019 |  | 09/16/2019 |  |  | 08/22/2019 | 09/16/2019 |  |  | 11/30/2020 |  | 09/16/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1385057 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [28 empty] | Initial submission |  |  |  |  | 08/22/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella 2.5. during support, the patient endured a cerebrovascular accident as a result of a thrombus traveling to the patient's brain. the physician could not rule out the impella device as a causal relationship to this event. the patient was withdrawn from care due to a due not resuscitate order and expired. the physician ruled the patient's death as the result of the original disease and not related to the impella device. | [5 empty] | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37692 | 1035166-2019-00078 | 9072545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9072545&pc=DYB | 08/17/2019 | 01/10/2020 |  | 09/16/2019 |  |  | 08/21/2019 | 09/16/2019 |  |  | 06/01/2022 |  | 09/16/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-17243 |  | 0052-3000 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup | Followup |  |  |  | * | Invalid/Incomplete |  |  | It was reported that patient was presented for acute myocardial infarction and cardiogenic shock (ami/cgs), while inserting the introducer, there was difficulty reported and an access site bleed resulted. initial placement was difficult due to resistance on vascular level. (b)(6) 2012 f dilator have been used. blood transfusion and surgical repair was required to control the blood loss. approximately 1500 ml blood was loss. bleeding occurred upon arrival at icu, after insertion. physician cannot determine if the bleeding was caused by the introducer or the patient's anatomy. patient outcome is death. patient was withdrawn from care on (b)(6) 2019 due to multiple organ failure. death is unrelated to the device. device is discarded by user. | [5 empty] | Corrected section: b4,g4 in follow up 1 medwatch b4 and g4 date should be 01/08/2020 not 01/08/2019. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | The device used in treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure abiomed introducer sheath in-process and final inspections: visual inspection : using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. slight discoloration from tipping process is acceptable. verify proper tip shape according to drawing. insert a tipped dilator of appropriate size through the tipped sheath and check if dilator inserts without excessive force or obstruction. with the naked eye at a distance of 12" to 18", inspect where the sheath and dilator meet and check the gap is minimal and a smooth transition is present. per ifu: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. avoid subjecting the device to unusual stresses. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | The device used in treatment, the device was not returned for analysis, therefore, the clinical observation could not be confirmed. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure ab-51-001x-e abiomed introducer sheath in-process and final inspections: visual inspection : using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. slight discoloration from tipping process is acceptable. verify proper tip shape according to drawing. insert a tipped dilator of appropriate size through the tipped sheath and check if dilator inserts without excessive force or obstruction. with the naked eye at a distance of 12" to 18", inspect where the sheath and dilator meet and check the gap is minimal and a smooth transition is present. inspect 100% per ifu: never advance or withdraw guide wire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. avoid subjecting the device to unusual stresses. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
37693 | 1220648-2019-00187 | 9012305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9012305&pc=OZD | 08/17/2019 | 09/12/2019 |  | 09/12/2019 |  |  | 08/17/2019 | 09/12/2019 |  |  | 06/30/2021 |  | 09/12/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1425103 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | Yes | Material Fragmentation | Physical Resistance/Sticking | [5 empty] | Blood Loss | Device Embedded In Tissue or Plaque | [11 empty] | Other | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 08/17/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp. the physician reported resistance while inserting the 14frx 13cm oscor introducer which resulted in bleeding from the insertion site. following the placement of the introducer, the pump was inserted over the guide wire without issue, however, physician experienced resistance when removing guide wire, resulting guide wire breaking off inside the patient. the patient's bleeding was treated with a blood transfusion of at least 3 units. the guide wire fragment remained in the patient and the impella was used with positive hemodynamic support, however, patient expired after being withdrawn from care and the wire was not retrieved. the patient's cause of death was due to multiple organ failure and not related to the impella or the introducer. there was no abnormality or damage noted with either the wire or introducer, prior to usage. the patient's anatomy was noted as a non-factor with "good anatomical conditions." | [5 empty] | The impella cp, introducer, and guide wire were discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37694 | 1220648-2019-00188 | 9012338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=9012338&pc=OZD | 08/21/2019 | 09/12/2019 |  | 09/12/2019 |  |  | 08/21/2019 | 09/12/2019 |  |  | 12/31/2020 |  | 09/12/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE |  |  | GM |  | 1391180 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 08/21/2019 | * | Invalid/Incomplete |  |  | A (b)(6) year old male was admitted to a medical facility in (b)(6) in cardiac arrest. he was seen to be in ventricular fibrillation (vf) and recurring pulseless electrical activity (pea). he was treated with cpr and placed on ecmo and an iabp for hemodynamic and respiratory support. when this support was not optimal, the iabp was explanted and an impella 2.5 was placed. after over 10 days of support the team was noticing signs of hemolysis. specifically elevated ldh and decreased urine output. the impella was a probable cause, as was the ecmo. the team chose to explant the impella pump. upon explant a thrombus was observed to be attached to the pump. the team discarded the pump and the observation of the thrombus was not confirmed. the patient remained on support with the ecmo. | [5 empty] | The complainant in (b)(6) discarded the pump product. no testing of the product can be done. the investigation is on-going at this time. should more information be shared that leads to a root cause of the presumed hemolysis, a final report will be filed.
37695 | 1220648-2019-00186 | 8992260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8992260&pc=OZD | 08/15/2019 | 11/07/2019 |  | 09/11/2019 |  |  | 08/15/2019 | 09/11/2019 |  |  | 08/31/2020 |  | 09/11/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1368160 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Improper or Incorrect Procedure or Method | [6 empty] | Hematoma | Blood Loss | Vascular Dissection | [10 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 08/15/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white female presenting with acute myocardia infarction and cardiogenic shock for hemodynamic support. an impella cp device was inserted without any reported issues, however a hematoma formed and continued to swell throughout support. the device was explanted surgically, due to continued hematoma swelling and patient's high anticoagulation status. during removal, the vascular surgeon identified a dissection in the patient's artery as a result of steep angle insertion. as treatment, multiple units of blood products were delivered to treat blood loss from the hematoma and surgical repair of the vessel was performed. | [5 empty] | The impella cp was returned by customer and an investigation of device is underway. a supplemental mdr will be filed with the results of the investigation upon completion. | Since the initial report was filed, the investigation has completed. the returned pump was evaluated and the root cause of the vascular damage with subsequent access site hematoma was a high stick during access site preparation. there were no abnormalities observed with the returned device.
37696 | 1220648-2019-00185 | 8970524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8970524&pc=OZD | 08/11/2019 | 09/06/2019 |  | 09/06/2019 |  |  | 08/11/2019 | 09/06/2019 |  |  | 12/31/2019 |  | 09/06/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | CP C-7 NO ACCESS PUMP SET, US |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1327282 |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/11/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male had impella cp pump inserted. the patient had a thrombus that was confirmed by vascular surgery when the thrombectomy was completed. the surgeon had to surgically remove the thrombus. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37697 | 1220648-2019-00184 | 8967182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8967182&pc=OZD | 08/06/2019 | 10/10/2019 |  | 09/05/2019 |  |  | 08/06/2019 | 09/05/2019 |  |  | 12/31/2020 |  | 09/05/2019 |  | 2)TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1391180 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 51 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 08/06/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male presenting with cardiomyopathy for hemodynamic support with the impella 2.5 pump. the pump was inserted without any reported issues, however, patient exhibited signs of hemolysis via "reddish-brown" urine during support. as a treatment, the physician removed the device and hemolysis improved, thereafter. | The complainant reported additional information on 2019-09-17 revealing the patient had also endured an access site bleed. hemostasis was achieved successfully, however, 10 units of red blood cell delivery were required. |  |  |  |  | The impella 2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation. | New information received on 2019-09-17 pertaining to blood loss endured that required blood product delivery. new code added for blood loss.
37698 | 1220648-2019-00183 | 8949828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8949828&pc=OZD | 08/03/2019 | 08/30/2019 |  | 08/30/2019 |  |  | 08/04/2019 | 08/30/2019 |  |  | 04/30/2021 |  | 08/30/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1415058 |  | 0048-0034-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 28 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage, Cerebral | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/04/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male had impella cp pump inserted in the right axillary. the left ventricular thrombus was claimed to be absorbed by impella and part of the thrombus flowed to the brain causing a stroke. the inside of the impella, especially the inlet part was clogged up with thrombus, the suction alarm was not able to be cleared and the impeller was removed. the physician in charged commented that the left ventricular thrombus was a blood stagnation in the left ventricle (lv) due to dilated cardiomyopathy (dcm),that became a thrombus after impella insertion. anticoagulant therapy was provided to treat the cerebral infarction and the patient is improving. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37699 | 1035166-2019-00074 | 8946509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8946509&pc=DYB | 07/18/2019 | 08/29/2019 |  | 08/29/2019 |  |  | 08/06/2019 | 08/29/2019 |  |  |  |  | 08/29/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  |  |  | 0052-3000 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that the physician performed a femoral cut down to expose the artery and placed a suture at the site without placing a surgical graft. the introducer was then inserted into the incision and high risk percutaneous coronary intervention was initiated. during procedure, the access site bleed with an estimated blood loss of 1100 ml. as treatment, 4 units of red blood cells were administered. the introducer and impella lp2.5 were also removed. oozing identified from the sheath's access site. bleeding is related to cutting down the vessel for insertion of imeplla. procedure was completed with impella, but bleeding from access site of imeplla did not stop, so it was removed after pci. the condition of patient was stable. device was used in tortuous and calcified path. device is not available for evaluation. | [5 empty] | Device used in treatment. the device was not returned for evaluation. there was no specific performance related failure reported by the user. no further investigation is required. the lot number of this device was not provided, therefore neither a review of the device history record nor complaint history could be performed. inspection procedures require any oscor product pass all in-process and qa final inspections before shipping to the customer. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. no further follow-up is required.
37700 | 1220648-2019-00182 | 8946944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8946944&pc=OZD | 08/05/2019 | 08/29/2019 |  | 08/29/2019 |  |  | 08/05/2019 | 08/29/2019 |  |  |  |  | 08/29/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | LITERATURE | [11 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Thrombosis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 08/05/2019 | * | Invalid/Incomplete |  |  | A single-center publication was discovered that pertained to impella use from april 2015 to april 2018. in the publication there were 13 impella patients in which there was thrombosis at the axillary artery access. the impella pumps were both the cp and 5.0 pumps, the 5.0 was used in 2 case supports and 11 were the cp. the publication stated that 4 received thrombectomy and/or balloon angioplasty. the remaining 9 received solely medical management with heparin infusion. | [5 empty] | The publication was discovered by abiomed medical affairs. without product, data logs, or further patient specific clinical case reports the root cause of the thrombosis can not be discovered. no analysis is possible. the failure mode will be monitored and trended.
37701 | 1220648-2019-00180 | 8940678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8940678&pc=OZD | 07/31/2019 | 10/10/2019 |  | 08/28/2019 |  |  | 07/31/2019 | 08/28/2019 |  |  | 01/31/2021 |  | 08/28/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1400754 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | Hemolysis | Renal Failure | Hematuria | [9 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 07/31/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female had impella 2.5 pump inserted in the left femoral. the patient had hemolysis during pump support and was seen in hematuria values going up and was put on continuous hemodiafiltration (chdf) for renal function support. hemolysis persisted until explant of impella. | [5 empty] | The impella 2.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | New information was received that the patient had bacteremia at the insertion or other procedures (including from previous doctors). there was interventional treatment which was antibacterial drug administration, early impella removal, swan (sg) catheter removal, central venous (cv) line, and intubation tube replacement. the pump was returned for evaluation. there was no biomaterial found in the returned pump. the pump was run in a beaker of water and ran properly without any suction alarms. the returned pump was also tested for hemolysis and passed hemolysis testing. the pump was caught on the papillary muscle during the case but it is unknown if the pump was released from the papillary. hemolysis during the case may have been caused by improper positioning or by being caught on the papillary muscle. impella passed all sterilization inspection therefore, the cause of sepsis was unable to be determined since it is unknown when bacteria was introduced to the patient or if bacteria was transmitted by other equipment.
37702 | 1220648-2019-00181 | 8940807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8940807&pc=OZD | 05/13/2019 | 10/09/2020 |  | 08/28/2019 |  |  | 05/13/2019 | 08/28/2019 |  |  |  |  | 08/28/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED INC. |  |  | US |  | N/A |  |  |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report | [8 empty] | 41 YR | [5 empty] | Injury | No | No | Insufficient Information | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 05/13/2019 | * | Invalid/Incomplete |  |  | In an abiomed personnel's literature review, a case publication was discovered in which there was an adverse event. the publication cites a case in which the impella 5.0 was placed via transcaval access, rather than the usual femoral or axillary access. while on support the 5.0 was the presumed cause of the patient's limb ischemia. the 5.0 had been placed via a large 24fr sheath that was exerting pressure on the artery. the limb ischemia was treated with surgical intervention/revascularization. | [5 empty] | Correction: g(4): upon review of this complaint, it was determined the incorrect aware date was provided in the initial report. the aware date should be (b)(6)2019. | No corrective action is recommended at this time as the root cause was unable to be determined. there was a lack of information regarding this case. the only information available was that which was disclosed in the publication. failures of this type will continue to be monitored for trends. (b)(4).
37703 | 1220648-2019-00177 | 8934875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8934875&pc=OZD | 07/29/2019 | 11/11/2019 |  | 08/27/2019 |  |  | 07/29/2019 | 08/27/2019 |  |  | 04/30/2020 |  | 08/27/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1340134 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 07/29/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old asian male presenting with myocarditis for hemodynamic support with the impella lp5.0. during support, the patient became septic and required antibiotics as treatment. the physician alleged the impella device may have caused or contributed to this diagnosis. ultimately, the patient recovered and the device was removed. | [5 empty] | The device was returned for analysis with no abnormalities found. the device history record was reviewed and found the device had been properly sterilized. no manufacturing issues with this lot were identified. there is no indication that the root cause of the reported sepsis was associated with the impella. | Components from the impella lp5.0 were received and in investigation is underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37704 | 1220648-2019-00178 | 8936607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8936607&pc=OZD | 07/31/2019 | 10/18/2019 |  | 08/27/2019 |  |  | 07/31/2019 | 08/27/2019 |  |  | 03/31/2021 |  | 08/27/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1401373 |  | 005060 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 07/31/2019 | * | Invalid/Incomplete |  |  | The complainant reported a 56 years old white female had impella 5.0 pump inserted in the right axillary. sepsis was suspected during treatment and patient was given antibiotics. one physician believes that impella could be a possible cause of sepsis. | [5 empty] | The pump was returned for evaluation but it was cut at the 29cm mark and only the 29 hours of data was returned. hemolysis testing could not be conducted because the catheter was cut too short. the data did reveal outflow blocked alarms and suction alarm, the alarms coincided with the reported increase in hemolysis parameter. biomaterial was found in the cannula near the impeller, this was likely the reason for the outflow blocked and suction alarms. without hemolysis testing the failure mode could not be confirmed, therefore the cause of hemolysis was unable to be determined. sepsis was suspected during the case and was resolved with antibiotics. the physicians discussed theories that included impella as a cause. sterilization records were reviewed and no abnormalities were discovered. there is no evidence that would implicate impella as the cause of sepsis. | The impella 5.0 pump has not been returned by the customer. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37705 | 1220648-2019-00179 | 8936611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8936611&pc=OZD | 05/04/2019 | 10/09/2020 |  | 08/27/2019 |  |  | 05/04/2019 | 08/27/2019 |  |  | 02/28/2021 |  | 08/27/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1405497 |  | 0048-0032 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | USER F |  |  |  |  | 38 YR | [5 empty] | Injury | No | No | Insufficient Information | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [22 empty] | Initial submission | Followup |  |  |  | 05/04/2019 | * | Invalid/Incomplete |  |  | A (b)(6) old patient was admitted with cardiogenic shock and acute myocardial infarction during a weekend. access of the femoral artery for placement of an impella cp was complicated. multiple access attempts were made. once the pump was successfully placed, the patient was sent to the icu for monitoring. in the icu there was an observation made of diminished pulses, which is indicative of limb ischemia. the patient was taken to the operating department for exploration for a possible cutdown and axillary access placement of the impella. the femoral access was successfully preserved and the patient remained on support. | [5 empty] | The medical facility discarded the impella pump. the data logs were not retrieved, nor would they have pertained to the failure mode of limb ischemia. without the pump, further clinical information such as vessel diameter or vessel morphology, the root cause of the limb ischemia could not be determined. the field team member did report that the pump functioned according to specification during her presence in the case support. the failure mode will be monitored and trended. | Correction: g(4): upon review of this complaint, it was determined the incorrect aware date was provided in the initial report. the aware date should be (b)(6) 2019.
37706 | 1035166-2019-00072 | 8930300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8930300&pc=DYB | 07/18/2019 | 10/04/2019 |  | 08/26/2019 |  |  | 08/01/2019 | 08/26/2019 |  |  | 12/01/2021 |  | 08/26/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-15824 |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that patient was presented to physician for ejection fraction of 55% dropping to 15%. during high risk percutaneous coronary intervention, the physician made the decision to leave the peel away sheath in the patient's anatomy during support with the impella device. this was done to prevent risk of losing proper positioning of the impella. as a result, the patient's leg became ischemic and the introducer was removed. there was no particular malfunction reported with the device reported, just that patient endured ischemia due to introducer remaining in access site during impella support. to resolve ischemia, physician removed the introducer. introducer was removed and pump was ultimately removed. no new pump was inserted and patient continued support via inotropic medication. patient outcome is stable. 1 device was planned to introduce on the introducer. | [5 empty] | The device was used for treatment. there was no product performance issue reported. the device was not returned for evaluation,therefore, the clinical observation could not be confirmed. the investigation will focus on the documentation review. the device history records were reviewed to confirm that the devices passed all applicable in process and final inspections. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress, follow up report will be submitted if we find any further additional information.
37707 | 1220648-2019-00176 | 8931873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8931873&pc=OZD | 07/30/2019 | 08/26/2019 |  | 08/26/2019 |  |  | 07/30/2019 | 08/26/2019 |  |  | 03/31/2021 |  | 08/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1412907 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/30/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp. the device was inserted without issue, however, the patient developed an access site bleed. after unsuccessful attempts to achieve hemostasis using manual compression and a femostop, the physician made the decision to remove the device surgically and deliver 2 units of blood products. no injuries to the vessel were found. the physician alleged the device's repositioning sheath had not sealed the site effectively, leading to the bleed. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37708 | 1220648-2019-00174 | 8927502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8927502&pc=OZD | 07/28/2019 | 10/01/2019 |  | 08/23/2019 |  |  | 07/28/2019 | 08/23/2019 |  |  | 04/30/2021 |  | 08/23/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1421332 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 07/28/2019 | * | Invalid/Incomplete |  |  | At the medical center in (b)(6) a patient was admitted with cardiomyopathy. the impella cp was placed via the left femoral artery for hemodynamic support. upon removal of the introducer and advancement of the impella pump there was an observation made of an access site bleed. the team applied a pressure dressing but due to the significant amount of blood lost, the team infused back approximately 900-1200 cc's of blood in the form of 2 units of red blood cells. after the blood replacement, the patient was stable and the care plan was to transfer the patient to a tertiary center for possible escalation of care to an lvad. | [5 empty] | The investigation is on-going at this time. upon completion of the investigation a final medwatch will be filed. | Since the original medwatch was filed, the investigation has completed. the team in the eu did not return the impella pump or introducer sheath for investigation, only the purge cassette was returned. the purge cassette did not pertain to the failure mode of bleeding or vascular injury. the root cause of the bleeding is unknown. the team did not share further clinical details regarding the amount of blood lost, only an estimate was given of between 900 and 1200ml. the access site bleed was treated by pressure dressing and blood replacement products. the failure mode will be monitored and trended.
37709 | 1220648-2019-00175 | 8927865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8927865&pc=OZD | 08/01/2019 | 10/03/2019 |  | 08/23/2019 |  |  | 08/01/2019 | 08/23/2019 |  |  | 05/31/2021 |  | 08/23/2019 |  | TEMPORARY NONROLLER TYPED LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1421005 |  | 005040 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Cardiac Tamponade | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 08/01/2019 | * | Invalid/Incomplete |  |  | The customer in (b)(6) admitted a patient for cardiogenic shock and acute myocardial infarction. due to poor condition an impella 2.5 placed for hemodynamic support for the planned coronary angioplasty. the team had some issues repositioning the pump after removal of the introducer and placement of the pump's repositioning sheath. the team noted pump movement at the placement of the repositioning sheath. shortly afterwards the team noted a pericardial effusion. the team performed a periocardiocentesis to pull off a volume of 1.5 liters. the team assumed the pump caused the puncture but, the definitive cause was not known. | [5 empty] | The investigation is on-going at this time. at the conclusion of the investigation into the ventricular perforation, a final medwatch will be drafted. | Since the original medwatch was filed, the investigation has concluded. the team forwarded the pump product, imaging, and data logs for analysis. the pump was returned with damage to repositioning unit of the pump. this is indicative of the difficulty in positioning mentioned by the medical team. during the repositioning the pump could have moved position within the left ventricle, and caused the perforation, but this can not be definitively determined with the limited clinical information and imaging provided by the foreign medical team. the failure mode will be monitored and trended.
37710 | 1220648-2019-00173 | 8915291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8915291&pc=OZD | 07/24/2019 | 08/21/2019 |  | 08/21/2019 |  |  | 07/24/2019 | 08/21/2019 |  |  | 12/31/2020 |  | 08/21/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1393820 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/24/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6)year old asian male presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella lp2.5. during use of the device, approximately 4 days after insertion, the patient's access side developed a bleed. as treatment, multiple blood products were delivered to the patient and hemostasis was achieved successfully. | [5 empty] | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37711 | 8896322 | 8896322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8896322&pc=PBL | 07/19/2019 | 08/07/2019 | 08/07/2019 | 08/15/2019 | 08/07/2019 | 08/15/2019 |  | 08/15/2019 |  |  |  |  | 08/15/2019 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | PURGE CASSETTE; Y CONNECTOR; IMPELLA |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Yes | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 16790 DA |  |  | 86 |  |  | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Consequences Or Impact To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Patient had impella percutaneous left ventricular assist device (plvad) implantation in morning. 5% dextrose was set to run as impella purge solution (heparin had not been added to the 5% dextrose, yet). the purge solution leaked while running the impella console resulting in 1 to 1-1/2 hours that the patient / machine was not receiving the 5% dextrose. the representative was contacted, and a new replacement kit was sent. the faulty purge set-up was replaced with a new set-up and this new set-up performed as intended. no patient injury.
37712 | 1220648-2019-00172 | 8898456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8898456&pc=OZD | 07/20/2019 | 09/20/2019 |  | 08/15/2019 |  |  | 07/20/2019 | 08/15/2019 |  |  | 09/30/2020 |  | 08/15/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1374516 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 07/20/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old (b)(6) male had impella 5.0 pump inserted in the right axillary. hemolysis was confirmed after about five (5) hours after insertion, plasma free hemoglobins drawn twice within 24 hours, and as a result fresh frozen plasma (ffp) and blood erythrocyte component preservation solution mixture of mannitol, adenine, and phosphate (map) was administered. | [5 empty] | Since the initial report was filed, the investigation into the reported hemolysis in the japanese facility has concluded. the pump and data logs were returned for analysis. the data logs revealed the flows were in specification throughout support. a few occurrences of positioning alarms occurred but, each alarm event lasted under 5 minutes. the impella pump was returned with biomaterial adhered to it. the thrombotic burden was on the motor housing and inside the pump housing. when the biomaterial was removed and the pump was run in the testing lab the pump passed abiomed hemolysis testing. the root cause of the biomaterial is not known. the failure mode will be monitored and trended. | The impella 5.0 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37713 | 1220648-2019-00170 | 8895265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8895265&pc=OZD | 07/16/2019 | 08/14/2019 |  | 08/14/2019 |  |  | 07/16/2019 | 08/14/2019 |  |  | 10/31/2020 |  | 08/14/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1376920 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 51 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [28 empty] | Initial submission |  |  |  |  | 07/16/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella lp2.5 device. during support, a cerebral infarct was confirmed via a computerized tomography (ct) scan. there was no additional intervention aside from routine management and observation. the physician alleged a thrombus had passed through the impella device and into the brain, resulting in the reported cerebrovascular accident. the patient was successfully weaned off support, once hemodynamics stabilized. the patient remains in critical condition. | [5 empty] | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37714 | 1220648-2019-00171 | 8895275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8895275&pc=OZD | 07/19/2019 | 08/14/2019 |  | 08/14/2019 |  |  | 07/19/2019 | 08/14/2019 |  |  | 05/31/2020 |  | 08/14/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | CP C-7 NO ACCESS PUMP SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1343727 |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 46 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/19/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp device. during support, the patient exhibited signs of hemolysis via hemolysis blood test results. the device was repositioned, however, hemolysis continued. ultimately, the patient was delivered continuous renal replacement therapy to treat the suspected hemolysis. the patient's support continued, thereafter, until device was explanted. the patient had planned to receive a permanent left ventricular assist device, however, we were unable to gather further details after the device was explanted. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37715 | 1220648-2019-00167 | 8888951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8888951&pc=OZD | 05/08/2019 | 08/13/2019 |  | 08/13/2019 |  |  | 07/16/2019 | 08/13/2019 |  |  | 12/31/2020 |  | 08/13/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1391206 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/16/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male had impella 2.5 pump inserted in the right femoral without any problems. during pump support the patient had renal failure and was put on dialysis solely because of impella. the patient recovered and no mechanical support was needed. | [5 empty] | Device discarded by customer. the impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37716 | 1220648-2019-00168 | 8889472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8889472&pc=OZD | 07/18/2019 | 10/09/2019 |  | 08/13/2019 |  |  | 07/18/2019 | 08/13/2019 |  |  | 01/31/2021 |  | 08/13/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1400489 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 07/18/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male presenting for high risk percutaneous coronary intervention with urgent support using the impella cp. during support, patient displayed signs of hemolysis. although hemolysis was not confirmed through plasma-free hemoglobin testing, patient's labs were hemolyzed and drops in both hemoglobin and platelet count were identified. as treatment, blood products were delivered and the device was ultimately removed. | [5 empty] | The impella cp was returned and an investigation was completed. data logs were not returned for review. device testing was performed and unable to reproduce the reported failure mode of hemolysis. we are unable to definitively determine a root cause as the device operated within specification during lab testing. | The impella cp patient cable was returned by the customer and an investigation is underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37717 | 1220648-2019-00169 | 8889606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8889606&pc=OZD | 07/16/2019 | 08/13/2019 |  | 08/13/2019 |  |  | 07/16/2019 | 08/13/2019 |  |  | 01/31/2021 |  | 08/13/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | AACHEN, GERMANY |  |  | GM |  | 1397672 |  | 0048-0002 |  | Yes | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Hemostasis | Ischemia | Sepsis | [8 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/16/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male had impella cp pump inserted in the left femoral. on (b)(6) 2019 computerized tomography (ct) scans shows active bleeding at the insertion site into the leg and retroperitoneal, intervention of a vascular surgeon was needed to stop the bleeding. the impella cp pump was removed to fix the vascular lesion at the level of the insertion site. a minimum of two (2) blood concentrates (330 ml each) was administered before starting the surgical repair of the vessel at an hb level of 7g/dl and an additional four (4) blood concentrates (330 ml each) after surgical closure of the vessel and removal of the hematoma. large amounts of fluids, catecholamines because of sepsis, and antibiotic therapy was also given as treatment. the patient was in hemorrhagic shock, developed a septic component, and compartment syndrome of the left upper leg and the patient expired in the intensive care unit (icu) on (b)(6) 2019. | [5 empty] | Device not returned. the impella cp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37718 | 1035166-2019-00069 | 8880838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8880838&pc=DYB | 07/03/2019 | 10/31/2019 |  | 08/09/2019 |  |  | 07/15/2019 | 08/09/2019 |  |  | 04/01/2022 |  | 08/09/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-17047 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN |  |  |  |  | 84 YR | [5 empty] | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Consequences Or Impact To Patient | [41 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that patient was presented to physician for 3 vessel disease, ejection fraction <20%. during high risk percutaneous coronary intervention physician was removing 14fr dilator and bleeding was discovered at the introducer's valve. to resolve bleeding, the decision was made to advance the impella cp into the introducer. approximately 20ml blood was lost and no blood products or surgical intervention required to control the blood loss. repositioning sheath was inserted into the introducer. thereafter, manual compression was applied for 20 minutes. physician initially thought there was an arterial puncture, however, advised after the intervention that it was not a complication nor a vessel damage. micro puncture w/echo technique used for insertion. procedure completed successfully by introducing the repo unit into the 14fr introducer. there was no patient harm reported. | [5 empty] | The device was used in treatment. one 14f 13 cm long abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. visible blood was found on the sheath and inside the sideport tubing. upon initial inspection, the hemostatic valve was observed to be damaged. the valve was inspected in more detail under 10x magnification. the damage was found to be a tear that stemmed off of the valve split line. the introducer was first manually tested by introducing water through sideport of the sheath using a syringe. the introducer did not leak. second leak test was performed as per iso liquid leakage test standard. the introducer failed the leak test at 200kpa with leakage from the hemostatic valve. the cause of this failure could be potentially related to the damage observed to the hemostatic valve when checked under 10x magnification. returned device analysis reveals one 14f abiomed introducer sheath that is within manufacturing specifications. the tear in the hemostatic valve is the cause of the leakage/bleeding. the probable root cause of the tear could be off center dilator/device insertion during use out in the field. the reason for return was not confirmed, as no manufacturing defects were found. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure adelante s2s introducer sheath in-process and final inspection: vacuum leak test - perform pressure and vacuum leak test per procedure, latest revision. this test is performed by qa on 100% of the product. per procedure introducer set, silicone seal slitting: the operator is to visually inspect all the slit seals 100% with a microscope. per qa procedure leak test - introducer/destino with seal- electronic tester : leak test is performed on 100% of the introducer sheaths per ifu :precautions: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. directions for use: slowly retract the guidewire and dilator, leaving the sheath in position. the hemostasis valve will reduce the loss of blood and the inadvertent aspiration of air through the sheath. introduce the catheter through the hemostasis valve/sheath and advance it into position. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress ,follow up report will be submitted if we find any further additional information .
37719 | 1220648-2019-00165 | 8873337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8873337&pc=OZD | 07/11/2019 | 10/04/2019 |  | 08/08/2019 |  |  | 07/11/2019 | 08/08/2019 |  |  | 11/30/2020 |  | 08/08/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED. INC. |  |  | US |  | 1397086 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 46 YR | [5 empty] | Injury | Yes | Yes | Break | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 07/11/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old asian male presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella lp5.0. the device was inserted and operated without any reported problems. after successful support was achieved, the physician experienced resistance when removing the device through the surgical graft, followed by bleeding. in order to prevent further bleeding, the physician used forceps to pinch graft and pull the device, however, the impella became separated at the inlet cage. surgical intervention was required to remove the remainder of the impella device. | [5 empty] | The impella lp5.0 was returned and in investigation is currently underway. a supplemental mdr will be filed with the results of the investigation, upon completion. | The impella 5.0 was received. visual inspection confirmed the reported detachment. the side inlet was found bent and cage struts were deformed; evidence of the forceps being clamped on the proximal portion of the inlet during removal. evidence of excessive applied force was identified as the device's cannula was deformed and kinked. the investigation suggests the root cause of the reported issue was related to improper technique by using forceps for removal of the pump.
37720 | 1220648-2019-00166 | 8873340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8873340&pc=OZD | 07/12/2019 | 08/08/2019 |  | 08/08/2019 |  |  | 07/12/2019 | 08/08/2019 |  |  | 01/31/2021 |  | 08/08/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1412345 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/12/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old (b)(6) male had impella 5.0 pump was inserted in the right axillary. after the impella was explanted the patient had mitral regurgitation (mr) and as treatment mitra clips procedure was performed. | [5 empty] | Device discarded by customer. the impella 5.0 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37721 | 1220648-2019-00161 | 8865870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8865870&pc=OZD | 07/20/2019 | 08/06/2019 |  | 08/06/2019 |  |  | 07/20/2019 | 08/06/2019 |  |  | 01/31/2021 |  | 08/06/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED INC. |  |  | US |  | 1385796 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 58 YR | [5 empty] | Malfunction | No | Yes | No Apparent Adverse Event | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission |  |  |  |  | 07/20/2019 | * | Invalid/Incomplete |  |  | The patient was admitted for an electrophysiology procedure. the impella cp was placed for hemodynamic support during the planned ablation. after two hours of successful support, there was an observation made of suction and pump flows decreased. the team attempted to troubleshoot the pump by adjusting the pump level and checking that the pump was in correct position within the left ventricle. when troubleshooting failed to rectify the low flows, the pump was explanted and the team replaced the pump. when the pump was returned for a complaint investigation there was an observation made of broken impeller blades. the pump damage was not observed by the customer, rather in bench top testing. | [5 empty] | The impella pump was returned to investigate the low flows and suction alarms during the support. upon return the impeller blades were found to be broken. the cause of the blade damage can not be confirmed. the pump was returned, but no data logs. clinical details were shared. there was no reported tavr device that could have interacted with the impeller to cause the damage. the root cause of the blade breakage is unknown. the failure mode will be monitored and trended.
37722 | 1220648-2019-00162 | 8865991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8865991&pc=OZD | 03/15/2018 | 08/06/2019 |  | 08/06/2019 |  |  | 07/08/2019 | 08/06/2019 |  |  | 09/30/2018 |  | 08/06/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1264525 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/08/2019 | * | Invalid/Incomplete |  |  | We received a publication from hindawi, case reports in cardiology, volume 2019, article id 4591250, 6 pages that a (b)(6) years old caucasian male had impella 5.0 pump inserted in the left ventricle where the impella cp was inserted via the left femoral artery. the patient developed profound intravascular hemolysis, transient complete heart block, and pulseless electrical activity requiring cardiopulmonary resuscitation. a repeat echocardiogram revealed small left ventricle (lv) and right ventricle (rv) cavity sizes. the impella 5.0 caused suction of endocardial tissue, leading to hemolysis and increased vagotonia. this improved with volume resuscitation and blood transfusion with reduction in flow rate. the impella 5.0 was explanted in the operating room (or). | [5 empty] | Device evaluated by mfr: the device was not returned. the impella 5.0 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37723 | 1220648-2019-00163 | 8866066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8866066&pc=OZD | 03/14/2018 | 08/06/2019 |  | 08/06/2019 |  |  | 07/08/2019 | 08/06/2019 |  |  | 12/31/2019 |  | 08/06/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA RP PUMP |  |  | OZD |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1319854 |  | 004334 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 61 YR | [5 empty] | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/08/2019 | * | Invalid/Incomplete |  |  | We received a publication from hindawi, case reports in cardiology, volume 2019, article id 4591250, 6 pages that a (b)(6) caucasian male had impella rp pump inserted in the right femoral vein. on day 3, the rp pump was explanted uneventful and surgical repair of the right common femoral artery was uneventful, despite confirmation of antegrade flow beyond the femoral artery, fasciotomy of the anterior compartment of the right leg was performed because of elevated compartment pressure and clinical evidence of limb ischemia. there was evidence of hematoma and myonecrosis in the right thigh, which was treated with resection of the affected muscle. the patient slowly recovered and was discharged to inpatient rehabilitation on hospital day 30. | [5 empty] | The device was not returned. the impella rp pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37724 | 1220648-2019-00164 | 8866977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8866977&pc=OZD | 09/16/2019 | 09/18/2019 |  | 08/06/2019 |  |  | 07/11/2019 | 08/06/2019 |  |  | 09/30/2020 |  | 08/06/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1374528 |  | 002162-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 37 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 07/11/2019 | * | Invalid/Incomplete |  |  | The complainant reported a 37 years old japanese male had impella 5.0 pump inserted in the left axillary. the patient developed an infection. blood culture test was done, the result was +3, the bacteria was detected. the same (b)(6) was detected on the (b)(6) blood cultures and as a result vancomycin antibiotic was given. | [5 empty] | Device not returned. the impella 5.0 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed. | The 5.0 pump was returned for investigation. the cause of access site infection was unable to be determined. sterilization records were reviewed and no abnormalities were discovered. the pump passed all sterilization. there is no evidence that would implicate impella as the cause of infection.
37725 | 1220648-2019-00160 | 8861697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8861697&pc=OZD | 07/10/2019 | 09/10/2019 |  | 08/05/2019 |  |  | 07/10/2019 | 08/05/2019 |  |  | 02/28/2021 |  | 08/05/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1405012 |  | 005048-JP |  | Yes | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 07/10/2019 | * | Invalid/Incomplete |  |  | The team in (b)(6) placed an impella 2.5 in a critical patient with an aortic dissection and an inability to come off of bypass in the operating suite. the pump supported the patient for five days. during the support the patient had signs of hemolysis. the hemolysis was treated with dialysis and blood product infusion. | [5 empty] | The impella pump and data logs were returned from japan for analysis. since the original report was filed, the investigation has completed. the data logs revealed suction alarms that coincided with the noted low patient volume and atrial fibrillation. the pump was analyzed and was found to have biomaterial attached that could have contributed to the observed signs of hemolysis. the biomaterial was removed and the pump was run. when the pump was run in benchtop testing the pump passed abiomed internal hemolysis testing. the observed signs of hemolysis were treated with blood product infusion, however the root cause could not be determined. patient condition could not be ruled out. the failure mode will be monitored and trended. | The complainant in (b)(6) has returned product and the investigation is on-going at this time. upon completion, a final medwatch will be drafted and submitted.
37726 | 1220648-2019-00159 | 8850098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8850098&pc=OZD | 02/20/2017 | 10/09/2020 |  | 08/01/2019 |  |  | 04/24/2019 | 08/01/2019 |  |  |  |  | 08/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED UNKNOWN |  |  |  |  | UNKNOWN |  | UNKNOWN |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 04/24/2019 | * | Invalid/Incomplete |  |  | We received a publication from clinical toxicology, 2017 vol. 55, no. 9, 1014-1015 that a (b)(6)years old female had impella 2.5 pump inserted. the patient had impella in place for 11 hours, while removing the device a minor femoral arterial tear was noted and the patient subsequently underwent successful surgical vascular repair. | [5 empty] | The device was not returned. the impella 2.5 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. | Correction: g(4): upon review of this complaint, it was determined the incorrect aware date was provided in the initial report. the aware date should be (b)(6)2019.
37727 | 1220648-2019-00158 | 8823752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8823752&pc=OZD | 06/01/2019 | 07/24/2019 |  | 07/24/2019 |  |  | 06/17/2019 | 07/24/2019 |  |  |  |  | 07/24/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED INC. | [6 empty] | UNKNOWN |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | OTHER | [11 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Stroke/CVA | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 06/25/2019 | * | Invalid/Incomplete |  |  | A user initiated medwatch (mw5087217) was received via fda communication on june 17, 2019. a thorough review of case reports was done to locate a similar case report from our field representatives. no similar case report was found. the user documented a patient of unknown race and age was admitted on (b)(6) 2019, with myocardial infarction and a low cardiac output. the team placed an impella for hemodynamic support via the femoral artery. there was bleeding from the access site and so, the patient was taken to the operating room for surgical repair and blood product infusion. after the surgery the patient experienced a neurological event from a thrombosis. the cause of the stroke was thought to have been from a clot embolized by the use of impella. | [5 empty] | Despite all due diligence, the event and report from the user can not be investigated. the user is unknown, the case report was not located, and therefore no root cause for the access site bleed or neurovascular event could be determined. these types of failures will be monitored and trended.
37728 | 8815340 | 8815340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8815340&pc=PBL | 06/27/2019 | 07/01/2019 | 07/01/2019 | 07/23/2019 | 07/01/2019 | 07/23/2019 |  | 07/23/2019 |  |  |  |  | 07/23/2019 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA 5 |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 25915 DA |  |  | 78 |  |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Consequences Or Impact To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella 5 device leaking scant amount of purge fluid from cassette area onto floor. charge nurse notified. cassette found to be cracked, changed and purge fluid bag changed. no issues with flows pre or post event. no patient harm.
37729 | 1220648-2019-00156 | 8812216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8812216&pc=OZD | 01/10/2019 | 07/22/2019 |  | 07/22/2019 |  |  | 06/25/2019 | 07/22/2019 |  |  | 10/31/2020 |  | 07/22/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1375181 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/25/2019 | * | Invalid/Incomplete |  |  | The complainant reported a ((b)(6) years old (b)(6) male had impella 2.5 pump inserted in the right axillary. the patient had hemolysis during pump support and was treated with 20 units of blood transfusion. | [5 empty] | Device discarded by customer. the impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37730 | 1220648-2019-00157 | 8812252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8812252&pc=OZD | 04/01/2019 | 07/22/2019 |  | 07/22/2019 |  |  | 06/24/2019 | 07/22/2019 |  |  | 11/30/2020 |  | 07/22/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1378913 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/24/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old (b)(6) male had impella 2.5 pump inserted in the right femoral. the patient had cardiac tamponade during pump support and was treated with pericardial puncture. | [5 empty] | Device evaluated by mfr: device discarded by customer. the impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37731 | 1220648-2019-00151 | 8808012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8808012&pc=OZD | 06/20/2019 | 07/19/2019 |  | 07/19/2019 |  |  | 06/20/2019 | 07/19/2019 |  |  | 02/28/2021 |  | 07/19/2019 |  | : TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1405497 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/20/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male had impella cp pump inserted in the right femoral. then unable to doppler distal pulse on the right leg so antegrade sheath was placed in the right femoral artery (rfa) to aid in perfusion to the impella leg, the patient did have vein harvested from that leg earlier in the day, the leg was wrapped with an ace wrap, initially no pulses were dopplerable but after 6fr antegrade sheath placed to rfa the patient¿s pulses were dopplerable, dorsalis pedis (dp) absent. the patient had ischemia and as a result the patient¿s leg was removed a couple days after impella was explanted. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37732 | 1220648-2019-00152 | 8809168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8809168&pc=OZD | 05/28/2019 | 07/19/2019 |  | 07/19/2019 |  |  | 06/24/2019 | 07/19/2019 |  |  | 11/30/2020 |  | 07/19/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1383954 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Hematuria | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/24/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old japanese male had impella 2.5 pump inserted in the left femoral. the patient had hematuria and as a result 34 units of red blood cells (rbcs), 40 units of fresh frozen plasma (ffp), and 20 units of platelets was administered. the impella was explanted due to ischemia and intra-aortic balloon pump (iabp) was inserted. | [5 empty] | Device discarded by customer. the impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37733 | 1220648-2019-00153 | 8809178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8809178&pc=OZD | 06/23/2019 | 07/19/2019 |  | 07/19/2019 |  |  | 06/23/2019 | 07/19/2019 |  |  | 12/31/2020 |  | 07/19/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1396886 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 57 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/23/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male had impella cp pump inserted in the left femoral. the patient had hemolysis as plasma-free hemoglobin (pfhgb) was over 1000 and as treatment 8 units of packed red blood cells (prbc) and 4 fresh frozen plasma (ffp) were administered. | [5 empty] | The impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37734 | 1220648-2019-00154 | 8809180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8809180&pc=OZD | 06/24/2019 | 07/19/2019 |  | 07/19/2019 |  |  | 06/24/2019 | 07/19/2019 |  |  | 11/30/2020 |  | 07/19/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1378913 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 64 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | [12 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 06/24/2019 | * | Invalid/Incomplete |  |  | The complainant reported a patient enduring a cerebrovascular accident, during hemodynamic support with the impella device, which lead to the patient's death. the physician alleged a causal relationship between the device and the patient's death. | [5 empty] | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37735 | 1220648-2019-00155 | 8809183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8809183&pc=OZD | 06/24/2019 | 09/11/2019 |  | 07/19/2019 |  |  | 06/24/2019 | 07/19/2019 |  |  | 12/31/2020 |  | 07/19/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1387850 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Material Separation | Improper or Incorrect Procedure or Method | [5 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 06/24/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient presenting for high risk percutaneous coronary intervention (hrpci). the impella lp2.5 was selected for hemodynamic support during this procedure. the physician attempted to deliver the device into the patient's anatomy without the required 0.018" guide wire. subsequently, the device kinked between the inlet and outlet cages. attempts to unkink the device resulted in catheter separating inside the patient. as treatment, surgical removal of the remaining components was performed and patient continued hrpci with va ecmo support. the patient remained in stable condition. | [5 empty] | The device was returned and a thorough investigation was completed. damage to the distal portion of the cannula was confirmed through visual inspection and pictures provided. it is unknown how or when the rupture ultimately occurred, however, evidence suggests it was likely facilitated by the wireless insertion process. due to the lack of clinical information and imaging of the initial kink that occurred inside the patient's body, the root cause of the cannula separation cannot definitively be determined. a review of the dhr was conducted and found no manufacturing issues or reworks associated with this pump lot. | The impella cp was returned by the customer and an investigation of device is underway. a supplemental mdr will be filed with the results of the investigation.
37736 | 1035166-2019-00059 | 8805193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8805193&pc=DYB | 06/17/2019 | 08/29/2019 |  | 07/18/2019 |  |  | 06/19/2019 | 07/18/2019 |  |  |  |  | 07/18/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 23F INTRODUCER KIT FOR IMPELLA |  |  | DYB | [5 empty] | OSCOR INC. |  |  | US |  | UNK |  |  |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | No Code Available | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The representative reported the patient expired on (b)(6)2019 , they think the family withdrew care. the representative does not have the pump or introducers to return to corporate. as far as the cause of death, the representative does not know what the official cause of death was determined to be. there were multiple physicians on the case and they are not sure which physician made the final decision on the cause of death. | On (b)(6) 2019, the patient a (b)(6) year old, hispanic, female was presented to physician in shock post cabg (coronary artery bypass grafting). during placement of impella cp via the femoral artery the physician placed the 23fr in error to the femoral artery, rather than the 14fr introducer. patient's abdomen was rigid in the cath lab, decision was made to stabilize patient first then address the abdomen. the vascular surgeon came and the vessel was repaired by placing 3 self-expanding stents to the iliac. investigation is still in progress; additional information has been requested. |  |  |  |  | The following sections were updated in follow-up 1: b4, b5, g4, g7, h2, h6, and h10: the device was used in treatment. the device was not returned for evaulation. there was no specific performance related failure reported by the user. the lot number of this device was not provided, therefore neither a review of the device history record nor complaint history could be performed. inspection procedures require oscor product to pass all in-process and qa final inspections before shipping to the customer. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete.
37737 | 1220648-2019-00118 | 8799219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8799219&pc=OZD | 02/12/2019 | 02/22/2019 |  | 07/17/2019 |  |  | 02/22/2019 | 07/17/2019 |  |  | 09/30/2020 |  | 07/17/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1370748 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 27 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Other | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old (b)(6) female had impella 2.5 pump inserted in the left femoral. the impella 2.5 pump was inserted for cardiac support, however, air was found in the patient's left ventricle when the device was turned on. no other adverse effects or additional intervention was reported. | [5 empty] | The root cause of the air in the left ventricle was unable to be determine without further clinical details.
37738 | 1220648-2019-00149 | 8799986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8799986&pc=OZD | 06/23/2019 | 09/11/2019 |  | 07/17/2019 |  |  | 06/23/2019 | 07/17/2019 |  |  | 02/28/2021 |  | 07/17/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1413259 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Contamination | [6 empty] | Septic Shock | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 06/23/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patent presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support with the impella cp. the device was inserted without issue, however, patient endured a large liquid bowel movement that pooled around the impella device and purge system, breaching the purge system and pump. the rn staff denied any tubing disconnection. the patient's white blood cell count rose and patient endured septic shock, thus, prophylactic antibiotics were administered as treatment and pump was removed. | [5 empty] | The impella cp was received and an investigation is underway. a supplemental mdr will be filed with the results of the investigation upon completion. | The impella cp was returned by the customer and an investigation was completed. visual inspection of the pump confirmed fecal matter on the exterior surfaces of both the pump handle and the side arms, however no fecal matter was observed inside the purge lumen or pressure lumen. data logs were analyzed and confirmed the pump had operated to specification, with no change in purge flow. the root cause of the patient injury was unable to be definitively determined as the patient had pre-existing sepsis and no fecal matter was observed on the interior of the purge lumen or pressure lumen.
37739 | 1220648-2019-00150 | 8799996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8799996&pc=OZD | 06/22/2019 | 09/11/2019 |  | 07/17/2019 |  |  | 06/22/2019 | 07/17/2019 |  |  | 07/31/2020 |  | 07/17/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1364190 |  | 005042 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 06/22/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male patient presenting post coronary artery bypass surgery with acute myocardial infarction and cardiogenic shock. the impella lp2.5 was selected for hemodynamic support and inserted without any issues reported. during support, the patient lost distal pulse and required device removal. post removal, the site was bled back in order to remove a suspected thrombus. the patient's leg became "mottled" and an arterial ultra-sound was performed which confirmed presence of a thrombus. as treatment a thrombectomy and fasciotomy were performed. patient remained in stable condition and did not require a second pump. | [5 empty] | The device was returned and investigation was completed. visual inspection identified biomaterial on the outflow cage of the pump. a review of the dhr was conducted and found no manufacturing issues or reworks associated with this pump lot. the root cause of the reported leg ischemia was attributed to thrombus formation at the access site, consistent with a lack of systemic heparin administration. | The impella lp2.5 was received and an investigation is underway. a supplemental mdr will be filed with the results of the investigation, upon completion.
37740 | 1220648-2019-00146 | 8790923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8790923&pc=OZD | 02/09/2019 | 07/15/2019 |  | 07/15/2019 |  |  | 06/18/2019 | 07/15/2019 |  |  | 07/31/2020 |  | 07/15/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1359038 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/18/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old (b)(6) male had impella 5.0 pump inserted in the right axillary. the patient had hemolysis during pump support and as a result eight (8) units of red blood cells (rbcs) was administered and also impella was repositioned to resolve the hemolysis. | [5 empty] | Device discarded by customer. the impella 5.0 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37741 | 1220648-2019-00147 | 8790947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8790947&pc=OZD | 06/17/2019 | 07/15/2019 |  | 07/15/2019 |  |  | 06/17/2019 | 07/15/2019 |  |  | 01/31/2020 |  | 07/15/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1331370 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 45 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/17/2019 | * | Invalid/Incomplete |  |  | The complainant reported a patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella lp5.0. the device was inserted in the right axillary artery without and reported issues. during treatment, the patient endured signs of hemolysis. no plasma-free hemoglobin levels were drawn to confirm hemolysis, however, multiple units of blood products were delivered and the device flow level was reduced, as treatment. | [5 empty] | The impella lp5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37742 | 1220648-2019-00148 | 8791105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8791105&pc=OZD | 06/18/2019 | 08/23/2019 |  | 07/15/2019 |  |  | 06/18/2019 | 07/15/2019 |  |  | 02/28/2021 |  | 07/15/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1410137 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 80 YR | [5 empty] | Malfunction | No | Yes | Detachment of Device or Device Component | Infusion or Flow Problem | No Apparent Adverse Event |  |  |  |  | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | Followup |  |  |  | 06/18/2019 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp. the device was inserted into the patient's left femoral artery without and reported issues. the device failed to achieve flows over 2.8 l/min, despite adequate filling pressures and multiple repositioning attempts. despite flow issue, successful hemodynamic support was achieved without replacing this device with no adverse impact to the patient. upon receipt and inspection of the device, damage to the device's impeller blades was identified. | [5 empty] | Correction: pma number was erroneously placed in the "ind number" section, when it should be placed in the "pma/501(k) number" section. entered pma number in it's correct spot. the impella cp and data logs were returned for evaluation. the device was visually inspected and noted to have a broken blade. review of the device's data confirmed multiple suction alarms with a momentary motor current spike. thereafter, the pump did not ramp up linearly as intended. simulated use testing in a basin of water successfully reproduced low pump flows reported. a review of the dhr was conducted and found no manufacturing issues or reworks associated with this pump lot. there are no other complaints related to this failure mode associated with pumps in this lot. the root cause of the low flow was a broken impeller blade. it is unknown how the blade was damaged. | The impella cp was received and an investigation is in process. a supplemental mdr will be filed with the results of the investigation.
37743 | 1035166-2019-00058 | 8786902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8786902&pc=DYB | 05/21/2019 | 09/12/2019 |  | 07/12/2019 |  |  | 06/13/2019 | 07/12/2019 |  |  | 01/01/2022 |  | 07/12/2019 |  | NTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC |  |  | US |  |  |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 64 YR | [5 empty] | Injury | Yes | Yes | Difficult to Insert | [6 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that patient as presented to physician with chest pain. patient went into ventricular fibrillation and crashed onto ecmo. during acute myocardial infarction / cardiogenic shock procedure physician was inserting the introducer into the patient right femoral artery, while inserting the introducer the physician experienced "a lot of resistance. peel away sheath removed and tapered sheath inserted into the groin. excessive bleeding from insertion site was identified. a hematoma developed around the insertion site. no blood products were required. manual pressure and sutures at site were applied. a new introducer was ultimately used and inserted without issue. it's unknown how much blood was lost. patient outcome is no harm. | [5 empty] | Our investigation is still in progress, follow up report will be submitted if we find any further additional information. | The device was used in treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. the investigation will focus on a review of product documentation. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure adelante s2s introducer sheath in-process and final inspection: visual inspection: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm. first 5 consecutive properly tipped parts, inspect 100% under 10x magnification verify: the tip is round. verify proper tip shape according to drawing. slight discoloration from tipping process is acceptable. no split score lines (inspect both score lines) along sheath body, cracks, gouges, tears. insert a tipped dilator of appropriate size through the hub of the sheath and check if dilator inserts without excessive force or obstruction. with the naked eye at a distance of 12" to 18", inspect where the sheath and dilator meet and check the gap is minimal and a smooth transition is present. per ifu: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. avoid subjecting the device to unusual stresses. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
37744 | 1220648-2019-00144 | 8781867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8781867&pc=OZD | 06/17/2019 | 07/11/2019 |  | 07/11/2019 |  |  | 06/17/2019 | 07/11/2019 |  |  | 12/31/2020 |  | 07/11/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1397692 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Insufficiency, Valvular | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | PCPS | [22 empty] | Initial submission |  |  |  |  | 06/17/2019 | * | Invalid/Incomplete |  |  | A patient was admitted to a community hospital in (b)(6), though as the patient condition deteriorated the patient was transferred to a tertiary hospital. upon transfer, the patient had the iabp (intra-aortic balloon pump) and pcps (percutaneous cardiopulmonary support) escalated to the placement of an impella. the impella 2.5 replaced the iabp. the impella supported the patient for 8 days. after removal of the impella the team observed aortic regurgitation. the patient was too ill to have the valve repaired or replaced. the echo revealed the valve leaflets to have prolapsed. | [5 empty] | The complainant in (b)(6) did not return the product nor share extensive clinical information or diagnostic imaging. the root cause of the aortic regurgitation could not be determined. the failure mode will be monitored and trended.
37745 | 1220648-2019-00145 | 8782559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8782559&pc=OZD | 09/13/2018 | 07/18/2019 |  | 07/11/2019 |  |  | 06/07/2019 | 07/11/2019 |  |  | 01/31/2020 |  | 07/11/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1324747 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 06/07/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old (b)(6) male had impella 2.5 pump inserted in the left femoral. the patient had acute renal failure due to hemolysis. the physician wanted to use impella to unload but hemolysis was getting worse, patient was hemodynamically stable, so the physician removed the device. the patient was put on dialysis solely for the signs of hemolysis attributed to impella use. | [5 empty] | Device discarded by customer. the impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed. | This was a duplicated report that was previously reported under 1220648-2019-00061.
37746 | 1220648-2019-00141 | 8775459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8775459&pc=OZD | 06/18/2019 | 08/29/2019 |  | 07/10/2019 |  |  | 06/25/2019 | 07/10/2019 |  |  | 01/31/2021 |  | 07/10/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1400754 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Dissection | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 06/25/2019 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) years old (b)(6) male had impella 2.5 pump inserted in the left femoral. emergency contrast-enhanced computerized tomography (ct) shows blood accumulation in the pericardium causing rupture of aortic dissection. patient was taken to the intensive care unit (icu) after the ct and stent graft placement was considered but the blood vessel tissue was severely damaged because of cardiac massage after rupture. the patient was already pulseless electrical activity (pea) on the electrocardiogram, and while the family watched, the patient¿s heart stopped, and treatment stopped. it was unclear whether the dissection progressed during pump insertion or percutaneous cardiopulmonary support (pcps) at the time of hospitalization. | [5 empty] | A correction was done on manufacturer name, city and state, manufacturing site name and address, initial reporter name and address. | Device discarded by customer. the impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37747 | 1220648-2019-00142 | 8777352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8777352&pc=OZD | 06/14/2019 | 07/10/2019 |  | 07/10/2019 |  |  | 06/14/2019 | 07/10/2019 |  |  | 01/31/2021 |  | 07/10/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1396842 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 57 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/14/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old asian male presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella lp2.5. during support, the device caused symptoms of hemolysis, although no plasma-free hemoglobin draws were taken to confirm this. as treatment, 20 units of red blood cells were delivered and the device as removed. the patient remained in stable condition. | [5 empty] | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37748 | 1220648-2019-00143 | 8778184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8778184&pc=OZD | 06/16/2019 | 09/10/2019 |  | 07/10/2019 |  |  | 06/16/2019 | 07/10/2019 |  |  | 03/31/2021 |  | 07/10/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, OZD |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1414235 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | Yes | Physical Resistance/Sticking | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Other | [28 empty] | Initial submission | Followup |  |  |  | 06/16/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old african american male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using an impella cp device. the pump was inserted successfully, however, when removing the 0.018" guide wire, there was resistance. although physician was successful in applying the necessary force to remove this component from the pump, the wire's coating was believed to have "sheared off" into the patient's leg. no intervention or extra measure were performed to remove any fragments. thereafter, the pump ran without any malfunctions. | [5 empty] | Corrections: initial report state the device was not received. the device was received and this section has been revised. the pump was not returned, however, the proximal portion of the guide wire was received. the wire was inspected and the reported wire fracture was confirmed. kinks were identified in the core of the wire towards the proximal end. due to lack of clinical data, and no returned pump, a root cause cannot be definitively determined. a review of the dhr was conducted and found no manufacturing issues or reworks associated with this pump lot. | The guide wire from the impella cp in question was received and an investigation is currently underway. a supplemental mdr will be filed with the results of the investigation.
37749 | 1220648-2019-00136 | 8771825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8771825&pc=OZD | 06/14/2019 | 08/01/2019 |  | 07/09/2019 |  |  | 06/14/2019 | 07/09/2019 |  |  | 02/28/2021 |  | 07/09/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1414879 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 06/14/2019 | * | Invalid/Incomplete |  |  | A patient was admitted to the medical facility in (b)(6) in cardiogenic shock. the team attempted to treat the multivessel coronary artery disease with stenting. the day after the angioplasty of the coronaries, the patient condition deteriorated and the team placed an impella cp to stabilize him for cardiac surgery the following week. shortly after the pump was placed, the patient developed an access site hematoma. the team performed a fluoroscopy which revealed a femoral artery dissection. the team consulted with vascular surgery and a femoral artery repair was done. the team also infused blood products to the patient. | [5 empty] | Since the filing of the initial medwatch the investigation has concluded. the pump was returned and no damage was observed that could have contributed to the blood loss at the access site. when reviewing the clinical account, there was an observation made of the patient's anticoagulation being in excess of abiomed recommendations. the ifu states to have an act (activated clotting time) of between 160 and 180 seconds during the impella support. the patient's act was noted to be 240 seconds and the aptts wer described as out of range. the blood loss and need for transfusion was due to excessive anticoagulation during impella support. the failure mode will be monitored and trended. | The medical facility in (b)(6) has not yet returned the product for evaluation. the product is expected. upon completion of the analysis a final medwatch will be filed.
37750 | 1220648-2019-00137 | 8771846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8771846&pc=OZD | 06/07/2019 | 07/09/2019 |  | 07/09/2019 |  |  | 06/10/2019 | 07/09/2019 |  |  | 10/31/2020 |  | 07/09/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1375181 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/10/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old (b)(6) male had impella 2.5 pump inserted in the left femoral. severe aortic regurgitation (ar) appeared after impella was inserted. it was unclear whether ar was caused impella or not so impella was removed. immediately after removal, patient became cardiopulmonary arrest (cpa), so percutaneous cardiopulmonary support (pcps) was inserted and emergency aortic valve replacement (avr) was implemented. | [5 empty] | Device discarded by customer. the impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37751 | 1220648-2019-00138 | 8772664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8772664&pc=OZD | 06/14/2019 | 07/09/2019 |  | 07/09/2019 |  |  | 06/14/2019 | 07/09/2019 |  |  | 01/31/2021 |  | 07/09/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 5.0 PUMP |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1404640 |  | 005066-JP |  | No answer provided | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemostasis | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/14/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old (b)(6) male had impella 5.0 pump inserted in the right axillary without any problems. the patient lost an estimated 500ml of blood and red blood cells (rbcs) was administered but the amount is unknown. hematoma was observed at the insertion site, so a surgical removal of the hematoma was made. | [5 empty] | Device not returned. the impella 5.0 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37752 | 1220648-2019-00139 | 8772699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8772699&pc=OZD | 06/14/2019 | 07/09/2019 |  | 07/09/2019 |  |  | 06/14/2019 | 07/09/2019 |  |  | 03/31/2021 |  | 07/09/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1411425 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemostasis | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/14/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old white male had impella cp optical pump inserted in the left femoral. the patient developed a very large hematoma inside the insertion site and the repositioning sheath migrated out causing bleeding. patient lost a lot of blood and as a result fifteen (15) units of red blood cells (rbcs) were administered. the patient was taken to the operating room (or) for impella removal that evening by a cardiothoracic vascular (cv) surgeon and closed the site the following day. patient was subsequently placed on extracorporeal membrane oxygenation (ecmo). | [5 empty] | Device discarded by customer. the impella cp optical was discarded by the customer therefore a failure analysis investigation cannot be completed.
37753 | 1220648-2019-00140 | 8773054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8773054&pc=OZD | 06/11/2019 | 07/09/2019 |  | 07/09/2019 |  |  | 06/11/2019 | 07/09/2019 |  |  | 03/31/2021 |  | 07/09/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, US |  |  | ABIOMED INC. |  |  |  |  | 1414303 |  | 0048-0003 |  | Yes | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | USER F |  |  |  |  | 64 YR | [5 empty] | Malfunction | No | Yes | No Apparent Adverse Event | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/11/2019 | * | Invalid/Incomplete |  |  | The complainant had a patient admitted to the emergency department in active cpr and coding. the patient was taken to the cardiac catheterization lab and found to have multivessel coronary artery disease. the team chose to place an impella cp for support during an angioplasty. the placement of the guide wire and cp pump to the left ventricle was successful, but upon attempted removal of the guide wire the pump was kinking and retaining the wire. troubleshooting to straighten the pump was unsuccessful, and so the pump and wire were snared out of the patient. once removed, the patient was stable and no further intervention was needed.
37754 | 1035166-2019-00057 | 8768658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8768658&pc=DYB | 09/25/2019 | 09/27/2019 |  | 07/08/2019 |  |  | 06/10/2019 | 07/08/2019 |  |  | 09/01/2021 |  | 07/08/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-15213 |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 73 YR | [5 empty] | Injury | Yes | No | Difficult to Advance | Adverse Event Without Identified Device or Use Problem | [5 empty] | Aortic Dissection | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup | Followup |  |  |  | * | Invalid/Incomplete |  |  | As per additional information, guidewire passed through dilator easily. resistance noticed inside the patient artery. patient was rescheduled for surgery and patient is doing fine. | It was reported that patient was presented for ventricular tachachardia. during ep/ablation physician was advancing the guidewire through the introducer. guidewire would not advance around ascending aorta. diagnostic pigtail was inserted, prior to the insertion of the 0.035" guidewire. once pigtail was inserted, 0.035 would not advance around ascending aorta. again, dye showed a lumen and tortuous artery. ep called for tee which showed a big flap from dissection of sfa to ascending aorta. impella was aborted and wires removed and perclose x 2 failed and surgical closure was done after manual pressure did not yield good homeostasis. procedure was aborted and patient outcome is stable. |  |  |  |  | The device was used for treatment. the device was not returned for analysis. the investigation will focus on a review of product documentation. the device history records and guidewire inspection sheet were reviewed to confirm that the devices passed all applicable in process and final inspections. there was no product performance issue reported. guidewire passed through dilator easily. because of tortuous artery guidewire was not able to advance around aorta. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | In follow up 1 mdr , date of this report date of report was left blank. it should be 09-25-2017. no change in date of event. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress ,follow up report will be submitted if we find any further additional information .
37755 | 1220648-2019-00135 | 8769393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8769393&pc=OZD | 06/05/2019 | 07/08/2019 |  | 07/08/2019 |  |  | 06/10/2019 | 07/08/2019 |  |  | 12/31/2020 |  | 07/08/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1391206 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/10/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old (b)(6) female had impella 2.5 pump inserted in the left femoral without any problem. there was bleeding from the insertion site and four (4) units of red blood cells (rccs) were administered as a result. the patient expired due to acute myocardial infarction (ami) that was unrelated to the impella device. | [5 empty] | Device discarded by customer. the impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37756 | 1220648-2019-00134 | 8758215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8758215&pc=OZD | 06/10/2019 | 07/03/2019 |  | 07/03/2019 |  |  | 06/10/2019 | 07/03/2019 |  |  | 03/31/2021 |  | 07/03/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1414706 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/10/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male caucasian presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support using the impella cp. the device was placed without issue, however, patient developed "severe hemolysis" during support. as treatment, fluid resuscitation and repositioning was attempted, however, did not resolve the issue. the decision was made to remove the device, at which point the patient's leg was found "cold and dusky in color" with absence of pulses. the physician opted to remove the pump surgically, as he did not feel comfortable removing the device and holding with manual pressure to achieve hemostasis. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37757 | 1035166-2019-00055 | 8760278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8760278&pc=DYB | 04/12/2019 | 07/03/2019 |  | 07/03/2019 |  |  | 06/07/2019 | 07/03/2019 |  |  |  |  | 07/03/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB | [5 empty] | OSCOR INC. |  |  | US | [5 empty] | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | On (b)(6) 2019 patient was presented to physician for right coronary artery occlusion with an ejection fraction of 30%. during high risk percutaneous coronary intervention physician had peeled away the introducer before being fully removed from the patient's body which cause bleeding. 6 unit blood transfusion performed as treatment and the patient recovered. there was no specific malfunction reported. bleeding started during removal of introducer from site. product is discarded by customer. device was used in tortuous path and there was surgical delay reported. procedure completed with impella lp2.5 device successfully. | [5 empty] | Device was used in treatment. the device was not returned for evaluation. there was no specific performance related failure reported by the user. the lot number of this device was not provided, therefore neither a review of the device history record nor complaint history could be performed. inspection procedures require any oscor product pass all in-process and qa final inspections before shipping to the customer. per instructions for use,introducer removal: flush the sheath with 5cc of saline immediately before peeling the sheath away in order to minimize backbleeding. first completely withdraw the sheath from vessel. then remove the sheath by sharply snapping the tabs of the valve housing in a plane perpendicular to the long axis of the sheath to split the valve and peel the sheath apart. in addition, as per ifu list these adverse effects and possible complications: air embolism, bleeding, hematoma formation, and vessel damage. based on the investigation, a capa is not required. the event will be re-evaluated if additional information becomes available. no further follow-up is required.
37758 | 1035166-2019-00056 | 8760516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8760516&pc=DYB | 05/28/2019 | 12/05/2019 |  | 07/03/2019 |  |  | 06/07/2019 | 07/03/2019 |  |  | 02/01/2022 |  | 07/03/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-16097 |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | On (b)(6) 2019, patient was presented to physician for cardiac arrest with intermittent return of spontaneous circulation during ems and er intervention. approximately 1 day after, during acute myocardial infarction / cardiogenic shock, impella implanted and introducer peeled away. as per reported event, introducer was peeled while still in patient's vessel, causing the bleed. hemostasis was achieved and patient was in stable condition. 3 units of blood products were given. procedure completed successfully with impella cp device. no new introducer was needed. blood products delivered approximately 1 day after introducer was removed. | [5 empty] | Device used in treatment. the device was not returned for analysis. the investigation will focus on a review of product documentation. the device passed all in-process and qa final inspection steps before shipping to the customer. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure adelante s2s introducer sheath in-process and final inspections: visual inspection: verify hub to be smooth, no cracks, short shots, sinking or unfilled areas, and check for fm. verify hub shape as per drawing. verify french size is properly embossed on hub. check for excessive flash (maximum allowable limit is 0.005") and tube pinch (maximum allowable limit is 0.005").using 10x microscope, look down into hub for irregularities. verify no splitting at break lines and a smooth transition between sheath and hub on inside of hub. along the entire length of the hub, verify that there is no break in the molded plastic where the extrusion is visible. verify presence of the two score lines spaced approximately 180 degree apart around the circumference of sheath tube. using 10x microscope, look down into the hub and verify score lines on sheath is in perfect alignment with break line of the hub. verify parts are free from contamination in the form of oil or residue. per a biomed adelante s2, instructions for use (ifu), it informs the user of these cautions: device is supplied sterile. do not use if package has been previously opened or damaged. prior to use, read all package inserts, warnings, precautions, and instructions. failure to do so may result in severe patient injury or death. procedure must be performed by trained medical personnel well versed in anatomical landmarks, safe technique, and potential complications. the product is designed for single use only. do not resterilize or reuse. do not alter the product in any way. in addition, they list these adverse effects and possible complications: air embolism, bleeding, hematoma formation, and vessel damage. based on the investigation, a capa is not required as there were no design, labeling or manufacturing related non-conformity found. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress; follow up report will be submitted if we find any further additional information.
37759 | 1220648-2019-00128 | 8748901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8748901&pc=PYX | 06/07/2019 | 09/03/2019 |  | 07/01/2019 |  |  | 06/07/2019 | 07/01/2019 |  |  | 09/30/2020 |  | 07/01/2019 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1386766 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | Yes | Physical Resistance/Sticking | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 06/07/2019 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old african american female presenting with right ventricular failure for hemodynamic support with the impella rp. successful placement of the device was achieved, however, advancing the pump proved difficult. the physician found the outflow cage appeared abnormal under fluoroscopy, as such, attempted to remove the device. during removal, the outflow cage and pigtail had separated from the catheter and was left in the right ventricle. a gooseneck snare was inserted and the remaining components of the device were removed successfully. a new rp device was opened and used without issue. | [5 empty] | The device was returned and an investigation was completed. visual inspection identified kinking in the device's cannula and a severed outflow cage/pigtail assembly; confirming the reported complaint. scratches, knicks, and gouges were also identified on the returned introducer, consistent with vessel calcification. the condition and size of the patient's vasculature are unknown. a review of the dhr was conducted and found no manufacturing issues or reworks for this pump lot. there are no other complaints related to this failure mode associated with pumps in this lot. the root cause of the pump damage is most likely excessive force used to deliver the pump resulting from difficult vasculature and placement technique. these factors placed an excessive strain on the outflow cage struts. | The impella rp was received from the customer and a device investigation is underway. a supplemental mdr will be filed with the results of the investigation.
37760 | 1220648-2019-00129 | 8748919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8748919&pc=OZD | 06/07/2019 | 07/01/2019 |  | 07/01/2019 |  |  | 06/07/2019 | 07/01/2019 |  |  | 04/30/2020 |  | 07/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1338126 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/07/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male patient presenting with acute myocardial infarction and cardiogenic shock for support with an impella lp2.5. the device was inserted without issue, however, the access site developed a bleed and required multiple units of blood products as treatment. the patient recovered and was successfully explanted, upon improved hemodynamic support. | [5 empty] | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37761 | 1220648-2019-00130 | 8748932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8748932&pc=OZD | 06/03/2019 | 07/01/2019 |  | 07/01/2019 |  |  | 06/03/2019 | 07/01/2019 |  |  | 12/31/2020 |  | 07/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1391294 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 82 YR | [5 empty] | Injury | No | No | No Apparent Adverse Event | [6 empty] | Aortic Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/03/2019 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) asian male patient presenting with "mild" aortic regurgitation for high risk percutaneous coronary intervention and an impella lp2.5 was inserted for hemodynamic support. during support, the patient's aortic regurgitation was exacerbated by the device, thus, prompted medical intervention to remove the device. | [5 empty] | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37762 | 1220648-2019-00131 | 8748938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8748938&pc=OZD | 06/03/2019 | 07/01/2019 |  | 07/01/2019 |  |  | 06/03/2019 | 07/01/2019 |  |  | 03/31/2020 |  | 07/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1412974 |  | 004334 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 40 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | IMPELLA RP | [22 empty] | Initial submission |  |  |  |  | 06/25/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient presenting with post cardiotomy cardiogenic shock. the impella rp was inserted for right side hemodynamic support, followed by an impella cp for left side support. during treatment, "severe hemolysis" with plasma-free hemoglobin levels of 250mg/dl were detected. blood products were delivered, the impella cp was removed, and hemolysis cleared. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37763 | 1220648-2019-00117 | 8749181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8749181&pc=OZD | 06/03/2019 | 06/03/2019 |  | 07/01/2019 |  |  | 05/20/2019 | 07/01/2019 |  |  | 11/30/2020 |  | 07/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1389598 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 54 YR | [5 empty] | Malfunction | No | Yes | No Apparent Adverse Event | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Hospitalization | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient was admitted for a planned high risk angioplasty of the coronary artery. for hemodynamic support, an impella cp was placed via the axillary artery. during the angioplasty the impella pump flows were low. the case was completed, and the pump was explanted and sent for product analysis. upon benchtop testing the engineering team discovered the impella pump to have cracked impeller blades. this blade damage is not observed by the complainant operator. | [5 empty] | The low pump flows observed by the operator was due to the damaged impeller blades. low pump flows can be caused by changes to the impeller shape. the cause of the blade damage was unable to be determined. the failure mode will be monitored and trended.
37764 | 1220648-2019-00132 | 8750237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8750237&pc=OZD | 06/03/2019 | 08/21/2019 |  | 07/01/2019 |  |  | 06/03/2019 | 07/01/2019 |  |  | 02/28/2021 |  | 07/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1405009 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 06/03/2019 | * | Invalid/Incomplete |  |  | The complainant in (b)(6) had a patient admitted in cardiogenic shock and experiencing an acute myocardial infarction have an impella 2.5 placed for hemodynamic support during the angioplasty procedure. the patient was then admitted to the icu for medical followup and monitoring. after over 49 hours of support the impella was re-positioned to remedy the signs of hemolysis. when that failed, the pump was explanted. the hemolysis was also treated with the infusion of blood products. | [5 empty] | The complainant in japan did not return the pump product for evaluation and analysis. the data logs were returned. data log analysis revealed early in the case there were pump position unknown alarms that resolved. at the last day of the support, the logs again show position unknown and position wrong alarms shortly before pump removal. the cause of the hemolysis that prompted pump removal and blood replacement is unknown. hemolysis testing can not be done without the pump product. the failure mode will be monitored and trended. | To date the team in (b)(6) has not replied to all clinical questions. pending more clinical detail, the root cause of the hemolysis is unable to be determined. when the investigation has completed, the final report will be drafted.
37765 | 1220648-2019-00133 | 8750241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8750241&pc=OZD | 06/03/2019 | 08/07/2019 |  | 07/01/2019 |  |  | 06/03/2019 | 07/01/2019 |  |  | 09/30/2020 |  | 07/01/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED INC. |  |  | US |  | 1370748 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 86 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 06/03/2019 | * | Invalid/Incomplete |  |  | The complainant in (b)(6) had a patient have a an impella 2.5 placed for hemodynamic support for cardiogenic shock and acute myocardial infarction. the pump supported the patient for over 65 hours, was weaned and explanted. months later the complainant informed abiomed that during support there was hemolysis. the hemolysis was attributed to the use of impella. the patient's hemolysis was treated by an infusion of blood products. the hemolysis was not confirmed by serial blood draws. | [5 empty] | Since the filing of the initial medwatch the investigation into the complaint from japan has concluded. the japanese medical team did not return for investigation the pump nor the data logs. therefore, no in-house hemolysis testing could be conducted. the claim of hemolysis, during impella support, was not confirmed with definitive lab testing. the physician report did state the hemolysis was "non-serious". with the limited information shared a root cause of the hemolysis could not be determined. | The complainant in (b)(6) has not returned the product for in-house hemolysis testing, as the pump was discarded. further clinical information has not yet been returned for analysis. should more information be shared, and a root cause of the hemolysis be determined, a final medwatch will be filed.
37766 | 1220648-2019-00127 | 8745357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8745357&pc=OZD | 05/31/2019 | 06/28/2019 |  | 06/28/2019 |  |  | 05/31/2019 | 06/28/2019 |  |  | 08/31/2020 |  | 06/28/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1371035 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 47 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Insufficiency | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/31/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old white male presenting with acute myocardial infarction and cardiogenic shock. an impella lp5.0 was selected for hemodynamic support and inserted without issue. the patient's heart function ultimately returned, thus, the device was removed after approximately 103 hours of support. upon removal, the physician noted 2 ruptured mitral valve chordae and attributed this to the impella device. as treatment, surgical repair was performed. the patient was in stable condition, thereafter. | [5 empty] | The impella lp5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37767 | 1220648-2019-00116 | 8734410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8734410&pc=OZD | 05/17/2019 | 07/17/2019 |  | 06/26/2019 |  |  | 05/28/2019 | 06/26/2019 |  |  | 11/30/2020 |  | 06/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1378913 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) years old (b)(6) male had an introducer 13fr inserted. the impella pump was stuck in the sheath, blood leaked and the patient got a shock temporary (systolic blood pressure (sbp)<70) and blood transfusion was made due to blood leaking. | [5 empty] | The pump and the introducer were returned to abiomed for failure investigation. the root cause of the pump unable to pass through the introducer was due to excess of epoxy at the pump housing resulting in an interference fit and the inability to insert the pump through the introducer. | The introducer 13 fr was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37768 | 1220648-2019-00125 | 8734771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8734771&pc=OZD | 05/31/2019 | 08/01/2019 |  | 06/26/2019 |  |  | 05/31/2019 | 06/26/2019 |  |  | 12/31/2020 |  | 06/26/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA LD AIC PUMP SET US |  |  | ABIOMED INC |  |  | US |  | 1397071 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR |  |  | 33 |  |  | Malfunction | No | Yes | Fracture | Difficult to Remove | [5 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 05/31/2019 | * | Invalid/Incomplete |  |  | A patient was admitted for a cabg (coronary artery bypass) and for hemodynamic support the team placed the impella cp via the right femoral. prior to the placement the team treated the arterial disease at the iliac to allow for the impella placement. the following day the impella was placed in surgical mode and the patient had her cabg procedure. upon completion of the cabg, the team attempted to restart the pump. when the pump failed to restart the team tried to pull back and reposition the pump. with the repositioning the physician felt great resistance, and then the pump broke in half at the outflow portion of the pump. at this time there was a portion of the pump remaining within the patient's aorta, and part was removed. | [5 empty] | Since the original medwatch was filed, the investigation into this event has concluded. the product was never returned for investigation. the data logs do not pertain to the failure of a detachment. according to physician, the detachment and subsequent death were unrelated to the use of the impella and likely due to the patient's multiple pre-existing comorbidities. the failure mode will be monitored and trended. | The medical facility has not yet released the product for analysis. upon completion of the investigation, the final medwatch will be filed.
37769 | 1220648-2019-00126 | 8734792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8734792&pc=OZD | 05/30/2019 | 08/21/2019 |  | 06/26/2019 |  |  | 05/30/2019 | 06/26/2019 |  |  | 01/31/2021 |  | 06/26/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1395453 |  | 005040 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | DISTRI |  |  |  |  | 0 YR | [5 empty] | Malfunction | No | Yes | No Apparent Adverse Event | Insufficient Information | [5 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 05/30/2019 | * | Invalid/Incomplete |  |  | The complainant in (b)(6) notified the distributor on may 30, 2019, of a product malfunction on (b)(6) 2019, involving the impella 2.5 pump. while advancing the impella to the left ventricle, the team encountered difficulty advancing beyond the aortic arch. the aorta was known to be heavily calcified and the physician felt friction and a need to use more than typical force. the pump became lodged within the aorta and was retrieved by a gooseneck snare. when snared from the aorta they observed the pump to be kinked and "knotted". the team placed a new impella via the contralateral access in the opposite leg and the support was successfully initiated. | [5 empty] | Since the filing of the initial medwatch report, the investigation has completed. the product was returned for analysis and as reported the impella shape was deformed. the team had reported a "knotted" pump, but upon return the pump was seen to be kinked. the catheter and cannula were kinked. imaging was not shared to observe the kink or knotting as reported. the pump was removed by some intervention, though the method of removal was not shared, whether it was snared or not. the pump deformation did not cause any lasting harm to the patient. the pump was replaced and hemodynamic support proceeded. the failure mode will be monitored and trended. | The distributor has not yet returned the impella pump product or the data logs. upon completion of the investigation, a final medwatch will be filed.
37770 | 1035166-2019-00051 | 8734935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8734935&pc=DYB | 05/06/2019 | 08/01/2019 |  | 06/26/2019 |  |  | 06/03/2019 | 06/26/2019 |  |  | 12/01/2020 |  | 06/26/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-13895 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 80 YR | [5 empty] | Injury | Yes | Yes | Difficult to Insert | [6 empty] | Calcium Deposits/Calcification | Perforation | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup | Followup |  |  |  | * | Invalid/Incomplete |  |  | It was reported that patient was presented to physician for triple vessel disease. during high risk percutaneous coronary intervention 25 cm failed to insert and 13cm sheath had difficulty in inserting. 13cm introducer was placed with difficulty. when event occur, user was attempting to pass pump through patient anatomy when heavily calcified arteries were found. after that femoral angiogram revealed perforation. because of multiple introducer insertion attempts vessel perforation occurred and surgical intervention required to repair injury. there were no specific technique used for insertion and 1 device was planned to be introduced in the device. device was used in tortuous path. the procedure was aborted, vascular surgery was performed, and procedure was successfully completed the following week without support of the impella. | [5 empty] | Our investigation is still in progress ,follow up report will be submitted if we find any further additional information . | Corrected data : follow up 1 mdr missing g7 information in h10 section. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Two 14f abiomed introducer sheaths were returned from the customer with five dilators. there were no other accessories. blood was found on the interior and exterior of the sheaths and the dilators. some of the devices were damaged upon receipt. upon evaluation of the returned devices, it was found that the 14f, 13cm sheath had damage around the circumference of the distal tip edge and scrape marks approximately 1cm from the distal tip. the 14f, 25cm sheath was not damaged. three dilators (12f 13cm, 14f 21cm and 14f 33cm) had deep scrape marks down the length of the dilator lumen. the 8 and 10f dilators looked normal. the 14f abiomed introducer sheath passed all in-process and qa final inspection steps as per dhr review. no manufacturing defects were found during returned device evaluation. multiple introducer insertion attempts into the vessel, tortuous patient anatomy and heavily calcified arteries mentioned in the event description could have been potential contributing factors to the sheath tip and dilators damage. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure adelante s2s introducer sheath in-process and final inspection: visual inspection: with naked eye at a distance of 12" to 18", verify the assembled sheath tip and body matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm. first 5 consecutive properly tipped parts, inspect 100%. visual inspection: under 10x magnification verify: the sheath tip is round. verify proper tip shape according to drawing. slight discoloration from tipping process is acceptable. no split score lines (inspect both score lines) along sheath body, cracks, gouges, tears. per ifu: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. avoid subjecting the device to unusual stresses. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
37771 | 1220648-2019-00124 | 8727739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8727739&pc=OZD | 05/28/2019 | 06/24/2019 |  | 06/24/2019 |  |  | 05/28/2019 | 06/24/2019 |  |  | 05/31/2020 |  | 06/24/2019 |  | 2)TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1342405 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/28/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male presenting with acute myocardial infarction and cardiogenic shock for support with the impella lp2.5 device. after support was successfully achieved, and device was explanted, an infection developed at the insertion site. as treatment, the physician administered antibiotics. the patient recovered. | [5 empty] | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37772 | 1035166-2019-00054 | 8728737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8728737&pc=DYB | 05/12/2019 | 07/02/2019 |  | 06/24/2019 |  |  | 06/05/2019 | 06/24/2019 |  |  | 10/01/2021 |  | 06/24/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-15234 |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | On (b)(6) 2019, the patient a (b)(6), white, female was presented to physician for coding, doing cpr (cardiopulmonary resuscitation). after the pccs (pulmonary critical care) procedure, while the patient was in icu (intensive care unit) it was noted there was bleeding at insertion site; the bleeding issue has not been identified. these multiple devices were also used during the procedure: intra-aortic balloon pump (iabp), impella and extracorporeal membrane oxygenation (ecmo). it is unknown how much blood the patient loss, however they were given one (1) unit of packed red blood cells (prbc). the procedure was completed with standard protocol; device was removed via vascular surgery cut down and the explant was due to heart recovery, not device failure or bleeding. it was reported the patient is still alive and in the hospital. the introducer will not be returned for evaluation, as it was discarded. | [5 empty] | Device used in treatment. the device was not returned for evaluation. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. there was no specific performance related failure reported by the user. the device passed all in-process and qa final inspection steps before shipping to the customer. per instructions for use (ifu), it informs the user of these cautions: device is supplied sterile. do not use if package has been previously opened or damaged. prior to use, read all package inserts, warnings, precautions, and instructions. failure to do so may result in severe patient injury or death. procedure must be performed by trained medical personnel well versed in anatomical landmarks, safe technique, and potential complications. the product is designed for single use only. do not resterilize or reuse. do not alter the product in any way. in addition, they list these adverse effects and possible complications: air embolism, bleeding, hematoma formation, and vessel damage. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. no further follow-up is required. | The following sections were updated in follow-up 1: b4, g4, g7, h2, h6, and h10: additional information h6 method codes: 4109 and 4110. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
37773 | 1220648-2019-00122 | 8724351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8724351&pc=OZD | 04/09/2019 | 06/21/2019 |  | 06/21/2019 |  |  | 05/28/2019 | 06/21/2019 |  |  | 11/30/2020 |  | 06/21/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1385057 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/28/2019 | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male had impella lp 2.5 pump inserted in the right femoral with no problem. during pump support on (b)(6) 2019, patient developed hemolysis and as a result fourteen (14) units of red blood cells concentrate (rbcs) were administered. on (b)(6) 2019 the impella device was explanted, patient status recovered, and mechanical support was no longer needed. | [5 empty] | The impella 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37774 | 1220648-2019-00123 | 8724377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8724377&pc=OZD | 05/30/2019 | 06/21/2019 |  | 06/21/2019 |  |  | 05/30/2019 | 06/21/2019 |  |  | 01/31/2021 |  | 06/21/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1411255 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 51 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/30/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female had impella cp pump inserted in the left femoral. the patient was going to the operating room (or) for a bilateral perfusion for the ecmo limbs and for a tap as they knew there was a need for pericardial drainage. upon draining, the anesthesia team noted fragrantly red blood draining in large and constant flow. after 15 minutes and over 200cc of drainage they decided to open the chest. at this time, they saw the white portion of the pigtail thru the apex of the heart. patient had surgical repair, a surgical left atrium (la) vent was placed by a cardiothoracic vascular (cv) surgeon and patient is still alive on ecmo. | [5 empty] | Device discarded by customer. the impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37775 | 1220648-2019-00121 | 8713305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8713305&pc=OZD | 05/24/2019 | 08/21/2019 |  | 06/19/2019 |  |  | 05/24/2019 | 06/19/2019 |  |  | 09/30/2020 |  | 06/19/2019 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA 2.5 PUMP SET, JAPAN |  |  | DANVERS, MA |  |  | US |  | 1373141 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 05/24/2019 | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian male presenting with myocarditis and the impella lp2.5 was selected for cardiac support. during treatment, the patient exhibited hemolysis. for treatment, the patient was infused with a total of 50 units of blood products. | [5 empty] | The impella lp2.5 was received and a failure investigation was completed. data logs identified pump had emitted positioning alarms, however, resolved by the following day. examination and testing of the pump found no problems and ran to specification. a review of the device history record found no manufacturing issues or reworks associated with this pump lot. the root cause of the hemolysis was unable to be determined due to lack of clinical information and pump had run within specification during use and testing. | The impella lp2.5 was received and an investigation is underway. a supplemental mdr will be filed with the results of the analysis, once completed.
37776 | 1035166-2019-00053 | 8713810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8713810&pc=DYB | 04/26/2019 | 06/19/2019 |  | 06/19/2019 |  |  | 05/24/2019 | 06/19/2019 |  |  |  |  | 06/19/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA, |  |  | DYB | [5 empty] | OSCOR INC. |  |  | US | [5 empty] | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiogenic Shock | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | On (b)(6) 2019, patient ((b)(6)) was presented to physician in cardiogenic shock. during (access) insertion at the left femoral artery, the physician noted he placed the 14fr introducer rather than the 13fr introducer, and then had difficultly controlling the blood loss. the patient had blood products infused, six (6) units of rbc (red blood cells) and six (6) units of ffp (fresh frozen plasma). the issue was resolved by blood product infusion, the patient is stable. there was no issue with the oscor 14fr introducer. the manufacturing lot number is unknown. the device will not be returned for evaluation. | [5 empty] | Device was used in treatment. the device was not returned for evaluation. there was no specific performance related failure reported by the user. the lot number of this device was not provided, therefore neither a review of the device history record nor complaint history could be performed. inspection procedures require any oscor product pass all in-process and qa final inspections before shipping to the customer. per instructions for use, it informs the user of these cautions: device is supplied sterile. do not use if package has been previously opened or damaged. prior to use, read all package inserts, warnings, precautions, and instructions. failure to do so may result in severe patient injury or death. procedure must be performed by trained medical personnel well versed in anatomical landmarks, safe technique, and potential complications. the product is designed for single use only. do not resterilize or reuse. do not alter the product in any way. in addition, they list these adverse effects and possible complications: air embolism, bleeding, hematoma formation, and vessel damage. the device was not returned and there was no specific performance related failure was reported by the user. based on the investigation, a capa is not required. the event will be re-evaluated if additional information becomes available. no further follow-up is required.
37777 | 1220648-2019-00054 | 8695282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8695282&pc=OZD | 05/20/2019 | 05/20/2019 |  | 06/13/2019 |  |  | 05/20/2019 | 06/13/2019 |  |  | 10/31/2020 |  | 06/13/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1388919 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male had impella cp pump inserted in the left femoral. patient had leg ischemia post removal, angiography of left leg showed clot in popliteal or anterior tibial region, cutdown of left femoral artery was done to remove the clot. the clot was successfully removed. | [5 empty] | The impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37778 | 1220648-2019-00026 | 8695283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8695283&pc=OZD | 05/14/2019 | 05/17/2019 |  | 06/13/2019 |  |  | 05/17/2019 | 06/13/2019 |  |  | 11/30/2020 |  | 06/13/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA LP5.0 PUMP |  |  | OZD |  |  | IMPELLA LP 5.0 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1395562 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male patient had impella lp5.0 pump inserted in the right axillary with no issue. patient had been septic after impella was implanted therefore patient was treated with antibiotics. | [5 empty] | The impella lp 5.0 pump was not returned by the customer therefore a failure analysis investigation cannot be completed. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37779 | 1220648-2019-00115 | 8695287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8695287&pc=OZD | 04/06/2019 | 05/15/2019 |  | 06/13/2019 |  |  | 05/15/2019 | 06/13/2019 |  |  | 12/31/2019 |  | 06/13/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA LP5.0 PUMP |  |  | OZD |  |  | IMPELLA LP 5.0 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1321702 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male had impella lp5.0 pump inserted in the right axillary. the patient had hemolysis which resulted in a blood transfusion of 6 units of red blood cells (rbc) and 10 units of platelets (pc) was administered during pump support. | [5 empty] | The impella lp 5.0 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37780 | 1035166-2019-00050 | 8692828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8692828&pc=DYB | 05/18/2019 | 10/03/2019 |  | 06/12/2019 |  |  | 05/20/2019 | 06/12/2019 |  |  | 02/01/2022 |  | 06/12/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-16097 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 75 YR | [5 empty] | Malfunction | No | Yes | Break | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that patient was presented to physician for cardiogenic shock post open heart surgery. intra-aortic balloon pump used prior to decision to place introducers/impella pumps. wired on 150cm 0.035 amplatz wire used for insertion. during access at left femoral guidewire kinked. there was no issue with introducer therefore, decision made to abort use of this access site due to dissection endured relating to previous iabp insertion. during access at right femoral physician forgotten that sheath has twist lock hub, so during procedure once the hub broke, sheath removed and replaced. there was 10 min delay reported and medical intervention is not related to sheath but rather underlying medical condition of patient. patient did have profound anemia with concern for rp bleed, but it was unconfirmed. patient died on (b)(6) 2019 because family withdrew care due to futility. no adverse event reported no additional information available. | [5 empty] | The following sections were updated in follow-up. b4,d10,g4,g7,h2,h3,h6,h10. the device used in treatment. one 14f, 13cm long abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. one of the orange split caps was detached from the sheath's hub and was not returned for evaluation. the other orange cap was still attached. visible blood was found on and inside the sheath and sideport tubing. when viewed under a 10x microscope, adhesive was present on both sides of the white hub. the bottom of the orange cap could not be checked for adhesive as it was not returned for evaluation. adhesive was identified on the hub, but it is difficult to quantify the amount of adhesive. physician was rocking the dilator back and forth when product issue identified. based on the available information it appears that the physician did now follow the ifu. no manufacturing defects were found. the device history records were reviewed to confirm that the device passed all applicable in process and final inspections. per procedure adelante s2s introducer sheath in-process and final inspection: "with naked eye at a distance of 12" to 18", verify the sheath hub and split caps are free of excess silicone oil applied between the seals during assembly. verify split caps are properly placed and secured onto sheath hub and free of cracks/damages or excessive adhesive." additionally, destructive testing is performed per sampling plan. inspector is to "insert dilator into sheath, lock dilator hub lock nut. hold dilator hub with one hand and one of the sheath hub wing with the other hand. bend the dilator hub to the side until dilator hub lock nut is disconnected from the sheath hub. split caps should remain in place. there should be no damage and fracture to split cap and tabs." manual break and peel test (sampling plan s4 1.0 reduced) using samples from assembly fit check as described in, manually break the sheath and hub and verify that the seals split easily without extreme elongation. also verify that the split cap and seals remain secure and do not break free or loosen from the sheath hub. per ifu : cautions: prior to use, read all package inserts, warnings, precautions, and instructions. failure to do so may result in severe injury or death. directions for use: insert vessel dilator into sheath and rotate the dilator cap over valve housing to secure the dilator onto sheath assembly. thread the dilator/sheath assembly over the guidewire. advance the dilator and sheath together with a twisting motion over the guidewire and into the vessel. fluoroscopic observation may be advisable. attaching a clamp or hemostat to the proximal end of the guidewire will prevent inadvertently advancing the guidewire entire into the patient. once assembly is fully introduced into the vascular system, separate the dilator cap from the sheath valve housing by rotating the dilator cap off the hub. corrective action has been completed to address the orange cap detachment issue. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress, follow up report will be submitted if we find any further additional information.
37781 | 8686551 | 8686551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8686551&pc=PBL | 04/12/2019 | 04/17/2019 | 04/17/2019 | 06/11/2019 | 04/17/2019 | 06/11/2019 |  | 06/11/2019 |  |  |  |  | 06/11/2019 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003 |  | Yes | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 22995 DA |  |  | 95 |  |  | Death | Yes | No | No Flow | [6 empty] | Death | [12 empty] | Death | [28 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | A patient was emergently brought into the cath lab for placement of the impella cp. chest compressions were being performed prior to insertion of the device. an impella cp was inserted into the patient's heart and little to no flow was observed through the device. this impella cp was removed and replaced with a second device. this resulted in slight improvement in flow. the patient ultimately expired in the cath lab. the physician stated that the lack of flow in the first impella may have been caused by the patients extremely low blood pressure, or perhaps a thrombus became lodged in the device preventing blood flow. it is unknown at the time of this report if the device malfunctioned or failed in any way.
37782 | 1220648-2019-00081 | 8683563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8683563&pc=OZD | 05/14/2019 | 05/14/2019 |  | 06/10/2019 |  |  | 05/14/2019 | 06/10/2019 |  |  | 02/28/2021 |  | 06/10/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1406778 |  | 0048-0024 |  | No | [5 empty] | No | No Information |  |  | Manufacturer report |  |  | HEALTH PROFESSIONAL | [5 empty] | 39 YR | [5 empty] | Injury | Yes | No | Material Twisted/Bent | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male presenting with acute myocardial infarction and cardiogenic shock. during insertion of cp, the pump failed to cross the aortic valve due to angulation. the device's pigtail "doubled back" and would not straighten. intervention was required with a snare, inserted through radial access, to straighten and remove the device. the device was removed and cardiac support was successfully achieved with a second device. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37783 | 1220648-2019-00106 | 8684279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8684279&pc=OZD | 09/22/2017 | 05/15/2019 |  | 06/10/2019 |  |  | 05/15/2019 | 06/10/2019 |  |  | 02/28/2019 |  | 06/10/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1280511 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 50 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Aortic Regurgitation | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | Via a social media posting, on (b)(6), abiomed was notified of an adverse event from two years prior. the patient published upon had two impella pumps, a cp and then escalated care to the impella 5.0. the patient had documented aortic regurgitation which prompted aortic valve replacement. the patient had been on impella 5.0 support for eighteen days and during those days did have multiple echo procedures done, though the true extent of the valvular complication was not noted til day sixteen of support. | [5 empty] | The complainant had returned the product and data logs in 2017 prior to the (b)(6) posting and publication. as such, and investigation could be done and the literature imaging, contained within the publication, could be reviewed. reviewing the clinical imaging, it is apparent that while mild aortic regurgitation (ar) possibly existed during impella cp support, there was severe aortic insufficiency during impella 5.0 explant. the clinical account states that there was ar four days prior to this event as well. the paper indicates that it was difficult to assess if there was ar under tte imaging, and only visible with ease under subsequent tee imaging. the returned video confirms that it would be difficult to assess given the quality of the tte taken the root cause of the severe ar cannot be determined. the failure mode will be monitored and trended.
37784 | 1220648-2019-00112 | 8680216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8680216&pc=OZD | 01/04/2019 | 05/10/2019 |  | 06/07/2019 |  |  | 05/10/2019 | 06/07/2019 |  |  | 06/30/2020 |  | 06/07/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1349665 |  | 1349665 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Use of Device Problem | Insufficient Information | [5 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) | [21 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The medical staff in (b)(6) placed an impella 5,0 for support in a patient admitted with cardiogenic shock and acute myocardial infarction as a supplement and vent for the previously placed ecmo. the impella supported the patient for 18 days before explant. months after explant the medical team informed abiomed that during support there was hemolysis treated with blood replacement and pump repositioning. | [5 empty] | The complainant in (b)(6) did return the pump product. the pump is currently undergoing hemolysis testing. upon completion of the hemolysis testing a final report will be sent. to date the conclusion available, drawn from clinical information shared by the medical team in japan, is that the pump was repositioned and rbc's were given to the patient in order to resolve the hemolysis, although it is unknown if these were effective. data logs confirm positioning alarms that could have been related to malposition or the use of ecmo. the failure mode will be monitored and trended. | The complainant did return the pump product. the clinical information shared by the medical team in japan, is that the pump was repositioned and rbc's were given to the patient in order to resolve the hemolysis, although it is unknown if these were effective. data logs confirm positioning alarms that could have been related to malposition or the use of ecmo. the in-house hemolysis testing passed. due to the fact that the pump passed hemolysis testing and that the hemolysis resolved with re-positioning, the root cause of the hemolysis was the malpositioning. the failure mode will be monitored and trended.
37785 | MW5087217 | 8683371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8683371&pc=PYX | 06/02/2019 | 06/05/2019 |  | 06/07/2019 |  |  |  | 06/07/2019 |  |  |  |  | 06/07/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 OR IMPELLA CP |  |  | PYX | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 60 YR |  |  | 95 |  |  | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Life Threatening | Other | Required Intervention | Disability | [25 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Pt had mi with low cardiac output. received impella 2.5 temporary assist device and transferred to our hosp. severe bleeding from femoral artery noted. the pt was taken to the operating room where the device was removed. the pt required repair of femoral artery as well as fasciotomy of left lower extremity. the pt required eight units of prbcs. pt noted to have massive stroke postoperatively (he was intact prior to being taken to the operating room). given the device is inserted from the femoral artery into the left ventricle, it is quite possible that clot(s) developed around the device embolized and caused massive stroke. please note we have multiple pts who have had massive stroke / died following impella use. we have notified mfr but have not heard back. this device is potentially serious health hazard and could cause death or other injury. fda safety report id# (b)(4).
37786 | 1220648-2019-00113 | 8666303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8666303&pc=OZD | 05/05/2019 | 05/06/2019 |  | 06/04/2019 |  |  | 05/06/2019 | 06/04/2019 |  |  | 10/31/2020 |  | 06/04/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA LP5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1388675 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | CONSUMER | HEALTH PROFESSIONAL |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Insufficiency | Rupture | [11 empty] | Life Threatening | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old caucasian male presenting with acute myocardial infarction and cardiogenic shock. during support, the patient decompensated and was intubated. the impella exhibited suction alarms and the pigtail was found "hooked" around the mitral chordae. damage to the chordae was identified and described as bruised, red, and ragged. as treatment, the p-level was reduced and emergent surgery was performed to repair the mitral valve. the device operated successfully for another 6-7 days until successful weaning/explant. since the impella's inflow was surrounding the damaged chordae, the physician suspected the impella was likely to blame for the damage. | [5 empty] | The impella lp5.0 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37787 | 1035166-2019-00047 | 8667046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8667046&pc=DYB | 05/07/2019 | 08/21/2019 |  | 06/04/2019 |  |  | 05/09/2019 | 06/04/2019 |  |  | 11/01/2021 |  | 06/04/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-15681 |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 68 YR | [5 empty] | Injury | Yes | Yes | Obstruction of Flow | [6 empty] | Ischemia | No Consequences Or Impact To Patient | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that an incident occurred during an outpatient procedure in the cath lab, for high risk pci (percutaneous coronary intervention) on (b)(6) 2019. patient had significant disease in her iliac artery and distal aorta but there was no prior angiogram that presented this before attempting to place the impella cp device across aortic valve. as the cp device was advanced over the .018 wire, it met resistance in the distal aorta. cp was removed and a pigtail catheter was placed in the aorta to take an angiogram. that revealed significant disease in left common iliac, left external iliac and that the oscor 14f x 13cm sheath was occluding flow to the left distal extremity. the occlusion lead to limb ischemia and the need for a cut down in the or with vascular repair and patch. patient's peripheral arteries are known to be calcified; outcome of patient is stable. device will not be returned as it was discarded. | [5 empty] | The following sections were updated in follow-up 1: b4, d4, g4, g7, h2, h6, h10 and h11 corrected data d4: expiration date 01-nov-2021 the device was used in treatment. the device will not be returned for analysis, therefore, the clinical observation could not be confirmed. the investigation will focus on review of product documentation. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per qa procedure preform occlusion test per procedure, latest revision. visual inspection: with the naked eye at a distance of 12" to 18", inspect sideport assembly for excessive glue. a bead of glue around the perimeter of the sideport tube to hub joint is acceptable. smeared glue on sideport assembly is not acceptable. the instructions for use (ifu) instructs the user: aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. aspirate all air from the sheath valve assembly by using a syringe connected to the sideport. flush the introducer through the sideport. if the introducer is to remain in place during catheter positioning and testing, flushing the introducer via the sideport periodically with saline is advised. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete.
37788 | 1220648-2019-00111 | 8668318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8668318&pc=OZD | 05/07/2019 | 05/07/2019 |  | 06/04/2019 |  |  | 05/07/2019 | 06/04/2019 |  |  | 09/30/2020 |  | 06/04/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED |  |  | US |  | 1374531 |  | 005-066-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 26 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hematoma | Mitral Regurgitation | Patient Problem/Medical Problem | [10 empty] | Hospitalization | Required Intervention |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) | [22 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | In (b)(6) a (b)(6) year old patient admitted with myocarditis had an impella 5.0 placed via the left femoral artery to supplement and vent the previously placed ecmo. due to the poor condition and deterioration of the patient, the team was considering an implanted vad device. after over 10 days of support, the team chose to explant the ecmo. once weaned off of the ecmo the team performed an echo and observed mitral regurgitation (mr). the team explanted the impella, performed a mitral valve replacement, and again placed ecmo for support. post explant of the impella the team also performed blood cultures. the blood culture was gram negative bacilli laden. the infection was thought to have been from a hematoma at the impella access site. the hematoma was surgically treated at the explant of the impella and the site was left open for vac (negative pressure wound therapy) at the medical facility in (b)(6). | [5 empty] | The complainant in (b)(6) has returned the product and data logs. the investigation is on-going at this time. there is a request out to the medical facility for further clinical updates. upon completion of the investigation, a final medwatch will be filed. | The investigation has completed since initial medwatch was filed. the data logs were not returned for analysis. the pump was returned and no damage was found, and no sharp edges. the visual inspection found no defects that may have increased the risk of damage to the chordae during any interaction. without more clinical data the root cause of the chordae damage is unknown. the root cause of the infectious hematoma is also unknown. it is suspected that the extended use of the access site could have caused the infection. the failure mode will be monitored and trended.
37789 | 1220648-2019-00095 | 8655263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8655263&pc=OZD | 05/03/2019 | 05/03/2019 |  | 05/30/2019 |  |  | 05/03/2019 | 05/30/2019 |  |  | 12/31/2020 |  | 05/30/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1407652 |  | 05042 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A patient with multiple cardiac comorbidities had an impella 2.5 placed for hemodynamic support for a high risk coronary angioplasty. there were difficulties in the placement of the pump's introducer sheath, though the pump was placed via the femoral artery, and support initiated. post angioplasty the pump remained on for support and the patient was transferred to the icu for monitoring. while in the icu the patient experienced diminished pedal pulses. the impella was explanted and vascular surgery consulted. the patient had a femoral endarterectomy done and was reported to be stable post surgery. | [5 empty] | The medical facility did not return all products used on the patient. the introducer sheath was returned, but not the impella pump. without the pump a complete investigation and root cause analysis for the limb ischemia could not be done. of note, the medical team knew the patient to have peripheral arterial disease, but still placed the impella pump for support via a heavily calcified artery. the impella 2.5 ifu does have in the contraindications the following: "severe peripheral arterial disease precluding placement of the impella system;". most likely patient condition contributed to the limb ischemia. the failure mode will be monitored and trended.
37790 | 1220648-2019-00100 | 8655295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8655295&pc=OZD | 02/16/2016 | 05/09/2019 |  | 05/30/2019 |  |  | 05/05/2019 | 05/30/2019 |  |  | 01/31/2017 |  | 05/30/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1182559 |  | 0048-0003 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | OTHER |  |  |  |  | 20 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Mitral Regurgitation | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | Via a social media post on may 5, 2019 manufacturer was made aware of an event occurring in 2016 in which post explant of an impella cp a patient was found to have chordae damage. via retrospective review the case was found in manufacturer database. the male patient was admitted in cardiogenic shock, and had an impella cp placed due to hypotension, deteriorating condition, and an unknown etiology of a low ejection fraction. the impella pump was placed and supported the patient for 4 days and was weaned and explanted. the field team was not informed of the chordae damage and mitral valve repair done post explant. | [5 empty] | The event occurred in 2016. product and data logs are not available for analysis. a root cause of chordae damage and mitral valve repair can not be established. should more information become available a final medwatch will be drafted. the failure mode will be monitored and trended.
37791 | 1220648-2019-00089 | 8650039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8650039&pc=OZD | 05/01/2019 | 05/01/2019 |  | 05/29/2019 |  |  | 05/01/2019 | 05/29/2019 |  |  | 11/30/2020 |  | 05/29/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1394648 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 57 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A patient presented in cardiogenic shock, and had an impella cp placed to hemodynamically support the patient. the access to place the impella was complicated by the amount of adipose tissue in the patient, of a larger body habitus. the team attempted to place the impella cp's repositioning sheath unit after removal of the 14fr x 30 cm introducer, and was unable to so. the team found the pump to be buckling in the longer adipose tissue track. the pump was explanted and the access site was bleeding while hemostasis was achieved by contralateral leg ballooning. in addition, the patient received blood products as a hematoma developed at the site. the patient was supported by an impella pump in the contralateral leg. | [5 empty] | The complainant did return the impella pump product for analysis. the analysis of the pump found ridges at the tip of the pump's repositioning sheath. this damage is consistent with the described difficult insertion, due to the patient's body habitus. this difficult exchange likely led to the bleeding observed. the failure mode will be monitored and trended for analysis.
37792 | 1220648-2019-00096 | 8650467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8650467&pc=OZD | 05/01/2019 | 05/01/2019 |  | 05/29/2019 |  |  | 05/01/2019 | 05/29/2019 |  |  | 01/31/2021 |  | 05/29/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1402213 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A patient was admitted with shortness of breath and a pulmonary embolism. the impella cp was placed via the left femoral artery to provide hemodynamic support. shortly after placement, the limb was cold due to limb ischemia. the team placed a antegrade sheath for an external femoral-femoral bypass. the bypass restored flow and on the next day the limb had pulses restored. the pump remained on for support for a total of 146 hours of support. this complaint was observed via a social media post. | [5 empty] | The complaint was received via a social media post. after reviewing the post, the investigation was initiated, but the product had already been discarded. the patient was on higher doses of vasopressors during the initial days of the impella support. the root cause of the leg ischemia was most likely the combination of small artery size and high dosage of vasopressors. the high dosage of vasopressors that further constricted the artery, led to the femoral artery blockage, and caused leg ischemia. the external femoral-femoral artery bypass and weaning from the vasopressors allowed for the limb to resolve. there was no lasting harm or injury.
37793 | 1220648-2019-00091 | 8650512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8650512&pc=OZD | 04/30/2019 | 04/30/2019 |  | 05/29/2019 |  |  | 04/30/2019 | 05/29/2019 |  |  | 10/31/2020 |  | 05/29/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1386387 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient was admitted in cardiogenic shock and acute myocardial infarction. upon admission the impella cp was placed to stabilize the patient and prepare for transfer to the tertiary medical center for coronary angioplasty. upon transfer to the tertiary center, there was bleeding observed from the impella cp access site. the patient had 1 unit of blood infused and a surgical consult to close the groin site and achieve hemostasis. the impella was explanted and escalated to the impella 5.0. | [5 empty] | The medical facility did not return for analysis the pump product. the root cause of the acces site bleed is not known at this time. should further details be share that point to a root cause or product malfunction, a final medwatch will be filed.
37794 | 1220648-2019-00110 | 8642517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8642517&pc=OZD | 04/12/2019 | 06/27/2019 |  | 05/24/2019 |  |  | 04/24/2019 | 05/24/2019 |  |  | 07/31/2020 |  | 05/24/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA LP5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1359038 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 39 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Other | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female asian presenting with myocarditis. the impella lp5.0 pump was inserted for cardiac support, however, air was found in the patient's left ventricle when the device was turned on. no adverse effects or additional intervention was reported, as a result. | [5 empty] | The impella lp5.0 was received and a failure investigation was completed. data log analysis and simulated use testing found pump had performed within specification. the cause of the reported air embolism was unable to be determined. a review of the device history record found no manufacturing issues or reworks related to this failure mode for any pump in this lot. there are no other complaints for this lot with other pumps associated with this failure mode. | The impella lp5.0 and console data logs were received by the customer and an investigation is underway. a supplemental mdr will be filed with the results of the analysis.
37795 | 1220648-2019-00090 | 8637651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8637651&pc=OZD | 04/27/2019 | 04/27/2019 |  | 05/23/2019 |  |  | 04/27/2019 | 05/23/2019 |  |  | 04/30/2020 |  | 05/23/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1338806 |  | 005060 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN | USER FACI |  |  |  | 64 YR | [5 empty] | Malfunction | No | Yes | Device Difficult to Setup or Prepare | Human-Device Interface Problem | [5 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant in the (b)(6) placed an impella 5.0 via the right axillary artery to support a patient admitted with an acute myocardial infarction and in shock. after placement of the pump, there was an observation made of bleeding from the pump. the bleeding prompted the pump to be replaced after their troubleshooting of surgically tying off failed. in addition, blood products were infused to the patient. the second pump was placed successfully and the hemodynamic support proceeded. | [5 empty] | The complainant returned, via the (b)(4) distributor, the impella 5.0 pump for analysis. the analysis of the pump found there to be a leak path originating at the repositioning unit of the impella pump. the leak path was found with benchtop pressure testing and was able to be reproduced. the leak path allowed for blood loss at the area of the hub and blue butterfly located on the pump. this leak path was most likely formed by physical stress applied to the system. the failure mode will be monitored and trended.
37796 | 1220648-2019-00107 | 8638002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8638002&pc=OZD | 12/26/2018 | 04/25/2019 |  | 05/23/2019 |  |  | 04/25/2019 | 05/23/2019 |  |  | 07/31/2020 |  | 05/23/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED |  |  | US |  | 1355322 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Renal Failure | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A patient admitted in cardiogenic shock had an impella placed for hemodynamic support during coronary revascularization in (b)(6). the pump was repositioned once due to concerns of hemolysis. the pump remained on for approximately 10 days of support, was weaned and explanted. the team later reported that the patient had experienced renal failure while on support with the impella that necessitated dialysis. | [5 empty] | Since the initial medwatch was filed, the investigation has completed. without the pump and data logs the definitive root cause of the renal failure could not be determined. | The team in (b)(6) had discarded the impella pump. analysis of it is not possible. the data logs were returned and will be analyzed. upon completion of the investigation, a final report will be submitted.
37797 | 1220648-2019-00108 | 8638005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8638005&pc=OZD | 02/15/2019 | 04/25/2019 |  | 05/23/2019 |  |  | 04/25/2019 | 05/23/2019 |  |  | 11/30/2020 |  | 05/23/2019 |  | TEMPORARY NON ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED |  |  | US |  | 1378913 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Hematoma | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A patient admitted in cardiogenic shock in (b)(6), had an impella placed for hemodynamic support during coronary revascularization. the pump was supporting the patient when there was an observation made of access site bleeding and hematoma development. the patient had 10 units of rbc and 10 units of ffp infused. the team explanted the pump and no further hemodynamic support was needed. | [5 empty] | Since the original medwatch was filed, the investigation has concluded. the team in japan never did return any product, nor answer any of the followup clinical questions. as such, the root cause of the hematoma at the access site could not be determined. the failure mode will be monitored and trended. | The complainant in (b)(6) has not yet answered clinical questions. the data logs have been returned for analysis, but the product was discarded. there will be a final report drafted upon completion of the investigation.
37798 | 1220648-2019-00109 | 8638274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8638274&pc=OZD | 04/26/2019 | 04/26/2019 |  | 05/23/2019 |  |  | 04/26/2019 | 05/23/2019 |  |  | 06/30/2020 |  | 05/23/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1352792 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Patient Problem/Medical Problem | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A (b)(6) year old asian male patient presented with acute myocardial infarction and cardiogenic shock. an impella lp2.5 was inserted without any reported issue, however, patient developed ischemia during the course of support. as treatment, the physician administered ecmo in the opposing femoral artery and the condition improved. | [5 empty] | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37799 | 1220648-2019-00105 | 8623546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8623546&pc=OZD | 04/23/2019 | 04/23/2019 |  | 05/20/2019 |  |  | 04/23/2019 | 05/20/2019 |  |  | 01/31/2021 |  | 05/20/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1394976 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemorrhage, Cerebral | Thrombosis | [11 empty] | Hospitalization | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant in the eu had an impella cp supporting a ventilated patient admitted with cardiogenic shock and acute myocardial infarction for over 171 hours, when they began to wean the patient in preparation of pump explant. during the support the patient suffered a cerebral vascular event of unknown causality. the stroke was confirmed with ct imaging. in addition the patient's sedation was ceased to evaluate the neurological status. upon pump explant there was an observation made of thrombotic burden adhering to the impella pump. | [5 empty] | The investigation is ongoing at this time. no root cause has yet been determined for the stroke and thrombus observed upon pump explant. a final report will be submitted upon the completion of the investigation. | Since the original medwatch was filed, the investigation has completed. the pump was returned and found to have biomaterial adhering to the inflow cage of the pump. when the biomaterial was removed the pump functioned appropriately without any suction alarms that were noted by the medical team and observed in the pump data logs. the suction alarms from low flows were due to biomaterial. the root cause of the neurological event though are still unknown. the failure mode will be monitored and trended.
37800 | 1220648-2019-00086 | 8620055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8620055&pc=OZD | 04/22/2019 | 04/22/2019 |  | 05/17/2019 |  |  | 04/22/2019 | 05/17/2019 |  |  | 11/30/2020 |  | 05/17/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1394462 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A patient with multiple pre-existing conditions was admitted to the hospital in cardiogenic shock. the patient was transferred to a tertiary medical center for care in critical condition. she presented to the tertiary center with av block, bradycardia, hypotension, and coronary thrombosis. the team in the cardiac catheterization lab placed an impella cp via the right femoral artery for hemodynamic support. the patient was noted to also be in right sided heart failure, but her poor prognosis did not prompt placement of right sided support. the team transferred the patient to the icu for continued monitoring on the impella pump. on the second day of support the access site of impella was oozing blood, which prompted 2 units of blood replacement. | [5 empty] | The medical facility did return the impella cp for analysis, though not the data logs from the console. the pump was returned in good condition. analysis under microscope found no damage that could have caused or contributed to the blood loss, and treatment with 2 units of blood. the blood loss, at the impella access site, was most likely a result of the patient's poor condition with dic (disseminated intravascular coagulation). in addition the physician does not attribute the outcome to the use of impella.
37801 | 1220648-2019-00104 | 8620970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8620970&pc=OZD | 04/18/2019 | 06/19/2019 |  | 05/17/2019 |  |  | 04/18/2019 | 05/17/2019 |  |  | 04/30/2020 |  | 05/17/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1345683 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | ECMO | ECMO | [21 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old asian female presenting with acute myocardial infarction and cardiogenic shock. the impella lp2.5 was inserted in the right femoral artery for cardiac support. during support, the limb in which the impella was inserted had become ischemic. as treatment, the device was removed and escalated to axillary insertion. the patient's ischemia had resolved and the patient was stable. | [5 empty] | The impella lp2.5 was returned for investigation by the customer. the investigation is ongoing and a supplemental mdr will be filed with the results of the analysis. | The impella lp2.5 was received and a failure investigation was completed. visual inspection found no damage, defects, or deficiencies that could have caused or contributed to the reported leg ischemia. a review of the device history record found no manufacturing issues or reworks related to this failure mode for any pump in this lot. there are no other complaints for other pumps in this lot or repositioning units in other kit lots that were associated with this failure mode.
37802 | 1220648-2019-00103 | 8617597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8617597&pc=OZD | 04/12/2019 | 06/24/2019 |  | 05/16/2019 |  |  | 04/18/2019 | 05/16/2019 |  |  | 11/30/2020 |  | 05/16/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1385393 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 89 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | ECMO | ECMO | [21 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old asian female patient presenting with acute myocardial infarction and cardiogenic shock. the impella lp2.5 was inserted for cardiac support. during support, the patient exhibited leg ischemia which prompted device removal and "retrograde perfusion" via blood products. the patient was in stable condition after explant. | [5 empty] | The impella lp2.5 was returned for investigation by the customer. the investigation is ongoing and a supplemental mdr will be filed with the results of the analysis. | The impella lp2.5 was received and a failure investigation was completed. visual inspection found no damage, defects, or deficiencies that could have contributed to the reported limb ischemia. a review of the device history record found no manufacturing issues or reworks related to this failure mode for any pump in this lot. there are no other complaints for other pumps in this lot or repositioning units in other kit lots that were associated with this failure mode. all introducers in this lot passed all inspection.
37803 | 1220648-2019-00102 | 8611732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8611732&pc=OZD | 04/19/2019 | 04/19/2019 |  | 05/15/2019 |  |  | 04/19/2019 | 05/15/2019 |  |  | 12/31/2020 |  | 05/15/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1397114 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 66 YR | [5 empty] | Malfunction | No | Yes | Detachment of Device or Device Component | Infusion or Flow Problem | No Apparent Adverse Event |  |  |  |  | No Information | [41 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male caucasian presenting for a coronary artery bypass graft. an impella cp was inserted for cardiac support. after surgery, the device was unable to achieve adequate flows and experienced multiple suction alarms. despite repositioning, the issue could not be resolved. the device was explanted and support was achieved successfully with a new device. the patient remained in stable condition. upon receipt and inspection of the device, damage to the device's impeller blades was identified. | [5 empty] | The impella cp and data logs were returned for evaluation. review of the device's data confirmed multiple temporary pump stops, followed by persistent suction alarms and low pump flows, thereafter. evaluation of the returned device found broken impeller blades. simulated use testing was successful in recreating low pump flow, after running device in a basin of water. device history record and found no manufacturing issues or reworks associated with this pump. of the 3 tumbled impeller lots, from which this component originated, one lot was scrapped for a concentricity nonconformity. there are no other complaints associated with pumps from this lot. the root cause of the low pump flow and suction was attributed to impeller blade damage.
37804 | 1220648-2019-00085 | 8605038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8605038&pc=OZD | 04/18/2019 | 04/18/2019 |  | 05/13/2019 |  |  | 04/18/2019 | 05/13/2019 |  |  | 02/28/2021 |  | 05/13/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1400928 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A patient admitted with cardiomyopathy needed his care escalated from an intra-aortic balloon pump (iabp) to the impella cp. on the second day of support, the team observed signs of hemolysis. pump position was checked via echo and the team decided to further escalate the care with placement of ecmo. at this time they also infused blood products for the hemolysis. the amount of blood infused was not shared for the investigation. the following day the team explanted and placed a bivad. | [5 empty] | To date no product or data logs have been shared for analysis, though both are expected. a final medwatch will be submitted upon completion of the investigation. | Since the original medwatch was filed, the product was returned. in house hemolysis testing and other investigative processes could be done. the pump was returned with clot on the pump housing. when the clot was removed the pump was run on the benchtop and passed in house hemolysis testing. the data logs were not returned for analysis. the clinical report did state that the pump was positioned in the papillary musculature and at the aortic valve, against recommendations. the root cause of the observed hemolysis was the combination of the malposition of the pump and the clot adhered to the pump. the failure mode will be monitored and trended.
37805 | 1220648-2019-00098 | 8605131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8605131&pc=OZD | 04/17/2019 | 04/17/2019 |  | 05/13/2019 |  |  | 04/17/2019 | 05/13/2019 |  |  | 01/31/2021 |  | 05/13/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1401830 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG | [10 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | VA ECMO | VA ECMO | [21 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | An impella cp was placed in (b)(6), for support of a patient in cardiogenic shock, that prompted coronary artery bypass surgery (cabg) and transfer to the tertiary medical center. during the support there was a question of hemolysis as the patient's plasma free hemoglobin levels were elevated on two lab draws. the team infused blood products to the patient and explanted the impella on the second day of support. the patient was stable post explant as he was also supported by va ecmo, veno-arterial extracorporeal membrane oxygenation. | [5 empty] | The complainant in the eu did return the impella pump for analysis. the hospital privacy policy does not allow the data logs to be analyzed. the investigation is on-going at this time. upon completion of the investigation, the final medwatch will be filed. | Since the original filing of the medwatch the investigation has concluded. the clinical report analysis found that during the impella case support high levels of plasma free hemoglobin confirmed true hemolysis was occuring. the fact that pump re-positioning did not remedy the issue suggests the pump was the cause of the hemolysis. the returned pump was inspected and run through benchtop hemolysis testing. during the testing, the sleeve bearing of the pump was found to have delamination of the epo-tek. the delamination can contribute to hemolysis. the failure mode will be monitored and trended.
37806 | 1220648-2019-00101 | 8605723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8605723&pc=OZD | 04/15/2019 | 06/06/2019 |  | 05/13/2019 |  |  | 04/15/2019 | 05/13/2019 |  |  | 11/30/2020 |  | 05/13/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1395292 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Material Twisted/Bent | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male presenting with post cardiotomy cardiogenic shock. during insertion of cp pump through aorta into the left ventricle, the cannula became kinked. as treatment, intervention with a snare was required to straighten the device. the device was removed and cardiac support was successfully achieved with a second device. | [5 empty] | The impella cp was received and a failure investigation is underway. a supplemental mdr will be filed with the results of the investigation, upon completion. | The impella cp was received and a failure investigation was completed. upon examination of the pump, and the fluoroscopy imagining provided, the catheter kink was confirmed. the location of the kink was identified directly above the outflow cage. the root cause of this failure was due to physician attempting to re-cross the aortic valve without a guide wire, after erroneously pulling the device out of the ventricle. a review of the device history record found no manufacturing issues or reworks associated with this pump lot. there are no other complaints associated with pumps from this lot.
37807 | 1220648-2019-00094 | 8600210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8600210&pc=OZD | 04/17/2019 | 04/17/2019 |  | 05/10/2019 |  |  | 04/17/2019 | 05/10/2019 |  |  | 05/31/2020 |  | 05/10/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED |  |  | US |  | 1351021 |  | 005042 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | Use of Device Problem | Insufficient Information | [5 empty] | Aortic Dissection | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A patient was scheduled for a high risk coronary angioplasty with the hemodynamic support of an impella 2.5 pump. the physician decided to insert the pump wirelessly despite recommendations not to do so. there was difficulty advancing the pump by the patient's previous intrahepatic portosystemic shunt (tips), though with effort the pump was advanced through the aorta. the physician was unable to cross the valve and place the impella within the left ventricle. the team removed the pump at this time and performed an aortogram and found there to be a infrarenal dissection in the aorta. the dissection closed itself without any intervention. the patient remained stable despite the dissection. the cause of the dissection is unknown. the patient was discharged home in stable condition. | [5 empty] | The investigation has completed since the initial medwatch was filed. the root cause of the dissection in the aorta was unable to be determined as the device, data, and imaging from the case was not returned despite all due diligence in retrieval attempts. the failure mode will be monitored and trended. | The medical facility did not return the product for analysis. the product was not used according to ifu recommendations of advancing to the left ventricle along a guidewire. as the ifu states: "advance the catheter through the hemostatic valve into the femoral artery (see figure 5.18) and along the placement guidewire and across the aortic valve using a fixed-wire technique. follow the catheter under fluoroscopy as it is advanced across the aortic valve, positioning the inlet area of the catheter 3.5 cm below the aortic valve annulus and in the middle of the ventricular chamber, free from the mitral valve chordae. be careful not to coil the guidewire in the left ventricle." the investigation is on-going at this time. whether a wireless insertion of the impella could have caused or contributed to the aortic dissection is not yet known at this time. a final medwatch will be filed at the investigation conclusion.
37808 | 1220648-2019-00079 | 8600465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8600465&pc=OZD | 04/16/2019 | 04/16/2019 |  | 05/10/2019 |  |  | 04/16/2019 | 05/10/2019 |  |  | 01/31/2021 |  | 05/10/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1400066 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A patient was having an emergent cardiac catheterization and angioplasty of the right coronary artery. during the intervention the patient condition deteriorated and so an impella cp was placed via the left femoral artery. after placement of the impella's repositioning sheath, the team observed a groin site bleed. the bleed was of a size that vascular surgery was consulted. the team chose to remove the impella and surgically repair the femoral artery with a viabahn covered stent. the team then placed another impella via the other leg, in the right femoral artery. | [5 empty] | The complainant did return the impella pump for analysis. the data logs did not pertain to the failure mode, and were not evaluated. when the pump was analyzed, there was no damage observed on the pump, inclusive of the repositioning sheath. the root cause of the vessel damage and associated blood loss could not be determined. the failure mode will be monitored and trended.
37809 | 1220648-2019-00099 | 8596035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8596035&pc=OZD | 04/12/2019 | 04/12/2019 |  | 05/09/2019 |  |  | 04/12/2019 | 05/09/2019 |  |  | 11/30/2020 |  | 05/09/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED |  |  | US |  | 1385393 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A patient admitted in cardiogenic shock and acute myocardial infarction in (b)(6) had an impella 2.5 placed via the left femoral artery. while being supported hemodynamically by the impella, the patient had coronary angioplasty completed successfully. the impella remained on for support while the patient was admitted to the icu for continued monitoring. while on the impella support the patient experienced limb ischemia. the team attempted to bypass the impella limb with an external bypass via a second introducer placed retrograde. the limb was not successfully perfused. during this time, the team also placed pcps ecmo and observed that there was an occlusion at the superficial femoral artery and that the patient was suffering from nonocclusive mesenteric ischemia (nomi). the patient expired with multiorgan failure. | [5 empty] | Since the original medwatch was filed, the investigation has completed. the complainant in japan never did return the pump or further clinical information, as requested. the console data logs do not pertain to the failure mode. the cause of the limb ischemia could not be determined. the failure mode will be monitored and trended. | The complainant in (b)(6) did not return the product for testing. the investigators are still awaiting a full clinical narrative as well as the console data logs. a final medwatch, with the failure investigation findings, will come upon completion of the investigation.
37810 | 1220648-2019-00092 | 8596211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8596211&pc=OZD | 04/05/2019 | 07/10/2019 |  | 05/09/2019 |  |  | 04/05/2019 | 05/09/2019 |  |  | 10/31/2020 |  | 05/09/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1383791 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 68 YR | [5 empty] | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female caucasian had impella cp pump inserted in the right femoral. on (b)(6) 2019 the patient had an antegrade sheath in the right leg to perfuse the limb but then patient had no pedal pulse in the morning as antegrade sheath was not effective so vascular surgeon took patient to operating room (or) to explant the impella and this resolved the ischemia. | [5 empty] | Corrected device evaluated by mfr from yes to no as the device was not returned for evaluation. a definitive root cause was not determined as the product was not returned. | The impella cp was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37811 | 1220648-2019-00088 | 8586226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8586226&pc=OZD | 04/12/2019 | 04/12/2019 |  | 05/07/2019 |  |  | 04/12/2019 | 05/07/2019 |  |  | 01/31/2021 |  | 05/07/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1397851 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 88 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | At the medical facility in (b)(6) a patient had an impella cp placed via the left femoral artery for a high risk coronary angioplasty. the angioplasty was completed, and the patient was sent to the icu on continued hemodynamic support with the impella cp. the weaning to explant was begun earlier than planned due to bleeding from the access site of the impella. the heparin concentration/anticoagulation was adjusted as well due to an elevated ptt (partial thromboplastin time) drawn bedside. the medical facility shared the ptt lab values and they are listed. the team additionally infused a total of 9 units of blood products to the patient. the impella pump was explanted and the patient was discharged home stable. the access site bleed needed no further intervention. | [5 empty] | Since the original medwatch was filed the investigation has completed. the pump analysis found no defects or damage that could have caused the access site bleeding. the failure mode will be monitored and trended. | The complainant in the eu has returned the product for analysis. the pump data logs have yet to be returned. the investigation is ongoing at this time. when the investigation completes, a final medwatch will be filed.
37812 | 1220648-2019-00093 | 8586446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8586446&pc=OZD | 04/08/2019 | 06/17/2019 |  | 05/07/2019 |  |  | 04/08/2019 | 05/07/2019 |  |  | 02/28/2021 |  | 05/07/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP OPTICAL PUMP |  |  | OZD |  |  | IMPELLA CP OPTICAL PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1400928 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 38 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female black or african-american had impella cp optical pump inserted. on (b)(6) 2019 the device was explanted, thrombus was fund in the impella leg so patient was taken for vascular surgery to remove it. | [5 empty] | The root cause of the leg ischemia was likely due to the clot observed in the patient leg because the leg ischemia was observed after pump removal. a definitive root cause was unable to be determined due to insufficient clinical details. | The impella cp optical was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37813 | 1220648-2019-00097 | 8587654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8587654&pc=OZD | 04/09/2019 | 04/09/2019 |  | 05/07/2019 |  |  | 04/09/2019 | 05/07/2019 |  |  | 01/31/2021 |  | 05/07/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1402854 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) asian female presenting for postpartum cardiomyopathy support. the impella cp was placed without any reported issues, however, the access site began to bleed approximately 30 minutes after arriving in the intensive care unit. the source of the leak was found at the repositioning sheath's blue cuff, as it was inserted just under the skin. treatment included adjusting the repositioning sheath, application of a femstop, and an additional suture. additional intervention was required in the form of blood products. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37814 | 1035166-2019-00032 | 8569110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8569110&pc=DYB | 04/05/2019 | 05/01/2019 |  | 05/01/2019 |  |  | 04/08/2019 | 05/01/2019 |  |  | 12/01/2021 |  | 05/01/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA, |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-16186 |  | 0052-3000 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN | [10 empty] | Injury | Yes | Yes | Structural Problem | Deformation Due to Compressive Stress | [5 empty] | Death | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that patient was presented to physician in cardiogenic shock. patient has very complex anatomy and calcification. during access of the right femoral artery for placement of the impella cp, 14fr peelaway sheath burst for a few millimeters but no bleeding problems occurred. user left the peelaway sheath to prevent bleeding problems form the puncture. user placed 13fr merit sheath 2 cm above and implella 2.5 placed successfully. there was no bleeding problem or leg perfusion problem noticed but patient died next day because of multi organ failure. date of death is not available. product was discarded. as per additional information, patient has very difficult anatomy, physician thinks the he kinked the peelaway sheath and he don't remember where it does exactly kinks and burst. he remember that the burst must be roughly 5mm. no other additional information is available. | [5 empty] | The device was used for treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. the investigation will focus on a review of product documentation. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. as per procedure adelante s2s introducer sheath in-process and final inspection: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. per ifu: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
37815 | 1220648-2019-00077 | 8569489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8569489&pc=OZD | 04/05/2019 | 06/03/2019 |  | 05/01/2019 |  |  | 04/05/2019 | 05/01/2019 |  |  | 11/30/2020 |  | 05/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA LP 2.5 PUMP |  |  | OZD |  |  | IMPELLA LP 2.5 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1385393 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male had impella lp 2.5 pump inserted in the right femoral. on (b)(6) 2019 the patient developed limb ischemia in lower limb, as a result an antegrade sheath was inserted. the next day the color of the lower limbs improved. | [5 empty] | The impella lp 2.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The root cause of the leg ischemia could not be determined because the introducer was not returned for evaluation.
37816 | 1220648-2019-00073 | 8569494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8569494&pc=OZD | 04/03/2019 | 04/03/2019 |  | 05/01/2019 |  |  | 04/03/2019 | 05/01/2019 |  |  | 10/31/2020 |  | 05/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA LP 5.0 PUMP |  |  | OZD |  |  | IMPELLA LP 5.0 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1387941 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | DISTRI |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Information | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) female caucasian had impella lp 5.0 pump inserted. on (b)(6) 2019, the patient had bleeding at multiple femoral sites including the groin, as a results prr - 29 with inr >2, fresh frozen plasma (ffp), vitamin k, platelets (plts), and 2 units of packed red blood cells (prbc) was administered. | [5 empty] | The impella lp 5.0 pump has not been returned by the customer. if the pump is returned a failure analysis investigation will be underway. a supplemental mdr will be filed at the completion of the device's investigation.
37817 | 1220648-2019-00078 | 8560115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8560115&pc=OZD | 04/04/2019 | 04/04/2019 |  | 04/29/2019 |  |  | 04/04/2019 | 04/29/2019 |  |  | 10/31/2020 |  | 04/29/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1387323 |  | 0048-0032 |  | Yes | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 57 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A patient was admitted for a high risk percutaneous angioplasty in the cardiac catheterization lab. the patient had a multivessel, inclusive of the left main coronary artery, angioplasty successfully completed with the hemodynamic support of an impella cp, which was inserted via the left femoral artery. after the angioplasty was completed, the patient was sent with the impella for continued hemodynamic support to the icu. during the overnight the patient developed a hematoma at the imepella access site. the team held manual pressure and applied a femostop. the pump remained on for support, but the groin site continued to ooze blood. due to the blood loss and hematoma the team infused 1 unit of blood products. on the fourth day of support, the pump was weaned and explanted. there was no further intervention needed for the blood loss and hematoma. | [5 empty] | To date the medical facility has yet to return to abiomed the pump product or data logs. the investigation is on-going at this time. should the facility return the product and a device analysis is done, a supplemental medwatch will be submitted. during the investigation the failure mode will be monitored and trended. | Since the original medwatch was filed the investigation has concluded. the product was returned for analysis. analysis found no product defects that could have injured the blood vessel. the root cause of the hematoma could not be determined. the failure mode will be monitored and trended.
37818 | 1220648-2019-00075 | 8561687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8561687&pc=OZD | 03/25/2019 | 04/02/2019 |  | 04/29/2019 |  |  | 04/02/2019 | 04/29/2019 |  |  | 08/31/2020 |  | 04/29/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1360977 |  | 0048-0003 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 38 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male had impella cp pump inserted in the right femoral. patient was transferred to another hospital, upon arrival reposition sheath was bleeding heavily. as a result, patient had a surgical repair of the femoral artery, suture was used, a cut down was done to access the artery and a 4-0 or 5-0 prolene was used for repair. some blood was lost but was given back through the cellsaver. | [5 empty] | The impella cp pump was discarded by the customer; therefore, a failure analysis investigation cannot be completed.
37819 | 1035166-2019-00035 | 8561739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8561739&pc=DYB | 03/29/2019 | 09/12/2019 |  | 04/29/2019 |  |  | 04/04/2019 | 04/29/2019 |  |  | 12/01/2021 |  | 04/29/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-15975 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN |  |  |  |  | 62 YR |  |  | 80 |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | Blood Loss | [41 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that on 29th march patient was presented for cardiogenic shock. during acute myocardial infarction procedure physician was unable to pass pump through introducer as introducer was kinked and preventing passage of impella cp. site was dilated, introducer was placed, and pump insertion was attempted. the physician advised that no excessive force was used when implanting the introducer. patient lost 100-200 ml of blood. however, no blood product or surgical intervention needed. kinked introducer was removed and new introducer (other brand) was used to complete surgery. as per additional information, introducer kinked during insertion. for insertion physician used dilator and no dilation performed. there was 10 -15 min delay reported to change introducer. device was not used in tortuous path. procedure completed with abiomed 14fr x 13cm introducer, no impella was use. | [5 empty] | The device used in treatment. one 14f abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. blood was observed on and inside the sheath and sideport tubing. the sheath was kinked approximately 9cm from the distal tip. the sheath body was also flattened and creased down the lumen approximately 7cm from the distal tip. there is a 3mm split in the distal tip of the sheath. returned device analysis reveals the 14f abiomed introducer sheath is within manufacturing specifications. potential contributing factors to the damage is if the device was advanced through an area of resistance (as stated in the precautions of the ifu), or if the device was subjected to unusual stresses (as stated in the handling and storage section of the ifu). no manufacturing defects were found. the device history records were reviewed to confirm that the device passed all applicable in process and final inspections. per procedure abiomed introducer sheath in-process and final inspection, the devices in this lot were inspected as follows: visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. this is performed by qa on the first 5 consecutive properly tipped parts. with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. this is performed by qa on 100% of the sheaths. per ifu : never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. no further follow-up is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress ,follow up report will be submitted if we find any further additional information.
37820 | 1220648-2019-00087 | 8557763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8557763&pc=OZD | 03/27/2019 | 03/27/2019 |  | 04/26/2019 |  |  | 03/27/2019 | 04/26/2019 |  |  | 09/30/2020 |  | 04/26/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED |  |  | US |  | 1387425 |  | 005042 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A patient with cancer and mutlivessel coronary artery disease was having a high risk angioplasty in the cardiac catheterization lab, and had an impella 2.5 pump placed for hemodynamic support via the left femoral artery. the pump supported the patient for the angioplasty but, once admitted to the icu for continued care, he developed limb ischemia in the impella limb. the pulses were lost to the left foot. the pump was explanted, vascular surgery performed a ax-fem bypass, and his hemodynamic care was escalated to the impella 5.0 via the axillary artery. | [5 empty] | The root cause of the limb ischemia could not be definatively determined. the product was never returned and the clinical report from the medical team did not provide any information as to what the cause of the ischemia was. the root cause most likely was the pre-existing peripheral disease that was exacerbated by the addition of the vasoconstricting medication, levophed. the failure mode will be monitored and trended. | The investigation is on-going at this time. no product, data logs, or further clinical report has been shared. upon completion of investigation, a final medwatch will be submitted.
37821 | 1035166-2019-00033 | 8544259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8544259&pc=DYB | 02/28/2019 | 07/24/2019 |  | 04/24/2019 |  |  | 04/02/2019 | 04/24/2019 |  |  | 08/01/2021 |  | 04/24/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA. |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-15174 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | Yes | Leak/Splash | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that on (b)(6) 2019 patient was presented to physician with low ejection fraction (~20%). during protected percutaneous coronary intervention (pci) for triple vessel disease (tvd) required rotablation a 14 fr introducer being used. after inserting the pump into the sheath and starting impella, physician discovered that the hemostatic valve of the peel away sheath is oozing blood. as per physician re-po sheath is slightly in touch with the sheath during impella insertion but it shouldn't cause the oozing. physician is expecting to complete the revasculation with impella support. due to blood loss and dropped bp, he aborted the procedure and blood transfusion performed. they removed the peel away sheath and inserted the repositioning sheath to stop the bleeding. physician completed one vessel and they stage the other 2 vessels, but they completed re-vas before patient discharged. adrenaline and amiodarone used as medical intervention. amount of blood loss is unknown and patient outcome is stable. | [5 empty] | Our investigation is still in progress, follow up report will be submitted if we find any further additional information. | One 14f 13 cm long abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. the sheath was received already peeled apart. both orange split caps were attached to the sheath's hub. visible blood was found on the sheath and inside the sideport tubing. according to the complaint , after inserting the pump into the sheath and starting impella, the physician discovered that the hemostatic valve of the peel away sheath was oozing blood. one of the helical cuts on the hemostatic valve was torn on the sideport half of the sheath. the introducer sheath looked like it peeled apart normally. since the introducer sheath was received back from the field to the customer already peeled apart, a leak test could not be performed. potential contributing factors to the damages of the sheath (resulting in leakage) is if the device was advanced through an area of resistance, or if the device was subjected to unusual stresses or the device was not inserted through center of valve membrane causing inadvertent puncture. no manufacturing defects were found. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure adelante s2s introducer sheath in-process and final inspection: vacuum leak test - perform pressure and vacuum leak test per procedure. this test is performed by qa on 100% of the product. per ifu: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. infusion through the sideport can be done only after all air is removed from the unit. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
37822 | 1220648-2019-00083 | 8534218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8534218&pc=OZD | 03/28/2019 | 05/16/2019 |  | 04/20/2019 |  |  | 03/28/2019 | 04/20/2019 |  |  | 08/31/2019 |  | 04/20/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1311673 |  | 005042 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Blood Loss | [10 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male caucasian presenting with acute myocardial infarction and cardiogenic shock for cardiac support with the impella lp2.5. the device was placed without issue, however, oozing from the insertion site developed during support and a hematoma formed. additionally, patient showed signs of hemolysis via hemolyzed labs. as treatment, blood products were delivered and the device was removed. the patient remained in stable condition and a 2nd pump insertion was not required. | [5 empty] | The impella lp2.5 was received and a failure investigation was completed. data logs found pump to have operated to specification. simulated use testing and visual inspection found no defects. a review of the device history record found no manufacturing issues or reworks associated with this pump lot. there are no other complaints associated with pumps or repositioning sheaths from this lot. the root cause of low flows and hemolysis were unable to be definitively determined. | The impella lp2.5 has not been returned from the customer, therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37823 | 1220648-2019-00082 | 8534219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8534219&pc=OZD | 12/07/2018 | 04/03/2019 |  | 04/20/2019 |  |  | 04/03/2019 | 04/20/2019 |  |  | 09/30/2020 |  | 04/20/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1370693 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) exhibiting signs of hemolysis via bloody urine and increased ldh values, as a result of suction from device. as treatment, 30 units of blood products were delivered. the device was removed and replaced with an intra-aortic balloon pump. the patient remained in stable condition. | [5 empty] | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37824 | 1220648-2019-00084 | 8534220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8534220&pc=OZD | 04/06/2019 | 06/12/2019 |  | 04/20/2019 |  |  | 04/06/2019 | 04/20/2019 |  |  | 10/31/2020 |  | 04/20/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1388157 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | Yes | Material Twisted/Bent | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male caucasian presenting with acute myocardial infarction and cardiogenic shock for cardiac support with the impella cp. during support, the device fell out of the left ventricle and physician was unable to re-cross the aortic valve for proper repositioning. the device's cannula had kinked, requiring intervention via snare in order to unfold and explant the device. patient received cardiac support with insertion of a second pump. patient did not suffer any lasting harm as a result. | [5 empty] | The impella cp was received and 2 kinks were confirmed on the pump's catheter. no further damage or deficiencies were observed. a review of the device history record found no manufacturing issues or reworks associated with this pump lot. there are no other complaints related to this failure mode associated with pumps from this lot. the root cause of these kinks were attributed to the clinical account attempting to wirelessly re-access the left ventricle, which is contradictory to the instructions for use for this device. | The impella cp has been received from the customer and an investigation is underway. a supplemental report will be filed with the results of the investigation.
37825 | 1220648-2019-00080 | 8533292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8533292&pc=OZD | 11/25/2018 | 06/10/2019 |  | 04/19/2019 |  |  | 04/08/2019 | 04/19/2019 |  |  | 04/30/2020 |  | 04/19/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1347537 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male asian patient exhibiting aortic regurgitation (ar) during cardiac support with the impella lp2.5 pump. the physician alleged the device had possibly migrated towards the aorta, causing the ar. as treatment, the device was removed. a second device was not required as patient was hemodynamically stable. | [5 empty] | The impella lp2.5 was received and a failure investigation was completed. visual inspection found no significant damage or biomaterial. the clinical account had went against the impella ifu recommendations to remove the device during extended pump stoppage, thus, likely contributing to the aortic regurgitation reported. simulated use testing could be definitively determined, as the pump stopped immediately during testing. a review of the device history record found no manufacturing issues or reworks related to this failure mode for any pump in this lot. there are no other complaints for other pumps in this lot associated with this failure mode. | The impella lp2.5 was received and a failure analysis investigation is underway. a supplemental report will be submitted upon completion of the evaluation.
37826 | 1220648-2019-00072 | 8527565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8527565&pc=OZD | 03/07/2019 | 04/03/2019 |  | 04/18/2019 |  |  | 04/03/2019 | 04/18/2019 |  |  | 07/31/2020 |  | 04/18/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1354010 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Use of Device Problem | [6 empty] | Insufficiency, Valvular | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A patient transferred from the community hospital in (b)(6), to a tertiary medical center in (b)(6), had an impella 2.5 placed due to post cardiotomy cardiogenic shock. the patient had mild aortic valve regurgitation prior to impella placement. during the placement of the impella the pump retracted back into the aorta and was re-positioned wirelessly back across the valve. the impella pump supported the patient for approximately 31 hours and was then removed. after the pump was removed, the patient was observed to have severe aortic regurgitation. the patient was sent to the or for valve repair. | [5 empty] | The clinical account states that the pump was re-advanced across the aortic valve wirelessly. it is likely that this wireless insertion led to increased valve damage. the valve was already known to be failing. the severe valve regurgitation prompted valve repair. the root cause can not be definitively determined as no product was returned, incomplete data logs, and further clinical details were not shared by the (b)(6) team. the failure mode will be monitored and trended.
37827 | 1220648-2019-00074 | 8527702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8527702&pc=OZD | 02/04/2019 | 04/01/2019 |  | 04/18/2019 |  |  | 04/01/2019 | 04/18/2019 |  |  | 06/30/2020 |  | 04/18/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA LP 2.5 PUMP |  |  | OZD |  |  | IMPELLA LP 2.5 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1352046 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 46 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese female had impella lp 2.5 pump inserted in the left femoral. on (b)(6) 2019 patient developed limb ischemia, as a result femoral angioplasty was used. | [5 empty] | The impella lp 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37828 | 1220648-2019-00076 | 8524858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8524858&pc=OZD | 12/21/2018 | 04/03/2019 |  | 04/17/2019 |  |  | 04/03/2019 | 04/17/2019 |  |  | 09/30/2020 |  | 04/17/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA LP 2.5 PUMP |  |  | OZD |  |  | IMPELLA LP 2.5 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1370693 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male had impella lp 2.5 pump inserted in the left femoral. on (b)(6) 2018 patient developed limb ischemia in lower left thigh and color got worse so physician removed the impella device. | [5 empty] | The impella lp 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37829 | 1035166-2019-00031 | 8499284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8499284&pc=DYB | 03/19/2019 | 04/10/2019 |  | 04/10/2019 |  |  | 03/20/2019 | 04/10/2019 |  |  | 07/01/2021 |  | 04/10/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-15070 |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 50 YR | [5 empty] | Injury | Yes | Yes | Obstruction of Flow | [6 empty] | Ischemia | Cardiogenic Shock | Blood Loss | [10 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that an incident occurred during a surgical procedure for placement of an impella on (b)(6) 2019. after the procedure was over (amifostine, cardiogenic shock impella insertion), while the patient was in icu, the patient's leg was cold and the pulse was absent. the patient went back to the cath lab for the contralateral bypass for limb ischemia. the patient was draped in standard fashion. the oscor sheath 14f x 25cm was inserted into the right femoral artery, the impella was placed through that. oozing occurred at the insertion site, preventing us from peeling away. continuous oozing at the sheath site resulted in transfusion; two units of red blood cells was given and an estimated amount of blood loss 4 grams of hemoglobin. the oscor sheath 14f x 25cm was occlusive (not allowing blood flow to pass and down the leg). we then placed a terumo rabbi sheath that is 45cm long in the left femoral artery, advanced the distal tip past the oscor sheath and down the right common femoral artery to provide distal flow to the lower leg. blood loss was resolved after impella and sheath were removed, perclose deployed. left femoral artery procedure took approximately 1 hour. outcome of patient is alive and doing well. device will not be returned as it was discarded. | [5 empty] | The device was used in treatment. the device will not be returned for analysis, therefore, the clinical observation could not be confirmed. the investigation will focus on a review of product documentation. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per qa adelante-s2s introducer sheath in process and final inspection procedure: preform occlusion test per procedure, per latest revision. visual inspection: with the naked eye at a distance of 12" to 18", inspect sideport assembly for excessive glue. a bead of glue around the perimeter of the sideport tube to hub joint is acceptable. smeared glue on sideport assembly is not acceptable. the instructions for use (ifu) instructs the user: aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. aspirate all air from the sheath valve assembly by using a syringe connected to the sideport. flush the introducer through the sideport. if the introducer is to remain in place during catheter positioning and testing, flushing the introducer via the sideport periodically with saline is advised. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type.
37830 | 1220648-2019-00071 | 8501285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8501285&pc=OZD | 03/14/2019 | 03/14/2019 |  | 04/10/2019 |  |  |  | 04/10/2019 |  |  | 10/31/2020 |  | 04/10/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA LP 2.5 PUMP |  |  | OZD |  |  | IMPELLA LP 2.5 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1377030 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male had impella lp 2.5 pump inserted in the left femoral at the time of rotablator from left main trunk (lmt) to left anterior descending (lad). patient had low heart function after cardia arrest and suspected ischemic heart disease. the patient had lower limb ischemia, the pump was removed, and surgical suture was applied. the color of the patient's feet improved after pump removal. | [5 empty] | The impella lp 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37831 | 1220648-2019-00066 | 8496804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8496804&pc=OZD | 03/14/2019 | 03/14/2019 |  | 04/09/2019 |  |  | 03/14/2019 | 04/09/2019 |  |  | 12/31/2020 |  | 04/09/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1397934 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Material Integrity Problem | Patient Device Interaction Problem | [5 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient was experiencing post cardiotomy cardiogenic shock after have coronary artery bypass and aortic and mitral valve replacements. due to the shock, the team chose to place an impella cp for support. upon transfer from the operating room to the icu, there was a hematoma that developed at the impella access site. the team observed the bleeding at the access site during the first 24 hours of support. on the second day of impella support the team placed a femostop and infused blood products. after 89 hours of impella support the patient condition improved, the pump was weaned and explanted. | [5 empty] | The medical facility did return for analysis the impella pump, but not the access site introducer nor the case data logs. upon inspection of the pump there was an observation made of a slight bend and buckling of the pump's reposistioning unit. the cause of the defect was unknown. the buckling of the unite likely caused a leak path to form for bleeding to occur at the access site. it is unknown what caused the buckling to occur. the cause of the access site bleeding and hematoma was most likely the buckling. the failure mode will be monitored and trended.
37832 | 1220648-2019-00070 | 8492614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8492614&pc=OZD | 03/10/2019 | 05/30/2019 |  | 04/08/2019 |  |  | 03/11/2019 | 04/08/2019 |  |  | 12/31/2020 |  | 04/08/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA LP 5.0 PUMP |  |  | OZD |  |  | IMPELLA LP 5.0 PUMP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1388547 |  | 005060 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 29 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old caucasian female had impella lp 5.0 pump inserted and the patient increasingly developed signs of hemolysis so the doctor pulled the pump. the patient's lactate dehydrogenase (ldh)>200, hemolysis index and bilirubin were also done, there was renal injury anuria and as a result the patient received eight (8) units of blood from (b)(6) to (b)(6) 2019. patient is recovering on ecmo support. | [5 empty] | The impella lp 5.0 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The root cause of hemolysis is biomaterial. the pump failed hemolysis test during failure investigation testing with biomaterial attached to the impeller. the pump passed the hemolysis test after the biomaterial was removed.
37833 | 1220648-2019-00056 | 8482390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8482390&pc=OZD | 03/08/2019 | 04/24/2019 |  | 04/04/2019 |  |  | 03/08/2019 | 04/04/2019 |  |  | 09/30/2020 |  | 04/04/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1380173 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 50 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old caucasian male morbidly obese with patient medical history (pmhx): hypertension (htn), diabetes mellitus (dm), coronary artery disease (cad) who was recently diagnosed with flu 2 weeks ago had impella cp pump inserted in the right femoral artery without issues. patient went into ventricular tachycardia (vt) arrest due to reperfusion and was shocked once. a large hematoma was noted at the impella site and bleeding as a result a topical thrombin was placed, and patient was taken to the intensive care unit (icu). | [5 empty] | The impella cp was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The antiplatelet drug: aggrastat may have limited hemostasis at the access site. a definite cause for bleeding around the repositioning sheath was unable to be determined.
37834 | 1220648-2019-00068 | 8477790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8477790&pc=OZD | 01/31/2019 | 03/07/2019 |  | 04/03/2019 |  |  | 03/07/2019 | 04/03/2019 |  |  | 09/30/2020 |  | 04/03/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1370693 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 84 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old asian male presenting with acute myocardial infarction and cardiogenic shock for cardiac support with the impella lp2.5 pump. during support, a bleed from the access site had formed. as treatment, the device was removed, 10 units of blood products were delivered, and an intra-aortic balloon pump was used as a replacement cardiac support device. | [5 empty] | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37835 | 1220648-2019-00069 | 8477811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8477811&pc=OZD | 01/29/2019 | 03/08/2019 |  | 04/03/2019 |  |  | 03/08/2019 | 04/03/2019 |  |  | 05/31/2020 |  | 04/03/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1361441 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG | [10 empty] | Injury | Yes | No | Use of Device Problem | [6 empty] | Mitral Insufficiency | [12 empty] | Other | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported an male patient presenting with acute myocardial infarction and cardiogenic shock for cardiac support with the impella cp pump. support commenced and ventricular function had improved. post explant, however, mitral insufficiency was exhibited via a newly formed flail leaflet that was identified. the cause of the flail leaflet was "most probably" due to a chordae rupture. the impella device may have contributed to this condition, but it was not conclusively determined. note: impella instructions for use provide clear indication on how to avoid chrodal trauma, when placing the impella within the subannular mitral apparatus. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation. impella 2.5 and impella cp circulatory support systems user manual, 0048-9006 rev a, page 11: "advance the catheter through the hemostatic valve into the femoral artery and along the placement guidewire and across the aortic valve using a fixed-wire technique. follow the catheter under fluoroscopy as it is advanced across the aortic valve, positioning the inlet area of the catheter 3.5 cm below the aortic valve annulus and in the middle of the ventricular chamber, free from the mitral valve chordae. be careful not to coil the guidewire in the left ventricle."
37836 | 1220648-2019-00057 | 8478364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8478364&pc=OZD | 02/19/2019 | 07/17/2019 |  | 04/03/2019 |  |  | 03/07/2019 | 04/03/2019 |  |  | 09/30/2020 |  | 04/03/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA LP 2.5 PUMP |  |  | OZD |  |  | IMPELLA LP 2.5 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1367870 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hypoxia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) female had impella lp 2.5 pump inserted in the left femoral before coronary artery bypass graft (cabg). patient had spontaneous coronary artery dissection (scad), dissociation of left anterior descending (lad) and circumflex (cx) was observed during coronary angiography (cag) examination and dissociation progressed with contrast agent. on (b)(6) 2019 patient had cerebrovascular accident (cva) and stroke. the hypoxic encephalopathy due to cardiopulmonary arrest. ischemic condition was found with mri check. | [5 empty] | The impella lp 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed. | We received new information that the cerebrovascular accident (cva) was due to hypoxic encephalopathy and this happened during cardiopulmonary arrest (cpa) which was before impella insertion. this event has no relation to impella device.
37837 | 1220648-2019-00061 | 8479718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8479718&pc=OZD |  | 03/06/2019 |  | 04/03/2019 |  |  | 03/06/2019 | 04/03/2019 |  |  | 01/31/2020 |  | 04/03/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1324747 |  | 005048-JP |  | No | [5 empty] | No | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Hemolysis | Renal Failure | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The medical team in (b)(6) placed an impella 2.5 for support for a patient who had been admitted in shock. the pump remained on for support for over 18 hours when it was explanted due to hemolysis concerns. months later the medical team published that the patient suffered renal failure and treated the patient with dialysis | [5 empty] | The medical team did not save for analysis the impella pump. no in-house hemolysis testing could be done. the publication and case report were reviewed for analysis. hemolysis was reported during support and presented as hematuria and elevated ldh. the hemolysis was not definitively diagnosed. the publication stated that the patient suffered renal failure attributed to the hemolysis that was treated with hemodialysis. the pump was noted to be "facing the posterior wall." improper positioning of the device may have contributed to hemolysis, but could not be evaluated as the data logs, pump, and imaging were not provided for this investigation.
37838 | 1035166-2019-00028 | 8475868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8475868&pc=DYB | 02/27/2019 | 04/02/2019 |  | 04/02/2019 |  |  | 03/08/2019 | 04/02/2019 |  |  | 10/01/2021 |  | 04/02/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-15224 |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | OTHER | [5 empty] | 53 YR | [5 empty] | Injury | Yes | Yes | Leak/Splash | [6 empty] | Ischemia | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that patient was admitted pre-operative placement of the impella prior to coronary artery bypass grafting surgery. introducer was used for insertion at the right femoral artery for an impella cp. the physician had to leave 14fr oscor in for 2 days from 25th february to 27th february, with external bypass placed to reduce limb ischemia, on 27th february blood loss was noticed at the introducer which saturated the groin site dressing. the blood loss was occurring through the valve of the introducer sheath. the amount of blood loss is unknown. the patient received 1 unit of blood replacement. bleeding required introducer explant and advancement of the impella repositioning sheath. during explant of introducer vessel damage was observed, as per physician vessel damage is related to patient movement. patient was taken to hybrid or to surgically repair impella cp access site in the right leg. right femoral, and right profunda surgical patch and right lower extremity thromboembolectomy and removal of atherosclerotic plaque was performed. patient maintained hemodynamic stability through the surgical procedure and was taken to icu. procedure completed with another introducer and pump from other leg. no additional information is available and no other adverse event reported. | [5 empty] | Device was used in treatment. the device was not returned for analysis. therefore, the clinical observation could not be confirmed. the investigation will focus on a review of product documentation. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. as per procedure adelante s2s introducer sheath in-process and final inspection: vacuum leak test - perform pressure and vacuum leak test per procedure. this test is performed by qa on 100% of the product. per instructions for use (ifu): when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. as per ifu bleeding is possible complications/adverse effects. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. no further follow-up is required.
37839 | 1220648-2019-00055 | 8465296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8465296&pc=OZD | 03/01/2019 | 04/12/2019 |  | 03/29/2019 |  |  | 03/01/2019 | 03/29/2019 |  |  | 11/30/2020 |  | 03/29/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA LP 2.5 PUMP |  |  | OZD |  |  | IMPELLA LP 2.5 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1378913 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 57 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese male had impella lp 2.5 pump inserted in the left femoral for high risk percutaneous coronary intervention (hrpci) with three (3) branch lesions. at night, bleeding occurred from the site which a vascular access catheter was inserted and the patient condition changed emergently to hemorrhagic shock. as treatment, ten (10) units of rbc and three (3) units of fresh frozen plasma (ffp) was administered along with manual compression to help. | [5 empty] | The cause of hemorrhagic shock was because of the patient's inability to achieve hemostasis at the access site due to the high dose of heparin that was administered at 50iu/ml and the repositioning sheath not reaching the femoral artery due to patient's thick adipose tissue. | The impella lp 2.5 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37840 | 1220648-2019-00053 | 8465302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8465302&pc=OZD | 01/05/2019 | 03/01/2019 |  | 03/29/2019 |  |  | 03/01/2019 | 03/29/2019 |  |  | 08/31/2020 |  | 03/29/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA 2.5 PUMP |  |  | OZD |  |  | IMPELLA LP 2.5 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1370701 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese male had impella lp 2.5 pump inserted in the right femoral without any problem. the impella was withdrawn on the fourth day due to mitral regurgitation (mr) at that time the mr was not severe. treatment was diuretic and dob (dobutamine hydrochloride) was used, heart failure management was done and confirmation of severe mr of valve leaflet on lateral side after transesophageal echocardiography (tee) and a 3d image that the chordae rupture was present. | [5 empty] | The impella lp 2.5 pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37841 | 1220648-2019-00052 | 8461685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8461685&pc=OZD | 03/01/2019 | 03/01/2019 |  | 03/28/2019 |  |  | 03/01/2019 | 03/28/2019 |  |  | 08/31/2020 |  | 03/28/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1380160 |  | 005082 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Activation, Positioning or Separation Problem | Patient Device Interaction Problem | [5 empty] | Mitral Regurgitation | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A patient had an impella ld placed due to shock post cardiotomy for a bypass surgery. the patient had four bypasses placed. the patient was sent to the icu with an open chest and the impella running for hemodynamic support. after 20 days of support, there was a diagnosis made of severe mitral valve regurgitation (mr). the team performed a surgical repair to the valve. subsequent to the valve surgery, the team reviewed echo imagining and found that the impella had been malpositioned up against the posterior leaflet of the valve for 7 days. this malposition had not been observed and recognized when the echo was done. | [5 empty] | Since the original medwatch was filed, the investigation has completed. the product and data logs were analyzed. the data logs revealed that pump flows were within specification throughout the case. there were positioning and suction issues that the medical team resolved by repositioning. the pump analysis found no irregularities, all pump surfaces were smooth. the cause of the mitral valve damage was most likely improper pump positioning. the failure mode will be monitored and trended. | The investigation is on-going at this time. a final report will be submitted at the conclusion of the investigation.
37842 | 1220648-2019-00067 | 8462002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8462002&pc=OZD | 02/26/2019 | 02/26/2019 |  | 03/28/2019 |  |  | 02/26/2019 | 03/28/2019 |  |  | 08/31/2020 |  | 03/28/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1361824 |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female presenting for cardiac support with an impella cp as a result of cardiomyopathy. during support, patient exhibited hematuria, therefore, device was removed. it's not clear if any tests for plasma free hemoglobin were measured in order to confirm hemolysis. as treatment, packed red blood cells (blood products) were administered to the patient over the course of several days. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37843 | 1220648-2019-00065 | 8453675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8453675&pc=OZD | 10/13/2018 | 04/02/2019 |  | 03/26/2019 |  |  | 02/27/2019 | 03/26/2019 |  |  | 02/29/2020 |  | 03/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1331385 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male presenting for cardiac support with the impella lp 2.5. the device was inserted without issue, patient received successful cardiac support, and device was explanted. approximately 12 days after explant, before patient was to be discharged from the hospital, the impella access site began to bleed. the blood loss was controlled and 6 units of red blood cells and fresh frozen plasma were delivered. the physician attributed the cause of this to the impella device, patient condition, and combination therapy. | [5 empty] | Additional information was received from the physician who performed the procedure with this device. the patient's specific ailment was the development of a hematoma at the insertion site of the impella, however, new information provided indicates physician deemed no relationship between this hematoma and the impella device. therefore, the original assessment has changed from reportable to not reportable as this injury was not related and device was not inserted at the time of occurrence. | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37844 | 1220648-2019-00062 | 8448004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8448004&pc=OZD | 11/07/2018 | 02/25/2019 |  | 03/25/2019 |  |  | 02/25/2019 | 03/25/2019 |  |  | 08/31/2020 |  | 03/25/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1363564 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese male presenting for impella cardiac support during high risk percutaneous coronary intervention. during support, a hematoma had formed at the device's access site. as treatment, blood products and surgical vascular repair was administered. the patient's condition improved, thereafter. | [5 empty] | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37845 | 1220648-2019-00063 | 8448028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8448028&pc=OZD | 06/12/2018 | 03/27/2019 |  | 03/25/2019 |  |  | 02/25/2019 | 03/25/2019 |  |  | 04/30/2019 |  | 03/25/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1300756 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old japanese female presenting for impella cardiac support due to acute myocardial infarction and cardiogenic shock. during placement of device, the access site bled, requiring 5 units of blood product administered. during support, the patient's leg, in which the device was inserted, became ischemic. as treatment, non-cardiac surgery was performed. thereafter, patient recovered from these ailments and received successful support from the impella device. | [5 empty] | Updated information received from complaint source (physician): new information was received from the physician involved in this complaint: there was no relationship between the ischemia and the impella device, rather, patient's original condition. non-device related ischemia was treated by adjusting the repositioning sheath. in regards to the original report of access site bleeding: the original cause was stated as "procedure or surgery" however, new information from the physician reveals this was not related to the impella device. as a result, the new information provided will change the assessment of this complaint from reportable to not reportable, as no adverse events were related to the impella device, as stated by the physician involved. | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37846 | 1220648-2019-00064 | 8448710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8448710&pc=OZD | 02/28/2019 | 02/28/2019 |  | 03/25/2019 |  |  | 02/28/2019 | 03/25/2019 |  |  | 01/31/2021 |  | 03/25/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1393531 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A (b)(6) year old female caucasian presented with post cardiomyopathy cardiogenic shock. an impella cp was inserted for cardiac support, without issue. the following day, patient exhibited loss of pulses in the ulnar and radial arteries. as treatment, the impella was removed and a new impella device was inserted in the opposing arm. patient endured no permanent injury and ischemia resolved, thereafter. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37847 | 1220648-2019-00048 | 8437421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8437421&pc=OZD | 02/24/2019 | 02/24/2019 |  | 03/20/2019 |  |  | 02/24/2019 | 03/20/2019 |  |  | 10/31/2020 |  | 03/20/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED |  |  | US |  | 1377030 |  | 005048-JP |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant in (b)(6) had a patient admitted in cardiogenic shock. the plan was to place an impella for hemodynamic support. the access of the patient's left femoral artery was complicated due to peripheral arterial disease. the right superficial femoral artery (sfa) was totally occluded and the left sfa had a stenosis. the team did insert the impella via the left limb access, but there was difficulty visualizing whether the pump was inserted via the sfa or deep femoral artery (dfa). the patient had multivessel coronary artery disease and was stented. the angioplasty was a success and the patient was sent to the icu with the impella pump. the following day, the left limb was seen to be ischemic and the impella pump was explanted and the site was surgically closed. | [5 empty] | Since the initial filing of the medwatch the investigation has completed. the root cause of the limb ischemia was most likely due to the small vascular size and known arterial stenosis. this could not be definitively determined as the pump was not returned. the failure mode will be monitored and trended. | The investigation is on-going at this time. no conclusion is yet available. a final medwatch will be filed upon investigation completion.
37848 | 1220648-2019-00058 | 8437442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8437442&pc=OZD | 11/17/2018 | 02/25/2019 |  | 03/20/2019 |  |  | 02/25/2019 | 03/20/2019 |  |  | 08/31/2020 |  | 03/20/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED |  |  | US |  | 1367668 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 86 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Blood Loss | Pseudoaneurysm | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | An (b)(6) year old male in (b)(6) with multiple cardiac comorbidities had an impella 5.0 placed via the right axillary artery. the access to the left ventricle was complicated for even diagnostic catheters, but eventually the impella 5.0 was placed for hemodynamic support. the pump was not optimally positioned, as it was posterior, however the pump was left in the location for the coronary angioplasty. during the coronary angioplasty the axillary tissue/site was seen to be bleeding, and blood products were infused. the angioplasty was successful, and the impella pump was weaned and explanted after just over 22 hours of support. months after the explant, the medical team noted to the abiomed field team in (b)(6) that there had been a pseudoaneurysm that had developed prior to pump explant. the pseudoaneurysm had been treated with blood products and a manual hold. | [5 empty] | The team in (b)(6) discarded the product before it could be sent for analysis. the investigation is on-going at this time. a supplemental final medwatch will be filed upon completion of the investigation. | Since the original filing of the medwatch was submitted, the investigation has completed. the product was not returned from japanese customer for investigation. the data logs from the console were not returned for analysis. the root cause of the access site bleed and pseudoaneurysm was unable to be determined. further clinical information was requested but, not shared for investigation purposes. the failure mode will be monitored and trended.
37849 | 1220648-2019-00060 | 8428006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8428006&pc=OZD | 02/20/2019 | 02/21/2019 |  | 03/18/2019 |  |  | 02/21/2019 | 03/18/2019 |  |  | 12/31/2020 |  | 03/18/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1382208 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 44 YR | [5 empty] | Injury | Yes | Yes | Break | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old caucasian male presenting with acute myocardial infarction and cardiogenic shock, while enduring ventricular tachycardia (vt) during vt ablation. an impella cp was inserted for cardiac support. while physician was completing the insertion process, the pump had dislodged from the ventricle. attempts were made to re-advance the device through the aortic valve, however, the device's cannula became kinked and would not straighten. the decision was made to remove the device. upon explant, it was discovered that the distal section of the device separated from the proximal and remained in the femoral artery. surgical intervention was required to remove the fragmented component. thereafter, a new impella cp was inserted via the same surgically prepared site and ran successfully. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37850 | 8433146 | 8433146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8433146&pc=PYX | 02/21/2019 | 03/13/2019 | 03/13/2019 | 03/14/2019 | 03/13/2019 | 02/21/2019 |  | 03/14/2019 |  |  | 02/29/2020 |  | 03/14/2019 |  | IMPELLA RP |  |  | IMPELLA RP |  |  | PYX |  |  | 004334 |  |  | ABIOMED |  |  |  |  | 21 155174 17 200229 |  |  |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 49 YR |  |  | 89 |  |  | Injury | Yes | Yes | Therapeutic or Diagnostic Output Failure | [6 empty] | Thrombus | Abdominal Distention | Patient Problem/Medical Problem | [10 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/13/2019 | Y | HOSPITAL |  |  | A (b)(6) y/o male on ecmo due to ischemic cardiomyopathy and acute on chronic cardiogenic shock. h/o (b)(6) and anal ca s/p xrt / chemo. planned for insertion of lvad and impella rp. during the procedure, there was successful lvad placement. impella rp placed under fluoroscopic guidance. for majority of procedure ventricular function remained stable with both stable rvad and lvad flows. pt acutely decompensated following chest closure with evidence on tee of significant rv distension and lv underfilling. this coincided with low lvad flows and extremely high pi. the impella rp was emergently removed, the chest re-opened, and a surgical rvad was placed with good decompression of the left ventricle. the chest was left open with a sternal bar in place. upon inspection of the impella rp, there was clot coating the inflow impeller. device malfunction reported to company rep (present in operating room) and to family at conclusion of procedure. the rvad was able to be removed the following day and chest closure was performed on 02/25. the pt continues to recover in the icu.
37851 | 1035166-2019-00023 | 8416422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8416422&pc=DYB | 02/04/2019 | 07/08/2019 |  | 03/13/2019 |  |  | 02/14/2019 | 03/13/2019 |  |  | 10/01/2021 |  | 03/13/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-15257 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [11 empty] | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that during high risk percutaneous coronary insertion, physician failure to advance pump due to patient tortuous anatomy, the pump became stuck in the introducer. when removing the pump physician noticed kink in introducer and damage to distal tip. additional force was used to remove pump; however, pump began to separate. therefore, introducer and pump were removed together, resulting in 300cc of estimated blood loss. as treatment, 1 unit of blood products was provided to patient. as per additional information, when physician pulled the sheath and pump out as unit, the inlet was stuck in the sheath and the cannula was tearing apart. initially 13cm introducer was placed in right axillary artery via cut down. cp was then inserted, but became lodged in sheath; kink was visible. trying to pull out cp, it began to pull apart. while tugging, everything pulled out of artery, but enough wire remained to keep access. the second introducer (25cm) was placed. it kinked at same spot. however, physician was able to complete the procedure successfully with second introducer (25cm). no additional information is available no other adverse event reported. | [5 empty] | Our investigation is still in progress. follow up report will be submitted if we found any further additional information. | One half of a 14f 13 cm long abiomed introducer sheath was returned from the customer without the dilator, and without the 25 cm long sheath and dilator. there were no other accessories. blood was found on the sheath. a kink was observed in the sheath tube approximately 2.5cm from the distal tip. there was also damage found around the circumference of the distal tip edge. the introducer sheath was already peeled apart upon receipt of the product. only the non-sideport half of the sheath was returned for evaluation. according to the complaint description, the device was used in a tortuous path which could account for the damage done to the sheath. the tip damage could also have been due to interaction with the impella cp as excessive force was used in attempt to remove device from peel-away sheath. no manufacturing defects were found. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. as per procedure adelante s2s introducer sheath in-process and final inspection: 12.2 visual inspection: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. per instructions for use (ifu): never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
37852 | 1220648-2019-00059 | 8417026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8417026&pc=OZD | 02/13/2019 | 02/13/2019 |  | 03/13/2019 |  |  | 02/13/2019 | 03/13/2019 |  |  | 09/30/2020 |  | 03/13/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP 2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1367729 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old asian female presenting with acute myocardial infarction and cardiogenic shock. an impella lp2.5 pump was inserted for cardiac support. during treatment, the leg distal to the impella insertion point became ischemic. as treatment, the pump was removed and ischemia resolved. | [5 empty] | The impella lp2.5 was received and a failure investigation was completed. examination of the pump revealed scratches, abrasions, and cracks on the wound closure tip. a review of the device history record found no manufacturing issues or reworks associated with this pump lot. there are no other complaints associated with pumps from this lot. the root cause of the limb ischemia was unable to be confirmed.
37853 | MW5084847 | 8414241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8414241&pc=OZD | 03/01/2019 | 03/08/2019 |  | 03/12/2019 |  |  |  | 03/12/2019 |  |  | 03/31/2019 |  | 03/12/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA |  |  | OZD | [5 empty] | ABIOMED |  |  | US | [5 empty] | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 68 YR |  |  | 90 |  |  | Injury |  | Yes | Display or Visual Feedback Problem | Pumping Stopped | Therapeutic or Diagnostic Output Failure |  |  |  |  | No Known Impact Or Consequence To Patient | [12 empty] | Hospitalization | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Css staff responded to an emergent page from cv recovery, stating that the patient's pump had stopped and was not turning back on. css staff immediately went down to analyze the issue. css staff also brought an additional console for a potential switch. once css staff arrived the patient was responsive. the console was alarming "impella pump off". css staff notified the impella rep in regards to the situation and explained that the pump had been off for longer than 5 min. the rep suggested to try and restart it. css attempted a restart but the pump stopped after the first attempt. css staff made preparation to switch the console but dr. (b)(6) determined that the catheter may be clotted so he ordered css to grab the 50cc balloon for a bedside insertion. after grabbing the iabp console and arriving to the room, another decision was made to take the patient directly to the operating room for emergent 5.0 insertion. css staff removed the impella consoles from the room and made preparation for the 5.0 case in the cv operating room.
37854 | 1220648-2019-00049 | 8407656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8407656&pc=OZD | 02/12/2019 | 02/12/2019 |  | 03/11/2019 |  |  | 02/12/2019 | 03/11/2019 |  |  | 10/31/2020 |  | 03/11/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1377030 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male had presented for cardiac support with the impella lp2.5 pump. during support, patient endured ischemia in the leg which the impella was inserted. as treatment, a surgical bypass was performed, positioning of the device was adjusted, and the patient's ischemia resolved. the patient had later expired, however, this was related to a bleed that had occurred which was not related to the impella device. | [5 empty] | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37855 | 1220648-2019-00050 | 8407689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8407689&pc=OZD | 02/13/2019 | 02/13/2019 |  | 03/11/2019 |  |  | 02/13/2019 | 03/11/2019 |  |  | 07/31/2020 |  | 03/11/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 172185A |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old asian male presenting for high risk percutaneous coronary insertion and received cardiac support with the impella cp. patient's access site developed a hematoma. as treatment, 3 units of packed red blood cells were administered and site was held with manual compression. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37856 | 1220648-2019-00051 | 8407721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8407721&pc=OZD | 02/13/2019 | 02/13/2019 |  | 03/11/2019 |  |  | 02/13/2019 | 03/11/2019 |  |  | 10/31/2020 |  | 03/11/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1382266 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 43 YR | [5 empty] | Injury | Yes | No | Unintended Movement | [6 empty] | Tissue Damage | Injury | [11 empty] | Other | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old caucasian male presenting with acute myocardial infarction and cardiogenic shock. an impella cp optical was inserted for cardiac support. during procedure, patient decompensated. as treatment, device was removed, patient was placed on ecmo support, an endured mitral valve repair. the patient ultimately expired of myocardial infarction. the physician attributed no relationship between the device and the patient's death, however, advised the impella may have exacerbated the patient's previously poor valve condition, as it was caught on the mitral apparatus while on support. | [5 empty] | The impella cp optical was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37857 | 1220648-2019-00037 | 8409802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8409802&pc=OZD | 02/17/2019 | 02/17/2019 |  | 03/11/2019 |  |  | 02/17/2019 | 03/11/2019 |  |  | 09/30/2020 |  | 03/11/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1380513 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hematoma | Patient Problem/Medical Problem | Vascular Dissection | [10 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A patient was transferred from the outlying medical center to the tertiary medical center for care. prior to the transfer the patient's acute myocardial infarction (ami) was treated with thrombolytic drugs, specifically tenecteplase, and arrested. during his arrest the patient received multiple defibrillating shocks to restore rhythm. upon arrival to the tertiary medical center the impella cp was placed for hemodynamic support in the cardiac catheterization lab. the introducer sheath was removed, and upon advancement of the impella, the patient developed a hematoma at the access site. despite all efforts, the hematoma and associated bleed could not be stopped with manual pressure, and so the team performed an angiography of the site which revealed a femoral artery dissection at the impella access site. the team explanted the impella pump, placed a new 14fr introducer, and the patient was transferred to the operating room. while in the operating room he had coronary artery bypasses placed. there was no repair to the femoral artery as hemostasis had been achieved with an introducer alone. | [5 empty] | The impella pump was returned for analysis. the introducer sheath was not returned for analysis. upon benchtop testing of the pump, there was no observed damage found to the pump's repositioning unit. the unit was tested for leaks and none were found. the trending analysis found no other complaints of a similar failure mode. the cause of the hematoma and femoral artery dissection were unable to be determined. the failure mode will be monitored and trended.
37858 | 1220648-2019-00032 | 8405102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8405102&pc=OZD | 11/11/2018 | 02/13/2019 |  | 03/08/2019 |  |  | 02/13/2019 | 03/08/2019 |  |  | 04/30/2020 |  | 03/08/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1347814 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 50 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hematoma | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | PCPS EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) | [22 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A (b)(6) patient with a history of myocarditis was admitted with a possible flare up/return of fluminant myocarditis and a reduced ejection fraction. the patient had pcps/ecmo placed to the right femoral artery. the patient had an impella 5.0 placed via the left axillary artery, despite difficulty with tortuous anatomy, and a cut down done by the surgeon to placed the graft. the impella supported the patient for a month with success. on (b)(6) 2018 the impella pump was explanted. during support the patient condition improved, and the pcps/ecmo was removed. two months later the (b)(6) medical team notified abiomed that the patient had a hematoma, at the access site of the impella, that had necessitated blood product infusion, in the form of 10 units of blood. | [5 empty] | The complainant in (b)(6) did not return the product for review and analysis. no imaging was shared of the axillary anatomy. the root cause of the hematoma could not be determined as no product was returned. the failure mode will be monitored and trended.
37859 | 1035166-2019-00017 | 8399278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8399278&pc=DYB | 02/02/2019 | 03/07/2019 |  | 03/07/2019 |  |  | 02/06/2019 | 03/07/2019 |  |  | 09/01/2021 |  | 03/07/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA, |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC |  |  | US |  | C1-15213 |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 68 YR | [5 empty] | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemostasis | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that patient was presented to the physician in cardiogenic shock stemi. during the procedure sheath placed to left femoral artery with much difficulty and there was noticeable kink in 14 fr 25cm long sheath. physician try to attempt impella but unable to pass, impella was removed and short 14 fr 13cm sheath was exchanged. there was noticeable bleeding around sheath. to control the blood loss 6 unit of prbc's given. staff held manual compression over bleeding area, contralateral balloon tamponade was done to create hemostasis after sheath was removed. surgical repair of the artery was perform. there was 2 hours of delay reported. procedure completed without impella support. patient outcome is stable and no other adverse event reported. | [5 empty] | The device was used for treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. the investigation will focus on a review of product documentation. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. as per procedure adelante s2s introducer sheath in-process and final inspection: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. per ifu : never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
37860 | 1035166-2019-00022 | 8401049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8401049&pc=DYB | 02/08/2019 | 03/07/2019 |  | 03/07/2019 |  |  | 02/11/2019 | 03/07/2019 |  |  |  |  | 03/07/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB | [5 empty] | OSCOR INC. |  |  | US | [5 empty] | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | Blood Loss | [11 empty] | Other | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that patient was presented to the physician for ami w/shock-impella post-pci. during the procedure sheath placed to left femoral artery, when long sheath (14f) was peeled away, bleeding happen. there was 300cc blood lost and rbc 2u transfusion was made. no additional surgery or medical intervention required to control the blood loss. bleeding was treated with manual compression hemostasis and inserted a sheath for insertion. there was unknown time of delay reported because of blood transfusion. there was no specific device performance failure reported by the user, the peel away sheath peeled all the way through. impella support was successful and patient outcome is stable. no additional information is available and no other adverse event reported. | [5 empty] | Device used in treatment. the device was not returned for analysis. there was no specific performance related failure reported by the user. the lot number of this device was not provided, therefore neither a review of the device history record nor complaint history could be performed. inspection procedures require any oscor product pass all in-process and qa final inspections before shipping to the customer. as per ifu: directions for use: introducer removal flush the sheath with 5 cc of saline immediately before peeling the sheath away in order to minimize back bleeding. first, completely withdraw the sheath from the vessel. then remove the sheath by sharply snapping the tabs of the valve housing in a plane perpendicular to the long axis of the sheath to split the valve and peel the sheath apart . as per ifu: bedding is possible complications/adverse effects. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. no further follow-up is require.
37861 | 1220648-2019-00047 | 8394448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8394448&pc=OZD | 02/08/2019 | 02/08/2019 |  | 03/06/2019 |  |  | 02/08/2019 | 03/06/2019 |  |  | 12/31/2020 |  | 03/06/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1385511 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A patient in cardiogenic shock was admitted and had an impella cp placed for hemodynamic support in (b)(6). the patient had been found collapsed for an unknown amount of time in the field and was in poor condition, as he was intubated and in shock. the impella cp was placed, coronary angioplasty completed, and shortly thereafter the patient exhibited signs of limb ischemia. the patient had an antegrade access made and after a short amount of time (an estimated 30 seconds), the leg perfusion improved. pulses returned. the team escalated the patient's care with placement of ecmo (extracorporeal membrane oxygenation). the patient did expire with multi-organ failure after 4 days of impella support. | [5 empty] | The cause of the reported limb ischemia could not be determined as the product was not returned for analysis. the data logs were returned but, they do not show any abnormality that is indicative of a root cause for limb ischemia. no imaging was available for review. the lot analysis revealed no other complaints pertaining to the same failure mode. the failure mode will be monitored and trended.
37862 | 1035166-2019-00021 | 8396865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8396865&pc=DYB | 02/04/2019 | 05/15/2019 |  | 03/06/2019 |  |  | 02/05/2019 | 03/06/2019 |  |  | 06/01/2021 |  | 03/06/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-14498 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN |  |  |  |  | 65 YR | [5 empty] | Malfunction | No | Yes | Break | [6 empty] | Myocardial Infarction | Multiple Organ Dysfunction Syndrome | [40 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported the patient was presented to the physician in acute mi and cardiogenic shock on (b)(6) 2019. during the procedure the sheath was placed in the right femoral artery. the representative reported it was not possible to place the competitor's 14fr sheath, due to wire (guidewire) interaction. damage to sheath occurred during wire (guidewire) removal. physician attempted to place another competitor's 14fr peel-away sheath (length size is unknown) but was not successful. it was reported there was delay of a few minutes in the surgical procedure during exchange of introducer. the patient received several dilations prior to introducer placement. it was reported the angulation to place the sheath was done in the normal way. the patient had no peripheral disease, no obesity, but it was very difficult under ongoing cpr to puncture the arteria. patient outcome was death. the physician reported cause of death as multi organ failure, patient expired on (b)(6) 2019. | [5 empty] | The device was used for treatment. one 14f abiomed introducer sheath was returned from the customer with the dilator. there were no other accessories. visible blood was found on and inside the sheath and dilator. the sheath was split up one of the scorelines starting at the distal tip and continuing approximately 6cm from the distal tip. there is another split in the scoreline starting at the distal tip and continuing approximately 4mm from the distal tip. there was also damage found around the circumference of the distal tip edge. there were two splits in the sheath tube and there was also damage found around the circumference of the distal tip edge. it is stated in the event description that the damage was due to pulling the wire out of the sheath. another potential cause of the damage is if the sheath tip came in contact with the guidewire (impact damage with the guidewire). other potential contributing factors to the damage is if the device was advanced through an area of resistance (as stated in the precautions of the instructions for use (ifu), or if the device was subjected to unusual stresses (as stated in the handling and storage section of the ifu). the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per abiomed introducer sheath in-process and final inspection procedure, the devices in this lot were inspected as follows: 9.2.2 visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. slight discoloration from tipping process is acceptable. verify proper tip shape according to drawing. the instructions for use (ifu) informs the user never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete.
37863 | 1220648-2019-00044 | 8390088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8390088&pc=OZD | 02/04/2019 | 02/05/2019 |  | 03/05/2019 |  |  | 02/05/2019 | 03/05/2019 |  |  | 10/31/2020 |  | 03/05/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1381868 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 40 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old hispanic female had impella cp pump inserted without difficulty. patient was constantly bleeding around repo sheath so daily blood products was given. the patient is obese and lost at least 200 cc of blood in seven (7) days. | [5 empty] | The impella cp pump was discarded by the customer therefore a failure analysis investigation cannot be completed.
37864 | 1220648-2019-00046 | 8384217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8384217&pc=OZD | 02/04/2019 | 02/04/2019 |  | 03/02/2019 |  |  | 02/04/2019 | 03/02/2019 |  |  | 10/31/2020 |  | 03/02/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1382142 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | (b)(6) year old caucasian male presented with acute myocardial infarction and cardiogenic shock. an impella cp was placed for cardiac support without any reported issues. the following day, the leg in which the impella was inserted became ischemic. as treatment, the pump was removed and ischemia resolved. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37865 | 1220648-2019-00033 | 8383529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8383529&pc=OZD | 02/04/2019 | 02/04/2019 |  | 03/01/2019 |  |  | 02/04/2019 | 03/01/2019 |  |  | 12/31/2020 |  | 03/01/2019 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1382244 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | An impella cp was placed in (b)(6) to support a patient for a high risk coronary angioplasty. the patient had mulitvessel coronary artery disease that was treated successfully. upon transfer from the cardiac catheterization suite to the icu for monitoring, the team observed blood loss from the impella access site. the team attempted to apply pressure to achieve hemostasis. when the manual hold failed the team explanted the impella. to treat the blood loss the patient received 3 units of blood products. in conversation with the field representative and physician afterwards, the thought was there was improper placement of the impella repositioning sheath, as it was too deep the puncture site allowed for blood loss. in addition, access was made at a steeper than recommended angulation. | [5 empty] | Since the original medwatch was filed, the investigation of the access site bleeding has concluded. according to the physicians, the impella pump's repositioning unit was placed too deep into the left femoral artery at a steeper than recommended insertion angle. the root cause of the bleeding was most likely due to the steep insertion angle, but without the pump returned this can not be definitively determined. the failure mode iwll be monitored and trended. | To date no data or product have been returned for analysis. more clinical details have been requested. should further information or product be shared that will lead to a root cause of the blood loss and associated intervention, a supplemental medwatch will be submitted. the failure mode will be monitored and trended.
37866 | 1220648-2019-00043 | 8377687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8377687&pc=OZD | 02/02/2019 | 03/28/2019 |  | 02/28/2019 |  |  | 02/02/2019 | 02/28/2019 |  |  | 09/30/2020 |  | 02/28/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1379921 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | Yes | Break | Material Fragmentation | Difficult to Remove | Failure to Advance | Detachment of Device or Device Component | Adverse Event Without Identified Device or Use Problem |  | Death | Foreign Body In Patient | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old caucasian male had impella cp pump inserted. while doctor was advancing the impella in aorta the device and wire became locked together and he was unable to advance or remove the impella over the wire. impella pulled back and buckled in aorta. doctor attempted to pull back and device broke off above the motor housing inside the patient. vascular surgeon attempted to snare but patient coded and expired despite everyone's best efforts. | [5 empty] | The root cause of catheter fracture leaving the pump inside the patient was unable to be determined due to the fractured portion of the pump not being returned for review. the reported guidewire fracture may have caused the pump and guidewire to become "locked" together and get stuck at the introducer during pump removal. to remove the pump from the patient, excessive force was applied, causing the catheter to fracture, and the pump housing to be left in the patient. corrected data: was changed from 02-08-2018 to 02-08-2019. | The impella cp was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37867 | 1220648-2019-00034 | 8370890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8370890&pc=OZD | 01/28/2019 | 01/28/2019 |  | 02/26/2019 |  |  | 01/28/2019 | 02/26/2019 |  |  | 10/31/2020 |  | 02/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1384654 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A patient presented to the outside hospital and was transferred to the tertiary medical facility for escalation of care. an impella pump was placed via the right femoral artery for hemodynamic support. three days later the team added va-ecmo for additional support and due to possible acute respiratory distress. after seven days of support the impella limb showed signs of ischemia. the impella was explanted and upon observation of the pump there was thrombus visualized. the team performed a 4 compartment fasciotomy of the limb and thromboendarterectomy. | [5 empty] | The returned product was inspected and found to have thrombus and biomaterial. the data logs did not pertain to the failure mode. the root cause of the leg ischemia was most likely inadequate anticoagulation resulting in thrombosis and the necessary interventions to treat the limb. per the ifu: "after insertion of the catheter (and until explant), act should be maintained at 160 to 180 seconds." in this patient support the act was below the recommended range throughout the case. the failure mode will be monitored and trended.
37868 | 1220648-2019-00040 | 8370971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8370971&pc=OZD | 01/18/2019 | 02/19/2019 |  | 02/26/2019 |  |  | 01/18/2019 | 02/26/2019 |  |  | 10/31/2020 |  | 02/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1382552 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) caucasian female presenting for cardiac support. post cardiology cardiogenic shock (pccs). patient had multiple risk factors and had underwent a multiple vessel coronary artery bypass graft (cabg). an impella cp was inserted without issue. during support, the device's motor had stopped. multiple attempts to re-start device were not successful; therefore, device was explanted and a new device was inserted and patient continued receiving support successfully. it should be noted that explant was not possible bedside, due to patient's small anatomy. subsequently, patient was taken to catheter laboratory for explant. between the time the device had stopped and a new device was inserted, patient had coded and required intubation and ecmo placed. | Patient's code was specifically due to "pea arrest." prior to being placed on va ecmo support, patient was placed on cardiopulmonary bypass and new impella was placed. patient experienced "temporary periods of stability followed by worsening respiratory response." after transitioning from cardiopulmonary bypass and placed on va ecmo, patient gained stability. |  |  |  |  | The impella cp was returned an investigation was completed. the clinical account indicates that the pump was run for over 11 days, which is beyond the indicated duration of use of the device. analysis of the data logs confirmed a pump stop had occurred. physical examination of the device's outflow cage found the a gap between the sleeve bearing and the impeller. a review the device history record was conducted and found no manufacturing issues or reworks associated with this pump lot. there are no other complaints associated with pumps from this lot. the root cause of the pump stop was sleeve bearing wear, due to duration of use. | Additional information was received via a user facility medwatch (b)(4) on 2/19/2019. the impella cp was returned and an investigation was completed. the clinical account indicates that the pump was run for over 11 days, which is beyond the indicated duration of use of the device. analysis of the data logs confirmed a pump stop had occurred. physical examination of the device's outflow cage found a gap between the sleeve bearing and the impeller. a review the device history record was conducted and found no manufacturing issues or reworks associated with this pump lot. there are no other complaints associated with pumps from this lot. the root cause of the pump stop was sleeve bearing wear, due to duration of use.
37869 | 1220648-2019-00042 | 8371299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8371299&pc=OZD | 01/31/2019 | 02/01/2019 |  | 02/26/2019 |  |  | 02/01/2019 | 02/26/2019 |  |  | 09/30/2020 |  | 02/26/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED |  |  | US |  | 1373141 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The medical team in (b)(6) placed an impella 2.5 for a patient suffering from myocarditis. during the support the patient exhibited signs of hemolysis that prompted the team to adjust the pump position, infuse volume, and eventually put the patient on dialysis and explant the pump. the patient remained on dialysis after pump explant. the patient had received impella support for just over 42 hours. | [5 empty] | The team in (b)(6) did not save for analysis the impella pump. no in-house hemolysis testing could be done. the data logs were returned for analysis and do show that the impella pump was out of position for extended periods of time and exhibited wrong position alarms. the ifu does alert the operator to the possibility of hemolysis as: "positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis." without benchtop testing for hemolysis the root cause of the hemolysis, supposed renal failure, and need for dialysis can not be determined. the failure mode will be monitored and trended.
37870 | 1035166-2019-00012 | 8366170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8366170&pc=DYB | 01/15/2019 | 02/25/2019 |  | 02/25/2019 |  |  | 01/28/2019 | 02/25/2019 |  |  | 08/01/2021 |  | 02/25/2019 |  | INTROUDCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-14552 |  | 0052-3025 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 89 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that on (b)(6) 2019, an (b)(6) male was having a high risk percutaneous coronary insertion procedure. during the procedure, the physician used a 14 french, 25 cm length oscor peel-away introducer. the peel-away sheath was left in access site during the procedure. an echo revealed the device's position was too shallow, and therefore the peel-away was removed by the physician to have better stability. shortly after removal, a slight ooze developed into a growing hematoma. they physician suspected a tear in the patient's artery was to blame, caused by sheath removal. however, it should be noted that the arterial tear was only suspected and never confirmed. no blood products or additional intervention was required to remedy the bleed, besides early removal of the impella cp pump and site closure. patient is in stable condition and did not endure any lasting harm. the peel-away sheath was discarded. | [5 empty] | Device used in treatment. the device was not returned for analysis. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. there was no specific performance related failure reported by the user. the device passed all in-process and qa final inspection steps before shipping to the customer. no further investigation is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. no further follow-up is required.
37871 | 1220648-2019-00030 | 8362794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8362794&pc=OZD | 01/30/2019 | 01/30/2019 |  | 02/22/2019 |  |  | 01/30/2019 | 02/22/2019 |  |  | 06/30/2020 |  | 02/22/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA LP 2.5 PUMP |  |  | OZD |  |  | IMPELLA LP 2.5 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1352017 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 88 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Intermenstrual Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) female presenting with pre shock and ejection fraction (ef) 30%, left main trunk (lmt) lesions and non-invasive blood pressure (nibp) 80-90 units. the impella lp 2.5 pump was inserted in the right femoral. the doctor sent the patient for a percutaneous coronary intervention (pci). hematoma and oozing was noted at the insertion site with blood pressure lowering and two (2) units of red blood cells (rbc) were administered. | [5 empty] | The facility did not return the impella lp 2.5 pump product for investigation; therefore, the root cause of the bleeding was unable to be determined.
37872 | 1220648-2019-00031 | 8362798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8362798&pc=OZD | 01/25/2019 | 03/18/2019 |  | 02/22/2019 |  |  | 01/25/2019 | 02/22/2019 |  |  | 09/30/2020 |  | 02/22/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1378951 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 14 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemostasis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A (b)(6) patient with a extensive medical history, inclusive of leukemia and many cardiac comorbidities, had an impella cp placed for hemodynamic support. while in the intensive care unit (icu) for continued monitoring, the patient began to have a blood loss from the impella access site at the femoral artery. the medical team treated the patient with platelets and a fresh frozen plasma infusion. in addition the team placed additional sutures and held manual pressure to the access site to remedy the blood loss. when these measures failed, the team placed a femstop and continued with infusion of platelets. | [5 empty] | The impella cp is being returned by the customer and a failure analysis investigation will be underway. a supplemental mdr will be filed at the completion of the device's investigation. | The root cause of the patient injury was likely due to a positioning issue of the pump being too deep in the left ventricle (lv). the positioning issue may have occurred if the locking anchor was not tightened appropriately to hold the catheter in place.
37873 | 1220648-2019-00039 | 8363178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8363178&pc=OZD | 01/28/2019 | 01/28/2019 |  | 02/22/2019 |  |  | 01/28/2019 | 02/22/2019 |  |  | 10/31/2020 |  | 02/22/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1382804 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Low Blood Pressure/ Hypotension | Ischemia | Diaphoresis | [10 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient with a known history of drug abuse, peripheral vascular disease, and cardiac comorbidities was transferred to the tertiary medical center for care escalation. the team decided to place an impella cp for support during a high risk coronary artery intervention. the 14fr introducer was placed in the left femoral artery. when the team placed the impella through the introducer the patient condition deteriorated. the patient's limb became ischemic. the patient further became hypotensive and diaphoretic. due to these reasons, the impella pump was removed. the team closed the femoral access and sent the patient to surgery. | [5 empty] | The hospital team did not forward the product for analysis, as it was discarded. the placement of the impella in a patient with known peripheral disease is mentioned within the contraindications of the impella ifu. the presence of "severe peripheral arterial disease" is noted in the ifu, and this patient did have known peripheral disease, which could have contributed to the limb ischemia observed after the placement of the impella pump in the femoral artery. as the product was not returned however the root cause of the ischemia and the associated hypotension and diaphoresis can not be determined. the failure mode will be monitored and trended.
37874 | 1220648-2019-00035 | 8358015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8358015&pc=OZD | 01/21/2019 | 03/07/2019 |  | 02/21/2019 |  |  | 01/23/2019 | 02/21/2019 |  |  | 06/30/2020 |  | 02/21/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA LP 5.0 PUMP |  |  | OZD |  |  | IMPELLA LP 5.0 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1349665 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Infarction, Cerebral | Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male had impella lp 5.0 pump inserted smoothly. on the eleventh day after the insertion while during weaning, cerebral infractions happened. the physician thought that the patient's cerebral infractions were related to the impella because he thinks some thrombus was blown out by impella. the patient expired after explant of impella because the patient did not want to have active treatment anymore. | [5 empty] | The impella lp 5.0 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The root cause of the cerebral infarction is unknown. no evidence relating impella to the infarction was identified. the motor current recorded during the case indicates the impella was providing hemodynamic support within specification.
37875 | MW5084269 | 8366150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8366150&pc=OZD | 11/14/2018 | 02/16/2019 |  | 02/21/2019 |  |  |  | 02/21/2019 | 02/21/2019 | 02/21/2019 |  |  | 02/21/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP | IMPLANTABLE PACEMAKER PULSE-GENERATOR | IMPLANTABLE, CARDIOVERTER, DEFIBRILLATOR (NON-CRT) | IMPELLA PUMP | PACEMAKER | AICD VISIA AF MRI SURESCAN DEFIBRILLATOR | OZD | DXY | LWS |  |  | DVFB1D4 | UNKNOWN | UNKNOWN | MEDTRONIC INC. | [6 empty] | Yes | Yes | Yes |  |  |  | Invalid/Incomplete | No Information | No Information | No Information | Voluntary report | [8 empty] | 34 YR | [5 empty] | Injury | Yes | Yes | Insufficient Information | [6 empty] | Fatigue | Impotence | Respiratory Distress | Seizures | Anxiety | Depression | [7 empty] | Hospitalization | Life Threatening | Required Intervention | Disability | [25 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | I had an impella pump put in on (b)(6) and a pacemaker put in on (b)(6) and have had continued health problems since then. seizures, fatigue, lack of breath, impotence, depression, anxiety, etc.
37876 | 1220648-2019-00038 | 8354419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8354419&pc=OZD |  |  |  | 02/20/2019 |  |  | 01/23/2019 | 02/20/2019 |  |  | 06/30/2020 |  | 02/20/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED |  |  | US |  | 1349665 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A (b)(6) male with cardiomyopathy was admitted to the medical facility in (b)(6) and had an impella 5.0 placed for support. the 5.0 was an escalation from the previously placed intra aortic balloon pump (iabp). the plan was to leave the 5.0 in as the patient was being prepped for a durable vad device. during the support the patient had presumed hemolysis and received 2 units of blood. after 15 days of support, the 5.0 was explanted and the heartmate ii vad was placed. | [5 empty] | The complainant in (b)(6) did not return the product for analysis. the clinical report is lacking in details at this time. more questions have been relayed to the team in (b)(6), but to date no reply has been made to assist in the investigation. the data logs were returned for analysis, and that analysis is ongoing. should more details into the cause of the hemolysis be shared and a root cause be determined, a supplemental medwatch will be completed. | Since the filing of the initial medwatch the investigation has completed. the complainant in japan never forwarded the product for investigation, nor the complete pump data logs. the partial data logs that were returned revealed only the first two hours of support when both placement and suction alarms were occurring. the clinical account stated that the echo was not used to confirm positioning. hemolysis was possibly related to improper positioning, but the root cause of the presumed hemolysis can not be determined without product and complete data logs. the failure mode will be monitored and trended.
37877 | 1220648-2019-00029 | 8351437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8351437&pc=OZD | 01/25/2019 | 03/07/2019 |  | 02/19/2019 |  |  | 01/25/2019 | 02/19/2019 |  |  | 10/31/2020 |  | 02/19/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1382671 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a 73 year old caucasian female presenting with st-segment elevation myocardial infarction (stemi) with ejection fraction (ef) 50%. the impella cp pump was inserted in the right femoral, patient taken to the operating room (or) for coronary artery bypass graft x 3 (cabg). impella cross clamped while on pump. after cabgx3, cross clamp removed and impella turned back to p1 while still on pump. impella pigtail noted to be through ventricle. the pigtail was sticking out of the left ventricle (lv), so the doctor had to surgically patch the lv hole. there was blood loss and many units were given. | [5 empty] | The impella cp was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation. | The root cause of the ventricular perforation was unable to be determined without further clinical information.
37878 | 1220648-2019-00036 | 8351749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8351749&pc=OZD | 01/22/2019 | 03/11/2019 |  | 02/19/2019 |  |  | 01/23/2019 | 02/19/2019 |  |  | 09/30/2020 |  | 02/19/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1379792 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 47 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old caucasian female presenting with acute myocardial infarction and cardiogenic shock. prior to presenting, patient endured cardiac arrest while at home and was administered 10 minutes of cardiopulmonary resuscitation by emergency management technicians. patient arrested a second time and return of spontaneous circulation was achieved via defibrillation. an impella cp was inserted for cardiac support, as treatment thereafter, and patient's condition notably improved by the following day. approximately 1 day after improvement was noted, physician identified pulse in impella leg as "questionable," and found blood flow as monophasic. patient was diagnosed with compartment syndrome. as treatment, a below knee amputation (bka) was conducted (due to necrotic tissue at calf) and impella was weaned and explanted. despite requiring a bka, cardiac circulatory support was successfully achieved with this device. | [5 empty] | The impella cp was returned and an investigation is underway. once completed, a supplemental mdr will be filed with the results of this investigation. | The impella cp was received and a failure investigation was completed. examination of the pump did not identify any issues that would likely attribute to an ischemic event. a review of the clinical account found no allegation of improper impella performance. a review of the device history record found this device had undergone a post-sterile re-work, however, no other complaints associated with pumps from this lot have been reported. the root cause could not definitively be determined.
37879 | 1220648-2019-00021 | 8339934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8339934&pc=PYX | 01/22/2019 | 01/22/2019 |  | 02/14/2019 |  |  | 01/22/2019 | 02/14/2019 |  |  | 08/31/2020 |  | 02/14/2019 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1379927 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 84 YR | [5 empty] | Malfunction | No | Yes | Break | No Apparent Adverse Event | [5 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | An (b)(6) year old male was having coronary artery bypass surgery, and had decompensated when he was taken off surgical bypass. the patient's right ventricle was failing, and so the team placed with difficulty a swan and then attempted to insert the impella rp. the team attempted multiple times and in different orientations to place the rp. upon removal of the rp, to try placement of a new rp catheter, the team noted the rp had become fractured. the team snared the part of the rp that had detached. after removal of the fractured rp, a new rp catheter was placed successfully via the same access at the right femoral vein. | [5 empty] | The impella rp was returned for analysis. the data logs from the case were not. the deformation of the rp was most likely created by the force exerted by the physician in the attempted placement of the rp. the force was due to the difficult insertion into the vasculature and over the guidewire, which had slack. the root cause for the rp pump separation was force applied to the outflow cage struts and welds during the difficult delivery. the failure mode will be monitored and trended.
37880 | 1220648-2019-00028 | 8331895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8331895&pc=OZD | 01/02/2019 | 03/05/2019 |  | 02/12/2019 |  |  | 01/14/2019 | 02/12/2019 |  |  | 06/30/2020 |  | 02/12/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA LP 5.0 PUMP |  |  | OZD |  |  | IMPELLA LP 5.0 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1353056 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male presenting with amifostine (ami)/ cardiogenic shock (cgs). patient had putamen bleeding and blood transfusion was performed. | [5 empty] | The cause of the putamen bleeding and iliolumbar is unknown. impella functioned properly throughout the duration of the case. | The impella lp 5.0 pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37881 | 1220648-2019-00027 | 8323896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8323896&pc=PYX | 01/12/2019 | 03/21/2019 |  | 02/08/2019 |  |  | 01/12/2019 | 02/08/2019 |  |  |  |  | 02/08/2019 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | IMPELLA AIC CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1198722 |  | 0042-0000-US |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | Yes | Unexpected Shutdown | [6 empty] | Cardiac Arrest | [12 empty] | Life Threatening | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old caucasian male presenting with right ventricular failure, post left ventricular assist device implantation. the impella rp was inserted for cardiac support. during the procedure, the automated impella controller (aic) console emitted a controller failure alarm, resulting in an unexpected shutdown. subsequently, the patient became unstable and coded. as treatment, patient was administered medical intervention via inotopes and pressors (medication). pump was switched to a back- up aic console, and patient became stable. the procedure was completed successfully, thereafter, with no lasting harm. | [5 empty] | The impella aic console was received and a failure investigation was completed. data logs confirmed the reported controller failure alarm. simulated use testing replicated the controller alarm with subsequent unexpected shut down. further investigation concluded the problem was caused by a faulty esd suppressor on the impellatronic assembly. after replacement of this component, the device ran without issue. a review of the device history record found no manufacturing issues with this device and applicable systems remain effective. the root cause of the controller failure alarm resulting in an unexpected shut down was a failed esd suppressor. | The device was received and an investigation is underway. should analysis identify a failure of the device to operate to specification, a supplemental mdr will be filed with the results of the investigation.
37882 | 8361320 | 8361320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8361320&pc=PBL | 01/18/2019 | 01/31/2019 | 01/31/2019 | 02/08/2019 | 02/01/2019 | 01/31/2019 |  | 02/08/2019 |  |  |  |  | 02/08/2019 |  | MECHANICAL CIRCULATORY SUPPORT DEVICE |  |  | IMPELLA CP |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 68 YR |  |  | 68 |  |  | Injury | Yes | No | Failure to Pump | Infusion or Flow Problem | Pumping Problem |  |  |  |  | Left Ventricular Failure | Heart Failure | [11 empty] | Life Threatening | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 01/21/2019 | Y | HOSPITAL |  |  | Patient requiring mechanical circulatory support via an impella cp. on (b)(6) 2019 the patient¿s impella device stopped functioning while the patient was in intensive care unit requiring emergent surgery. (b)(6) female with cad, type 2 diabetes and history of myocardial infarction and cabg. patient was admitted on (b)(6) 2018 due to worsening cardiac and respiratory symptoms. due to her low cardiac output and severe left ventricular systolic failure it was decided that she would be a candidate for mechanical circulatory support. on (b)(6) 2019 she underwent placement of the impella cp. the impella cp was positioned without complication trans-aortic into the left ventricle and left ventricular support was initiated. on (b)(6) 2019 it was discovered that the impella cp had an acute malfunction in the intensive care unit with complete stoppage of the pump and inability to restart the pump. the patient was taken immediately to the hybrid room for impella removal and replacement. on arrival to the hybrid room the patient became unstable and went into pea arrest. the patient was placed on cardiopulmonary bypass. the impella was removed. another impella was placed. there was multiple attempts to wean the patient from cardiopulmonary bypass with temporary periods of stability followed by worsening respiratory response. the patient was transitioned from cardiopulmonary bypass to ecmo. the patient was transferred on va ecmo support, impella support to the intensive care unit in stable condition. the impella was picked up by representatives of abiomed for interrogation of the device to determine the cause of the malfunction.
37883 | 1220648-2019-00025 | 8298015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8298015&pc=OZD | 01/10/2019 | 01/11/2019 |  | 02/01/2019 |  |  | 01/11/2019 | 02/01/2019 |  |  | 10/31/2020 |  | 02/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1383060 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male caucasian presenting with acute myocardial infarction and cardiogenic shock. the impella was used successfully; however, a thrombus was identified at the access site, post explant. as treatment, patient required re-intubation, pressors, vascular intervention to remove clot. | [5 empty] | The impella cp was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37884 | 1220648-2019-00024 | 8298076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8298076&pc=OZD | 12/04/2018 | 01/07/2019 |  | 02/01/2019 |  |  | 01/07/2019 | 02/01/2019 |  |  | 08/31/2020 |  | 02/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1363700 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | No Apparent Adverse Event | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A (b)(6) male presented with acute myocardial infarction and cardiogenic shock. an impella lp2.5 was inserted for cardiac support. during treatment, patient exhibited hemolysis. as treatment, patient received a blood transfusion. | [5 empty] | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37885 | 1220648-2019-00023 | 8298101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8298101&pc=OZD | 11/29/2018 | 02/04/2019 |  | 02/01/2019 |  |  | 01/07/2019 | 02/01/2019 |  |  | 09/30/2020 |  | 02/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1370693 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 86 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) asian male patient presenting for high risk percutaneous coronary insertion via an impella lp2.5 cardiac support pump. during support, the patient developed a hematoma at the device's insertion site. as a result, patient's hospital stay was prolonged and a 10 unit blood transfusion was administered. patient is in recovery and stable. | [5 empty] | New information was received from reporting source. the physician has now concluded that there was no causal relationship between the device and the originally reported event (hematoma requiring blood products). rather, the event was the result of "combination medicine" that was used during the procedure. in light of this new information, we respectfully retract the original reporting decision to not reportable. | The impella lp2.5 was discarded by the customer and therefore, investigation of device was not possible. should any new information be returned, a supplemental mdr will be filed with the results of the investigation.
37886 | 1035166-2019-00006 | 8298432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8298432&pc=DYB | 01/04/2019 | 02/28/2019 |  | 02/01/2019 |  |  | 01/07/2019 | 02/01/2019 |  |  | 08/01/2021 |  | 02/01/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-15174 |  | 0052-3000 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN | [10 empty] | Injury | Yes | No | Break | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that during access at femoral artery introducer break into 4 pieces which promoted vascular repair. patient survived after surgery. no other additional information is available. no other adverse event reported. | As per additional information, there was longitudinal spilt, tip and end intact observed in introducer. impella placement was abandoned and patient was supported by va ecmo. |  |  |  |  | Our investigation is still in pogress, follow up report will be submitted if we found any further addtional informtion. | The device was used for treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure abiomed introducer sheath in-process and final inspection 9.2.2 visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. slight discoloration from tipping process is acceptable. verify proper tip shape according to drawing. per ifu: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
37887 | 1220648-2019-00009 | 8298564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8298564&pc=OZD | 01/04/2019 | 01/07/2019 |  | 02/01/2019 |  |  | 01/07/2019 | 02/01/2019 |  |  | 11/30/2020 |  | 02/01/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1375988 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG | [10 empty] | Malfunction | No | Yes | Material Frayed | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A patient was admitted to facility in (b)(6) and the team attempted to place the 14fr introducer to allow for placement of the impella cp. after placement of the 14fr introducer, the team noted there was damage to the introducer and the patient was sent to surgery for vascular repair. | [5 empty] | To date no product or data have been shared. clinical questions have been requested but, to date no reply has been shared. upon completed investigation, a final medwatch will be filed. | Since the initial medwatch was filed, the investigation has completed. the cause of the introducer damage which lead to a vascular repair could not be determined. the complainant in the eu did not return any product for analysis. sufficient clinical information was also not shared. the failure mode will be monitored and trended.
37888 | MW5083629 | 8297108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8297108&pc=OZD | 01/25/2019 | 01/30/2019 |  | 01/31/2019 |  |  |  | 01/31/2019 |  |  |  |  | 01/31/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED EUROPE GMBH | [8 empty] | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 85 YR |  |  | 87 |  |  | Malfunction |  | Yes | Application Program Problem | Component or Accessory Incompatibility | [5 empty] | Underdose | [41 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Our institution recently changed to an impella heparin purge solution concentration of 20 units/ml, based off of the live, online user manual on abiomed¿s website. there are a variety of programmable concentrations of heparin from 0-50 units/ml included in their user manual. we chose 20 units/ml based on the concentration being listed as an available concentration, as well as abiomed representative review and input. this patient was the first patient with our new order set and when the concentration was being programmed into the impella console it was discovered that 20 units/ml was unable to be selected. upon further review and calls to the company, the impella software does not support this (and other) concentrations listed in their user manual. as a temporary work around our nurses were instructed to use the 25 unit/ml concentration that was available in the console. once discovered, pharmacy provided a heparin bag that matched this concentration. however, until the bag was replaced, this patient received less systemic heparin exposure than anticipated due to the fact that the user manual was not up to date with their current software.
37889 | MW5083437 | 8286331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8286331&pc=PBL | 01/04/2019 | 01/25/2019 |  | 01/28/2019 |  |  |  | 01/28/2019 |  |  | 07/31/2020 |  | 01/28/2019 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP SET |  |  | PBL |  |  | 01-00813502010879101357925 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 66 YR |  |  | 64 |  |  | Injury | No | No | Peeled/Delaminated | Failure to Power Up | [5 empty] | Low Blood Pressure/ Hypotension | Dizziness | Stenosis | [10 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | A (b)(6) year old male presenting with hypotension (63/41), dizziness requiring pressor support. ti 23.7 increasing to 37.5, severe cardiomyopathy with ef 20-25%. taken to cl for nstemi and treatment of cardiogenic shock. found to have critical left main stenosis with severe left circumflex obtuse marginal branch stenosis and severe to critical rca stenosis. an impella was placed. the sheath was exchanged for a 14-french sheath and then was able to cross the impella device across the aortic valve under fluoroscopic guidance and confirmed adequate placement. peel -away sheath was then taken and removed with insertion sheath. i then put in place at the arteriotomy site the impella insertion sheath. maintain hemostasis . however, the impella device simply would not turn on despite all adequate checks. multiple attempts then were made at maintaining arteriotomy site access including floating a 9-french, as well as a 10-french sheath over the insertion sheath via the impella catheter itself. however , i could not get access into the arteriotomy site. in light of this , i then maintained that the impella was pulled back into the level of the descending aorta. the patient was hemodynamically stable . however, given the fact that we could not maintain access site if it was removed, the decision was to leave the impella in the body in the level of the descending aorta, maintaining hemostasis with the impella insertion sheath in place, maintaining hemostasis at the level of the arteriotomy site. patient was transferred to a facility with critical pci/cabg capability . they were unable to get the impella to work, so the impella was removed and patient placed on iabp (occurred at another facility). on inspection, it was noted that the cheater had not been completely removed from the device before insertion which shut down the impella. no apparent patient injury.
37890 | 1220648-2019-00020 | 8278356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8278356&pc=OZD | 12/31/2018 | 12/31/2018 |  | 01/25/2019 |  |  | 12/31/2018 | 01/25/2019 |  |  | 08/31/2020 |  | 01/25/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1357060 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) hispanic female presenting with acute myocardial infarction, cardiogenic shock, receiving continuous dialysis, and receiving support from a mechanical ventilator. the impella cp was inserted and used for cardiovascular support. during the procedure, there was an absence in distal pulses in the patient's limb. as treatment, the device was removed to prevent lasting harm. | [5 empty] | The impella cp has been requested for return; however, not yet received. should device return and any new information pertaining to product deficiency be identified, a supplemental mdr will be filed with the results of the investigation.
37891 | 1220648-2019-00004 | 8279477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8279477&pc=PYX | 12/28/2018 | 12/28/2018 |  | 01/25/2019 |  |  | 12/28/2018 | 01/25/2019 |  |  | 01/31/2020 |  | 01/25/2019 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1323607 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 43 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A patient was placed on the impella rp for support due to right heart failure post-pulmonary embolism. after eleven hours of support, the team removed the rp due to a venous access site bleed, which troubleshooting in the form of both the placement of a femostop and stitch failed. the team infused units of blood product back to the patient. it is unknown if the blood infusion was 1 or 2 units. | [5 empty] | For the investigation the rp pump was returned for analysis, though it was returned cut by the user. the pump was returned for review with both the catheter and repositioning unit cut. the introducer was not returned for analysis. no damage was observed to the returned portion of the device, the root cause of the access site bleed was unable to be determined without the complete product. should the introducer be returned for analysis, a supplemental medwatch will be completed inclusive of any new findings.
37892 | 1220648-2019-00022 | 8279692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8279692&pc=OZD | 09/26/2018 | 01/02/2019 |  | 01/25/2019 |  |  | 01/02/2019 | 01/25/2019 |  |  | 08/31/2019 |  | 01/25/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA LP5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1307296 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | HEALTH PROFESSIONAL | USER FACI |  |  |  |  | 72 YR |  |  | 111 |  |  | Injury | Yes | Yes | Disconnection | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male presenting with cardiomyopathy. an impella lp 5.0 was inserted for temporary cardiovascular support. the patient was receiving physical therapy, during which point the pump's cable connector became separated from the power cable, and pump motor stopped. attempts to reconnect the cable were not successful and patient became hypotensive. as treatment to stabilize patient, epinephrine was restarted and norepinephrine dosage was increased. damage to the connector prongs was reported. a new pump was inserted and treatment continued successfully. | [5 empty] | The impella lp5.0 was received and a failure investigation was performed. broken lemo pins were confirmed, however, we are unable to definitively determine a root cause. a review the device history record was completed and found no manufacturing issues or reworks associated with this pump lot. there are no other complaints associated with pumps from this lot.
37893 | 8270041 | 8270041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8270041&pc=OZD | 12/20/2018 | 12/21/2018 | 12/21/2018 | 01/23/2019 | 12/21/2018 | 01/23/2019 |  | 01/23/2019 |  |  |  |  | 01/23/2019 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | 0042-0000-US |  | Yes | [5 empty] | Invalid/Incomplete | Health Professional |  |  | User Facility report | [14 empty] | Malfunction |  | Yes | No Display/Image | Protective Measures Problem | [5 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella pump was on patient and device failed. screen went blank, and then received message "attempting to boot a hard drive" followed by a page of computer code. abiomed customer service number immediately called. cardiologist notified, and able to secure an alternate device from the cath lab to replace malfunctioning pump. manufacturer response for abiomed impella controller, (brand not provided) (per site reporter). not known at this time.
37894 | 1220648-2019-00019 | 8270307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8270307&pc=OZD | 12/31/2018 | 02/22/2019 |  | 01/23/2019 |  |  | 12/31/2018 | 01/23/2019 |  |  | 09/30/2020 |  | 01/23/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1378000 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | Yes | Output Problem | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old caucasian male in cardiogenic shock (cgs). doctor attempted to repo prior to leaving the cardiac cath lab (ccl) and the device jumped out of left ventricle (lv). the doctor attempted to push back across the atrioventricular node (av). the impella then bent back. the decision was made to explant and re-cross with a new device. while pulling the device through the sheath the inflow detached entirely from the cannula and snare was unsuccessful. the doctor cut down in the left leg to retrieve the device as it separated from the inflow cage to pigtail. | [5 empty] | The polyurethane was ruptured near the inflow cage, consistent with the clinical report that the cannula bent at this location during attempts to re-advance the pump. the root cause cannot be definitively determined as the introducer and the detached inflow cage were not returned for review. proximal to the delaminated region, the cannula was free of kinks and there were no signs of deformation. the end of the polyurethane was torn, and the step in the coating was visible. failures of this type will continue to be monitored for trends. | The polyurethane was ruptured near the inflow cage, consistent with the clinical report that the cannula bent at this location during attempts to re-advance the pump. the root cause cannot be definitively determined as the introducer and the detached inflow cage were not returned for review. proximal to the delaminated region, the cannula was free of kinks and there were no signs of deformation. the end of the polyurethane was torn, and the step in the coating was visible. failures of this type will continue to be monitored for trends.
37895 | 8266665 | 8266665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8266665&pc=OZD | 12/29/2018 | 01/17/2019 | 01/17/2019 | 01/22/2019 | 01/17/2019 | 01/22/2019 |  | 01/22/2019 |  |  |  |  | 01/22/2019 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA AUTOMATED CONTROLLER *(AIC) |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0000-US |  | No answer provided | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 6570 DA |  |  | 66 |  |  | Malfunction | No | Yes | Use of Device Problem | Obstruction of Flow | [5 empty] | Patient Problem/Medical Problem | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Per vad coord: the high purge pressure and purge flow blocked alarms around 0800. patient vitals and impella flows remained stable and unchanged. map >65 flow 3.8. purge pressure 1000-1300, purge flow 0-1 ml/hr. line de-aired and alarms briefly resolved. rn and team captain were preparing a new purge cassette (per impella online troubleshooting recommendations), when vad coord. called back saying not to change purge, but to do a maneuver that had been done on the patient last week, per impella rep. the maneuver was performed by md and np. impella flows dropped to p0 x 1 second and then changed back to pt baseline p6. pt was asymptomatic and tolerated procedure very well. no change in vital signs. alarms and issue with purge pressures/blockage resolved. the rep communicated the procedure to vad coord. communicated with ct icu and the primary team. this is our current practice. this is not a trend. the vad team is tracking the impella issues, and there is no trend. each one has been different. this patient was (b)(6) and therefore could not receive heparin. the purge solution that is delivered by the purge cassette is supposed to have heparin in it, but we could not do that because of his hit. initially, the team ordered argatroban to be put in the purge solution because that was an anticoagulant that the patient could tolerate but it was unclear if the argatroban was making the fluid too viscous to flow through the tubing properly, so the icu team decided to take it out and run d5w without any anticoagulant. this is not recommended by the manufacturer because of the risk of the impella clotting. clinically we did not have a choice, and it was a decision made by our physicians. the situation that was described here was likely caused by the fact that there was no anticoagulation in the purge fluid. appropriate procedure was carried out.
37896 | 1220648-2019-00016 | 8269029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8269029&pc=OZD | 12/26/2018 | 12/26/2018 |  | 01/22/2019 |  |  | 12/26/2018 | 01/22/2019 |  |  | 09/30/2020 |  | 01/22/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1375524 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 57 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) caucasian male with no history of cardiogenic shock and ef 35. the patient was brought to cardiac lab (cath) from the emergency room (er), impella cp was placed through right femoral artery (rfa) without incident. on (b)(6) 2018 the doctor tried to reposition the device due to hemolysis above 100 level and low flows caused by device being wrapped up in the mitral valve (mv). after multiple attempts and accidentally pulling device out of the heart and unable to re-cross, the decision was made to take the patient to the cardiac catheterization laboratory (ccl) to replace. upon set up of a new device, the patient began to decompensate and doctor was able to re-advance the current pump back across the atrioventricular node (av). the pump was explanted on (b)(6) 2018.
37897 | 1220648-2019-00008 | 8269030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8269030&pc=OZD | 12/23/2018 | 02/04/2019 |  | 01/22/2019 |  |  | 12/25/2018 | 01/22/2019 |  |  | 08/31/2020 |  | 01/22/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1371792 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 39 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male with past medical history (pmh) of smoking. on (b)(6) 2018, patient presented with myocardial infarction (awmi). the impella cp was placed in the right femoral artery (rfa) in order to allow the heart to rest and echocardiogram (echo) showed good position. patient received two (2) units of blood due to hemoglobin (hb) dropping from 9.9 and on (b)(6) 2018 to 7.9. the impella was explanted on (b)(6) 2018, and was within target range for weaning or recovery. | [5 empty] | The returned repositioning unit was inspected and found to be free of damage and kinks. minimal blood was found in the tuohy-borst fitting indicating that the sleeve was properly secured to the catheter. evidence of blood at the back of the blue suture pad was found. the bleeding probably occurred due to act levels reported in the clinical account to be above the recommended range. per (b)(4), "after insertion of the catheter (and until explant), act should be maintained at 160 to 180 seconds." the root cause of the excessive bleeding at the insertion site was most likely due to excessive anti-coagulation therapy. | The impella cp pump was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37898 | 1220648-2019-00007 | 8269036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8269036&pc=OZD | 12/23/2018 | 01/30/2019 |  | 01/22/2019 |  |  | 12/23/2018 | 01/22/2019 |  |  | 10/31/2020 |  | 01/22/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1380546 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 21 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) caucasian female with diagnosis of juvenile cardiomyopathy. patient was diagnosed at (b)(6). an impella cp was placed. the doctors called and stated the patient has hemolysis above 100 level despite multiple repositions and elevated plasma-free hemoglobin (pfhgb). given the severity of the disease and the patient's age the decision was made to change the pump. there was multiple plasma-hemolysis labs done and when they switched to a new pump the hemolysis went away. | [5 empty] | The returned pump was inspected and found to be free of damages, with no biomaterial noted on the pump. the pump passed hemolysis testing, indicating that the failure was not related to a defect of the device. the root cause of the hemolysis was unable to be determined because the failure mode was unable to be reproduced. | The impella cp optical was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37899 | 1220648-2019-00018 | 8261927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8261927&pc=OZD | 12/23/2018 | 12/23/2018 |  | 01/18/2019 |  |  | 12/23/2018 | 01/18/2019 |  |  | 10/31/2020 |  | 01/18/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL WITH SMART ASSIST |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1378379 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A patient with extensive cardiovascular history and comorbidities had an impella cp placed at the right femoral artery for support during a high risk coronary angioplasty. at the conclusion of the angioplasty, the team noted a blood loss at the access site that prompted a manual hold and sutures. the patient was then transferred to the icu with the pump remaining on for continued hemodynamic support. during the overnight hours, the patient's hemoglobin levels decreased and the team gave 2 units of blood and applied a femostop to the impella site. the impella was successfully weaned and explanted. at the explant the physician did comment that perhaps the bleeding at the impella site could have been contributed to by the significant scarring and calcification at the femoral artery. | [5 empty] | Since the initial medwatch was filed the investigation has concluded. the product was returned for analysis and upon review no damage was observed. based upon the lack of device damage and the clinical account, the patient's peripheral vascular disease most likely the cause of the access site bleeding. the physician had stated the patient had significant scarring and native calcification in the area of the vascular access that could have caused the bleed. the failure mode will be monitored and trended. | The product was returned for analysis. the data logs were also returned for analysis.at this time the investigation is on-going and so a supplemental medwatch will be filed at investigation conclusion.
37900 | 1220648-2019-00017 | 8257818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8257818&pc=OZD | 12/21/2018 | 12/21/2018 |  | 01/17/2019 |  |  | 12/21/2018 | 01/17/2019 |  |  | 09/30/2020 |  | 01/17/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1376563 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old caucasian female with hypotension, decreased ejection fraction (ef) 35%, and 4+ mitral regurgitation (mr). the impella cp was placed easily in the right axillary and immediately provided flows of 3.4l/min. on (b)(6) 2018, the patient's radial pulse was absent in the impella arm and fingers turned black. the impella was explanted for loss of pressure and flow and patient was given milrinone drip. patient was stable and no other device was implanted.
37901 | MW5083215 | 8260736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8260736&pc=OZD | 12/01/2018 | 01/15/2019 |  | 01/17/2019 |  |  |  | 01/17/2019 |  |  |  |  | 01/17/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 76 YR |  |  | 39 |  |  | Injury |  | Yes | Fluid/Blood Leak | Protective Measures Problem | [5 empty] | No Information | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Impella 2.5 device inserted for patient's unstable condition. after insertion, purge pressure alarm noted. device was leaking and then failed. new device inserted.
37902 | 1220648-2019-00014 | 8248136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8248136&pc=OZD | 12/19/2018 | 12/19/2018 |  | 01/15/2019 |  |  | 12/19/2019 | 01/15/2019 |  |  | 07/31/2020 |  | 01/15/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED |  |  | US |  | 1359859 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Stroke/CVA | [12 empty] | Hospitalization | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A patient with multiple cardiovascular comorbidities was admitted in cardiogenic shock and had an impella placed via the right axillay artery and graft. the impella supported the patient for 4 days until explant. during the support and subsequent hospital stay the patient was noted to have suffered a parietal stroke. the cause of the stroke is unknown, and the use of impella could not be excluded by the medical team as a contributor to the neurological event. | [5 empty] | The product was not returned for analysis, nor were imaging studies. the data logs were retuned for analysis. the root cause of the parietal stroke was unable to be definitively determined as the pump was not returned from the field, but may have related to the reported carotid disease and repeated delivery attempts. no corrective action is recommended as the root cause was unable to be determined. failures of this type will continue to be monitored for trends.
37903 | 1035166-2019-00003 | 8248288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8248288&pc=DYB | 12/28/2018 | 05/29/2019 |  | 01/15/2019 |  |  | 01/03/2019 | 01/15/2019 |  |  |  |  | 01/15/2019 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | CR-02060 |  | 0052-3021 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 43 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup | Followup | Followup |  |  | * | Invalid/Incomplete |  |  | The customer reported the facility only returned the impella pump, the introducer was discarded. | It was reported the patient was presented to the physician with right ventricular failure post pulmonary artery embolism post-operative to an orthopedic surgery. during access at right femoral vein for the placement of the impella rp, the site was seen to be oozing blood. the ooze was treated with 1-2 units of blood product and vascular surgery to stop the bleeding, after mattress "stich" and femostop failed to achieve hemostasis. the team made a decision to remove the product in the operating room, and afterwards patient was noted to be stable. to date, no problems noted by hospital. the product is expected to be returned for evaluation. |  |  |  |  | The device was not returned for evaluation as it was discarded; there was no specific device related failure reported by the user. the information was requested on the cause of the reported blood loss but was not available. the device history record could not be located, however, inspection procedures require any oscor product to pass all in-process and qa final inspections before shipping to the customer. as per procedure adelante s2s introducer sheath in-process and final inspection: vacuum leak test - perform pressure and vacuum leak test per procedure. this test is performed by qa on 100% of the product. as per instructions for use (ifu): directions for use: introducer removal 16. flush the sheath with 5 cc of saline immediately before peeling the sheath away in order to minimize backbleeding. 17. first, completely withdraw the sheath from the vessel. 18. then remove the sheath by sharply snapping the tabs of the valve housing in a plane perpendicular to the long axis of the sheath to split the valve and peel the sheath apart . as per ifu: bleeding is possible complications/adverse effects. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. no further follow-up is required. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available as the investigation is on-going. a follow-up will be submitted as soon as the investigation is complete.
37904 | 1220648-2019-00015 | 8250658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8250658&pc=OZD | 12/19/2018 | 12/19/2018 |  | 01/15/2019 |  |  | 12/19/2018 | 01/15/2019 |  |  | 09/30/2020 |  | 01/15/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1372258 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Hematoma | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The patient was admitted to the cardiac catheterization lab in cardiogenic shock and had the impella cp placed via the left femoral artery. the access was made and of note the access was complicated by obesity and tortuous native vascular anatomy. the cp was to support the patient during the coronary angioplasty. upon removal of the introducer, there was an apparent hematoma at the impella access site. due to the hematoma, the impella was weaned and removed. the patient received a vascular repair in the form of an arterial closure as well as infusion of blood products. | [5 empty] | Since the filing of the initial medwatch the investigation has completed. the hospital team did not return all the product used. as such, the root cause of the hematoma and access site bleeding was unable to be determined. the large patient size and tortuous vasculature might have contributed but, this can not be determined without full product investigation. the failure mode will be monitored and trended. | The root cause of the hematoma and access site bleeding was unable to be determined as the product was not returned for review. upon return of the product, or should more information be shared with abiomed but user facility, a supplemental report will be submitted with investigation findings.
37905 | 1220648-2019-00006 | 8235235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8235235&pc=OZD | 12/15/2018 | 02/01/2019 |  | 01/10/2019 |  |  |  | 01/10/2019 |  |  | 09/30/2020 |  | 01/10/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1378000 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old caucasian male with history of coronary artery disease (cad), implantable cardiac monitor (icm), non-st-elevation myocardial infarction (nstemi) in march, atrial fibrillation (afib), chronic kidney disease (ckd), and diabetes mellitus (dm). patient had a ventricular fibrillation (vfib) arrest, defibrillated, return of spontaneous circulation (rosc). the patient also had gastrointestinal bleeding (gi bleed). patient was found to have three-vessel disease (3vd) and stent to left anterior descending (lad). the patient continued to worsen, brought back to the cath lab and impella cp was placed easily in the left groin. patient had oozing from every access site, decision was made to leave peel away in place and patient also has a large body habitus and panis. the echocardiogram (echo) and the pump was close to or on the mitral valve (mv) and the patient developed an mv rupture, the decision was made to wean the pump and remove it, removal of the pump was successful though peel away sheath. the patient now has significant mitral valve disease and required another pump to be replaced. patient received three (3) units of blood. the patient will need some kind of mv repair when patient is stable enough. | [5 empty] | The mitral valve injury and regurgitation was could have been due to improper pump position, abiomed was unable to determine as the pump and data were not returned and clinical imaging was not provided.
37906 | 1220648-2019-00013 | 8236715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8236715&pc=OZD | 12/18/2018 | 12/18/2018 |  | 01/10/2019 |  |  | 12/18/2018 | 01/10/2019 |  |  | 06/30/2020 |  | 01/10/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1367241 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG | [10 empty] | Injury | Yes | Yes | Physical Resistance/Sticking | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Other | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a male patient presenting with acute myocardial infarction and cardiogenic shock. an impella cp was inserted successfully, however, when removing the 0.018" guide wire, there was resistance. although physician was successful in applying the necessary force to remove this component from the pump, the wire's coating was "shaved" off and wire was torn. no intervention or extra measure were performed to remove any fragments. thereafter, the pump ran without any malfunctions. | [5 empty] | The guide wire was returned for analysis, however, pump was not returned. field images show coil and core of wire were fractured. the returned guidewire fracture surface confirmed a ductile fracture occurred. bends and kinks in the core existing only in the distal third of the core were also found. device history record and found no manufacturing issues or reworks associated with this pump lot. all guide wires in this lot passed all inspection. there were no other complaints related to the failure mode in this guidewire lot. the root cause of the guidewire fracture was most likely excessive force used in attempting to remove the wire, however, need for excessive force is unable to be determined.
37907 | 1035166-2019-00002 | 8237479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8237479&pc=DYB | 12/14/2018 | 01/10/2019 |  | 01/10/2019 |  |  | 12/19/2018 | 01/10/2019 |  |  | 07/01/2021 |  | 01/10/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-15069 |  | 0052-3000 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN | [10 empty] | Injury | Yes | Yes | Break | Crack | [5 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that during placement of impella cp, sheath broke to 4 pieces in the patient. as per additional information, sheath was torn and there were 4 or 5 longitudinal cracks. patient loss of 2 liter blood in cath lab, as result 14 units of blood transfusion performed. to prevent the blood loss access site was treated by 25 minute of manual compression and after that vascular surgeon sutured the puncture site. the procedure is aborted and patient was supported by va ecmo. no additional information is available and no other adverse event was reported. | [5 empty] | The device was used for treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure abiomed introducer sheath in-process and final inspection 9.2.2 visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. slight discoloration from tipping process is acceptable. verify proper tip shape according to drawing. per ifu never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. oscor will continue to monitor this event type and risk.
37908 | 1220648-2019-00012 | 8230960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8230960&pc=OZD | 09/11/2018 | 12/14/2018 |  | 01/09/2019 |  |  | 12/14/2018 | 01/09/2019 |  |  | 02/29/2020 |  | 01/09/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED |  |  | US |  | 1331385 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A (b)(6) year old female in cardiogenic shock was admitted to the facility in (b)(6) for cardiac care. during the coronary angioplasty the patient condition deteriorated and so the impella 2.5 was placed. on the 3rd day of support there was an infusion made of 4 units of blood to treat presumed hemolysis. the pump remained on for support for another 4 days until explant. | [5 empty] | The impella pump and data logs were not returned for analysis. the logs are expected as are more clinical details. the investigation of the event has yet to begin. upon completion of the investigation, results will be shared via a supplemental medwatch.
37909 | 1035166-2019-00001 | 8226575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8226575&pc=DYB | 12/24/2018 | 01/08/2019 |  | 01/08/2019 |  |  | 12/28/2018 | 01/08/2019 |  |  | 08/01/2021 |  | 01/08/2019 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-15174 |  | 0052-3000 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN | [10 empty] | Injury | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Other | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that patient was admitted for high risk angioplasty. during the access at the femoral artery for placement of the impella cp, 150cc of blood lost noted which prompted placement of a new 14fr x13cm introducer. there was no surgical intervention or blood products required to control the blood loss. angioplasty procedure was done with impella support. however, there was no product problem reported. no additional information is available and no other adverse event was reported. | [5 empty] | Device used in treatment. the device was not returned for analysis. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. there was no specific performance related failure reported by the user. the device passed all in-process and qa final inspection steps before shipping to the customer. no further investigation is required. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. no further follow-up is required.
37910 | 1220648-2018-00178 | 8227067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8227067&pc=OZD | 11/26/2018 | 12/13/2018 |  | 01/08/2019 |  |  | 12/13/2018 | 01/08/2019 |  |  | 07/31/2020 |  | 01/08/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1354010 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male presenting with acute myocardial infarction and cardiogenic shock. during the procedure, patient endured limb ischemia, as a result of using the device. as treatment, device was removed and patient's condition improved. | [5 empty] | The impella lp2.5 was discarded by the customer and therefore, no investigation of device was not possible. should any new information be returned, we a supplemental mdr will be filed with the results of the investigation.
37911 | 1220648-2019-00010 | 8228538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8228538&pc=OZD | 10/02/2018 | 02/08/2019 |  | 01/08/2019 |  |  | 12/12/2018 | 01/08/2019 |  |  | 04/30/2020 |  | 01/08/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1347537 |  | 005048-JP |  | No | [5 empty] | No | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 50 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | No Code Available | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male presenting with myocardial infarction and cardiogenic shock. an impella lp2.5 device was inserted, thereafter. the patient was diagnosed with thrombocytopenia, developing from sepsis. as a result, the patient was treated with antibiotics. | [5 empty] | The impella lp2.5 was not received as it was discarded by the customer. should new information be received after this report, a final (b)(6) will be filed at that point. | Erroneously referred to regulatory report as "yoshiki" in initial mdr. statement should read "should new information be received after this report, a final mdr will be filed at that point." the impella lp2.5 discarded by the customer, therefore, unable to be returned for investigation. a failure analysis investigation was performed to review the device history record and found no manufacturing issues or reworks associated with this pump lot. there are no other complaints associated with pumps from this lot.
37912 | 1220648-2019-00002 | 8228622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8228622&pc=OZD | 10/25/2018 | 10/25/2018 |  | 01/08/2019 |  |  | 10/24/2018 | 01/08/2019 |  |  | 06/30/2020 |  | 01/08/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1363451 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 33 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) caucasian male that had collapsed on the street, resuscitation unclear at the time, there was intervention of right coronary artery (rca) and left coronary artery (lca) with 3 drug -eluting stent (des) and impella was implanted. later that evening patient was unstable with hemodynamic and ecmo implanted. a day later on (b)(6) 2018, patient had leg ischemia on the impella side and the impella was removed. after the explantation the leg recovered. | [5 empty] | The device was not received. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37913 | 1220648-2019-00003 | 8224278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8224278&pc=OZD | 12/10/2018 | 12/10/2018 |  | 01/07/2019 |  |  | 12/10/2018 | 01/07/2019 |  |  | 08/31/2020 |  | 01/07/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1374663 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 70 YR | [5 empty] | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old caucasian male was transferred with intra-aortic balloon pump (iabp) in cardiogenic shock (cgs) on levophed and dopamine drips. the iabp was removed and the 14 fr sheath was inserted and then impella cp was quickly inserted with no issues. there was an estimated amount of blood loss at approximately 100cc and 1 unit of blood was given. heparin was used as anticoagulation medication. | [5 empty] | The device was not received. if device is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37914 | 1220648-2018-00176 | 8219312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8219312&pc=OZD | 12/08/2018 | 12/08/2018 |  | 01/04/2019 |  |  | 12/08/2018 | 01/04/2019 |  |  | 10/31/2020 |  | 01/04/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP OPTICAL |  |  | ABIOMED, INC. |  |  | US |  | 1376352 |  | 0048-0024 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male presenting in an unstable condition. decision was made to implant an impella cp optical device in the patient's left axillary artery. during the course of treatment the following day, patient exhibited ischemia in their left arm. as treatment, the device was removed and the ischemia improved. no new pump was inserted and no additional intervention was required. | [5 empty] | The impella cp was discarded by the customer and therefore, no investigation of device was not possible. should any new information be returned, we a supplemental mdr will be filed with the results of the investigation.
37915 | 1220648-2018-00177 | 8219384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8219384&pc=OZD | 12/09/2018 | 12/10/2018 |  | 01/04/2019 |  |  | 12/10/2018 | 01/04/2019 |  |  | 08/31/2020 |  | 01/04/2019 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA LP5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1374252 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Mitral Regurgitation | Injury | [10 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male presenting as a post heart transplant recipient. an impella lp5.0 pump was inserted into the patient's right axillary artery. during implant, there was a moment of resistance that resolved and pump achieved correct positioning within the heart. overnight, patient exhibited signs of hemolysis and subsequent echocardiogram confirmed mitral regurgitation due to a damaged mitral valve. it was believed that the impella device may have damaged the mitral valve during insertion, leading to the aforementioned adverse events. as treatment, the impella device was removed. hemolysis was resolved via plasmapheresis and dialysis, while mitral valve was surgically repaired. | [5 empty] | The impella lp5.0 was discarded by the customer and therefore, no investigation of device was not possible. should any new information be returned, we a supplemental mdr will be filed with the results of the investigation.
37916 | 1220648-2018-00173 | 8206908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8206908&pc=OZD | 12/04/2018 | 12/04/2018 |  | 12/31/2018 |  |  | 12/04/2018 | 12/31/2018 |  |  | 08/31/2020 |  | 12/31/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1365883 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) caucasian female patient presenting with myocardial infarction, flash pulmonary edema and the beginning states of cardiogenic shock with an ejection fraction of 20%. impella cp was inserted with no issues. during procedure, a hematoma developed at the device's insertion site that required 2 units of blood products delivered. device was successfully weaned, pressure was applied, and site was closed successfully via combination of perclose and angioseal. the following day, access site remained stable. | [5 empty] | The impella cp was not returned by the customer, therefore, investigation of device not possible. data logs were received and will be analyzed to ensure proper device functionality. should there be any deficiencies in data logs, or new information received, a supplemental mdr will be filed.
37917 | 1220648-2018-00175 | 8206914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8206914&pc=PYX | 12/11/2019 | 02/07/2019 |  | 12/31/2018 |  |  | 12/05/2018 | 12/31/2018 |  |  | 12/31/2019 |  | 12/31/2018 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1321488 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A complainant reported a (b)(6) year old male presenting with right ventricular failure. an impella rp was inserted without issue. during procedure, insertion site was found bleeding. as treatment, 1 unit of blood products were administered and bleed has since resolved. patient continued to receive treatment with this device and has since been weaned successfully. | [5 empty] | The impella rp was returned for failure analysis investigation. device was observed to have rough edges and minor scratches. a review the device history record and found no manufacturing issues or reworks associated with this pump lot. all repositioning units in this lot passed incoming inspection. there were no complaints associated with other units in this lot. the cause of the minor damage to re-positioning sheath and access site bleed was unable to be definitively determined. | The impella rp was returned by the customer and a failure analysis investigation is underway. a supplemental mdr will be filed at the completion of the device's investigation.
37918 | 1220648-2018-00174 | 8206922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8206922&pc=OZD | 12/06/2018 | 12/06/2018 |  | 12/31/2018 |  |  | 12/06/2018 | 12/31/2018 |  |  | 09/30/2020 |  | 12/31/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP OPTICAL |  |  | ABIOMED, INC. |  |  | US |  | 1371125 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 87 YR | [5 empty] | Malfunction | No | Yes | Detachment of Device or Device Component | Infusion or Flow Problem | No Apparent Adverse Event |  |  |  |  | No Information | [41 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male presenting for high risk percutaneous coronary insertion. an impella cp optical was placed and initiated without issue. approximately 20 minutes into procedure, flows reduced and suction alarms ensued. device was repositioned, however, issues did not resolve. decision was made to explant device. a new impella cp was inserted and ran successfully. there were no adverse effects to the patient, and none expected if issue were to recur. | [5 empty] | The impella cp optical and data logs were returned for evaluation. review of the device's data logs confirm suction responses approximately 18 minutes into the case, following a spike in motor current. after approximately 22 minutes, 3 motor current spikes were observed along with a motor current drop. evaluation of the returned device found fractured impeller blades. simulated use testing was successful in recreating low pump flow, after running device in a basin of water. device history record and found no manufacturing issues or reworks associated with this pump. there are no other complaints associated with pumps from this lot. one impeller from these two lots was rejected for failure during the visual inspection test, however, all other impellers passed inspection. the root cause of the low pump flow was the broken impeller blade.
37919 | 1220648-2018-00170 | 8204417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8204417&pc=OZD | 11/30/2018 | 02/04/2019 |  | 12/28/2018 |  |  | 12/03/2018 | 12/28/2018 |  |  | 06/30/2020 |  | 12/28/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1356085 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 83 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Intermenstrual Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) year old caucasian male presenting with high risk percutaneous coronary intervention. impella cp pump was inserted on (b)(6) 2018 without any reported issues, however, on (b)(6) 2018 registered nurse stated the patient received a total of ten (10) units of packed red blood cells (prbc) for bilateral groin site bleeding along with suspected heparin induced thrombocytopenia (hit), primarily from impella site. there was two (2) units of blood given on (b)(6) 2018, four (4) units of blood given on (b)(6) 2018, two (2) units of blood given on (b)(6) 2018, one (1) unit of platelets was given on (b)(6) 2018, and two (2) units of blood given on (b)(6) 2018. hit panel came back negative. | [5 empty] | The failure mode of access site bleeding was confirmed. the cause of the access site bleeding was most likely due to arterial calcification, but was unable to be definitively determined without clinical details. | The device was received but has not been evaluated. a supplemental mdr will be filed once device evaluation is completed.
37920 | 8193716 | 8193716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8193716&pc=KFM | 09/26/2018 | 11/02/2018 | 11/20/2018 | 12/24/2018 | 11/20/2018 | 12/24/2018 |  | 12/24/2018 |  |  |  |  | 12/24/2018 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA |  |  | KFM | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 26280 DA |  |  | 111 |  |  | Injury |  | Yes | Break | Device Dislodged or Dislocated | [5 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Life Threatening | Required Intervention | [27 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Elderly male with past medical history (pmh) of cardiac and pulmonary sarcoidosis got admitted for cardiogenic shock with complaints of progressively worsening dyspnea, lower extremity edema as well as 15 pound weight gain. diagnosed as combined systolic and diastolic cardiomyopathy with cardiogenic shock and oliguria. impella 5.0 was placed via right axillary artery as a bridge to decision regarding vad candidacy. the next day: pt was working with ot and rn. while returning to bed, impella 5.0 cable became dislodged, as a result pt became hypotensive, attending provider and cv surgery team notified, epinephrine restarted, norepinephrine dose increased. impella representative notified. new console and cables connected without return of function to pump. clinical team were at the bedside. pt stabilized with restarting epinephrine, increased norepinephrine. after cable disconnected, pump was no longer functioning. pt returned to cardiovascular operating room (cvor), and had new impella placed. prior to working with ot and during the session, position of securement device was verified by both ot and rn to ensure securement. it is unclear how cable became disconnected during repositioning. while connecting the device back, it appears the prongs inside of the gray connection cable was damaged. one side of the gray connection cable is connected to the impella inside of the patient and the other end is connected to the impella console. the prongs must have been damaged on the impella connection side because a new impella console cable was connected to the patient with no change and a new impella console was used as well.
37921 | 1220648-2018-00171 | 8191852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8191852&pc=OZD | 11/30/2018 | 11/30/2018 |  | 12/21/2018 |  |  | 11/30/2018 | 12/21/2018 |  |  | 11/30/2020 |  | 12/21/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1375988 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | Yes | Material Twisted/Bent | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male presenting with acute myocardial infarction and cardiogenic shock. during insertion of cp pump through aorta into the left ventricle, the cannula bent. a radial puncture was required, snare was inserted, and cannula was straighten for removal. a new impella was inserted and ran without issue. | [5 empty] | The device has not been received for evaluation as device was discarded by customer. data logs were obtained and the investigation is underway. if new information is obtained, a supplemental mdr will be filed.
37922 | 1035166-2018-00127 | 8184777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8184777&pc=DYB | 12/03/2018 | 01/25/2019 |  | 12/20/2018 |  |  | 12/05/2018 | 12/20/2018 |  |  | 03/01/2018 |  | 12/20/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 23F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-14074 |  | 0052-3021 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 72 YR | [5 empty] | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | Cardiogenic Shock | [41 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | As per additional information, there was no adverse event associated with this device failure. details about second introducer is unknown. there was delay to removal and attempted placement of second introducer. product was discarded. | The customer reported that an event occurred on (b)(6) 2018 when using the 23f abiomed introducer. the patient presented with cardiogenic shock with both right and left heart failure. the physician was attempting insertion at the femoral vein for placement of the impella rp. during use the physician stated that both 23f devices kinked when attempting to gain access at the femoral vein. when both 23 f failed, the impella placement was abandoned. additional information was reported on (b)(6) 2018. the patient was a (b)(6) female. the kink was due to inability to cross the femoral vein and torque appropriately and the placement was aborted. the patient was placed on comfort care and support was withdrawn resulting in patient death. the patient died on (b)(6) 2018. it was reported that the device was not a contributing factor to the patient's death. |  |  |  |  | . the device was used for treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. the device history records were reviewed to confirm that the device passed all applicable in process and final inspections. per procedure abiomed introducer sheath in-process and final inspections 12.2 visual inspection: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. per ifu: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | A preliminary review showed there were no manufacturing rejects or anomalies of this event type recorded in the device history record. a conclusion is not yet available as the investigation is on-going.
37923 | 1220648-2018-00169 | 8186229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8186229&pc=OZD | 07/17/2018 | 11/26/2018 |  | 12/20/2018 |  |  | 11/26/2018 | 12/20/2018 |  |  | 05/31/2020 |  | 12/20/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP 5.0 PUMP |  |  | OZD |  |  | IMPELLA LP 5.0 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 1342102 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male presenting with low cardiac function suspected of acute myocarditis. impella lp5.0 pump was inserted on (b)(6) 2018 without any reported issues, however, on (b)(6) 2018, hematoma was found during impella support. as treatment, doctor performed a surgical operation. | [5 empty] | The device nor data logs were received for evaluation. if device or data is received after this date, an evaluation will be performed and a supplemental mdr will be filed.
37924 | 1220648-2018-00166 | 8186278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8186278&pc=OZD | 12/10/2018 | 12/10/2018 |  | 12/20/2018 |  |  | 12/10/2018 | 12/20/2018 |  |  | 07/31/2020 |  | 12/20/2018 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1359103 |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | Insufficient Information | [5 empty] | Hematoma | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The physician placed an impella cp via the left femoral artery for support of an obese patient with cardiogenic shock. the support was successful during the coronary angioplasty, however the team did observe a hematoma at the access site. the hematoma was controlled by manual pressure, removal of the impella, and an infusion of two units of blood. when the access site bleed was controlled, a second cp was placed due to continued need of hemodynamic support. | [5 empty] | The product was discarded at the user facility and therefore an investigation of the product could not be performed. the data logs were returned, but do not pertain to the failure. the ifu does state an "evaluation prior to inserting the impella catheter... observation of obesity" should be made. whether the obesity and deep tissue track contributed to the hematoma and bleed, and the impella pump failing to achieve hemostasis at the access, is unknown at this time. should more information be shared and new investigation findings become available, a supplemental report will be submitted.
37925 | 1035166-2018-00104 | 8160019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8160019&pc=DYB | 11/17/2018 | 03/22/2019 |  | 12/13/2018 |  |  | 11/19/2018 | 12/13/2018 |  |  | 03/01/2021 |  | 12/13/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC |  |  | US |  | C1-14153 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN | [10 empty] | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Rupture | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | Visual inspection of return device shows that sheath was split up, one of the scorelines approximately 11.5cm from the distal tip. | It was reported that dilator lock mechanism on the sheath turns around its own axis and does not tighten when it's pushed back into the sheath. there was bleeding from groin. to stop bleeding physician advanced the introducer in femoral artery but it failed as dilator pushed out again. additionally the femoral artery ruptured. because hemodynamic support is urgently needed, puncture the contralateral groin and placed the impella catheter over a replacement introducer. rupture of the femoral artery was closed by intervision radiologists & vascular surgeons via stent implantation in a femoral artery. patient was presented to physician in cardiogenic shock. patient died on (b)(6) 2018 as care was withdrawn by family due to question of neurological status from prolonged cpr. no additional information is available. |  |  |  |  | Device used in treatment. one 14f abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. visible blood was found on and inside the sheath. the sheath was split up one of the scorelines approximately 11.5cm from the distal tip. according to the complaint, the end user reported that the dilator lock mechanism on the sheath turns around its own axis and does not tighten. upon evaluation of the returned product under a 10x microscope, it was found that the notches (locking mechanism) on the hub cap gouged/deformed. however, this hub was tested by the distributor prior to returning to oscor. the sheath was also split up one of the scorelines approximately 11.5cm from the distal tip and the circumference around the edge of the distal tip is deformed. the device was handled by the customer prior to shipping it back to oscor. it is unknown what kind of further damage was done to the hub cap locking mechanism by the customer by engaging a reference dilator through the sheath. no manufacturing defects were found. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure abiomed introducer sheath in-process and final inspections the following are inspected 100% unless otherwise stated: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. with naked eye at a distance of 12" to 18", verify the sheath hub and split caps are free of excess silicone oil applied between the seals during assembly. verify split caps are properly placed and secured onto sheath hub and free of cracks/damages or excessive adhesive. vacuum leak test: perform pressure and vacuum leak test per procedure. destructive testing: sampling plan ansi z 1.4, special level iv, aql 1.0 reduced (product used for destructive testing mustbe rejected and not included in the final production lot. assembly fit check: coat a dilator with silicone oil med400, 1000 cp using a foam. insert dilator into sheath, lock and unlock dilator hub lock nut for verification of proper fit. at a distance of 12-18, verify there are no gaps between the sheath tip and dilator. as per instructions for use ifu: 9. insert vessel dilator into sheath and rotate the dilator cap over valve housing to secure the dilator onto sheath assembly. 10. thread the dilator/sheath assembly over the guidewire. 11. advance the dilator and sheath together with a twisting motion over the guidewire and into the vessel. fluoroscopic observation may be advisable. attaching a clamp or hemostat to the proximal end of the guidewire will prevent inadvertently advancing the guidewire entirely into the patient. 12. once assembly is fully introduced into the vascular system, separate the dilator cap from the sheath valve housing by rotating the dilator cap off the hub. capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress, follow up report will be submitted if we found any further additional information .
37926 | 1220648-2018-00161 | 8157515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8157515&pc=OZD | 11/18/2018 | 11/18/2018 |  | 12/12/2018 |  |  | 11/18/2018 | 12/12/2018 |  |  | 08/31/2020 |  | 12/12/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1366126 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR |  |  | 136 |  |  | Injury | Yes |  | Difficult to Remove | Difficult to Advance | [5 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that a (b)(6) male patient weighing (b)(6) was admitted to this facility following chest pains suffered at home. the patient was found to have thick tissue and to have had previous stents, during which multiple femoral sticks, resulting in observed scar tissue. during patient treatment thrombi were found in the ramus and circumflex arterial stents placed 2 weeks previous. the thrombi were removed and new replacement arterial stents were placed. following the placement of the replacement stents patient circulation was restored. at this point an impella cp was placed. access site bleeding and a hematoma was observed at this time. numerous unsuccessful attempts to reposition and advance the guide-wire repositioning unit were made. the physician then decided to remove the impella device. a contralateral balloon tamponade was obtained to occlude and limit blood loss. vascular surgery was called and the patient was taken to operating room for open left femoral cut down closure procedure to surgically repair the left femoral artery. closure was successful and an impella lp5.0 was then successfully placed in the patient. following this event and the placement of an impella lp5.0 the patient was reported to be in stable condition, and greatly improved. | [5 empty] | The impella cp and the introducer were returned for evaluation. the guide-wire repositioning unit was examined and found to be deformed along the length of the sheath, with visible ribbing present along the sheath. the ribbing was consistent with damage due to buckling of the sheath. damage was also observed at the tip of the guide-wire. the returned guide-wire repositioning unit showed signs of buckling, with curvature of the sheath consistent with the physical described reported by the complainant. there was also damage to the tip of the sheath. based on the patient condition as described by the clinical account, it is likely that the tip of the guide-wire repositioning unit became stuck on an obstruction in the vasculature. excessive force was likely applied in attempts t0 overcome this, leading to the buckling and ribbing evident in the returned sheath. the guide-wire repositioning unit was tested for sheath kinking, buckling, and tip split forces to the required specifications, and was found to be within clinically acceptable range. the root cause of the access site bleeding was most likely excessive force used to overcome the difficult patient anatomy, leading to damage to the sheath allowing for the access site bleeding observed.
37927 | 1220648-2018-00165 | 8151398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8151398&pc=OZD | 11/19/2018 | 01/15/2019 |  | 12/11/2018 |  |  | 11/19/2018 | 12/11/2018 |  |  | 07/31/2020 |  | 12/11/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1368182 |  | 005042 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 82 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | Blood Loss | [11 empty] | Death | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a female patient presenting with hypertension, diabetes, multi-vessel coronary artery disease, obesity, and an ejection fraction of 20%. due to patient's size a 14fr x 30cm sheath was used in an attempt to insert the impella lp2.5, however, physician was unable to pass the device through patient's anatomy. against the clinical field representative's recommendations to use 30cm sheath, physician elected to use a shorter 13fr x 13cm merit introducer to reach patient's left femoral artery. pump was inserted without issue and physician elected to leave sheath in place. however, during radial intervention, severe blood loss was detected at the impella insertion site. as treatment, manual pressure was applied, blood products were delivered, protamine administered (to reverse heparin), and a vascular surgeon was contacted for device removal. patient lost pulsatility and vitals deteriorated. after 40 minutes of attempted resuscitation, patient care was withdrawn and patient expired. the cause of death was due to complications due to bleeding and congestive heart failure. bleed was attributed introducer and repositioning sheath not bleeding long enough to effectively reach the patients arteriotomy due to patient's size. | [5 empty] | No samples were returned from the customer, as they had been discarded after this event. a review of the device history record and found no manufacturing issues or reworks associated with this pump lot. there are no other complaints associated with pumps from this lot. all introducers in this lot passed all inspection. | The impella lp2.5 was discarded by the customer, however, a failure analysis investigation is underway and supplemental medwatch report will be submitted upon completion of the evaluation.
37928 | 1220648-2018-00160 | 8152270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8152270&pc=OZD | 11/12/2018 | 11/13/2018 |  | 12/11/2018 |  |  | 11/13/2018 | 12/11/2018 |  |  | 06/30/2020 |  | 12/11/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA LP5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1353068 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | On (b)(6) 2018 an impella lp5.0 was placed in a (b)(6) year old female patient. the patient's hemodynamics did not improve following placement of the lp5.0 pump, so an ecmo was also placed in the patient's lingual regions. following placement of the impella 5.0 and ecmo lower limb ischemia occurred. there was no reported treatment of the ischemia. | [5 empty] | The facility did not return the impella pump product for investigation; therefore the root cause of the limb ischemia was unable to be determined.
37929 | 1220648-2018-00158 | 8152505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8152505&pc=OZD | 09/04/2018 | 09/08/2018 |  | 12/11/2018 |  |  | 09/08/2018 | 12/11/2018 |  |  | 06/30/2020 |  | 12/11/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1356085 |  | 0048-0032 |  | No | [5 empty] | Invalid/Incomplete | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) african-american male presenting with cardiomyopathy. impella cp was placed without any reported issues. attempted weaning of device was not tolerated by patient and hemoglobin levels decreased. it was reported that access site bleeding was present and could have also contributed to the decrease in hemoglobin levels. as treatment, 1 unit of blood products were administered. | [5 empty] | No evaluation was performed as the impella cp and data logs were not returned by the customer. should these items be received ,a supplemental medwatch report will be submitted, thereafter.
37930 | 1220648-2018-00164 | 8148767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8148767&pc=OZD | 08/10/2018 | 01/24/2019 |  | 12/10/2018 |  |  | 11/16/2018 | 12/10/2018 |  |  | 01/31/2020 |  | 12/10/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1324747 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 91 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female patient presenting with heart failure, mitral regurgitation, and hemodynamic breakdown during balloon aortic valvuloplasty. an impella lp2.5 was prepped and inserted on (b)(6) 2018 in the right femoral artery with no issues. device was explanted on (b)(6) 2018 and no malfunctions regarding the operation of the device were reported. the following day, however, the physician identified limb ischemia in proximity to the impella insertion site. as treatment, the thrombus was removed and there were no lasting harm endured by the patient, as a result. although it was not confirmed, the physician alleged the device had either caused and/or contributed to the thrombus and limb ischemia. | [5 empty] | The root cause of the thrombus formation and subsequent ischemia was unable to be determined as the device was not returned due to customer discarding the device. datalogs were analyzed and revealed an instance of suction during the case. a failure analysis investigation was performed to review the device history record and found no manufacturing issues or reworks associated with this pump lot. there are no other complaints associated with pumps from this lot. | The impella lp2.5 has not been received from the customer, however, data logs have been received and a failure analysis investigation is underway. a supplemental medwatch report will be submitted upon completion of the evaluation.
37931 | 1220648-2018-00162 | 8139699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8139699&pc=OZD | 04/21/2018 | 04/15/2019 |  | 12/06/2018 |  |  | 11/09/2018 | 12/06/2018 |  |  | 01/31/2020 |  | 12/06/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1324050 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 35 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup | Followup |  |  |  | * | Invalid/Incomplete |  |  | A (b)(6) patient in (b)(6) was admitted in cardiogenic shock and had an impella 2.5 placed. during procedure, a bent pigtail was discovered, however, the pump supported the patient successfully with no operational issues. a thrombus in the proximity of the fourth intercostal space of the descending aorta was identified and removed following device explant. patient has since been discharged from the hospital. it should be known that physician had lowered act values below the recommended values set for the patient and may have contributed to the formation of the reported thrombus. | [5 empty] | This pump was not returned for analysis, however, data logs were returned and a failure investigation is underway. a supplemental report will be submitted upon completion of the investigation. | A definitive root cause was unable to be determined as the patient's clotting times were not provided and the device was not returned for the investigation. an analysis was performed to review the device history record and found no manufacturing issues or reworks associated with this pump lot. there are no other complaints associated with pumps from this lot. | Data logs were returned and reviewed. brief suction alarms were identified and had occurred multiple times during the first two days of support. impella position wrong alarms for the outlet being under the aortic valve were also identified.
37932 | 1220648-2018-00163 | 8139715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8139715&pc=OZD | 11/14/2018 | 11/14/2018 |  | 12/06/2018 |  |  | 11/14/2018 | 12/06/2018 |  |  | 06/30/2020 |  | 12/06/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1355357 |  | 0048-0032 |  | No | [5 empty] | Y | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old male presenting with acute myocardial infarction and cardiogenic shock. impella cp was inserted without any reported issues, however, during procedure patient endured bleeding around the insertion site. as treatment, 1 unit of blood was administered and control was achieved. device operated successfully and patient was explanted after 164.58 hours in stable condition. | [5 empty] | The impella lp cp was discarded by the customer, however, a failure analysis investigation is underway and supplemental report will be submitted upon completion of the evaluation.
37933 | 8124035 | 8124035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8124035&pc=PBL | 11/04/2018 | 11/26/2018 | 11/26/2018 | 12/03/2018 | 11/26/2018 | 12/03/2018 |  | 12/03/2018 |  |  |  |  | 12/03/2018 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA VAD MOBILE CONSOLE |  |  | PBL |  |  | 004603 |  |  | ABIOMED, INC |  |  | US | [5 empty] | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | No Information |  |  | User Facility report | [8 empty] | 19710 DA | [5 empty] | Malfunction |  | Yes | Display or Visual Feedback Problem | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | During shift change the impella alarmed "controller error" with a message that stated to use back up controller. it was exchanged immediately. device was sent to the vendor.
37934 | 1220648-2018-00157 | 8125329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8125329&pc=OZD | 11/05/2018 | 12/28/2018 |  | 12/03/2018 |  |  | 11/06/2018 | 12/03/2018 |  |  | 01/31/2020 |  | 12/03/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA LP5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1326543 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Regurgitation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old (b)(6) male presenting for treatment with ecmo. impella lp5.0 pump was inserted on (b)(6) 2018 without any reported issues, however, mitral regurgitation was identified. as treatment, pump was removed and condition improved. patient remains in stable condition. | [5 empty] | The device was not returned for analysis, however, data logs were retrieved and "impella outflow blocked" alarms were confirmed. the root cause of the mitral regurgitation was unable to be determined because product was not returned for review. a review of the device history record was completed and found no manufacturing issues or reworks associated with this pump lot. there are no other complaints associated with pumps from this lot. all introducers in this lot passed all inspection. | The impella lp5.0 was discarded by the customer, however, device data logs were received. a failure analysis investigation is underway and a supplemental mdr will be submitted upon completion of the evaluation.
37935 | 1220648-2018-00159 | 8125596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8125596&pc=OZD |  | 11/04/2018 |  | 12/03/2018 |  |  | 11/04/2018 | 12/03/2018 |  |  | 06/30/2020 |  | 12/03/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1370030 |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2018 a (b)(6) male patient was admitted to the hospital after having chest pains at home for a month. the physician placed an impella cp in the patient. stents were placed in the left anterior descending artery and in the circumflex artery. that evening bleeding issues occurred, and a fem stop was placed. the next morning a swan-ganz catheterization was performed in the patient's arm. in addition, the repositioning sheath was advanced in an attempt to help seal the femoral artery access site. several suction events occurred through the nights which were resolved with the administration of fluids. at 1:00am the patient's act measured 200 rather than the abiomed recommendation of 160-180. three hours later the patient's act had risen to 237. the patient was still experiencing bleeding. six units of blood were administered. at 9:00am the patient's act was at 300. platelets were then administered. the patient was reported to have been successfully supported for 17.93 hours, although the patient expired on (b)(6) 2018. there was no indication from the complainant that the impella cp caused or contributed to the patient outcome. | [5 empty] | The device was not returned to abiomed for evaluation, as the device was not removed following the patient expiring; therefore, an investigation was unable to be performed. consequently a root cause for the reported event was unable to be determined, and there is no indication that the impella cp caused or contributed to the event or patient outcome.
37936 | 1220648-2018-00146 | 8122693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8122693&pc=OZD | 10/31/2018 | 10/31/2018 |  | 12/01/2018 |  |  | 10/31/2018 | 12/01/2018 |  |  | 08/31/2020 |  | 12/01/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1366412 |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR | [5 empty] | Death | Yes | No | Suction Problem | [6 empty] | Complete Heart Block | [12 empty] | Death | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that an impella cp was placed on a (b)(6) male patient on (b)(6) 2018. during patient support suction alarms occurred. during the night suction alarms occurred when running the pump above p-2. an echo revealed that the inlet was 3.3 cm below the aortic valve, and the position was adjusted to resolve the suction. it was then confirmed with echo that there was good positioning and adequate patient volume, although the suction issue was not resolved. the patient's act was between 160-180 seconds for the duration of the case. the physician noted that there were no signs of patient hemolysis. because the patient was stable, the decision was made to leave the pump at p-2. no pressors were administered to the patient. on the morning of (b)(6) 2018 the patient had runs of ventricular tachycardia, and the patient developed a third degree heart block and quickly decompensated. stability could not be regained as the patient coded and expired due to the heart block while on impella support. the pump was checked for clots on explant and none were found, and the pump was then discarded. | [5 empty] | The impella cp was not returned to the manufacturer for evaluation, as the pump was discarded after use. the automatic impella controller logs were returned for analysis. the analysis confirmed this event, and showed that 3 hours into the case a sudden drop in motor current occurred despite no change in p-level. it is likely that an acute occlusion of the cannula occurred at this time, as the clinical account confirmed good positioning and adequate patient volume. the occlusion may have been due to a kink in the cannula or a biological obstruction, but this could not be determined as the pump was not returned for evaluation. there is no known history of impella contributing to heart block, and the development of this condition was most likely unrelated to the impella device.
37937 | 1035166-2018-00102 | 8121295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8121295&pc=DYB | 10/25/2018 | 02/13/2018 |  | 11/30/2018 |  |  | 11/08/2018 | 11/30/2018 |  |  | 03/01/2021 |  | 11/30/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3014 |  |  | OSCOR INC. |  |  | US |  | C1-14180 |  | 0052-3014 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 47 YR | [5 empty] | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Consequences Or Impact To Patient | Blood Loss | [40 empty] | Initial submission | Followup | Followup |  |  |  | * | Invalid/Incomplete |  |  | Patient was admitted for a high risk angioplasty of the coronary arteries. access at the left femoral artery for placement of the impella cp, hub leak observed with blood lost. 14fr introducer removed and replaced by 16fr cook. no blood products given to control blood loss. patient is stable and no extended hospitalization reported. no other additional information is available. | [5 empty] | The device was used in treatment. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. a review of complaints involving this lot number identified no additional reports. one 14f abiomed introducer sheath was returned from the customer. there were no other accessories. blood was found on and inside the sheath. the introducer was returned from the end user with both orange split caps detached from the hub and were not returned for evaluation. the hub was checked for a crack and the valve was checked for damage under a 10x microscope. no crack was found on or inside the hub. the valve looked normal and no damage or fatigue was found. there was a small amount of adhesive residue found on the hub. the customer reported that blood leaked from the hub during use of the introducer sheath. no crack was found on or inside the hub when viewed under a 10x microscope. the valve looked normal and no damage or fatigue was found. it was not possible to perform a leak test on the product as the valve and caps were already detached. the reason for return cannot be confirmed. no manufacturing defects were found. as per qa adelante s2s introducer sheath in-process and final inspection procedure, vacuum leak test - perform pressure and vacuum leak test. this test is performed by qa on 100% of the product. per the instructions for use (ifu): when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. infusion through the sideport can be done only after all air is removed from the unit. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. when injecting or aspirating through the sheath, use the sideport only. insert vessel dilator into sheath and rotate the dilator cap over valve housing to secure the dilator onto sheath assembly. thread the dilator/sheath assembly over the guidewire. once assembly is fully introduced into the vascular system, separate the dilator cap from the sheath valve housing by rotating the dilator cap off the hub. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Based on the investigation, cause of caps detachment could not be confirmed. however, the corrective action has been initiated to address the cap detachment issue. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is still in progress, follow up report will be submitted if we found any further additional information.
37938 | 1220648-2018-00145 | 8110701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8110701&pc=OZD | 10/29/2018 | 10/29/2018 |  | 11/28/2018 |  |  | 10/29/2018 | 11/28/2018 |  |  | 06/30/2020 |  | 11/28/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1356177 |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Use of Device Problem | Activation, Positioning or Separation Problem | Human-Device Interface Problem | Patient Device Interaction Problem |  |  |  | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that a (b)(6) year old diabetic female patient presented at the hospital with a non-st-elevation myocardial infarction (nstemi.) the patient had a history of left main artery (lm) disease and multi-vessel disease (mvd.) an impella cp was placed in the patient on (b)(6) 2018. md peeled a 14fr sheath, and then pushed the repositioning sheath with only 1/4 portion of the sheath into the vessel, rather than pushing the sheath in completely, as instructed in the abiomed "instructions for use" for the impella cp. the repositioning sheath was then sutured. the patient was then transferred to the unit where bleeding occurred at the patient's groin. the repositioning sheath was then cracked and pressure was applied to the bleeding site. it was reported that the patient was administered 2 bags of replacement blood. on (b)(6) 2018 it was decided to remove the pump and place an intra-aortic balloon pump in the patient. the physician cut the 9fr pump cannula, and tried to rewire the impella using .014/.018 wires, but was unsuccessful. the physician then placed a 12fr sheath over the catheter and sterile wrapped the 9fr catheter with tape. hemostasis was maintained throughout the procedure. in addition, the patient was hemodynamically stable throughout the procedure. the patient was transferred to another facility where the catheter was successfully removed by a vascular surgeon. the patient was reported to have been hemodynamically stable throughout patient support, and following removal of the device. | [5 empty] | The impella cp was discarded by the user facility, therefore a full investigation could not be conducted. a review of the clinical case revealed that the bleeding was most likely due to the handling of the repositioning sheath. the instructions for use advises the following: trim any excess graft and slide the repositioning sheath into the graft. note: the hub of the repositioning sheath should be at the skin level and the length of the remaining graft material should be just long enough to secure the graft around the repositioning sheath with all of the graft buried beneath the skin. using heavy silk suture, secure the graft around the hub of the repositioning sheath so that the position of the impella catheter can still be adjusted. remove the vascular clamp adjacent to the axillary artery. the wound should be closed over the trimmed graft with the end of the repositioning sheath clearly visible. anchor the repositioning sheath securely to the skin remove excess slack from the impella catheter and re-check position. tighten the tuohy-borst valve to prevent catheter migration.
37939 | 1220648-2018-00149 | 8110703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8110703&pc=OZD | 10/26/2018 | 12/28/2018 |  | 11/28/2018 |  |  | 10/30/2018 | 11/28/2018 |  |  |  |  | 11/28/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, PRODUCT |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1361035 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2018 an impella lp2.5 was placed in the left femoral artery of a (b)(6) year old female patient. on (b)(6) 2018 lower limb ischemia was seen on the patient's left lower extremity. the lower limb was then flexed from the side tube of the ecmo and the circulation of the lower then improved. it was reported that then lower extremity compartment syndrome developed. the patient then underwent a telesurgery to resolve the issue. the patient was supported for 146.42 hours until patient care was withdrawn and the patient expired. | [5 empty] | The impella 2.5 pump was not returned for evaluation. the data logs was received and evaluated, no further information was found that could show the impella device contributed to limb ischemia. | The impella lp2.5 has not yet been returned for evaluation, upon the return of the product and the completion of the evaluation a supplemental mdr will be submitted.
37940 | 1220648-2018-00144 | 8106117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8106117&pc=PYX | 11/01/2018 | 11/01/2018 |  | 11/27/2018 |  |  | 11/01/2018 | 11/27/2018 |  |  | 05/31/2020 |  | 11/27/2018 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1348530 |  | 004334 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 46 YR | [5 empty] | Injury | Yes | No | Human-Device Interface Problem | Insufficient Information | [5 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | An obese patient was admitted presenting with a 3 day old myocardial infarction, presenting as a stemi. it was determined that the patient had a septal infarct and had both right and left sided impella pumps placed. the patient "recieved" the impella rp and impella cp. in the rp access site, at the patient developed a hematoma. there was a question as too whether or not the patient's body habitus was too large and the tissue track was too deep for hemostasis by the rp repositioning sheath. the team called in vascular surgery to remedy the patient blood loss. the team achieved hemostasis and infused 3 units of blood products as well as 1 unit of platelets. within 24 hours, the rp was explanted and the patient was solely supported by the impella cp. | [5 empty] | The investigation is on-going at this time. should product or other information be shared, a supplemental report will be filed.
37941 | 1220648-2018-00155 | 8102652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8102652&pc=OZD | 10/31/2018 | 10/31/2018 |  | 11/26/2018 |  |  | 10/31/2018 | 11/26/2018 |  |  | 05/31/2020 |  | 11/26/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, IN. |  |  | US |  | 1351504 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A (b)(6) african-american male presented with acute myocardial infarction and cardiogenic shock. an impella cp was inserted without issue, however, bleeding and hematoma was identified at site. 2 units of blood products were delivered to patient and heparin administration was reduced, as treatment. the bleeding slowed and hematoma resolved, thereafter. patient has since been weaned off pump and discharged from hospital. | [5 empty] | The impella cp was returned by the customer and a failure analysis investigation was performed. simulated use testing found no deficiencies and the device operated as intended with no structural defects found. review the device history record found no manufacturing issues or reworks associated with this pump. there are no other complaints associated with pumps from this lot. based on the clinical account and the returned product, the sheath was most likely inserted at a steep angle, leading to the deformation of the guidewire repositioning unit observed and applying stress to the blood vessel.
37942 | 1220648-2018-00156 | 8096567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8096567&pc=PYX | 10/04/2018 | 11/05/2018 |  | 11/21/2018 |  |  | 11/05/2018 | 11/21/2018 |  |  | 04/30/2020 |  | 11/21/2018 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1341178 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | Yes | Material Deformation | Material Twisted/Bent | [5 empty] | Infiltration into Tissue | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old caucasian male presenting with right heart failure. an impella rp was inserted without issue. patient endured 4 sustained vt events, which were alleviated by way of patient's internal defibrillator. during procedure, a kink was discovered in the cannula and was described as "coiled." due to malformed cannula, surgical intervention was performed to remove the rp to avoid serious injury to patient during explant. a new rp pump was inserted and operated successfully. | [5 empty] | The impella rp was received and a failure investigation was conducted. kinks and general deformation was confirmed and support the reported allegation. a review of the data logs confirmed placement signal reduction. a review of the device history record found no manufacturing issues or reworks associated with this pump. there are no other complaints associated with pumps from this lot. the loss of position and cannula kinking on itself most likely occurred during the vt storms and defibrillation of the patient, but the definitive cause of the pump losing position and kinking is unknown.
37943 | 1220648-2018-00151 | 8088914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8088914&pc=OZD | 10/17/2018 | 10/22/2018 |  | 11/20/2018 |  |  | 10/22/2018 | 11/20/2018 |  |  | 07/31/2020 |  | 11/20/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA LP5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1359038 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 38 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting with myocarditis. impella lp5.0 was inserted in the right axillary site. a small incision was made in the left chest to expose the pulmonary artery, in order to perform a surgical procedure (blood removal from the right atrium and feeding blood to the pulmonary artery). the wound created by this surgical procedure produced a bleed that required surgical intervention to resolve. although the impella device was not the original cause of this bleed, it was reported that the device's use of heparin contributed to the bleed. after surgical intervention, patient's bleed subsided. | [5 empty] | The impella lp5.0 was received and evaluated. a failure analysis investigation was performed and no issues were found with the device. a review the device history record and found no manufacturing issues or reworks associated with this pump. there are no other complaints associated with pumps from this lot. data logs confirm the device operated within design specification during use. the root cause of the bleeding was most likely related to the patient condition and surgical procedures performed during the case and is unrelated to impella. it is likely that the heparin contributed to promoting bleeding after it initially occurred. it should be noted that thoracic drains are contraindicated for uses in which the blood's ability to clot is impaired.
37944 | 1220648-2018-00150 | 8092060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8092060&pc=OZD | 10/21/2018 | 10/22/2018 |  | 11/20/2018 |  |  | 10/22/2018 | 11/20/2018 |  |  | 07/31/2020 |  | 11/20/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1355322 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hematoma | Hemolysis | Blood Loss | [10 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported that following a pci procedure an impella lp2.5 was placed in a (b)(6) male patient. during the pump support hemolysis occurred. a ct revealed that the pump was located close to the aortic valve. the impella's position was adjusted under echo. the pump was removed following 2.50 hours of support. | [5 empty] | The impella cp was returned for evaluation. in addition, the automatic impella controller's data logs were also returned for analysis; however, the logs were found to be corrupted which disallowed performing an analysis of the logs. the 13fr introducer was not returned for evaluation. an inspection was performed on the returned impella cp. there were no damages found on the outflow cage or inflow cage. in addition, there was no damage to the impeller and the impeller was appropriately concentric to the pump housing. the impella was cleaned and tested for hemolysis, and passed all specifications. the clinical account received from the complainant reported that the repositioning sheath was fixed to the skin with sutures, but the clinical narrative suggests that the fixation may not have been adequate and may have allowed for movement of the device, as the pump was found to be out of position with the outlet cage too close to the aortic valve. patient hemolysis and then hematuria was reported, with a decrease in the patient's hemoglobin values from 13 to 7g/dl. the patient was reported to have been administered 8 units of replacement blood products. the returned pump passed hemolysis testing. based on this clinical account and the condition of the returned device, the hemolysis was most likely due to the improper positioning of the device. without the data logs, the improper positioning cannot be definitively verified. a review of the device history record revealed that there were no issues related to this failure mode found during the manufacture, inspection or testing with this pump. in addition, there have been no other complaints relevant to this failure mode for other pumps in this lot. | The impella lp2.5 and console data have been returned. the evaluation is currently in process. upon the conclusion of the evaluation the results will be reported on a supplemental mdr.
37945 | 1220648-2018-00152 | 8087151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8087151&pc=OZD | 08/09/2018 | 11/15/2018 |  | 11/19/2018 |  |  | 11/15/2018 | 11/19/2018 |  |  | 04/30/2020 |  | 11/19/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA LP5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1342588 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2018 an impella lp5.0 was placed in a (b)(6) male patient. following placement slight bleeding occurred at the access site. a sand bag was applied, and the heparin was reduced for 3 hours in an effort to stem the bleeding. the following day the patient was brought to the operating room due to a hematoma that developed at the access site. in addition, serious drainage (bleeding) occurred. the patient was given 5 units of fresh frozen plasma (ffp), 2 units of platelets, and the patient also received a third unit of packed red blood cells. the patient was reported to have been successfully supported for 76-65 hours. the device was removed on (b)(6) 2018 and the patient was transferred to another facility were another impella lp5.0 was placed in the patient. there were no other adverse patient effects and the patient was in stable condition upon the removal of the first impella lp5.0. it is unclear whether these blood products were given because of the bleeding at the access site, the hematoma, or a pre-existing patient condition. | [5 empty] | The previous unique identifier (udi) # provided was inaccurate. updated unique identifier (udi) # from (b)(4) to (b)(4). | The impella lp5.0 was not returned for evaluation, as it was discarded following patient use and removal. in conclusion, a root cause was unable to be determined because product and data were not returned. there are no other complaints for pumps in this lot.
37946 | 1220648-2018-00137 | 8081802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8081802&pc=OZD | 10/19/2018 | 10/19/2018 |  | 11/16/2018 |  |  | 10/19/2018 | 11/16/2018 |  |  | 07/31/2020 |  | 11/16/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1361211 |  | 0048-0032 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 81 YR | [5 empty] | Malfunction | No | Yes | Detachment of Device or Device Component | Infusion or Flow Problem | No Apparent Adverse Event |  |  |  |  | Missing Value Reason | [41 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that the clinicians placed an impella cp in an (b)(6) male patient. immediately following placement low flow and suction alarms occurred upon activation. the pump was explanted, replaced with a new cp pump which successfully supported the patient. the patient was reported to be in stable condition, with no adverse patient effects as a result of the reported issue. | [5 empty] | The impella cp was returned for evaluation. the cause of the low flow was found to be a fracture of an impeller blade. the impeller blades are capable of contacting a hard material ingested through the inlet cage or through the exit of the outlet cage. hard material such as calcium is capable of fracturing an impeller blade. the impeller blade geometry is critical to achieving appropriate flow rates. during manufacturing testing, the pump ran at p9 with a motor current of 852 ma, which is a higher current than when the pump ran at p9 in the field with a motor current of 784 ma indicating that the impeller blade fracture most likely occurred in the field. the pump was re-run in the lab at p9 and experienced motor currents lower than in the field at 729 ma. the low flows and suction alarms experienced in the field at setting p9 were reproduced on the returned product in the lab. the pump never achieved appropriate flow rates in the field indicating that a crack or fracture of the impeller blade occurred at the start of use. a hard material may have contacted the rotating impeller blade during activation. in conclusion the root cause of the low flow was a fracture of an impeller blade which was unable to be determined. the pump may have ingested a hard material that came in contact with the blade, but no hard material was found upon inspection.
37947 | 1220648-2018-00147 | 8075282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8075282&pc=OZD | 07/11/2018 | 10/19/2018 |  | 11/15/2018 |  |  | 10/19/2018 | 11/15/2018 |  |  | 01/31/2020 |  | 11/15/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP. |  |  | IMPELLA :LP5.0 |  |  | OZD |  |  | IMPELLA LP5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1331370 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [5 empty] | ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION) | [22 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The impella lp5.0 was placed in the patient on (b)(6) 2018. an ecmo was also placed in the patient, and the impella and an ecmo were used in combination to treat the patient's low cardiac output. the patient was diagnosed with disseminated intravascular coagulation (dic), and it was reported that the patient's condition was complicated by a deterioration of his general condition. the patient did require a mass blood transfusion in response to hemorrhage and hemolysis as a result of the dic. during patient support the dic improved and the patient's hemodynamics became stable. the impella lp5.0 was removed from the patient on (b)(6) 2018 after 220 hours of successful patient support. | [5 empty] | An evaluation of the device could not be performed as the impella lp5.0 was not returned to the manufacture as it was discarded following use. a review of the device history record did not reveal any issues with the pump during manufacture. there were no other complaints reported for other pumps in this lot. the root cause of the blood loss was unable to be determined as the product was not returned and further clinical information was not provided. there was no malfunction of the device that we are aware may have occurred.
37948 | 1220648-2018-00148 | 8075700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8075700&pc=OZD | 09/27/2018 | 10/19/2018 |  | 11/15/2018 |  |  | 10/19/2018 | 11/15/2018 |  |  | 07/13/2020 |  | 11/15/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1355322 |  | 0005048-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hematoma | Mitral Regurgitation | Blood Loss | [10 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2018 an impella lp2.5 was placed in an (b)(6) year old male patient in hemodynamic failure. some difficulty was encountered during placement as a result of the patient's abdominal vascular tortuosity. following pump placement a hemorrhage from the impella puncture site occurred, causing a hematoma to develop. the physician performed manual compression hemostasis to stop the bleeding and the patient was administered replacement blood products (amount unknown). upon removal of the pump on (b)(6) 2018 and visualizing the patient it was suspected that chordendolysis had occurred during the insertion of the impella lp2.5 and there was the appearance of mitral regurgitation. there was no indication of any required performed treatment for the suspected chordendolysis and the appearance of the mitral regurgitation. it was reported that the patient was successfully supported for 101.60 hours, and was reported to be in stable condition when the device was removed. | [5 empty] | The impella lp2.5 and repositioning sheath were returned for evaluation. the automatic impella console data logs ere also returned for analysis. the data logs were found to have been corrupted, preventing an analysis of the logs. during the evaluation the pump and repositioning sheath, the wound portion of the sheath was found to have been extensively damaged. no damage or deformities were found on the pump. as reported by the complainant the patient's chordae tendineae was mechanically broken during the impella lp2.5 insertion, resulting in mitral regurgitation. the cause of the chordae tendineae breaking and subsequent mitral regurgitation was most likely the result of a difficult insertion. the impella may have interacted with the patient's mitral chordae during the difficult insertion, as the pigtail was noted to be turned towards the back wall of the ventricle. this cannot be confirmed without log data or imagery. without the data and/or imagery from the event a definitive root cause of the mitral chordae breaking or the subsequent mitral regurgitation cannot be definitively determined. the impella lp2.5 was also evaluated for the reported access site hematoma that occurred. as detailed above in this evaluation report the wound closure portion of the repositioning sheath had been extensively damaged. the identified damage is consistent with interaction with calcifications. according to the clinical team the repositioning sheath was not damaged on insertion, but had been damaged upon explant of the pump. the root cause of the hematoma may have been related to the extended time taken for insertion, although this cannot be definitively determined. a review of the device history record revealed that there were no issues related to these failure modes found during the manufacture, inspection or testing with this pump. in addition, there have been no other complaints relevant to the failure modes for other pumps in this lot. | The device has been returned to the manufacturer; however the evaluation has not yet been completed. upon the completion of the evaluation of the impella cp a supplemental mdr will be submitted.
37949 | 1220648-2018-00134 | 8070716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8070716&pc=OZD | 10/18/2018 | 10/19/2018 |  | 11/14/2018 |  |  | 10/18/2018 | 11/14/2018 |  |  | 08/31/2020 |  | 11/14/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1363818 |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | Yes | Insufficient Information | [6 empty] | Hemolysis | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2018 a (b)(6) male patient presented at the facility's cath lab with a stemi, and in respiratory distress. testing in the cath lab revealed severe multi-vessel disease. an impella cp was placed in the patient; the physician then placed 7 drug eluding stents in the patient with great angiographic results. following this procedure the 14fr sheath was removed and the repositioning sheath was inserted. shortly thereafter the patient developed a large hematoma of the groin. manual pressure was held, and then a fem stop was placed. the patient was administered packed replacement blood cells in response to the loss of blood experienced by the patient. the impella cp was removed from the patient on (b)(6) following 44.73 hours of successful hemodynamic support. the patient was reported to be in stable condition and doing well following the removal of the pump. | [5 empty] | The device was nor returned for evaluation as it was discarded by the user following patient support. there is no known malfunction that we are aware of. if the manufacturer becomes aware of any additional information, a follow-up will be submitted.
37950 | 1220648-2018-00135 | 8070721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8070721&pc=OZD | 10/19/2018 | 10/19/2018 |  | 11/14/2018 |  |  | 10/19/2018 | 11/14/2018 |  |  | 08/31/2020 |  | 11/14/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1360977 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Ischemia | Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that a (b)(6) year old male patient presented at the hospital on (b)(6) 2018. the patient was found to have severe multi-vessel disease of left anterior descending artery, mid-circumflex, obtuse marginal, and chronic total occlusion of the right coronary artery. an impella cp was placed in the patient's right femoral artery to stabilize the patient. the 14fr peel-away sheath was left in place in the patient's anatomy. after 44.75 hours of patient support it was decided to remove the impella cp and escalate to an impella 5.0, but as the clinicians went to remove the impella cp they noticed that the patient had no pedal pulses. vascular surgery removed the pump at which point they noticed a clot. vascular surgery repaired the vessel in the leg, at which time the surgeon remarked that the patient had advanced peripheral artery disease. following the surgical repair of the vessel perfusion to the distal leg improved. | [5 empty] | The impella cp was returned for analysis; however it was reported that the 14fr peel-away sheath had been discarded following removal from the patient. the inspection of the device revealed a clot on the impeller. the evaluation of the returned impella cp found that the device was within specification. a review of the device history record did not reveal any issues with the pump during manufacture. the 14fr introducer from this lot passed all incoming inspections. there have been no other complaints associated with introducer kits from this lot. the clinical account of this event stated that the patient had severely diseased vasculature, including that of the peripheral arteries. the clinical account also stated that the 14fr peel away introducer was never removed from the access site, and remained in place for the duration of the case. the diseased vasculature was later repaired via vascular surgery, after which perfusion to the distal leg improved. this indicates that the ischemia was most likely due to the condition of the patient's vessels. the pump was found to be within specification and most likely did not contribute to the condition. in conclusion, the root cause of the leg ischemia was most likely the patient's diseased vasculature. the root cause was unable to be definitively determined because the 14fr introducer was not returned for evaluation.
37951 | 1220648-2018-00142 | 8046312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8046312&pc=OZD | 09/07/2018 | 10/10/2018 |  | 11/07/2018 |  |  | 10/10/2018 | 11/07/2018 |  |  | 04/30/2020 |  | 11/07/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1341340 |  | 0048-0003 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombus | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) caucasian male presenting with nstemi. an impella cp was inserted into the left femoral artery and no issues were noted during use. approximately 1 hour post explant, patient's leg was found cold. a blood clot was found at the insertion site. as treatment, clot was surgically removed and pulse to right leg was noted as "improved" the following day. | [5 empty] | The impella cp was discarded by the customer, therefore, unable to be returned for investigation. a failure analysis investigation was performed to review the device history record and found no manufacturing issues or reworks associated with this pump. there are no other complaints associated with pumps from this lot. all introducers in this lot passed all inspection.
37952 | 1220648-2018-00143 | 8046391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8046391&pc=OZD | 09/29/2018 | 10/11/2018 |  | 11/07/2018 |  |  | 10/11/2018 | 11/07/2018 |  |  | 05/31/2019 |  | 11/07/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1348042 |  | 0048-0014 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | CONSUMER | FOREIGN | HEALTH PROFE |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Life Threatening | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male caucasian presenting with acute myocardial infarction and cardiogenic shock. patient's initial therapy was received via intra-aortic balloon pump and escalated to the impella cp pump via left femoral artery. during treatment, patient's leg became "cooler" and it was suspected that a thrombotic event was occurring and had led to closure of femoral artery. pump was removed, however, it was reported that patient's leg required amputation. patient's was placed on alternate therapy via ecmo and patient was noted in stable condition. | [5 empty] | The impella cp was discarded by the customer, therefore, the root cause of the limb ischemia and amputation was unable to be determined. a review the device history record found no manufacturing issues or reworks associated with this pump. there are no other complaints associated with pumps from this lot. all introducers in this lot passed all inspection. no corrective action is recommended at this time as the root cause was unable to be determined and the failure mode was determined to have a low risk index. failures of this type will continue to be monitored for trends.
37953 | MW5081158 | 8050941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8050941&pc=OZD | 10/29/2018 | 10/30/2018 |  | 11/07/2018 |  |  |  | 11/07/2018 |  |  | 08/30/2020 |  | 11/07/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | CP 00480003 |  |  | ABIOMED INC. |  |  |  |  | 172712 |  |  |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 72 YR |  |  | 84 |  |  | Injury | No | Yes | Obstruction of Flow | Difficult to Advance | [5 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Stemi pt taken to cath lab for ptca. md placed impella cp to as. difficulty crossing av and device quit working. placed a second device which functioned as expected. (b)(6) 2018 pci of rca: initial pda revealed flow and then placed guide pda and plb were occluded. opened pda and did ptca.
37954 | 1220648-2018-00128 | 8040306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8040306&pc=OZD | 10/07/2018 | 10/07/2018 |  | 11/05/2018 |  |  | 10/07/2018 | 11/05/2018 |  |  | 06/30/2020 |  | 11/05/2018 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1356949 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Death | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | Complainant reported that a (b)(6) year old male patient presented with a stemi on (b)(6) 2018. the patient was hypotensive and examination in the cath lab revealed that the patient had a total occlusion of the proximal left anterior descending artery (plad). the impella cp was placed in the patient on (b)(6) 2018. a lucas device was also started, in addition a tandem heart with generator was placed. the patient was successfully supported with the impella cp for 8.9 days. the physician reported that there was bleeding from the impella insertion site. the patient was administered at least 4 units of blood before the pump was removing from the patient. the physician also reported that the patient was supported by multiple devices during this time. it was also reported that vascular surgery repaired both vessels in the leg in order to control bleeding. | [5 empty] | The facility did not return the impella cp pump for investigation; therefore, the root cause of the access site bleeding was unable to be determined.
37955 | 1220648-2018-00141 | 8028505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8028505&pc=OZD |  |  |  | 11/01/2018 |  |  | 10/09/2018 | 11/01/2018 |  |  | 07/31/2020 |  | 11/01/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1355322 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male presenting with acute myocardial infarction and cardiogenic shock. impella lp2.5 pump was inserted on (b)(6) 2018 without any reported issues, however, on (b)(6) 2018, ischemia developed on the insertion side of the patient's leg. as treatment, pump was removed and condition improved. patient remains in stable condition. | [5 empty] | The impella lp2.5 was discarded by the customer; however, a device history review was performed. no manufacturing failures applicable to the failure mode for this pump. there were no other complaints relevant to the failure mode in this pump lot. there were no other complaints in either the introducer lot or the repositioning unit lot related to this failure mode. the introducer lot passed all incoming inspection. the failure mode was determined to have a low risk index based on moderate severity and very low occurrence (usage: 6227; occurrence: 5; rate: 0.08%). no further corrective actions are being taken at this time; however, the failure mode will be monitored and trended.
37956 | 1220648-2018-00131 | 8019281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8019281&pc=PYX | 10/02/2018 | 10/02/2018 |  | 10/30/2018 |  |  | 10/02/2018 | 10/30/2018 |  |  | 07/31/2019 |  | 10/30/2018 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1301319 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Labelling, Instructions for Use or Training Problem | Use of Device Problem | Insufficient Information |  |  |  |  | Cardiopulmonary Arrest | No Code Available | [11 empty] | Hospitalization | Life Threatening | Required Intervention | [26 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A (b)(6) patient in cardiogenic shock and an acute myocardial infarction, was brought to the cardiac cath lab for placement of the impella 2.5, without a representative from abiomed on site. some time later, the physician removed the impella as it had come out of the left ventricle due to an unknown cause. when the physician placed a second impella, the pump placed to the left ventricle was not another 2.5, a left sided pump, but rather the right heart sided, rp device. the team noticed this after the pump was running, removed the pump, and began cpr. the patient did expire despite efforts. | [5 empty] | The product and data logs have been returned for analysis. the investigation is on-going at this time. a supplemental will be drafted upon completion of the investigation. | Since the original filing the investigation has completed. mistaken for a left ventricular assist device, the impella rp was placed in the patient's left ventricle and turned on. the emergent nature of the case and inexperienced staff may have contributed to the misuse. the impella rp was appropriately labeled and programmed as an impella rp. the impella rp is not indicated for left ventricular support. the clinical staff was retrained on setup and insertion on the impella 2.5, cp, 5.0 and rp on 2018-10-20. this failure mode will continue to be monitored for trends.
37957 | 1220648-2018-00139 | 8011337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8011337&pc=OZD | 09/28/2018 | 11/28/2018 |  | 10/28/2018 |  |  | 09/28/2018 | 10/28/2018 |  |  | 06/30/2020 |  | 10/28/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1350196 |  | 0048-0020 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 45 YR | [5 empty] | Injury | Yes | Yes | Device Dislodged or Dislocated | Unintended Movement | [5 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) african american male patient presenting with cardiomyopathy and cardiogenic shock. a decision was made to place the impella cp pump for post ecmo decannulation in the right axillary site. insertion was completed without incident. patient was administered a procedure in a different area of the hospital, and when returned to the complainant, the device's motor current was flat and positioning indicated as 10 cm back from previous spot. upon taking patient the cardiac cath lab for replacement of device, patient coded. the pump was found partially dislodged from the insertion site. the outflow cage was outside of the body while the inflow cage was inside the patient. patient endured blood loss, blood products were administered, and hemostasis was achieved. a new impella cp was placed without incident. | [5 empty] | The device has not yet been received from the customer, however, a failure analysis investigation is underway and supplemental medwatch report will be submitted upon completion of the evaluation. | Revised d(2) and g(3) to align to correct manufacturing site. neither device nor data logs were returned for evaluation. the device history record was reviewed and found no manufacturing issues or reworks associated with this pump. there are no other complaints associated with pumps from this lot. all introducers in this lot passed all inspection. although device nor data logs were received, the root cause of the positioning issue was attributed to patient movement. the pump was allowed to change positions because of the removal of, and non-replacement of the statlock by a member of the radiology team post procedure. the cause of the bleeding was the pump being extremely out of position, across the hemostatic valve of the introducer.
37958 | MW5080855 | 8009264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=8009264&pc=PBL | 10/23/2018 | 10/24/2018 |  | 10/25/2018 |  |  |  | 10/25/2018 |  |  |  |  | 10/25/2018 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP CATHETER |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  |  |  | 01008135022010879101356397 |  |  |  | Yes | [5 empty] | Invalid/Incomplete | No Information |  |  | Voluntary report | [8 empty] | 80 YR | [6 empty] | No | No | Defective Component | Device Displays Incorrect Message | [5 empty] | No Consequences Or Impact To Patient | [41 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Motor failure on impella device, device alarmed. impella cath removed due to cath malfunction. balloon pump was put back in. no harm to pt but the balloon had to be reinserted due to equipment malfunction.
37959 | 1220648-2018-00138 | 7999897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7999897&pc=OZD | 09/26/2018 | 09/27/2018 |  | 10/24/2018 |  |  | 09/27/2018 | 10/24/2018 |  |  | 05/31/2020 |  | 10/24/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1342405 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 43 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that the impella lp 2.5 was successfully placed in a (b)(6) male patient on (b)(6) 2018. the patient was successfully supported by the impella lp2.5 pump for 4.7 days. upon removal of the pump on (b)(6) 2018 a thrombus was found on the pump's pigtail. it was also found that the patient had developed limb ischemia. a percutaneous transluminal angioplasty was performed, which successfully resolved the reported ischemia. the physician was unable to determine the cause of the ischemia, and if it was the result of the stenosis of the limb vessel from the patient's dm (diabetes mellitus) or thrombosis occurring as a result of the impella support. the physician further reported that the patient may have been prone to thrombosis due to the dm. the patient support was reported to have been successful, and the patient was in stable condition. | [5 empty] | The device was not returned to the manufacturer for evaluation as it was discarded by the user following use. additional information regarding this event was obtained that reported that this patient had a percutaneous transluminal angioplasty (pta) and arterial switch operation (aso) procedures prior to impella placement which might result in stenosis of the limb vessel. it was unknown if the limb ischemia was caused by stenosis of the limb vessel or thrombus. the automatic impella controller data logs were returned for analysis. the analysis revealed that the pump functioned appropriately during the entire case. a review of the device history record was performed that revealed that there were no issue related to this failure mode for any pump in this lot. a lot review revealed that there were no complaint for other pumps in this lot relevant to this failure mode. in conclusion, this patient had ventricular fibrillation and was on ecmo which are risks for thrombus formation. although the thrombus was found on the pigtail upon explant, it is likely that it was from the patient condition and not caused by the impella. there was no evidence the impella can create thrombus. this patient also underwent the percutaneous transluminal angioplasty (pta) and arterial switch operation procedures prior to impella placement which may result in stenosis of the limb vessel and cause ischemia. as the device was not returned for review, the root cause of ischemia was unable to be definitively determined.
37960 | 1220648-2018-00119 | 7992628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7992628&pc=OZD | 10/01/2018 | 10/01/2018 |  | 10/22/2018 |  |  | 10/01/2018 | 10/22/2018 |  |  | 05/31/2020 |  | 10/22/2018 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1342102 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 37 YR | [5 empty] | Injury | Yes | No | Off-Label Use | [6 empty] | Inflammation | Regurgitation, Valvular | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | (b)(6) year old patient in (b)(6) was admitted in cardiomyopathy and had an impella 5.0 placed via the axillary artery. in addition, the patient was supported by ecmo. after over 2 months of successful support, the pump was removed and replaced. this duration was beyond the indicated use per usa ifu. the team replaced the pump with another impella 5.0 for 16 more days of successful support. the team then explanted the impella to transition the patient to the heartmate ii device. while in the surgical suite for the heart mate ii placement, the team observed damage to the valve. there was aortic regurgitation that was remedied with the surgical patch. | [5 empty] | The root cause of the damage to the aortic valve was unable to be determined because the product and data logs were not returned for review. the product was used beyond the indicated duration of use. the failure mode will be monitored and trended.
37961 | 1220648-2018-00117 | 7985039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7985039&pc=OZD | 09/20/2018 | 09/24/2018 |  | 10/19/2018 |  |  | 09/24/2018 | 10/19/2018 |  |  | 03/31/2020 |  | 10/19/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1334259 |  | 0048-0003 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hematoma | Thrombosis | Blood Loss | [10 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient was admitted with chest pain and sent to the cardiac catheterization lab for evaluation. the team placed an impella cp for support via the right femoral artery while cardiothoracic surgery was consulted. the patient was transferred on the next day to the tertiary medical center for continued care and monitoring. the patient received a coronary intervention while on support with the impella and subsequently the patient weaning from the impella began. the patient had heparin removed from systemic infusion in preparation for pump removal. the pump remained on overnight during this time of no systemic anticoagulation, the patient developed limb ischemia. the medical team then infused anticoagulation medication to address the clotting and associated limb ischemia. at this time the patient developed a bleed at the access site and a hematoma developed. the team applied manual pressure to the site and infused blood products back to the patient. | [5 empty] | The patient developed leg ischemia after the impella pump was removed. according to the clinical account, heparin was removed from the patient in anticipation of removal, but the pump remained in place overnight. it is unknown whether the act was maintained as recommended after systemic anticoagulation was discontinued. the pump was unable to be evaluated as it was not returned. the lack of systemic anticoagulation for the final 23 hours of support was most likely the cause of the clot development. the root cause of the access site bleed was most likely the anticoagulation administered to address the ischemia. the failure mode will be monitored and trended. no corrective action is warranted at this time.
37962 | 1035166-2018-00099 | 7985526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7985526&pc=DYB | 08/03/2018 | 02/05/2019 |  | 10/19/2018 |  |  | 10/12/2018 | 10/19/2018 |  |  | 01/01/2021 |  | 10/19/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 23F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-13864 |  | 0052-3021 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 74 YR | [5 empty] | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | Death | [18 empty] | IMPELLA RP | [22 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | As per additional information, there was difficulty getting access to the anatomy when placing impella and for that reason multiple insertions were made with various wires. | The patient was presented with a right heart failure post pulmonary embolism. during the placement of the impella rp, the tip split during repeated insertion attempts. the introducer was not in use when the patient expired. the patient expired due to family withdraw of care, the patient had a dnr (do not resuscitate). the customer reported the device will be returned for analysis. |  |  |  |  | Device used in treatment. one 23 fr abiomed introducer sheath was received in a used condition without the dilator. no other accessories were received with this device. blood was found on the exterior and interior of the sheath and sideport tubing. upon evaluation of the returned product, it was found that there were splits in the tip of the sheath and the distal tip circumference was dented inward and damaged. as stated in the event description, there was difficulty getting access to the anatomy with the impella rp, repeated attempts were made to place the impella and multiple insertions were made with various wires. this could have been a contributing factor to the damage done to the distal tip of the sheath. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. returned device analysis reveals one 23f abiomed introducer sheath that is within manufacturing specifications. no manufacturing defects were found. potential contributing factors to the damage is if the device was advanced through an area of resistance as stated in the precautions of the ifu, or if the device was subjected to unusual stresses as stated in the handling and storage section of the ifu. as per procedure abiomed introducer sheath in-process and final inspection, the devices in this lot were inspected as follows: visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. insert a tipped dilator of appropriate size through the tipped sheath and check if dilator inserts without excessive force or obstruction. with the naked eye at a distance of 12" to 18", inspect where the sheath and dilator meet and check the gap is minimal and a smooth transition is present. the instructions for use ifu instructs the end user: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available as the investigation is on-going. a follow-up will be submitted as soon as the investigation is complete.
37963 | 1220648-2018-00132 | 7986946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7986946&pc=OZD | 09/29/2018 | 11/30/2018 |  | 10/19/2018 |  |  | 09/29/2108 | 10/19/2018 |  |  | 02/29/2020 |  | 10/19/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1336295 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | CONSUM |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | Yes | Material Separation | [6 empty] | No Consequences Or Impact To Patient | Blood Loss | Device Embedded In Tissue or Plaque | [10 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | Patient endured blood loss from impella insertion site and required 2 units of blood products administered. | The complainant reported a (b)(6) year old caucasian male patient presenting with acute myocardial infarction with cardiogenic shock. cp pump was placed successfully in left femoral artery. when explanting, device separated at the outflow cage when attempting to pull device pack into the aorta. a snare was inserted in the right femoral artery and separated component of device was successfully removed. there was no adverse effects endured by patient, as a result. |  |  |  |  | The impella cp, pump cable, and purge cassette were returned for investigation. a failure analysis investigation was performed and found outer diameter of the pump stretched, however, within specification. stretched out diameter of pump and cannula detachment was deemed most likely due to excessive force required to remove catheter. bleeding was most likely due to increased levels of anti-coagulation, as bleeding subsided after heparin concentration was decreased. a review the device history record and found no manufacturing issues or reworks associated with this pump. there are no other complaints associated with pumps from this lot. all introducers in this lot passed all inspection. | The impella cp was not returned by the customer, however, a failure analysis investigation is underway and supplemental medwatch report will be submitted upon completion of the evaluation.
37964 | 1220648-2018-00127 | 7974049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7974049&pc=OZD | 09/26/2018 | 11/30/2018 |  | 10/17/2018 |  |  | 09/27/2018 | 10/17/2018 |  |  | 01/31/2020 |  | 10/17/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA LP5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1324277 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 49 YR | [5 empty] | Injury | Yes | No | Off-Label Use | [6 empty] | Death | [12 empty] | Death | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) male patient presenting with heart failure due to hypertrophic cardiomyopathy. impella lp 5.0 pump was inserted without issue via right axillary. the pump remained implanted for a period of 994.33 hours until it stopped. it should be known that patient was not deemed a heart transplant candidate and there was "no end strategy." it should also be known that physicians and family were aware that when/if pump had stopped, patient would expire. | [5 empty] | The impella lp 5.0 has been returned for evaluation. the evaluation is underway and a supplemental mdr will be submitted upon completion of the evaluation. instructions for use state impella 5.0 "in conjunction with the automated impella® controller, are temporary ventricular support devices intended for short term use" and are designed for usage time of "< 6 days." | The impella lp 5.0 was returned and a failure analysis investigation was completed. data logs confirm a pump stop had occurred beyond design specification. physical examination and simulated use testing found the root cause of this failure related to coil delamination. a review the device history record found no manufacturing issues or reworks associated with this pump. there are no other complaints associated with pumps from this lot.
37965 | 1220648-2018-00130 | 7969496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7969496&pc=OZD | 09/21/2018 | 09/21/2018 |  | 10/16/2018 |  |  |  | 10/16/2018 |  |  | 06/30/2020 |  | 10/16/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA LP5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1353492 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 36 YR | [5 empty] | Injury | No | Yes | Difficult to Insert | Difficult to Remove | [5 empty] | No Consequences Or Impact To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) year old female patient presenting with post-partum cardiomyopathy. an impella lp5.0 pump was inserted into the right axillary via a hemashield platinum graft that was sewn on the artery. patient required right side support, therefore, an rp pump was placed and patient was noted as "doing much better." planned removal of rp was conducted without issue, however, lp5.0 required surgical intervention for removal. removal was unable to be conducted, as intended, due to patient's physical anatomy. subsequently, physician intentionally arrested the heart, cut open aorta, cut impella pump wire above motor housing, and pulled through the aortic incision to remove. there was no noticeable damage to the pump, other than what had been intentionally cut. lvad was successfully implanted and patient has since been discharged home. it should be noted that patient anatomy also made pump advancement difficult, upon implant and required use of a snare to successfully place device. no adverse effects were endured by patient, as a result of the event reported. | [5 empty] | The impella lp5.0 was discarded by the customer, however, a failure analysis investigation is underway and supplemental medwatch report will be submitted upon completion of the evaluation. | The impella lp5.0 was discarded by the customer, therefore, unable to be returned for investigation. a failure analysis investigation was performed to review the device history record and found no manufacturing issues or reworks associated with this pump. there are no other complaints associated with pumps from this lot. the root cause of inability to explant the impella 5.0 was most likely due to patient's anatomy. the root cause cannot be definitively determined as the product was not returned for the failure investigation.
37966 | 1035166-2018-00097 | 7960202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7960202&pc=DYB | 09/20/2018 | 11/30/2018 |  | 10/12/2018 |  |  | 09/21/2018 | 10/12/2018 |  |  | 05/01/2021 |  | 10/12/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA 23F |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC |  |  | US |  | C1-14283 |  | 0052-3006 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 73 YR | [5 empty] | Injury | Yes | No | Leak/Splash | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that during insertion at the axillary artery for placement of the impella 5.0, the valve of the axillary sheath had a slit that resulted in blood loss estimated at 500-1000 cc. as per additional information no surgical technique required to control blood loss. surgeon applied jaw clamp at graft to stop bleeding. patient is stable and still on support with impella pump. no additional information is available and no other adverse event was reported. | [5 empty] | Our investigation is still in progress, follow up report will be submitted if we found any further additional information. | Device used in treatment. one 23f abiomed axillary introducer sheath was returned from the customer. there were no other accessories. blood was found on and inside the sheath. the sheath was returned already peeled in half. the hemostatic valve was damaged (torn). the rest of the sheath looked normal. no manufacturing defects were found. upon evaluation of the returned product, it was found that the hemostatic valve was torn when viewed under a 10x microscope. the sheath was already torn in half when received back from the customer, so a leak test could not be performed. returned device analysis reveals one 23f abiomed axillary introducer sheath that is within manufacturing specifications. according to the complaint report, the patient experienced bleeding from the valve of the sheath. the sheath was found to have a torn hemostatic valve which would account for the blood loss. the valve may have been torn as a result of off center insertion due to use error/mishandling in the field. per qa procedure vacuum leak test - perform pressure and vacuum leak test per procedure. this procedure is performed by qa 100%. per instructions for use ifu: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
37967 | 7960316 | 7960316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7960316&pc=PBL | 05/04/2018 | 09/28/2018 | 09/28/2018 | 10/12/2018 | 09/28/2018 | 10/12/2018 |  | 10/12/2018 |  |  |  |  | 10/12/2018 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 21535 DA | [5 empty] | Malfunction | No | Yes | Fluid/Blood Leak | Dull, Blunt | Material Twisted/Bent |  |  |  |  | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Just before 0400, the rn and pca had completed a patient bath post bowel movement. while cleaning the patient, the patient was moving her legs and squirming but the staff was able to keep her on her side for a sheet change. there were no observed incidents involving the impella catheter at this time. patient was returned supine with pillow under right side. the rn began a sterile dressing change at about 0405. while the rn was taking down the old dressing, the impella controller alarmed "air in line". the rn observed a kink in the tubing and the alarm was resolved. at approximately 0409, the controller alarmed "purge system open". at this time, the rn assessed the tubing again and noticed purge fluid leaking from the translucent yellow connection which was connected by the blue disc of the glucose filter module. the rn manager was called to the bedside and the issue was assessed. noticed that the male portion near the translucent yellow connection appeared snapped and a piece of the clear connector was lodged in the female portion near the blue disc. abiomed support line called and spoke with rep who instructed us to cut off the glucose filter and connect the dark yellow to a blunt tip 18-20 gauge needle. we cut the one inch line between the 0.2 micron filter and inserted the needle into the remaining portion of the one inch tubing. impella flow, purge flow, and purge pressure returned to the pre-incident level. per abiomed rep, he will report it to the team and they will follow up today.
37968 | 1220648-2018-00124 | 7957894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7957894&pc=OZD | 09/17/2018 | 11/23/2018 |  | 10/11/2018 |  |  | 09/17/2018 | 10/11/2018 |  |  | 03/31/2020 |  | 10/11/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1353646 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 45 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Injury | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The eu complainant reported a (b)(6) female presenting with cardiogenic shock after enduring a heart stop at home, during the night. patient was resuscitated and impella cp was implanted. patient was admitted to the intensive care unit where suction alarms were emitted. procedure was successful; however, a tear in the arterial wall was identified by surgeon, upon closure. surgical repair of artery was performed and patient was noted to be in stable condition. | [5 empty] | The impella cp has been returned for evaluation. the evaluation is underway and supplemental medwatch report will be submitted upon completion of the evaluation. | The impella cp was returned by the customer and a failure analysis investigation was completed. we are unable to confirm the cause of the reported vascular damage, as no abnormalities with the device that would have caused this were identified. a review the device history record and found no manufacturing issues or reworks associated with this pump. there are no other complaints associated with pumps from this lot. all introducers in this lot passed all inspection.
37969 | 1220648-2018-00126 | 7951907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7951907&pc=OZD | 09/16/2018 | 09/16/2018 |  | 10/10/2018 |  |  | 09/16/2018 | 10/10/2018 |  |  | 06/30/2020 |  | 10/10/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1355357 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | CONSUMER | HEALTH PROFESSIONAL |  |  |  |  | 64 MO | [5 empty] | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2018 the clinicians placed an impella cp in a (b)(6) male patient. the impella placed without issue in the left femoral artery. the impella was assessed by echo and found to be in good position. it was reported that following pump placement the patient was given a rapid infusion of 5 liters of intravenous fluids. the patient moved to the ctu where he received another 2 liters of intravenous fluids and 6 units of packed red blood cells. the patient was successfully supported for 10.22 hours. on (b)(6) 2018 during the removal of the impella cp the physician experienced difficulty gaining access to the femoral artery. a vascular physician was called. the patient was found to have developed a large hematoma and required more blood products and fluid to maintain arterial pressure. upon obtaining access it was observed that the repositioning sheath was not in the femoral artery. it was determined by the vascular surgeon and ic fellow, that as the patient was given double digits of fluids, the distance from the artery to the skin surface grew greater and the sheath slid out of the artery. this resulted in bleeding that continued to worsen as they gave fluids. the ic fellow did remark that no hematoma was noted until well after a large amount of fluids had been administered to the patient. | [5 empty] | The impella cp has not yet been returned for evaluation, preventing any evaluation of the suspect this device at this time. in conclusion, a root cause was not "yet been" able to be determined because product was not returned.
37970 | 1220648-2018-00125 | 7951909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7951909&pc=OZD | 09/23/2018 | 09/23/2018 |  | 10/10/2018 |  |  | 09/23/2018 | 10/10/2018 |  |  | 06/30/2020 |  | 10/10/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA LP5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1354183 |  | 005062 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | [6 empty] | Rupture | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that a (b)(6) year old male patient presented with ischemic cardiomyopathy and right ventricle failure. an impella lp5.0 pump was placed through a previously used graft used. pump caused "gross hemolysis" during use. upon explant the mitral valve ruptured. as treatment, the valve was repaired; however, the repair delayed the placement of a planned heartware hvad system. | [5 empty] | The device has not yet been returned for evaluation. upon receipt, and completion of device evaluation a supplemental mdr will be submitted.
37971 | 1035166-2018-00095 | 7948826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7948826&pc=DYB | 09/14/2018 | 10/09/2018 |  | 10/09/2018 |  |  | 09/20/2018 | 10/09/2018 |  |  |  |  | 10/09/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB | [5 empty] | OSCOR INC. |  |  | US | [5 empty] | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [11 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that introducer access site had blood loss and require 1 unit of blood. patient presented to emergency room with congestive heart failure exacerbation. subsequently patient went into pulses electrical activity arrest and was successfully resuscitated. patient was in icu and continued to decline hemodynamically. decision was made to implant impella for heart failure. patient outcome was stable and no issue identified with device. no additional information is available and no other adverse event was reported. | [5 empty] | The device was not returned for analysis. a review of the device history record could not be performed as the lot number is unknown. there was no specific performance related failure was reported by the user, therefore an investigation could not be conducted. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. the device passed all in-process and qa final inspection steps before shipping to the customer.
37972 | 1220648-2018-00123 | 7949824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7949824&pc=OZD | 09/14/2018 | 12/26/2018 |  | 10/09/2018 |  |  | 09/14/2018 | 10/09/2018 |  |  |  |  | 10/09/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, OZD |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 80 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported an (b)(6) male patient presenting with congestive heart failure and cardiomyopathy. patient was successfully resuscitated and admitted to the intensive care unit, where and impella cp pump was placed. while accessing the artery with a 14f x13 cm oscor introducer, the patient endured blood loss. as treatment, 1 unit of blood was delivered. pump was successfully inserted and patient's heart was said to have "greatly improved." | [5 empty] | The impella cp has not been returned for evaluation. the investigation is underway and a supplemental medwatch report will be submitted upon completion of the failure analysis investigation. | Neither the device nor data logs were received for failure investigation. a failure analysis investigation was performed to review the device history record and found no manufacturing issues or reworks associated with this pump. there are no other complaints associated with pumps from this lot.
37973 | 1220648-2018-00110 | 7935990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7935990&pc=OZD |  |  |  | 10/04/2018 |  |  | 09/12/2018 | 10/04/2018 |  |  | 02/29/2020 |  | 10/04/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1336708 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Malfunction | No | Yes | Break | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A patient was taken to the catheterization lab for angioplasty of two coronary arteries. an impella cp was placed via axillary artery due to the femoral arteries being calcified and small in diameter. as the impella cp was supporting the patient, the introducer was seen to have damage. the orange hub broke off and this prompted blood loss of an estimated 500cc. a second introducer was placed and the intervention continued. | [5 empty] | The introducer product was returned for analysis. the broken orange hub cap of the 14fr x 13 cm introducer was confirmed. during inspection of the returned sheath, half of the orange hub cap was detached from the sheath. there were no traces of glue on either the orange hub or the white sheath. the remaining hub half still attached to the sheath was easily removed by hand; it was not glued either. a supplier corrective action request (scar) # (b)(4) was generated to address this failure mode. the failure will continue to be monitored and trended.
37974 | 1220648-2018-00122 | 7936358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7936358&pc=OZD | 08/07/2018 | 08/08/2018 |  | 10/04/2018 |  |  | 08/08/2018 | 10/04/2018 |  |  | 06/30/2020 |  | 10/04/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1349611 |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 83 YR | [5 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Ischemia | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported on (b)(6) 2018 an impella cp was placed in an (b)(6) female patient. the patient was successfully supported during a stenting and ballooning procedure. following the completion of the procedure it was noticed that there was groin bleeding from the left femoral artery at the pump's insertion site. because of the patient's religious prohibitions only dextrose had been added to the purge solutions. a fem stop was applied to the groin site, and the bleeding was controlled. by the morning of (b)(6) 2018 the patient's hemoglobin was down to 4 mg/dl, and the left leg was found to be pulseless. the patient's lactate was also noted to be very high and increasing. as a result of the patient's frail condition, and the patient's religious restriction of not receiving replacement blood products the family withdrew care. on (b)(6) 2018, after 28.18 hours of impella cp support the patient expired. | [5 empty] | The impella cp was discarded subsequent to use, preventing any evaluation of the suspect this device. a review of the device history record was performed. there were no issues identified with this device during manufacture or inspection. there have not been any other complaints with pumps from this lot number. in conclusion, a root cause was unable to be determined because product was not returned.
37975 | 1220648-2018-00115 | 7930085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7930085&pc=OZD | 09/06/2018 | 09/06/2018 |  | 10/03/2018 |  |  | 09/06/2018 | 10/03/2018 |  |  | 02/29/2020 |  | 10/03/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1330451 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that a (b)(6) male patient was being administered a stress test and collapsed due to myocardial infarction. impella cp device was implanted without complication. the following day, a plasma free hemoglobin result was reported to exceed 3000 mg/dl, indicating hemolysis was present. after repositioning and treatment with crrt, pump was explanted and patient was placed on ecmo; however, hemolysis worsened. ecmo has since been removed and an iabp has been placed. it should be known that, while impella pump was inserted, 8 units of blood was transfused. | [5 empty] | The device has been received and the investigation is underway. a supplemental report will be filed upon completion of evaluation. | The product was received and a failure investigation was performed. the pump had biomaterial near the base of the impeller and in the purge gap. this material could have contributed to the hemolysis. biomaterial was removed, pump was cleaned, and device passed simulated use testing for hemolysis. data logs confirm an auto-suction response early after support begins, followed by elevated motor current throughout support, possibly indicating thrombus ingestion. both alarms resolved within a few minutes. a device history review was conducted and found no manufacturing issues were identified. due to pump passing failure analysis testing, we're unable to definitively determine a root cause of the reported hemolysis.
37976 | 1220648-2018-00112 | 7930086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7930086&pc=OZD | 09/04/2018 | 09/04/2018 |  | 10/03/2018 |  |  | 09/04/2018 | 10/03/2018 |  |  | 03/31/2020 |  | 10/03/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP | [5 empty] | GM |  | 1339807 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | CONSUMER | FOREIGN | HEALTH PROFE |  |  |  | 49 YR | [5 empty] | Injury | Yes | No | Suction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported inserting a (b)(6) male patient presenting to the intensive care unit with cardiogenic shock on (b)(6) 2018. patient was sedated, ventilated, and an impella cp pump was inserted into the right femoral artery without any complications. during procedure, an alarm for loss of placement signal occurred and was resolved. in addition, multiple suction alarms were present. echo was performed and correct positioning of the device was achieved. despite troubleshooting, suction alarms continued. hematuria was visible during procedure and physician confirmed hemolysis present, however, no measurement of plasma free hemoglobin was taken for definitive diagnosis. due to continued suction alarms, the pump was removed. patient was treated with continuous veno-venous hemodialysis and an intercooler. physician believes impella pump responsible for hemolysis. the signs of hemolysis and suction were not treated with any volume per the ifu as recommended: "suction may occur if the blood volume available for the impella® catheter is inadequate or restricted. suction limits the amount of support that the impella® catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure." | [5 empty] | The device has been received and the investigation is underway. a supplemental report wll be filed upon completion of evaluation. | The product was received and a failure investigation was performed. a small amount of biomaterial was discovered on the pump. biomaterial was removed, pump was cleaned, and device passed simulated use testing for hemolysis. data logs confirm suction alarms and pump out of position. hemolysis may have been caused by biomaterial identified and/or pump housing sitting too close to aortic valve; either situation could result in excess shear stress on red blood cells. a device history review was conducted and found no manufacturing issues were identified. due to pump passing failure analysis testing, we're unable to definitively determine a root cause of the reported hemolysis.
37977 | 1220648-2018-00114 | 7930097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7930097&pc=OZD | 04/01/2018 | 09/07/2018 |  | 10/03/2018 |  |  | 09/07/2018 | 10/03/2018 |  |  | 11/30/2019 |  | 10/03/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1320543 |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 37 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tissue Damage | Rupture | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | On (b)(6) 2018, a (b)(6) male patient presented with acute myocardial infarction and cardiogenic shock. the impella cp heart pump was placed successfully. despite patient arresting shortly after placement, cpr was performed and patient was resuscitated. placement signal and motor current were checked and operating as intended. femostop was administered due to groin bleeding. the following day, femostop was removed and repositioning sheath was resutured. while repositioning with a c-arm, the device "caught" on the papillary muscle, however, "freed up nicely." a single suction alarm emitted, however, corrected itself. on september 7th, 2018, during a fellows educational session, a physician had reported this patient to have endured "severe" mr 2 days post explant of the device. this event required "immediate surgery" and an the impella, used during procedure, was alleged to have been responsible. it was not clear if posterior-medial or antero-lateral pap muscle had ruptured. | [5 empty] | An investigation is underway. upon completion, a supplemental report will be filed. | The product was discarded by the customer, therefore, evaluation of actual device could not be completed. imaging was not provided and device and data logs were not returned to abiomed for evaluation. a device history review was conducted and found no manufacturing issues were identified in relationship to this failure mode. we cannot confirm the relationship of the impella to the reported papillary muscle rupture.
37978 | 1220648-2018-00108 | 7931779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7931779&pc=OZD | 09/05/2018 | 09/05/2018 |  | 10/03/2018 |  |  | 09/05/2018 | 10/03/2018 |  |  | 11/30/2018 |  | 10/03/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA LP5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1269703 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | CONSUMER | HEALTH PROFESSIONAL |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | No | Suction Problem | [6 empty] | Hemolysis | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported a (b)(6) white male presenting with acute myocardial infarction and cardiogenic shock. patient's o2 saturation dropped and ecmo was placed. impella lp5.0 pump was prepped and inserted, to assist unload of left ventricle, without issue. patient was subsequently returned to intensive care unit. during the course of treatment, there were no suction issues or alarms. however, at one point the pump was found to be close to the mitral valve and required re-positioning. the pump was eventually removed due to "multiple positioning issues" and hemolysis in which the device was implicated. additionally, there were multiple hematomas found at the impella and ecmo insertion sites. as treatment, patient received "numerous" blood transfusions due to retro-peritoneal bleed and ecmo. a new pump was exchanged and patient is in stable condition. | [5 empty] | The device has been received and a failure analysis investigation has begun. upon completion of the evaluation, a supplemental report will be completed. | The pump was returned for investigation. a large crack in the epo-tek near the impeller was discovered upon visual inspection. the clinical account also pointed out positioning issues during the case. hemolysis was not mentioned as a symptom until the very end of the case, just before the pump was explanted. the abiomed representative mentioned that there were repositioning issues, and hemolysis was still present after the pump was repositioned. he also believed that the patient's anatomy was preventing good placement of the pump. the physician originally planned to explant the pump and re-insert it, but decided to replace the pump instead. the data logs returned confirm the positioning issues experienced during the case. the pump was tested in our lab for hemolysis and failed, reproducing the failure mode. since the failure mode was reproduced in the lab with neutral positioning, the crack in the epo-tek is determined to be the root cause for the hemolysis. the origin of the crack is unknown. at 530 hours, the pump was run beyond product claims. the root cause of the failure is the cracked epo-tek near the impeller.
37979 | 1220648-2018-00111 | 7927731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7927731&pc=OZD | 09/06/2008 | 09/07/2018 |  | 10/02/2018 |  |  | 09/07/2018 | 10/02/2018 |  |  | 02/29/2020 |  | 10/02/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1353284 |  | 0048-0020 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient arrived in cardiogenic shock and had an impella cp placed after approximately 25 minutes of active cpr. the patient had a coronary angioplasty performed while on the imeplla support. the patient was then transferred to the (b)(6) medical center for followup and monitoring. while on support at the second tertiary facility, there was bleeding from the access sites and lungs. the patient needed 5 units of blood product replacement. there was communication with the medical affairs department at abiomed that the bleeding was observed from the impella pump's repositioning sheath. | [5 empty] | The root cause of the access site bleeding was unable to be determined as the product was not returned for analysis and testing. no corrective action is warranted at this time. the failure mode will be monitored and trended.
37980 | 1220648-2018-00120 | 7928414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7928414&pc=OZD | 09/05/2018 | 09/05/2018 |  | 10/02/2018 |  |  | 09/05/2018 | 10/02/2018 |  |  | 06/30/2020 |  | 10/02/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1352792 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2018 an impella lp2.5 was placed in a (b)(6) male patient. on (b)(6) 2018 the patient's urine became bloody, and the patient's ldh increased to over 5000. in response to the patient's hemolysis, 4 units of replacement blood products were administered to the patient, and the device was repositioned. the patient hemolysis improved, the blood urine resolved. the patient support was successfully continued. the patient was reported to have been successfully supported for 237.84 hours (9.91 days). the patient was in stable condition upon the removal of the impella cp. | [5 empty] | The impella lp2.5 was discarded subsequent to use, preventing any evaluation of the suspect impella lp2.5. the automatic impella controller data logs were also not returned for analysis. in conclusion, a root cause was unable to be determined because product and data were not returned.
37981 | 1220648-2018-00118 | 7920219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7920219&pc=OZD | 08/28/2018 | 08/29/2018 |  | 09/28/2018 |  |  | 08/29/2018 | 09/28/2018 |  |  | 12/31/2019 |  | 09/28/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, PRODUCT |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1329330 |  | 005042 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Brain Injury | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that a (b)(6) male patient was found to have stents that were poorly opposed to the arterial wall. as a result the patient was transferred to another facility, and was administered tpa prior to transport. following transport the patient's stents were under-dilated and an nc balloon was used to expand the stents. the patient's motor skills were tested in the cath lab, but the patient gave inappropriate verbal responses. in addition, an in-stent thrombus was noted. heparin given in cath lab, the patient's act was 382. there was no heparin added to the purge or systemic in unit due to the patient's high act. three ct's were performed within a few hours, and only the last ct showed a bleed. the patient was then transferred out to another facility. patient impella support continued, and the pump was still supporting the patient when care was withdrawn by the family, and the patient expired. | [5 empty] | The impella lp2.5 was discarded subsequent to use, preventing any evaluation of the suspect impella lp2.5. the automatic impella controller data logs were also not returned for analysis. in conclusion, a root cause was unable to be determined because product and data were not returned, although there is no indication that the impella lp2.5 caused or contributed to the cerebrovascular accident. contrary to the instructions for use of this product, heparin was not added to the purge solution. in addition there were multiple exchanges during patient care that may have contributed to the event.
37982 | 1035166-2018-00064 | 7920303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7920303&pc=DYB | 05/12/2018 | 11/05/2018 |  | 09/28/2018 |  |  | 09/27/2018 | 09/28/2018 |  |  | 11/01/2020 |  | 09/28/2018 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-13849 |  | 0052-3021 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 65 YR | [5 empty] | Death | Yes | No | Break | Crack | Deformation Due to Compressive Stress |  |  |  |  | Death | Myocardial Infarction | Shock | [10 empty] | Death | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The patient was presented with ami (acute myocardial infarction) with shock, impella post pci (percutaneous coronary intervention). after second attempt at impella rp insertion via 23fr sheath, distal tip of introducer sheath was misshaped (wavy) and crack/break was noted on introducer hub. when md was able to remove impella rp, new impella rp pitched and inserted into patient. there were no reported issues noted with the product prior to the procedure. the patient was not obese, no excessive force was used to deliver this introducer. the rep also reported that there was significant bleeding at the reported crack/ break; more than 2 units of replacement blood products were given to the patient. the family withdrew support and the patient subsequently expired. | [5 empty] | During our investigation it was discovered, the initial mdr 1035166-2018-00064 report had failed submission. based upon our findings, we have resubmitted this initial mdr report. concomitant medical products: impella rp 004343, serial number: (b)(4), lot number 1314688. date: (b)(6) 2018. conclusion not yet available as the investigation is on-going. | Device used in treatment. for bleeding more than 2 units of replacement blood products were given to the patient. the family withdrew support and the patient subsequently expired. one 23 fr abiomed introducer sheath was received from the customer without the dilator. no other accessories were received with this device. blood was found on and inside the sheath. the distal tip circumference edge was damaged. the hemostasis valve was heavily damaged. upon evaluation of the returned product, it was found that there was impact damage (waviness) around the circumference of the distal tip edge as stated in the event description. the hemostasis valve was heavily damaged/torn in a few places. the introducer was then tested using the iso liquid leakage test. the device was connected at the distal tip and pressurized to less than 1.5kpa before the sheath started leaking at the crack in the hub and at the damaged hemostasis valve. the crack in the hub was confirmed when viewed under a 10x microscope and when leak tested. the customer complaint also indicated that there were several kinks in the sheath. no kinks in the sheath were found. returned device analysis reveals one 23f abiomed introducer sheath that is within manufacturing specifications. no manufacturing defects were found. potential contributing factors to the damages of the sheath is if the device was advanced through an area of resistance (as stated in the precautions of the ifu), or if the device was subjected to unusual stresses (as stated in the handling and storage section of the ifu). based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. as per procedure abiomed introducer sheath in-process and final inspection, the devices in this lot were inspected as follows: 9.2.2 visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. per section 12.2, visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages . the leak test is performed by qa on 100% of the sheaths. per ifu: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane.
37983 | 7913719 | 7913719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7913719&pc=OZD | 08/26/2018 | 09/11/2018 | 09/12/2018 | 09/27/2018 | 09/12/2018 | 09/27/2018 |  | 09/27/2018 |  |  |  |  | 09/27/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Device was returned to manufacturer | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 25915 DA |  |  | 93 |  |  | Death | Yes | Yes | Break | Failure to Power Up | Contamination /Decontamination Problem | Infusion or Flow Problem |  |  |  | Death | [12 empty] | Death | [28 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Pump 1 had flows less than expected. material was found in the pump. it's currently being soaked in solution to remove and determine what the material is. pump 2 did not start. the pump housing/outlet cage and impeller were crushed. engineers are determining how this could have occurred. this is the first time this occurred within last year. i have heard of past problems after f/u with cardiac cath manager. manufacturer response for impella, (brand not provided) (per site reporter). the two impella pumps that were used during the recent case were returned to abiomed. i spoke with one of our engineers yesterday. their initial finding were: pump 1. had flows less than expected. material was found in the pump. it's currently being soaked in solution to remove and determine what the material is. pump 2. did not start. the pump housing / outlet cage and impeller were crushed. engineers are determining how this could have occurred. manufacturer response for impella, impella (per site reporter). the two impella pumps that were used during the recent case were returned to abiomed. i spoke with one of our engineers yesterday. their initial finding were: pump 1. had flows less than expected. material was found in the pump. it's currently being soaked in solution to remove and determine what the material is. pump 2. did not start. the pump housing / outlet cage and impeller were crushed. engineers are determining how this could have occurred.
37984 | 1220648-2018-00116 | 7911921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7911921&pc=OZD | 08/30/2018 | 08/30/2018 |  | 09/26/2018 |  |  | 08/30/2018 | 09/26/2018 |  |  | 05/31/2020 |  | 09/26/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1352382 |  | 0048-0032 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 46 YR | [5 empty] | Injury | Yes | No | Malposition of Device | [6 empty] | Mitral Perforation | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient was admitted to the hospital in cardiogenic shock and had an impella cp placed for hemodynamic support, in addition to ecmo. the patient was subsequently transferred, on support, to the tertiary medical center for continued care. while at the tertiary medical center there was the ability to wean down from the ecmo support. after 10 days of impella support there was a transesophageal echo done (tee) that showed the impella position was not optimal. the team had improperly positioned the impella and the mitral valve had been damaged. the patient was sent for a mitra clip procedure and the impella was explanted. | [5 empty] | To date no product or imaging have been returned from the medical facility. imaging from the tee and data logs from the impella console have been requested, and are awaiting return. should more details be shared, the supplemental medwatch will be completed and filed.
37985 | MW5080098 | 7910023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7910023&pc=DWA | 09/02/2018 | 09/24/2018 |  | 09/25/2018 |  |  |  | 09/25/2018 |  |  | 03/31/2019 |  | 09/25/2018 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | IMPELLA CONTROLLER |  |  | DWA |  |  | 0042-0000-US |  |  | ABIOMED INC. |  |  | US | [5 empty] | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 61 YR |  |  | 78 |  |  | Injury | No | Yes | Loss of Power | [6 empty] | No Code Available | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | The patient was moving from the surgical table to the bed and all of a sudden, the impella console lost its signal and afterwards the impella software shutdown.
37986 | 1035166-2018-00036 | 7902918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7902918&pc=DYB | 03/26/2018 | 09/24/2018 |  | 09/24/2018 |  |  | 03/28/2018 | 09/24/2018 |  |  | 12/01/2020 |  | 09/24/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC |  |  | US |  | C1-13895 |  | 0052-3025 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [7 empty] | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | On (b)(6) 2018, during a hrpci procedure, the dilator perforated a l-iliac artery. the physician observed the incident during the actual procedure. after the physician had used the abiomed .018 wire and the benson wire, it was noted that the dilator perforated l- iliac. a balloon used to control bleeding and the perforation was repaired. a 7f introducer was used above perforation to do intervention and the impella was placed without issue inside r-iliac. the physician used balloon and stents to rca and kissing balloons and stents to lad, lm and circ. preclose x 2 was also utilized. patient was reported to be stable. | On (b)(6) 2018, during a hrpci procedure, the dilator perforated a l-iliac artery. the physician observed the incident during the actual procedure. after the physician had used the abiomed .018 wire and the benson wire, it was noted that the dilator perforated l- iliac. a balloon used to control bleeding and the perforation was repaired. a 7f introducer was used above perforation to do intervention and the impella was placed without issue inside r-iliac. the physician used balloon and stents to rca and kissing balloons and stents to lad, lm and circ. preclose x 2 was also utilized. patient was reported to be stable. |  |  |  |  | H6: a preliminary review showed there were no manufacturing rejects or anomalies of this event type recorded in the device history record. conclusion not yet available as the investigation is on-going. h11: during our final investigation, it was discovered the initial mdr 1035166-2018-00036 (b)(4) report had failed submission. based upon our findings, we have corrected the reported failure in this initial mdr report. oscor inc., is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc., which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor, inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor, inc., or its employees, that the report constitutes an admission that the device, oscor, inc., or its employees caused or contributed to the event described in this report. | This report is being submitted retrospectively per fda request during our final investigation, it was discovered the initial mdr cmp (b)(4) report had failed submission. based upon our findings, we have corrected the reported failure in tmdr follow up 2 mdr submitted on 09/26/2018. oscor inc., is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc., which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor, inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor, inc., or its employees, that the report constitutes an admission that the device, oscor, inc., or its employees caused or contributed to the event described in this report.
37987 | 1035166-2018-00088 | 7902932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7902932&pc=DYB | 09/12/2018 | 11/05/2018 |  | 09/24/2018 |  |  | 09/13/2018 | 09/24/2018 |  |  | 01/01/2021 |  | 09/24/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14 F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-13998 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 63 YR | [5 empty] | Injury | Yes | No | Break | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | It was reported that removal of dilator the md noticed the orange piece had broken off. within a minute there was excessive bleeding and patient lost 500cc blood. admitted for a high risk angioplasty to the cardiac cath lab. axillary access site chosen due to small and diseased femoral arteries. patient was reported to be stable. based on additional information , md swapped for 14x30 in the kit which inserted without problems, impella weaned and removed. percloses deployed, then extravasation noted on angiogram, ballooned for 10 minutes, then deployed covered stent. hemostasis achieved with the long introducer and 1 unit of blood was infused to the patient. the team percutaneously placed a covered stent - this was not done surgically. no additional information is available and no other adverse event was reported. | [5 empty] | Two 14f abiomed introducer sheaths (13cm and 25cm) were returned from the customer along with four dilators of varying lengths and french sizes. there were no other accessories. the customer handled this device prior to returning to oscor for evaluation, and had removed and placed back on, the orange split cap that is currently on the device. blood was found on all returned devices. one half of the orange cap from the 14f, 13cm introducer sheath was detached from the hub. the other half of the orange cap was still attached to the hub. the rest of the devices looked normal. ne of the orange split caps was detached from the sheath's hub. the other orange cap was still attached. when viewed under a 10x microscope, possible adhesive was found on white hub but it was difficult to verify. side 1 appears to have minimal adhesive and side 2 has slightly more. the bottom of the orange cap does not have much evidence of adhesive on either side of the tab. however, due to the location of where the adhesive is placed, it can be very difficult to properly identify adhesive. the ifu instructs the user to unscrew the dilator from the threaded caps prior to removing. capa has been previously opened to investigate the root cause and to implement corrective action to prevent recurrence of this issue and as per our preliminary investigation user error (deviation from the instruction for use) determined as probable cause. per procedure adelante s2s introducer sheath in-process and final inspection: "with naked eye at a distance of 12" to 18", verify the sheath hub and split caps are free of excess silicone oil applied between the seals during assembly. verify split caps are properly placed and secured onto sheath hub and free of cracks/damages or excessive adhesive." additionally, destructive testing is performed per sampling plan. inspector is to "insert dilator into sheath, lock dilator hub lock nut. hold dilator hub with one hand and one of the sheath hub wing with the other hand. bend the dilator hub to the side until dilator hub lock nut is disconnected from the sheath hub. split caps should remain in place. there should be no damage and fracture to split cap and tabs." per ifu: cautions: prior to use, read all package inserts, warnings, precautions, and instructions. failure to do so may result in severe injury or death. directions for use: insert vessel dilator into sheath and rotate the dilator cap over valve housing to secure the dilator onto sheath assembly. thread the dilator/sheath assembly over the guidewire. advance the dilator and sheath together with a twisting motion over the guidewire and into the vessel. fluoroscopic observation may be advisable. attaching a clamp or hemostat to the proximal end of the guidewire will prevent inadvertently advancing the guidewire entire into the patient. once assembly is fully introduced into the vascular system, separate the dilator cap from the sheath valve housing by rotating the dilator cap off the hub. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Our investigation is in progress; follow up report will be submitted when investigation is completed.
37988 | 1220648-2018-00107 | 7899811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7899811&pc=OZD | 08/24/2018 | 08/24/2018 |  | 09/21/2018 |  |  | 08/24/2018 | 09/21/2018 |  |  | 06/30/2020 |  | 09/21/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1356018 |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported implanting an impella cp pump on a (b)(6) male patient on (b)(6) 2018. while inserting the repositioning sheath, physician found patient's groin "blowing up" with a femoral bleed present. pressure was applied, pump explanted, femostop was administered, and a blood transfusion of approximately 1 to 2 units was completed. a vascular surgeon was consulted and agreed with treatment and no further intervention was performed. | [5 empty] | The pump was received and an investigation was performed to identify any possible malfunctions or deficiencies that could have played a roll in the reported femoral bleed. the pump was inspected and no abnormalities or deficiencies were noted. the device operated as intended with no issues.
37989 | 1220648-2018-00109 | 7900323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7900323&pc=OZD | 08/27/2018 | 08/27/2018 |  | 09/21/2018 |  |  | 08/27/2018 | 09/21/2018 |  |  | 08/31/2019 |  | 09/21/2018 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1306307 |  | 0048-0003 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 41 YR | [5 empty] | Injury | Yes | No | Material Rupture | Insufficient Information | [5 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A patient admitted with cardiogenic shock and acute myocardial infarction had an impella cp placed for hemodynamic support via the left femoral artery while undergoing active cpr. while on support the patient had a multivessel coronary artery intervention performed and was then admitted to the icu for monitoring. during her time in the icu the team was attempting to determine if she had disseminated intravascular coagulation (dic) as she was bleeding from all access sites. while in the icu the patient had an observed impella access site bleed, at the femoral artery, after being turned in the bed. the team controlled the bleeding only to have the patient re-bleed during the overnight. the patient was sent to the or for treatment of the bleed site, though she did expire in the operating room (or). | [5 empty] | The pump product was not returned for analysis or testing. the facility did return a photograph of the product showing a tear. the tear of the pump sheath most likely cased the access site bleeding, however without return of the product the definitive cause could not be determined. the failure mode will be monitored and trended.
37990 | 1035166-2018-00049 | 7893218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7893218&pc=DYB | 05/06/2018 | 09/20/2018 |  | 09/20/2018 |  |  | 05/10/2018 | 09/20/2018 |  |  | 11/01/2020 |  | 09/20/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-13850 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 71 YR | [5 empty] | Malfunction | No | Yes | Break | [6 empty] | Chest Pain | Ischemia | Myocardial Infarction | [39 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that on (b)(6) 2018, during the insertion of the impella pump the 14f 13cm (short sheath) hub broke open during dilator removal. there was a large blood release, and they scrambled to re-wire and remove the sheath. the device was being used for primary support during revascularization. the second sheath (long sheath) was then placed over the wire to be able to insert the impella cp. patient is reported to be stable. see mdr # 1035166-2018-00077 for 14f, 25cm sheath. | [5 empty] | Concomitant medical products: impella cp, model 0048-0026, serial: (b)(4), lot: 1324139. during our final investigation, it was discovered the initial mdr 1035166-2018-00049 cmp 5013 report had failed submission. based upon our findings, we have corrected the reported failure in this initial mdr report. one 14f, 13cm long abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. both orange split caps were detached from the sheath's hub. visible blood was found on and inside the sheath. the introducer sheath was received from the field to the customer with only 1 orange split cap detached. in a conversation with the customer, it was revealed that they removed the other orange split cap. when viewed under a 10x microscope, traces of adhesive were found on side 1 of the white hub and side 2 had a sufficient amount of adhesive. the bottom of the orange caps do not have much evidence of adhesive on the tab of side 1; side 2 has slightly more. the orange split caps were separated from the hub of device. the device appeared to have had minimal to sufficient amounts of adhesive on the hub and orange cap tabs. based on the returned device, it appears that the lack of glue on the hub/orange cap tabs could have contributed to this failure. capa p was previously opened to investigate the root cause and to implement corrective action to prevent recurrence of this issue. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. per procedure adelante s2s introducer sheath in-process and final inspection: with naked eye at a distance of 12" to 18", verify the sheath hub and split caps are free of excess silicone oil applied between the seals during assembly. verify split caps are properly placed and secured onto sheath hub and free of cracks/damages or excessive adhesive. additionally, destructive testing is performed per sampling plan. inspector is to "insert dilator into sheath, lock dilator hub lock nut. hold dilator hub with one hand and one of the sheath hub wing with the other hand. bend the dilator hub to the side until dilator hub lock nut is disconnected from the sheath hub. split caps should remain in place. there should be no damage and fracture to split cap and tabs. per ifu: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action.
37991 | 1220648-2018-00100 | 7893542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7893542&pc=OZD | 07/02/2018 | 08/23/2018 |  | 09/20/2018 |  |  | 08/23/2018 | 09/20/2018 |  |  | 05/31/2020 |  | 09/20/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1342102 |  | 005066-JP |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | No | Human-Device Interface Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | INTRA AORTIC BALLOON PUMP | PCPS, VA-ECMO | INTRA AORTIC BALLOON PUMP | PCPS, VA-ECMO | [19 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant in (b)(6) was utilizing an impella 5.0 for support, in combination with ecmo and an intra aortic balloon pump (iabp). the impella was placed via a graft to the axillary artery due to the inability to place femorally, as both femoral arteries were accessed by the ecmo and iabp. the insertion was the facilities first impella patient and there was an access site complication. the patient developed a hematoma at the axillary site. during the patient's 12 days of support, the patient received 54 units of blood replacement. after the 12 days of support the impella was explanted and the patient is noted to be stable. | [5 empty] | The investigation is on-going at this time. no product has been returned for analysis from the complainant in (b)(6). should more information be shared the supplemental report will be filed. | Since the original medwatch was filed, the product was returned and investigation completed. the reported hematoma was confirmed via the clinical account. the impella pump was returned intact. the root cause of the access site hematoma was unable to be determined as insufficient clinical information was provided.
37992 | 1220648-2018-00106 | 7893546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7893546&pc=OZD | 08/17/2018 | 09/10/2018 |  | 09/20/2018 |  |  | 09/10/2018 | 09/20/2018 |  |  | 06/30/2020 |  | 09/20/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1354035 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report | [8 empty] | 72 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Mitral Regurgitation | [12 empty] | Hospitalization | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient was admitted for a high risk angioplasty and placed on impella hemodynamic support. the patient had low flows with the impella, and so it was explanted and replaced by another impella cp. there was not any issue with the replacement. the second pump successfully supported the patient until the decision was made by the family to withdraw care. at that time the patient expired. | [5 empty] | The pump product and data logs were returned for analysis. the initial complaint was pump flows being low during patient support. during the investigation, and visit to the facility, further details were shared of possible mitral valve regurgitation caused by the use of impella. the most likely cause for the low flows was a low native volume within the patient's left ventricle. this low volume could be caused by mitral regurgitation that possibly could have been caused by the use of impella. when the pump was tested the low flows could not be reproduced. the pump performed as designed. there is a possibility that the pump interacted with the mitral valve; however this could not be confirmed. the failure mode will be monitored and trended.
37993 | 1220648-2018-00104 | 7887484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7887484&pc=OZD | 07/25/2018 | 08/20/2018 |  | 09/18/2018 |  |  | 08/20/2018 | 09/18/2018 |  |  | 10/31/2018 |  | 09/18/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPLELLA LP5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1273217 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 47 YR | [5 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2018 the physician placed an impella lp2.5 in a (b)(6) male patient who had a mechanical mitral valve previously placed. approximately one hour prior to the insertion of the impella a va-ecmo had previously been placed. on (b)(6) 2018 patient bleeding and hemolysis occurred, requiring the administration of 240 units of packed red blood cells and 180 units of fresh frozen plasma. the patient bleeding emanating from both the ecmo and impella lp2.5 insertion sites. the clinicians determined that the hemolysis was the result of the inlet of the impella touching the mechanical valve in the mitral position as a result of the abnormal shape of ventricle and the dilated cardiomyopathy (dcm). the physician reported that by re-positioning the impella lp2.5 the hemolysis improved in the next day. on (b)(6) 2018 the impella lp2.5 was removed from the patient, and an impella lp5.0 was placed in an effort to provide the patient with additional hemodynamic support. on (b)(6) 2018 the patient expired due to worsening condition. there was no indication from the complainant that any issue with the impella lp2.5 caused or contributed to the patient outcome. | [5 empty] | The impella lp2.5 was discarded subsequent to use, preventing any evaluation of the suspect impella lp2.5. the automatic impella controller data logs were also not returned for analysis. in conclusion, a root cause was unable to be determined because product and data were not returned.
37994 | 1220648-2018-00090 | 7887525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7887525&pc=OZD | 08/07/2018 | 08/07/2018 |  | 09/18/2018 |  |  | 08/07/2018 | 09/18/2018 |  |  | 05/31/2020 |  | 09/18/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1348316 |  | 0048-0032 |  | No | [5 empty] | No | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | On (b)(6) 2018 a (b)(6) male patient was admitted to the er feeling unwell. the patient had a history of diabetes and hypertension, as well as mitral valve disease and chronic total occlusion in the right coronary artery. the patient was transferred to the cath lab and an impella cp was placed. the pump was inserted through the right femoral artery using a 14fr x 25cm oscor introducer. later that night, the nurse reported that she was unable to get pulses in the patient's right foot and that the foot was cold. the decision ws made to remove the impella. tee was weaned over 30 minutes and smoothly removed. a perclose was deployed, but the patient still had no pulses. the patient was then taken to the or and an endarterectomy was performed, which restored pulses in the leg. the patient had been successfully supported for 16 hours when the event occurred, and patient support was reported as successful, despite the reported ischemia. in addition, the patient was reported in stable condition. | [5 empty] | The impella cp was not returned for evaluation, as it was discarded following patient use and removal. the patient was reported to have chronic total occlusion in the right coronary artery, as well as diabetes. chronic total occlusion (cto) causes a narrowing or blockage of the arteries due to atherosclerosis, which is a buildup of fatty deposits and plaque on the inner walls of arteries. both diabetes and hypertension are increased risk factors for cto. given that the patient had cto, it is likely that fatty deposits or plaque were already present in the vasculature. the insertion of the pump most likely caused a larger blockage of the vasculature in the right leg, leading to ischemia, although this could not be definitively determined, there are no other complaints associated with pumps from this lot relating to this failure mode. .
37995 | 1220648-2018-00101 | 7887554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7887554&pc=OZD | 08/22/2018 | 08/22/2018 |  | 09/18/2018 |  |  | 08/22/2018 | 09/18/2018 |  |  | 01/31/2020 |  | 09/18/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1342764 |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A cardiomyopathy patient was being supported by an impella cp due to cardiogenic shock. during the48 hours of support the patient's physician observed signs of hemolysis. the team chose to treat the patient with blood products, dialysis, and explant of the impella and placement of va ecmo. | [5 empty] | The facility in (b)(6) did not return the product for investigation and hemolysis testing. the root cause of the reported hemolysis was unable to be determined.
37996 | 1220648-2018-00099 | 7878738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7878738&pc=OZD | 04/03/2018 | 08/16/2018 |  | 09/14/2018 |  |  | 08/16/2018 | 09/14/2018 |  |  | 08/31/2019 |  | 09/14/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1324050 |  | 005048-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2018 an impella lp2.5 was placed in a (b)(6) female patient who was in cardiogenic shock. the patient was supported via the pump and ecmo during the performance of a balloon valvuloplasty. following pump placement a hematoma was found at the pump's insertion site. patient bleeding occurred requiring the administration of 12 units of packed red blood cells and blood products and 4 units fresh frozen plasma. the physician reported that the patient was supported for 92.12 hours with the impella lp2.5, and following the pump support the patient was reported to be in stable condition. | [5 empty] | The impella lp2.5 was discarded subsequent to use, preventing any evaluation of the suspect impella lp2.5. the automatic impella controller data logs were also not returned for analysis. in conclusion, a root cause was unable to be determined because product and data were not returned, although there was no allegation of the impella having contributed to the patient outcome. there are no other complaints for pumps from this lot associated with this failure mode.
37997 | 1220648-2018-00105 | 7879383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7879383&pc=OZD | 08/17/2018 | 08/17/2018 |  | 09/14/2018 |  |  | 08/17/2018 | 09/14/2018 |  |  | 07/10/2018 |  | 09/14/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1348507 |  | 005042 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Use of Device Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2018 an impella lp2.5 was placed in the left femoral artery of a (b)(6) male patient. optimal positioning and flows were achieved under fluoroscopy. it was observed that some oozing was observed at the insertion site. the physician mad the decision to leave the peel-away sheath in place. on (b)(6) 2018, arterial pulses were absent on the pump insertion side. it was decided to remove the impella, and peel-way sheath and increase pressers. the patient was taken to the or to re-vascularize the affected leg. following this event the patient was reported to be in stable condition. the instructions for the impella lp2.5 advises the user to: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." | [5 empty] | The impella lp2.5 was discarded subsequent to use, preventing any evaluation of the suspect impella lp2.5. the automatic impella controller data logs were also not returned for analysis. in conclusion, a root cause was unable to be determined because product and data were not returned. the instructions for the impella lp2.5 advises the user to: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site."
37998 | 1220648-2018-00102 | 7855179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7855179&pc=OZD | 08/06/2018 | 08/08/2018 |  | 09/07/2018 |  |  |  | 09/07/2018 |  |  | 05/31/2020 |  | 09/07/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1344654 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 86 YR | [5 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported that an impella cp was placed in an (b)(6) female patient. following pump placement hematoma developed and patient bleeding from the groin occurred. the impella cp was removed from the patient and vascular repair was completed. in addition the patient received 3 units of replacement blood products. patient support from the impella cp was reported as successful. following the vascular repair the patient was reported to be in stable condition. | [5 empty] | The impella cp was returned for evaluation. damage was observed to the tip of the wound closure of the repositioning unit. there is no evidence of flash removal aligning with the damages. the patient was noted to have calcification in their peripheral vasculature, which most likely led to the damage of the wound closure. the root cause of the access site bleeding was damage to the repositioning unit wound closure. the damage to the wound closure was most likely caused by the calcification in the patient's peripheral vasculature. there are no other complaints associated with pumps from this lot. there are no other complaints associated with repositioning units from this lot. there is one other complaint for a tapered wound closure from this lot relating to access site bleeding. | The device has been returned and is currently under evaluation. at the conclusion of the evaluation a supplemental medwatch report will be submitted.
37999 | 1220648-2018-00098 | 7855180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7855180&pc=OZD | 08/09/2018 | 08/10/2018 |  | 09/07/2018 |  |  | 08/10/2018 | 09/07/2018 |  |  | 04/30/2020 |  | 09/07/2018 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1349052 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 0 YR | [5 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | On (b)(6) 2018 an impella cp was placed in a male patient (age unknown). the patient was successfully supported for 9 hours with this pump. the following day leg ischemia was observed in the impella access leg. the impella cp was removed and the leg ischemia resolved. patient support was reported as successful, and the patient's condition was reported as stable. | [5 empty] | The facility did not return the impella pump product for investigation; therefore the root cause of the leg ischemia was unable to be determined. a manufacturing review found no other complaints for this failure mode made against this lot of impella pumps. failures of this type will be monitored and trended.
38000 | 1220648-2018-00093 | 7847704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7847704&pc=OZD | 08/07/2018 | 08/08/2018 |  | 09/05/2018 |  |  | 08/08/2018 | 09/05/2018 |  |  | 04/30/2020 |  | 09/05/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1340374 |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Patient Device Interaction Problem | Migration | [5 empty] | Perforation | Cardiogenic Shock | Pericardial Effusion | [10 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | An impella cp was placed for hemodynamic support of a patient who had been refused cardiac bypass surgery by the cardiothoracic surgeons. after placement of the impella, the patient needed further support, and so a temporary pacer was placed. after the pacer was placed, the patient condition worsened and there was a pericardiocentesis done. at this time the decision was made to transfer the patient to the tertiary medical center for care. the medical team sent the patient to surgery and observed the left ventricle had been punctured and the impella was observed outside of the ventricle. the impella was explanted and the patient was supported by ecmo. | [5 empty] | The facility did not return the impella pump product for investigation. fluoroscopic imaging from the procedure was shared for analysis. additionally clinical details were shared for analysis. the patient had an observed akinetic heart apex seen on echo prior to impella placement. the heart muscle wall might also be compromised and aneurysmal from the acute myocardial infarction 7 days prior. the investigation is still on-going at this time. when the analysis is completed, a supplemental report will be drafted. | Since the initial filing of the medwatch, the investigation has closed. no product has been returned to date. the root cause of the perforation of the left ventricle was most likely the advancement of the pigtail 10 minutes after initiating support. the left ventricle was more susceptible to perforation due to the akinetic and aneurysmal nature of the apex following a previous acute myocardial infarction. the failure mode will be monitored and trended.
38001 | 1220648-2018-00096 | 7837133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7837133&pc=OZD | 08/06/2018 | 08/06/2018 |  | 08/31/2018 |  |  | 08/06/2018 | 08/31/2018 |  |  | 04/30/2020 |  | 08/31/2018 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, PRODUCT |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1341730 |  | 0048-0020 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2018 a (b)(6) male patient with end-stage renal disease and hypertension presented to the hospital's emergency department complaining of chest pains. the patient was diagnosed with acute coronary syndrome with flash pulmonary edema. the patient was transferred to the icu for hemodynamic support, where he began to decompensate. the patient was then removed to the cath lab where an impella cp was placed in response to the patient's cardiogenic shock. following placement of the impella cp the patient was noted to have some oozing at the groin site requiring the application of surgicell and lidocaine. the following morning the oozing was reported to have subsided, and it was reported that the patient had received replacement blood products. on (b)(6) 2018 the patient was reported to be in stable condition, although additional replacement blood was administered due to continued oozing at the insertion site. the oozing at the site continued through (b)(6) 2018 with additional blood being administered to the patient. on (b)(6) 2018 the impella cp was removed from the patient after 156.77 hours of patient support. the support was reported as successful, with the patient in stable condition. | [5 empty] | The impella cp was discarded by the user facility, therefore a full investigation could not be conducted. the cause of the continued oozing at the patient's insertion site was unable to be determined.
38002 | 1220648-2018-00092 | 7834846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7834846&pc=OZD | 06/04/2018 | 08/06/2018 |  | 08/30/2018 |  |  | 08/06/2018 | 08/30/2018 |  |  | 01/31/2020 |  | 08/30/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1330032 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient in (b)(6) was admitted for a mitraclip procedure and decompensated in the surgical procedure. the impella 5.0 was placed as the patient went into cardiogenic shock. during the patient's icu stay there were signs of hemolysis while on support with the impella 5.0. the facility treated the patient with blood transfusions. the 5.0 pump remained on for hemodynamic support until the patient could be weaned off of support and the pump was explanted after six days. | [5 empty] | The facility did not return the product or data logs for analysis. the investigation has yet to reach a root cause. should the cause be discovered, a supplemental will be filed. the failure mode will be monitored and trended.
38003 | 1220648-2018-00089 | 7834876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7834876&pc=OZD | 08/04/2018 | 08/06/2018 |  | 08/30/2018 |  |  | 08/04/2018 | 08/30/2018 |  |  | 05/31/2020 |  | 08/30/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1346318 |  | 0048-0020 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | No | Use of Device Problem | Patient Device Interaction Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient had an impella cp placed for hemodynamic support after a myocardial infarction. during the support the patient was not therapeutically anticoagulated due to a subarachnoid bleed. as the patient condition improved, the impella was weaned and explanted. after explant from the femoral artery the patient's limb became ischemic. the patient had a surgical embolectomy and flow was restored. | [5 empty] | The investigation is on-going at this time. the product is discarded and no imaging has been shared, though requested. the need for embolectomy prompts reporting, and if more data is shared and a root cause determined, a supplemental will be filed.
38004 | 1035166-2018-00077 | 7824540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7824540&pc=DYB | 05/06/2018 | 12/18/2018 |  | 08/28/2018 |  |  | 08/01/2018 | 08/28/2018 |  |  | 11/01/2020 |  | 08/28/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-13850 |  | 0052-3025 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Leak/Splash | [5 empty] | Chest Pain | Ischemia | Myocardial Infarction | [16 empty] | 0052-3025 14F, 13 CM SHEATH LOT C1-13850 | SEE SECTION H10 | SEE SECTION H11 | [16 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | On august 1, 2018, the customer provided additional information concerning the second sheath used in this procedure, a 14f, 25cm long sheath. they reported that on (b)(6) 2018, the patient presented to the physician with chest pain. the ekg showed avr elevation and the patient had acute mesenteric ischemia and acute myocardial infarction. the product was being used for primary support during revascularization. during the insertion of the impella pump, a 14f 13cm (short sheath) hub broke open during dilator removal. there was a large blood release. they scrambled to re-wire and remove the sheath. (see mdr# 1035166-2018-00049 for 14f, 13 cm sheath). the long sheath was then placed over the wire to be able to insert the impella cp. the md stated that the 25 cm sheath exhibited a slow drip-leak until it was removed. the patient was stable reported as stable following the procedure. see mdr# 1035166-2018-00049 for 14f, 13 cm sheath. | On (b)(6) 2018, the customer provided additional information concerning the second sheath used in this procedure, a 14f, 25cm long sheath. they reported that on (b)(6) 2018, during the insertion of the impella pump a 14f 13cm (short sheath) the hub broke open during dilator removal. there was a large blood release, and they scrambled to re-wire and remove the sheath. the device was being used for primary support during revascularization. the second sheath (14f, 25cm long) was then placed over the wire to be able to insert the impella cp. the md stated this 25cm long sheath also leaked with a slow drip until it was removed. patient is reported to be stable. |  |  |  |  | H10: the following sections were updated in follow-up 2: b4, d11, g4, g7, h2, h10, and h11. h11: the following section was corrected: d11 impella cp, model 0048-0026, serial number: (b)(6), lot number 1324136; therapy date (b)(6) 2018 oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Section d11 information: impella cp, model 0048-0026 serial: (b)(6), lot: 1324139. one 14f, 25 cm long abiomed intruducer sheath was returned from the customer, peeled into two pieces, without the dilator. there were no other accessories returned. the end user shipped the device in two pieces to the customer. the customer cleaned the device, and removed the orange split caps from both sides of the product prior to return to oscor. a leak test could not be performed by oscor as the sheath was torn apart by the end user. the introducer sheath and hemostatic valve looked like it peeled apart normally. also, the pegs that hold the hemostatic valve in place looked normal. based on the condition the device was returned in, the cause for leakage is unknown. per procedure adelante s2s introducer sheath in-process and final inspection, vacuum leak test - perform pressure and vacuum leak test per procedure. this procedure is performed by qa 100%. as per instructions for use: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the reprt was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | This report is being submitted retrospectively per fda request during our final investigation, it was discovered the initial mdr 1035166-2018-00077 cmp (b)(6) report had failed submission. based upon our findings, we have corrected the reported failure in mdr initial mdr 1035166-2018-00077 cmp (b)(6) follow up 3 mdr submitted on 12/15/2018. oscor inc., is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc., which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor, inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor, inc., or its employees, that the report constitutes an admission that the device, oscor, inc., or its employees caused or contributed to the event described in this report. | This report is being submitted retrospectively per fda request. initial mdr (1035166-2018-00077) was submitted on august 23rd, 2018; however, submission failed as the d11 character was maxed out. follow-up mdr was submitted on august 27th, 2018. however, submission failed as the d11 character was maxed out. follow-up 2 mdr was submitted on august 28th, 2018, as remedial action to correct the initial and follow-up mdr errors. and this submission was passed. follow-up 3 mdr was submitted on december 18th,2018, and submission was passed. and on may 26th, we submitted the initial mdr, and the submission was passed. oscor inc., is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc., which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor, inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor, inc., or its employees, that the report constitutes an admission that the device, oscor, inc., or its employees caused or contributed to the event described in this report.
38005 | 1220648-2018-00087 | 7810929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7810929&pc=PYX | 07/25/2018 | 07/25/2018 |  | 08/23/2018 |  |  | 07/25/2018 | 08/23/2018 |  |  | 06/30/2019 |  | 08/23/2018 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1293940 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 74 YR | [5 empty] | Malfunction | No | Yes | Material Separation | [6 empty] | No Consequences Or Impact To Patient | No Code Available | [40 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2018 the physician was attempting to place an impella rp in the left femoral artery of a (b)(6) female patient. it was reported that when the impella's distal tip reached the poster anterior pulmonary artery the clinician noticed that the outflow cage and the pigtail were separated from the impella cannula. the remaining portion of the impella rp cannula was removed through the sheath, a snare was employed to pull the remaining outflow cage and pigtail through the sheath. another impella rp a used and that device was easily inserted without complication. the complainant reported that the patient was successfully supported with the second impella rp that was placed, and that there were no adverse effects to the patient as a result of the reported event. | [5 empty] | The impella rp was returned to the manufacturer for evaluation. during the visual examination of the returned pump, deformation and damage was found at the struts of the outflow cage, suggesting that a radically compressive force was applied to the struts at some point during the set-up and delivery of the pump to the pa. the cause of this load is unknown, but may have resulted from the tortuosity of the path from the femoral vein to the ivc. left femoral insertion is cautioned against in the impella rp's instructions for use. a review of the device history records for this device revealed that there had been no manufacturing issues for this pump related to this failure mode. there have been no other complaints for other pumps in this lot relevant to this failure mode.
38006 | 1220648-2018-00095 | 7801441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7801441&pc=OZD | 03/22/2018 | 08/15/2018 |  | 08/21/2018 |  |  | 03/22/2018 | 08/21/2018 |  |  | 01/31/2020 |  | 08/21/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1324139 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 80 YR | [5 empty] | Malfunction | No | Yes | Device-Device Incompatibility | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant was performing a tavr on a patient when the patient condition deteriorated and, so the impella cp was placed for hemodynamic support. shortly after implant of the cp pump, the pump's flow dropped and there were suction alarms. these alarms prompted the team to explant and place a new impella. the second pump supported the patient for a week until the pump could be weaned off and explanted. the patient is stable. | [5 empty] | The product analysis revealed that there was a possible interaction between the impella cp another device. the returned pump had both impeller blades broken. the pump did have suction when rerun in a water basin, confirming the failure mode. the root cause of the low flows was the broken blades. the failure mode will be monitored and trended. no corrective actions are warranted at this time.
38007 | 1220648-2018-00091 | 7801834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7801834&pc=OZD | 06/30/2018 | 07/26/2018 |  | 08/21/2018 |  |  | 07/26/2018 | 08/21/2018 |  |  | 09/30/2019 |  | 08/21/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1312629 |  | 005048-JP |  | No | [5 empty] | No | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemolysis | Patient Problem/Medical Problem | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient in (b)(6) had a syncopal episode and was found unconscious. the impella 2.5 was placed for support for the patient in cardiogenic shock. during the support, the patient had 10 units of blood product infused for treatment of signs of suspected hemolysis, per physician statement in the post marketing surveillance report (pms). the patient was on support for 68 hours and is reported to be recovered. the pump was explanted and no other support was needed. | [5 empty] | To date no product nor data have been returned. the event was reported in post marketing surveillance report. no other clinical details have been shared. should additional information be shared, that will point to the cause of the suspected hemolysis and blood product infusion, a supplemental report will be submitted.
38008 | 1220648-2018-00094 | 7788703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7788703&pc=OZD | 07/04/2018 | 07/21/2018 |  | 08/16/2018 |  |  | 07/21/2018 | 08/16/2018 |  |  | 03/31/2020 |  | 08/16/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1339788 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 83 YR | [5 empty] | Malfunction | No | Yes | Device-Device Incompatibility | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Hospitalization | [5 empty] | EDWARDS SAPIEN XT VALVE | [22 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | An impella cp was placed on a patient in the eu for support during a high risk cardiac angioplasty. after approximately 2 hours of support the pump flows diminished. pump remained on for support during angioplasty despite the flows being lower than expected. | [5 empty] | The (b)(6) facility returned the impella product and console data logs. the analysis of the data logs confirmed the drop in flows. upon investigation of the product, the pump was found to have a broken impeller blade, localized to one side. no other defects were found on the pump. the cause of the low pump flows was the broken impeller blade which was possibly caused by interaction with the tavr. no corrective action is warranted at this time. the failure mode will be monitored and trended.
38009 | 1220648-2018-00085 | 7773114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7773114&pc=OZD | 07/17/2018 | 07/17/2018 |  | 08/10/2018 |  |  | 07/17/2018 | 08/10/2018 |  |  | 06/30/2020 |  | 08/10/2018 |  | TEMPORARY NONROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1348368 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient was admitted for a cardiac angioplasty and had the impella cp placed via the left femoral artery. the patient developed a hematoma at the access site which was attributed to the cp re-positioning sheath. the patient had her impella cp removed and a second one placed to continue the patient support. in addition, the patient received four units of blood replacement. the second cp supported the patient successfully until wean and explant the following day. | [5 empty] | The facility returned the impella cp product for analysis. upon inspection there were two areas of damage observed at the 80 and 90cm marks. in addition, there was damage noted on the tip of the re-positioning sheath. the cause of hematoma was damage to the tip of the wound closure. the cause of this damage was unable to be determined. the failure mode will be monitored and trended.
38010 | 1220648-2018-00082 | 7773262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7773262&pc=OZD | 07/12/2018 | 07/13/2018 |  | 08/10/2018 |  |  | 07/12/2018 | 08/10/2018 |  |  | 11/30/2019 |  | 08/10/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1321428 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 51 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2018 an impella cp was successfully placed in a 51 year old male patient's femoral artery. following the placement of stents in the patient, the physician changed to the repositioning sheath, but heavy patient bleeding occurred at the side of the repositioning sheath. a vascular surgeon performed a tobacco-pouch suture technique to stem the bleeding. the patient was administered 3 to 4 units of replacement blood products. eventually the patient bleeding stopped and the patient continued to be supported by the impella cp. the pump was removed from the patient without issue on (b)(6) 2018 after a total of 93 hours of patient support. patient support was reported as successful. | [5 empty] | The impella cp has been returned for evaluation. the investigation has not yet begun; therefore, a supplemental medwatch report sill be submitted upon completion of the evaluation. | The impella cp, purge cassette, and white cable were returned to the manufacturer for evaluation. no leaks were found during the evaluation of the pump. a review of the device history record was performed. there were no other complaints reported that were associated with impella cp pumps from this lot, and no other complaints associated with repositioning units from this lot. the clinical account of this event described heavy bleeding from the side of the repositioning unit. this was confirmed during the inspection of the returned device, as blood was present around the hub at the juncture with the butterfly. integrin, a blood thinner that prevents clotting was administered. the impella cp's instructions for use recommends heparin usage in the purge solution. the instructions for use also recommends that act be maintained at 160 to 180 seconds after implant of the device. the leak was not reproduced in leak testing. in conclusion, the root cause of the access site bleeding was most likely as a result of improper anticoagulation related to the use of integrin and high resulting act, although this has been unable to be definitively determined.
38011 | 1220648-2018-00084 | 7757988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7757988&pc=OZD | 07/13/2018 | 07/13/2018 |  | 08/07/2018 |  |  | 07/13/2018 | 08/07/2018 |  |  | 07/31/2019 |  | 08/07/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1318240 |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report | [8 empty] | 81 YR | [5 empty] | Injury | Yes | No | Inadequacy of Device Shape and/or Size | Patient Device Interaction Problem | [5 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A patient was admitted in ventricular tachycardia with a history of a syncopal episode, and was found to have multivessel coronary artery disease. the team placed an intra-aortic balloon pump and called for bypass surgery. when the surgical team declined, the team placed an impella cp via the right femoral artery. due to the patient's obesity placement was very difficult and the patient developed a hematoma at the access site. the cardiologist asked for a vascular surgeon to consult on the hematoma, but the patient expired in pulseless electrical activity (pea) and active cpr before the hematoma could be addressed. | [5 empty] | The clinical report was analysed, but no product or data logs were returned for analysis. the clinical report did note the patient's obesity. obesity is listed as a caution for impella placement within the pump ifu. without more information the investigation findings and root cause of the groin complication are not available. should more details become available, a supplemental will be filed.
38012 | 1220648-2018-00080 | 7757990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7757990&pc=OZD | 07/05/2018 | 07/10/2018 |  | 08/07/2018 |  |  | 07/10/2018 | 08/07/2018 |  |  | 02/29/2020 |  | 08/07/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1336246 |  | 0048-0020 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Malfunction | Yes | No | Use of Device Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A critical patient was transferred with signs of hemolysis on impella cp support. at the second tertiary facility the patient had blood products, for the presumed gastrointestinal bleed, and dialysis for intervention to treat ongoing hemolysis. the patient was weaned off the impella support and the pump was explanted. at support's end the patient remained on cvvh dialysis. | [5 empty] | The product, data logs, and clinical account were analysed. the returned pump passed in-house hemolysis testing, and the failure mode could not be reproduced. since hemofiltration did not occur until the end of the case, the hemolysis may have persisted due to a lack of native kidney function. the failure mode will be monitored and trended.
38013 | 1220648-2018-00086 | 7759630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7759630&pc=OZD | 07/09/2018 | 07/09/2018 |  | 08/07/2018 |  |  | 07/09/2018 | 08/07/2018 |  |  | 04/30/2020 |  | 08/07/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP5.0 |  |  | OZD |  |  | IMPELLA LP5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1340134 |  | 005066-JP |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 37 YR | [5 empty] | Injury | Yes | No | Appropriate Term/Code Not Available | [6 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup | Followup |  |  |  | * | Invalid/Incomplete |  |  | On (b)(6) 2018 an impella lp5.0 was placed in a (b)(6) male patient. the patient was successfully supported for 8 days when the clinician found a sub-dermal hematoma located at the pump's right axillary insertion site. the patient was administered 4 units of red blood cells in response to an undisclosed amount of blood loss. there was no other intervention required for the hematoma, and no other issues reported with this device. the patient is still being successfully supported by this impella lp5.0. | [5 empty] | To date neither product, imaging, nor data logs have yet been returned for analysis, as it is reported that the patient is still under support of this impella lp5.0. a supplemental medwatch will be filed if the product and console logs are received and when investigation is completed. | The impella lp5.0, purge cassette and white cable were returned to the manufacturer for evaluation. during the inspection of the device blood was found around the cap of the repositioning sheath. the repositioning unit was cleaned and inspected for damages. no damages were observed to the distal tip of the sheath, although after cleaning damages and scratches were observed to the body of the repositioning unit sheath. the repositioning unit was tested for a leak in the leak test fixture. the tubing was filled with water and pressurized with the delta cal pressure transducer. the end of the end of the tube was fitted around the over molded region of the butterfly. the pump was tested at 120 mmhg and 180 mmhg. leak was checked with the sheath laying stationary in the fixture, with the butterfly bent to the side, and with the catheter bent to the side. no leaks were recorded in any configuration at either pressure. a review of the device history record did not identify any issues with the device during manufacture, inspection or testing. there have been other complaints associated with repositioning sheath units from this lot. in summary, the clinical account described heavy bleeding from the side of the repositioning unit was confirmed by the returned device, as blood was present around the hub at the juncture with the butterfly. integrin, a blood thinner that prevents clotting, was administered to the patient. the patient's act was noted to be extremely high. the heparin usage is unknown. the impella lp5.0 instructions for use (ifu) recommends heparin usage in the purge solution, as the impella lp5.0 has not been tested with other anticoagulants such as integrin. the ifu also recommends that act be maintained at 160 to 180 seconds after implant of the device. the returned device was cleaned and scratches and damages were found along the body of the wound closure of the repositioning unit, but none that extended past the surface level. the leak was not reproduced in leak testing. the root cause of the access site bleeding was most likely improper coagulation related to the use of integrin and the high resulting act. | This supplemental medwatch report is being submitted to correct supplemental medwatch report -01 which was submitted with erroneous evaluation results. this device was not returned to the manufacturer for evaluation as reported on supplement 01, as it was discarded by the user subsequent to use. the impella 5.0 was not returned for evaluation; therefore, the root cause of the access site bleeding was unable to be determined. review the device history record found no manufacturing issues or reworks associated with this pump. there are no other complaints associated with with pumps from this lot. review the device history record found no manufacturing issues or reworks associated with this pump. there are no other complaints with this failure mode associated with pumps from this lot.
38014 | 7713645 | 7713645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7713645&pc=PBL | 06/13/2018 | 06/25/2018 | 06/25/2018 | 07/24/2018 | 06/25/2018 | 07/24/2018 |  | 07/24/2018 |  |  |  |  | 07/24/2018 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PBL |  |  | NOT KNOWN |  |  | ABIOMED, INC. |  |  | US |  | NOT KNOWN |  |  |  | Device was returned to manufacturer | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 28470 DA | [5 empty] | Malfunction |  | Yes | Fluid/Blood Leak | Decrease in Pressure | [5 empty] | No Consequences Or Impact To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Patient's impella cp alarmed "low purge pressure". all connections both to the patient and to the pressure bag/purge fluid bag were checked to make sure there were no leaks or kinks. staff then noticed a small puddle under the impella controller, and upon opening the cassette flap noted that there was a leak in the cassette somewhere. it appeared to be leaking to the right of the cassette from the small circular membrane. staff quickly used the spare cassette and old purge fluid bag and switched out the cassette. no harm to the patient.
38015 | 1220648-2018-00073 | 7707061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7707061&pc=PYX | 06/22/2018 | 06/22/2018 |  | 07/20/2018 |  |  | 06/22/2018 | 07/20/2018 |  |  | 07/31/2019 |  | 07/20/2018 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1301336 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 59 YR | [5 empty] | Death | Yes | No | Appropriate Term/Code Not Available | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | On (b)(6) 2018 a (b)(6) year old male patient with post-cardiotomy cardiogenic shock had post aortic valve replacement surgery. the valve replacement surgery was difficult, and a va ecmo device was placed urgently. the ecmo was decannulated after supporting the patient for 24 hours, the patient was administered 2 units of packed red blood cells and an impella rp was placed in the patient. the physician did not use echo to evaluate for clotting prior to the rp insertion. it was reported that there was heparin in the purge solution from implant, but not systemically administered due to bleeding from the ecmo insertion site. during the night patient hemolysis occurred. the labs continued to indicate hemolysis, and the patient's urine output was black. under x-ray the pump appeared to be located just below the pulmonic valve. the staff was unable to get the heparin level up due to the hemolysis. the impella rp was unable to be repositioned, and on (b)(6) 2018 the impella rp was removed and a va ecmo was placed in the patient. patient care was withdrawn on (b)(6) 2018 and subsequently expired on that same date. | [5 empty] | The device was returned to abiomed for evaluation. the catheter was found to have been cut, which was done by the clinicians to maintain access for va ecmo insertion after the impella rp was explanted. this disallowed the ability to perform hemolysis testing of the returned impella rp. a thrombus was noted to be inside the pump housing on the impeller, but no other defects were found on the pump. the automatic impella controller's data logs were returned for this case. the analysis of the logs revealed that the pump had been run for a total of about 12 hours. there were no alarms that triggered related to the failure mode throughout patient support. there was elevated and changing placement signals throughout, which could correlate with the biomaterial ingestion. the changes in the placement signal and motor current during support indicated either changing position or interaction with patient vasculature or ingested biomaterial during this time. a review of the device history record revealed that there had been no issues found during the manufacture or inspection of the device that apply to this failure mode. in addition, there have been on complaints reported related to this failure mode. the root cause of the hemolysis was most likely thrombus ingestion, but this could not be definitively determined as hemolysis testing could not be completed on the pump. no corrective action is recommended at this time as the root cause of the hemolysis was unable to be definitively determined. failures of this type will continue to be monitored for trends.
38016 | 1220648-2018-00079 | 7703997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7703997&pc=OZD | 06/20/2018 | 06/20/2018 |  | 07/19/2018 |  |  | 06/20/2018 | 07/19/2018 |  |  | 01/30/2020 |  | 07/19/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1324651 |  | 005040 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 79 YR | [5 empty] | Death | Yes | No | Appropriate Term/Code Not Available | [6 empty] | Hemolysis | Blood Loss | [11 empty] | Death | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | Complainant reported that on (b)(6) 2018 an impella lp2.5 was placed in a (b)(6) patient who was in cariogenic shock and with unstable circulation. after impella lp2.5 placement the patient was reported to have bleeding problems from the eyes, ears and abdomen and hemolysis. the physician reported that there were discrepancies between act (activated clotting time) and the ptt (partial thromboplastin time) values. under high doses of heparin the patient had <200 sec act and ptt >200. on (b)(6) 2018 the patient received 7-8 units of replacement blood products at an hb (hemoglobin) reading of 5. the patient was bleeding massively in the abdomen, and expired on (b)(6) 2018. | [5 empty] | The impella lp2.5 was returned for evaluation as it was discarded by the user at the conclusion of patient support. as the product was not returned, no root cause can be determined.
38017 | 1220648-2018-00074 | 7704331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7704331&pc=OZD | 07/01/2018 | 07/01/2018 |  | 07/19/2018 |  |  | 07/01/2018 | 07/19/2018 |  |  | 04/30/2020 |  | 07/19/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1339655 |  | 0048-0003 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 78 YR | [5 empty] | Malfunction | No | Yes | Difficult to Insert | Inadequacy of Device Shape and/or Size | [5 empty] | Perforation | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A morbidly obese patient with a known cardiac history was transferred to the tertiary medical center for escalated care. the patient was actively arresting, and being treated by repeated shocks, when admitted to the cardiac catheterization lab. due to his condition the team attempted to place an impella cp via the right femoral artery and 14fr introducer. there was difficulty placing the introducer and upon inspection it was seen to be damaged with a kink and fraying. the patient's vasculature had been injured and so the team placed a covered stent to the right femoral artery and then placed the impella cp via the contralateral leg. in addition the team treated the patient's access site bleed with 3 units of blood replacement. the placement was successful on the left side. | [5 empty] | To date neither product, imaging, nor data logs have yet been returned for analysis. a supplemental medwatch will be filed when investigation has completed. | Since the initial medwatch was filed, the investigation h as completed. the product was not returned for investigation, only the clinical report was analyzed. due to the patient's body habitus, the placement was made at a steep angulation. the root cause of the re-positioning sheath damage, and associated blood loss, was most likely a combination of steep angulation, tortuous native anatomy, and operator error in the attempted force applied at insertion. no corrective action is warranted at this time. the failure mode will be monitored and trended for analysis.
38018 | 1220648-2018-00068 | 7699378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7699378&pc=OZD | 06/18/2018 | 06/18/2018 |  | 07/18/2018 |  |  |  | 07/18/2018 |  |  | 02/29/2020 |  | 07/18/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLLA CP |  |  | ABIOMED, INC |  |  | US |  | 1332857 |  | 0048-0003 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | Yes | Appropriate Term/Code Not Available | [6 empty] | Perforation of Vessels | Blood Loss | [11 empty] | Required Intervention | [5 empty] | ABBOTT VASCULAR PROGLIDE | [22 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that the physician were preparing for placement of an impella cp in a (b)(6) female patient. during the placement the left femoral artery was accessed. the impella cp was placed and ramped on. a femstop was applied to the left groin, and the oozing slowed. after about 10 more minutes the patient's pressure dropped and groin was again bleeding. manual pressure was applied. a distal aortagram was performed that revealed a patent vessel in the left groin common femoral with sluggish flow and no flow past the access site. the impella was removed due to no flow in the patient's left leg. the concern was that the proglide had failed and was occluding the artery and that a dissection was present. vascular surgery was consulted and the patient was taken to the or for a cut-down & vascular repair. the physician and staff reported that the patient would not have survived without impella support. | [5 empty] | The failure of the cannula detaching from the pump housing was unable to be confirmed because neither product nor console data logs were returned; no visual evidence was provided by the user facility of the failure for investigation. if additional information is obtained a supplemental report will be submitted.
38019 | 1220648-2018-00078 | 7693714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7693714&pc=OZD | 06/20/2018 | 06/20/2018 |  | 07/17/2018 |  |  | 06/20/2018 | 07/17/2018 |  |  | 08/31/2019 |  | 07/17/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1308677 |  | 005062 |  | No | [5 empty] | No | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemorrhage, Cerebral | Blood Loss | [11 empty] | Hospitalization | Required Intervention |  |  |  |  | ECMO | [22 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient was admitted in cardiogenic shock and supported by the impella for seven days, at which time the impella was explanted for cannulation of ecmo. eleven days later, the team placed a new, second impella via the axillary artery due to deteriorating condition. after four days the patient was bleeding from the axillary insertion site and showed neurological changes and right sided weakness. the team had the neurosurgeon intervene surgically to remove the clot. despite these efforts, the patient did expire when the family chose to withdraw care the following day. in review of the case the team did state the patient was on dual antiplatelet therapy for blood thinning. | [5 empty] | The impella pump was not returned for evaluation. there is no indication from the data logs that the use of the impella caused or contributed to the stroke. the report from the hospital team says that the stroke was caused by the combination of dual antiplatelet therapies as well as the simultaneous delivery of heparin. no corrective action is warranted. the failure mode will be monitored and trended.
38020 | 1220648-2018-00072 | 7693787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7693787&pc=OZD | 06/19/2018 | 06/21/2018 |  | 07/17/2018 |  |  | 06/21/2018 | 07/17/2018 |  |  | 03/31/2020 |  | 07/17/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1337540 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 57 YR | [5 empty] | Malfunction | No | Yes | Material Integrity Problem | [6 empty] | Low Blood Pressure/ Hypotension | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient was admitted for a high risk angioplasty (hrpci) and the team observed that at the access site there was excessive bleeding. the observation of the blood loss was at a split in the introducer placed at the left femoral artery. the introducer was placed as access for the impella cp. the team believes the patient to have lost 1 liter of blood. the blood loss prompted the ordering of blood replacement products and the treatment of hypotension with neo-synephrine boluses x 11. the patient remained on impella support, despite the blood loss, and the angioplasty was completed and the patient discharged from the catheterization lab without lasting harm or injury. | [5 empty] | The product was returned along with a photograph of the product in use at the hospital. the cause of the introducer leaking is a failure of the seal between the valve and the hub. the cause of the failure of this seal is unknown.
38021 | 1220648-2018-00081 | 7682118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7682118&pc=OZD | 06/12/2018 | 06/12/2018 |  | 07/12/2018 |  |  | 06/12/2018 | 07/12/2018 |  |  | 12/31/2019 |  | 07/12/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1325622 |  | 0048-0026 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | Human-Device Interface Problem | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient was admitted in cardiogenic shock, and while awaiting cardiac surgery he decompensated, which prompted the team to implant and impella cp via the left femoral artery. the support from the impella allowed the team to perform a valvuloplasty. during the valvuloplasty the patient condition was critical and so the cp pump was left in for support. on the second day of support the groin access site was oozing blood which prompted to the team to infuse 1 unit of blood products. the groin site continued to ooze but, the blood loss was remedied the following day when the impella was explanted. | [5 empty] | To date no product has been returned by the facility for investigation. no root cause of the groin access site bleed has been determined. should more information be shared, the manufacturer will file a supplemental medwatch report.
38022 | 1220648-2018-00076 | 7673027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7673027&pc=OZD | 06/12/2018 | 06/14/2018 |  | 07/10/2018 |  |  | 06/14/2018 | 07/10/2018 |  |  | 02/29/2020 |  | 07/10/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1341080 |  | 005042 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 46 YR | [5 empty] | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | An impella 2.5 was attempted to be inserted for support for hrpci. during insertion the pump kinked, requiring the use of a snare for removal of the device. | [5 empty] | The pump was returned for evaluation, however the investigation is still underway. | Since the initial filing of the medwatch, the investigation has completed. the kinking of the pump cannula was due to wireless access attempts by the user. the introducer then buckled, when the excessive force was applied by the user, to remove the kinked pump from the patient. the snare was then utilized to remove the product from the patient when it became stuck within the introducer. the failure is of very low occurrence and will be monitored and trended.
38023 | MW5078289 | 7673330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7673330&pc=OJE | 06/21/2018 | 07/02/2018 |  | 07/09/2018 |  |  |  | 07/09/2018 |  |  |  |  | 07/09/2018 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA R. P. DEVICE |  |  | OJE | [5 empty] | UNK | [8 empty] | Device was returned to manufacturer | [5 empty] | Y | No Information |  |  | Voluntary report | [8 empty] | 59 YR |  |  | 72 |  |  | Injury | Yes | No | Pumping Problem | [6 empty] | Aortic Valve Stenosis | Calcium Deposits/Calcification | Death | Hemolysis | Heart Failure | Multiple Organ Dysfunction Syndrome | [36 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Massive hemolysis, no right heart support despite flor readings of 3 liters per minute - worsening shock. pt was receiving from complex aortic root surgery and was on partial right heart support with ecmo (extra corporeal membrane oxygenation), for 18 hrs and slowly improving. impella rp inserted in the cath lab while ecmo was wearing. impella rp read flows of 2.8 to 3.1 liters per minute. however despite good lv function. the left ventricle was not filling (suggesting that impella rp was not working). lactate kept going up. blood showed hemolysis. ldh kept going up. urine turned black and stopped. as an emergency, the impella rp was taken out and ecmo re-instituted. when the impella rp device was shut off, and cut, there was no change to hemodynamics. everything got better with ecmo. however, 8 hours of profound hemolysis and negative effects on the right heart caused kidney, liver and multi - organ failure and death. complex aortic root reconstruction 2 days before impella rp came out, performed for a very heavily calcified bicuspid aortic valve with severe stenosis and a narrow aortic root. s/p aortic valve replacement and aortic root reconstruction. dates of use: 15.30 hrs, (b)(6) 2018; 0230 hrs, (b)(6) 2018. diagnosis or reason for use: right heart failure. "is the product compounded: no; is the product over-the-counter: no." event abated after use stopped or dose reduced: yes. abiomed, inc.
38024 | 1220648-2018-00077 | 7666938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7666938&pc=OZD | 07/04/2018 | 07/04/2018 |  | 07/06/2018 |  |  | 07/04/2018 | 07/06/2018 |  |  | 12/31/2018 |  | 07/06/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1272923 |  | 0048-0002 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 78 YR | [5 empty] | Death | Yes | No | Use of Device Problem | Human-Device Interface Problem | [5 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient was undergoing a coronary intervention, and due to an intraprocedural arrest, the patient had an impella cp placed via the right femoral artery. the patient did have known peripheral arterial disease, but the pump was placed. the cp was supporting the patient in the icu for hours, despite his respiratory distress, renal function being depressed, and beginning of multi-organ failure. during the overnight hours, the patient condition deteriorated further and a hematoma developed in the cp access side. the team began to infuse blood product, but the patient expired with a presumed retroperitoneal bleed and multi-organ failure, along with cardiogenic and hemorrhagic shock. | [5 empty] | The complainant in the eu did not return the product for analysis. when more data and information is shared with the manufacturer, a supplemental medwatch will be reported.
38025 | 1220648-2018-00067 | 7665216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7665216&pc=OZD | 06/07/2018 | 06/07/2018 |  | 07/05/2018 |  |  | 06/07/2018 | 07/05/2018 |  |  | 09/30/2019 |  | 07/05/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1312629 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Appropriate Term/Code Not Available | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2018 an impella lp2.5 was placed in a (b)(6) male patient. the following day patient hemolysis occurred. the impella was repositioned under tte, but the hemolysis did not improve. plasma free hemoglobin was drawn twice confirming the hemolysis; therefore an ecmo was placed. the physician performed repositioning of the pump after ecmo implant, but the patient's hemolysis was not improved. the patient was administered replacement red blood cells (amount of blood administered undisclosed) on both (b)(6) 2018. the impella lp2.5 was removed on (b)(6) 2018 after 44.55 hours of patient support. the patient's hemolysis resolved and the patient was implanted with a vad. | [5 empty] | The impella lp2.5 and aic console data logs were returned for evaluation. the evaluation has not yet been completed. upon the completion of the device evaluation and analysis of the data logs a supplemental medwatch will be submitted. | The returned console data logs and impella lp2.5 were returned for analysis. the analysis of the data logs revealed that the pump was out of position when patient support was first initiated. the position was adjusted and the pump ran without issue for the first day of support. hemolysis was first reported on the following day. no suction alarms were triggered prior to the first observation of hemolysis. the log analysis revealed that a high motor current pump stop occurred on the third day of support. the logs show that the pump migrated out of position leading to the pump stop, at which point biomaterial may have been ingested. after the pump was restarted, the support level was lowered and ecmo support was added. after this change in support, the pump exhibited atypical motor current behavior, with the motor current appearing non-pulsatile with intermittent spikes. the spikes were aligned with spikes in rpm. this is likely attributable to biomaterial present in the pump after the high motor current pump stop. the evaluation of the impella lp2.5 revealed biomaterial around the impeller of the pump. it's likely that this biomaterial was ingested leading to the high motor current pump stop which occurred after the first report of hemolysis. the impeller was slightly eccentric n the pump housing, but no scratching to the housing was observed. the pump was tested twice for hemolysis, and failed to meet specification in both tests. further inspection of the device then revealed delamination of the epoxy from the sleeve bearing. there is a known correlation between epoxy delamination and hemolysis. in conclusion, the root cause of the hemolysis was the delamination of the epoxy from the sleeve bearing. contributing factors to this condition include the biomaterial build up around the impeller, and the eccentricity of the impeller in the pump housing.
38026 | 1220648-2018-00063 | 7665221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7665221&pc=OZD | 06/06/2018 | 06/06/2018 |  | 07/05/2018 |  |  | 06/06/2018 | 07/05/2018 |  |  | 04/30/2020 |  | 07/05/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1340185 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Missing Value Reason | [6 empty] | Hemolysis | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | The complainant reported that following the removal of an impella cp (medwatch report number 1220648-2018-00054) another impella cp was placed at the right femoral artery in a (b)(6) female patient. during this support the physician reported concerns from patient hemolysis, resulting in a ct scan being performed which revealed a retroperitoneal hematoma. due to the continued need for the administration of blood products (11 units) as a result of issues of bleeding from the ecmo site and a hematocrit reading of 21 the patient was taken to the or for exploration of bilateral iliac cut-downs from a previous surgery. the physician remained concerned with the patient's hemolysis and the patient's increased needs for pressors. as a result the impella cp was removed from the patient after 126.10 hours of support and the site was surgically closed. an impella lp5.0 was then placed in the patient without issue. the patient was reported to be successfully supported with the impella lp5.0. | [5 empty] | The device has been returned for evaluation, the evaluation is in process. upon completion of the evaluation, a supplemental medwatch report will be submitted. | The device was returned for review, and showed no indication of damage or biomaterial buildup. the pump was tested for hemolysis and was found to be within specification. the data logs show that the pump experienced suction and improper positioning alarms throughout the case, indicating that positioning issues occurred intermittently throughout the case. positioning issues resulting in suction alarms are most likely the cause of the hemolysis.
38027 | 1220648-2018-00075 | 7656543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7656543&pc=OZD | 12/20/2017 | 06/21/2018 |  | 07/02/2018 |  |  | 06/21/2018 | 07/02/2018 |  |  | 04/30/2019 |  | 07/02/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | 138129 |  |  | ABIOMED, INC. |  |  | US |  | 1300756 |  | 005048-JP |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 83 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An adverse event was reported to abiomed through the postmarket surveillance report ((b)(6)). an impella 2.5 was inserted for support into a patient with heart failure for pccs in (b)(6) 2017. the patient experienced hemolysis during the case, and received blood replacement products. the patient expired. | [5 empty] | The product was not returned from the user facility for evaluation, therefore a full investigation was unable to be conducted. failures of this type will continue to be monitored for trends.
38028 | 1220648-2018-00064 | 7652620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7652620&pc=OZD | 06/05/2018 | 06/08/2018 |  | 06/29/2018 |  |  | 06/08/2018 | 06/29/2018 |  |  | 09/30/2019 |  | 06/29/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1312629 |  | 005048-JP |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | After three days of successful support the impella 2.5 was removed from the patient in the (b)(6) facility. upon removal the patient had a manual hold to achieve hemostasis. after an hour the limb was noted to be cold. the patient was treated by a fogarty embolectomy and thrombus was removed. the flow to the limb was restored and there as no lasting harm or injury. | [5 empty] | The pump product was not returned for analysis from the (b)(4) facility. the data logs were also not returned, nor was imaging as requested, sent for analysis. the root cause of the reported limb ischemia, that prompted embolectomy, could not be determined. the failure mode will be monitored and trended. should more information be shared, a supplemental will be drafted.
38029 | 1220648-2018-00069 | 7652753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7652753&pc=OZD | 06/07/2018 | 06/07/2018 |  | 06/29/2018 |  |  | 06/07/2018 | 06/29/2018 |  |  | 02/29/2020 |  | 06/29/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1330432 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 73 YR | [5 empty] | Malfunction | Yes | No | Difficult to Insert | Kinked | [5 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A patient post-operative from a amplatz plug placement for embolization, had an impella cp placed via the right femoral artery. shortly after placement, the pump came out of position in the left ventricle. during re-positioning attempts the cannula of the pump kinked. the team attempted to re-wire and remove the kinked pump. during these attempts the patient lost blood and had one unit of blood replacement. the team was unable to access again via the right femoral artery, and instead placed the impella in the left femoral and support was successful. support was successful for over 244 hours. | [5 empty] | The impella pump product was returned for evaluation and analysis. the clinical account from the physician indicates that the outflow cage was positioned outside the body while attempting to insert a new guidewire to regain access. this allows a clear path for bleeding from the site, resulting in the loss of one unit of blood via the positioning of the outflow cage outside of the body during attempted rewiring. no corrective action is warranted and the failure mode will be monitored and trended.
38030 | 7645251 | 7645251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7645251&pc=KFM | 05/29/2018 | 06/25/2018 | 06/25/2018 | 06/28/2018 | 06/25/2018 | 06/28/2018 |  | 06/28/2018 |  |  |  |  | 06/28/2018 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS |  |  | IMPELLA |  |  | KFM |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 155522 |  |  |  | Device was returned to manufacturer | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 25185 DA | [5 empty] | Malfunction | No | Yes | Disconnection | Device Inoperable | [5 empty] | Low Blood Pressure/ Hypotension | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Patient was ambulating back to bed from the chair when the impella power cord came disconnected at the red/grey connector site with arrows. staff noticed right away, informed the patient's nurse and quickly reconnected the two connections - however when interviewed after the event the staff was not absolutely certain that the arrows were aligned. the patient was placed into bed and his pressure immediately began to drop, requiring titration of his pressors. upon assessment by the team, the impella was not working. the patient continued to decline and ultimately patient was intubated and cannulated for ecmo as a result of the impella not functioning as intended and staff inability to find a solution in the moment.
38031 | 1035166-2018-00065 | 7648109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7648109&pc=DYB | 06/21/2018 | 06/28/2018 |  | 06/28/2018 |  |  | 06/25/2018 | 06/28/2018 |  |  | 10/01/2019 |  | 06/28/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC |  |  | US |  | C1-12828 |  | 0052-3000 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 79 YR | [5 empty] | Malfunction | No | Yes | Crack | [6 empty] | No Consequences Or Impact To Patient | [41 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | It was reported that an incident which involved a 14fx13 cm introducer occurred during a planned high risk percutaneous coronary intervention procedure on (b)(6) 2018. during insertion at the left femoral artery, after dilations with 8f, 10f, and 12f the 14fr introducer cracked. the physician conducted the first three dilations, put the 14fx13 introducer over the .035 stiff wire, it cracked, was removed and a new 14fx13 introducer was placed. per the reporter, the tip of the sheath is what split down and along the length to the middle portion. there was roughly 30 cc of blood lost during the switch to 2nd replacement 14fx13 oscor sheath. patient is reported to be stable and no extended hospitalization was required. the product was discarded. | [5 empty] | A review of device history record (dhr) was conducted and there were no manufacturing rejects or anomalies recorded in the dhr affecting released product. the introducer sheath passed all in-process and qa final inspection steps before shipping to the customer including visual, mechanical and dimensional testing. a review of complaints involving the reported lot identified one additional report for a similar event. per the quality assurance abiomed introducer sheath in-process and final inspection procedure it states: using 10x magnification, look down into hub for irregularities; verify no splitting at break lines and a smooth transition between sheath and hub on inside of hub. verify score lines are spaced evenly around the circumference of sheath tube; look down into the hub and verify score lines on sheath align with break line of the hub. using 10x magnification, verify the tip is round, no flash, no splitting, cracks or other damages. insert a tipped dilator through the tipped sheath and check if dilator inserts without excessive force or obstruction. verify split caps are properly placed and secured onto sheath hub and free of cracks/damages or excessive adhesive. perform pressure and vacuum leak test and perform occlusion test. the abiomed adelante s2 hemostatic peel-away introducer system with infusion sideport instructions for use (ifu) lists the following: precaution - never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. the ifu lists the following adverse effects and possible complications: air embolism, bleeding, hematoma formation and vessel damage. as the device was not returned for analysis, the allegations and clinical observations cannot be confirmed and a root cause of the reported failure could not be determined. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
38032 | 1220648-2018-00055 | 7641561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7641561&pc=OZD | 05/27/2018 | 05/29/2018 |  | 06/27/2018 |  |  | 05/29/2018 | 06/27/2018 |  |  | 12/31/2019 |  | 06/27/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1321944 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 58 YR | [5 empty] | Malfunction | No | Yes | Appropriate Term/Code Not Available | [6 empty] | Thrombus | [41 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that an impella cp was used to support a (b)(6) male patient for ami/cgs on (b)(6) 2018. the pump was inserted without issue and functioned properly. on (b)(6) 2018 the impella cp was explanted without complication. after the device was explanted 2 large thrombus were removed from patient's right femoral artery (rfa) prior to closure. the first thrombus was located at the insertion site, with the 2nd located proximal to the insertion site. no clots or material seen on the catheter or the pump itself. both thrombi were successful aspirated from the vessel by the surgeon prior to site closure. strong pulses where confirmed in right leg/foot with doppler post-procedure. patient support was reported as successful, with no adverse effects to the patient. | [5 empty] | The automatic impella data logs were returned for analysis. the analysis of the logs showed that the pump functioned properly without issue for the entire case. the logs revealed that heparin was not used in the purge in the first 5 hours of the pump's patient support. the heparin was then added to the purge with 50 ui/ml in d5. approximately 4.5 hours later the heparin dose decreased to 25 ui/ml, and remained in the purge system with that dose for 7 hours. the next, heparin dose increased to 40 ui/ml for 5.5 hours, and was then increased to 50 ui/ml, and stayed on that level to the rest of the case. the root cause of the reported thrombus has been determined to be due to the patient's condition, and the inability to maintain therapeutic act levels to the recommended range. there were no complaints related to this failure mode for other pumps in this pump lot.
38033 | 7637383 | 7637383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7637383&pc=KFM | 12/19/2017 | 05/31/2018 | 06/13/2018 | 06/26/2018 | 06/13/2018 | 06/26/2018 |  | 06/26/2018 |  |  |  |  | 06/26/2018 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA |  |  | KFM | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 23360 DA |  |  | 87 |  |  | Malfunction | No | Yes | Occlusion Within Device | [6 empty] | Thrombosis | [41 empty] | Initial submission | [5 empty] | Y | OPERATING ROOM |  |  | Impella 5.0 was placed approximately 7 months ago. it required replacing by the afternoon the following day due to thrombosed catheter. a new 5.0 impella was placed via same entry--right axillary.
38034 | 1220648-2018-00066 | 7638013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7638013&pc=OZD | 05/26/2018 | 05/28/2018 |  | 06/26/2018 |  |  | 05/28/2018 | 06/26/2018 |  |  | 04/30/2019 |  | 06/26/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LP2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1300756 |  | 005048-JP |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | Yes | Appropriate Term/Code Not Available | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | Lower limb ischemia occurred after the implant of the impella lp2.5 via the right femoral artery. the impella lp2.5 was removed from the patient due to the lower limb ischemia. the ischemia did not improve after the pump's removal from the patient. the patient subsequently expired due to myonephropathic metabolic syndrome (mnms.) | [5 empty] | The impella lp2.5 was not returned for evaluation. a review of the device history did not reveal any nonconformities.
38035 | 1220648-2018-00053 | 7638470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7638470&pc=OZD | 05/29/2018 | 05/29/2018 |  | 06/26/2018 |  |  | 05/29/2018 | 06/26/2018 |  |  | 03/31/2020 |  | 06/26/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1333928 |  | 0048-0026 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Human-Device Interface Problem | [6 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention |  |  |  |  | INTRA-AORTIC BALLOON PUMP | [22 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A patient was admitted for an aortic valvuloplasty and had the impella cp placed for hemodynamic support. upon placement of the cp the patient's previously placed intra-aortic balloon pump was shut off and the patient was sent to the icu for monitoring. while admitted to the icu the patient decompensated. the medical team attributed the decline to a possible retroperitoneal bleed. the bleed was thought to have occured from the impella access, but the team did not consult vascular surgery to diagnose nor treat the bleed. the patient later expired. | [5 empty] | The product was not returned for analysis. the data logs do not pertain to the failure mode of a bleed. the root cause of the suspected retroperitoneal bleed can not be determined.
38036 | 1220648-2018-00054 | 7638518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7638518&pc=OZD | 05/29/2018 | 05/31/2018 |  | 06/26/2018 |  |  | 05/31/2018 | 06/26/2018 |  |  | 02/29/2020 |  | 06/26/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1336295 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Human-Device Interface Problem | [6 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The patient had a mitraclip procedure for mitral valve disease, decompensated, then an impella cp was placed emergently. after implant the patient was admitted to the icu for monitoring, at this time the patient had a hematoma develop. due to the patient's instability, the decision was made to bring her back to the or to repair the groin access site and infuse 4 units of blood product to the patient. | [5 empty] | The impella product was returned for analysis. the data logs were not returned but, they would not pertain to the failure mode and analysis. the impella was inspected and no defect was found. the root cause of the hematoma was unable to be determined. the returned repositioning unit showed no signs of damage. the manufacturing review found no similar complaints made against the lot of pumps. the failure mode will be monitored and trended.
38037 | 1220648-2018-00061 | 7638539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7638539&pc=OZD | 06/05/2018 | 06/07/2018 |  | 06/26/2018 |  |  | 06/07/2018 | 06/26/2018 |  |  | 04/30/2020 |  | 06/26/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1340623 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 79 YR | [5 empty] | Malfunction | No | Yes | Break | Difficult to Remove | Material Separation | Device Handling Problem |  |  |  | Blood Loss | Patient Problem/Medical Problem | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The patient admitted in cardiogenic shock had an impella placed for support as he coded during a gi study. the patient was supported successfully and upon explant of the pump, the catheter broke. the vascular surgery team was called to remove the portion of the pump that had remained within the vasculature. the vascular surgeon noted calcified peripheral vasculature that he believed may have contributed to the breakage of the pump. | [5 empty] | The pump product in its totality was returned for analysis. both portions of the pump were analyzed. the root cause of the cannula detachment was excessive force applied to remove the device. the reason for the need for excessive force was unable to be determined. there was no information relayed regarding where the device stuck. the failure mode will be monitored and trended. no corrective action is warranted at this time. medwatch mw5077873, (b)(4).
38038 | 1035166-2018-00060 | 7633882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7633882&pc=DYB | 06/01/2018 | 11/08/2018 |  | 06/25/2018 |  |  | 06/12/2018 | 06/25/2018 |  |  | 01/01/2021 |  | 06/25/2018 |  | INTRODUCER, CATHETER, |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC |  |  | US |  | C1-13998 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | [11 empty] | Injury | Yes | Yes | Break | [6 empty] | Myocardial Infarction | Cardiogenic Shock | Blood Loss | [10 empty] | Required Intervention | [28 empty] | Initial submission | Followup | Followup |  |  |  | N | Invalid/Incomplete |  |  | An incident was reported that occurred on (b)(6) 2018 during an angioplasty, also called percutaneous coronary intervention (pci) on the rca (right coronary artery). during placement of the impella cp, while advancing the oscor 14f x 25 sheath into the groin the valve broke, causing significant patient bleeding, resulting in the administration of replacement blood products. no additional device was used as once sheath peeled away, groin was okay. patient was reported to be in stable condition. additional information has been requested, however not received as of date of this report. | [5 empty] | A preliminary review showed there were no manufacturing rejects or anomalies of this event type recorded in the device history record. a conclusion is not yet available as the investigation is on-going. additional information has been requested, however not received as of date of this report. | One 14f, 13cm long abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. both of the orange split caps were detached from the sheath's hub. visible blood was found on and inside the sheath. both of the orange split caps were detached from the sheath's hub. when viewed under a 10x microscope, the following results were found: hub side a: very minimal adhesive. hub side b: minimal sporadic adhesive. orange cap 1. some adhesive placed incorrectly. it is flaky up the side wall and there is more adhesive on the tab portion. orange cap 2: very minimal adhesive. the orange split caps were separated from the hub of device. the device appeared to have had a minimal to sporadic amount of adhesive on the hub and orange cap tabs. the device history record was reviewed to confirm that the device passed all applicable in-process and final inspections. based on the returned device, it appears that the lack of glue on the hub/orange cap tabs could have contributed to this failure. capa was opened to investigate the root cause and to implement corrective action to prevent recurrence of this issue. per procedure adelante s2s introducer sheath in-process and final inspection: with naked eye at a distance of 12" to 18", verify the sheath hub and split caps are free of excess silicone oil applied between the seals during assembly. verify split caps are properly placed and secured onto sheath hub and free of cracks/damages or excessive adhesive. additionally, destructive testing is performed per sampling plan. inspector is to "insert dilator into sheath, lock dilator hub lock nut. hold dilator hub with one hand and one of the sheath hub wing with the other hand. bend the dilator hub to the side until dilator hub lock nut is disconnected from the sheath hub. split caps should remain in place. there should be no damage and fracture to split cap and tabs. per ifu : never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | The following sections were updated in follow-up. date of report,pma/510k,if follow-up, what type. one 14f, 13cm long abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. both of the orange split caps were detached from the sheath's hub. visible blood was found on and inside the sheath. note: the introducer sheath was received from the field to the customer with only 1 orange split cap detached. in a conversation with the customer, it was revealed that they removed the other orange split cap. based upon this, we cannot determine which orange split cap was the one that the end user had come off of the device.varying amounts of glue were observed, but due to the area where the glue is placed, and due to handling prior to oscor receiving the product, it is difficult to quantify the amount of glue. additionally, the end user did not state how the dilator was removed. if the dilator was not twisted off of the sheath, it is possible that use error could have contributed to this failure mode. the ifu instructs the user to unscrew the dilator from the threaded caps prior to removing. capa has been previously opened to investigate the root cause and to implement corrective action to prevent recurrence of this issue and as per our preliminary investigation user error (deviation from the instruction for use) determined as probable cause. per ifu : cautions: prior to use, read all package inserts, warnings, precautions, and instructions. failure to do so may result in severe injury or death. directions for use: 9. insert vessel dilator into sheath and rotate the dilator cap over valve housing to secure the dilator onto sheath assembly. 10. thread the dilator/sheath assembly over the guidewire. 11. advance the dilator and sheath together with a twisting motion over the guidewire and into the vessel. fluoroscopic observation may be advisable. attaching a clamp or hemostat to the proximal end of the guidewire will prevent inadvertently advancing the guidewire entire into the patient. 12. once assembly is fully introduced into the vascular system, separate the dilator cap from the sheath valve housing by rotating the dilator cap off the hub. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
38039 | 1220648-2018-00059 | 7613535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7613535&pc=OZD | 05/27/2018 | 05/27/2018 |  | 06/19/2018 |  |  | 05/27/2018 | 06/19/2018 |  |  | 02/29/2020 |  | 06/19/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1336246 |  | 0048-0020 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 42 YR | [5 empty] | Injury | Yes | No | Difficult To Position | [6 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A patient experiencing a myocardial infarction had an impella cp placed. later, after having a ct scan a hematoma was observed at the cp insertion site. as the team was moving the patient, the cp moved out of position and the team therefore replaced the pump. in addition to a new pump, the team infused blood product to the patient to treat the hematoma. the patient had blood infusion and manual compression to the site, and the patient's hematoma resolved. | [5 empty] | The root cause of the hematoma was unable to be determined. when the impella was returned there was no signs of damage to the pump. (b)(4).
38040 | 1220648-2018-00070 | 7614725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7614725&pc=OZD | 05/26/2018 | 05/26/2018 |  | 06/19/2018 |  |  | 05/26/2018 | 06/19/2018 |  |  | 04/30/2020 |  | 06/19/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1337765 |  | 0048-0032 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | CONSUMER | HEALTH PROFESSIONAL |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella cp was inserted for support for pre-op cabg. three hours after insertion of the device, the patient developed a large hematoma at the access site. a vascular cut down was performed to remove the device, and a 16 fr introducer was inserted to control the bleeding. the patient received 1 unit of blood replacement products two days after the device was removed. | [5 empty] | The impella cp was discarded by the user facility, therefore a full investigaion could not be conducted. a review of the clinical case revealed the injury was due to damage of the external iliac artery near the inguinal ligament. the cause of the injury is unabe to be determined because the device was unable to be evaluated. a review of the device history did not reveal any nonconformaties.
38041 | 1220648-2018-00071 | 7614798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7614798&pc=OZD | 05/23/2018 | 05/23/2018 |  | 06/19/2018 |  |  | 05/23/2018 | 06/19/2018 |  |  | 01/31/2020 |  | 06/19/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1328713 |  | 0048-0032 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella cp was inserted for support for acute myocardial infarction and cardiogenic shock. while on support the patient 's distal pulses were absent in the leg with impella access, requiring opening the vessel with a balloon to improve flow. the patient was reported to be in stable condition following explant of the device. | [5 empty] | The impella cp was discarded by the user, therefore a complete investigation could not be conducted. should the manufacturer receive additional information a follow up report will be sent.
38042 | 1035166-2018-00056 | 7611196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7611196&pc=DYB | 05/18/2018 | 07/03/2018 |  | 06/18/2018 |  |  | 05/22/2018 | 06/18/2018 |  |  | 03/01/2021 |  | 06/18/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3014 |  |  | OSCOR INC |  |  | US |  | C1-14098 |  | 0052-3014 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | [5 empty] | 55 YR | [5 empty] | Malfunction | No | Yes | Crack | [6 empty] | No Consequences Or Impact To Patient | [41 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | It was reported that during a hrpci (high risk percutaneous coronary intervention) procedure, on (b)(6) 2018, an introducer cracked after insertion. the issue was identified during the procedure as follows: upon removal of the dilator, bleeding was noted from around the sheath. after removing the sheath, a crack in the 14f 13cm introducer was found. amount of blood loss is unknown at this time. patient is reported to be stable, no extended hospitalization was required. the devices are not available to return as they were discarded. although requested, additional information has not been received as of date of this report. | [5 empty] | A preliminary review showed there were no manufacturing rejects or anomalies of this event type recorded in the device history record. a conclusion is not yet available as the investigation is on-going. | A review of device history record (dhr) was conducted and there were no manufacturing rejects or anomalies recorded in the dhr affecting released product. the released device passed all in-process and qa final inspection steps before shipping to the customer. a complaint review of the reported lot found no additional reports involving this lot number. per the adelante s2s introducer sheath in-process and final inspection procedure, it states the following regarding visual inspection: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded foreign material or any other damages. in the instructions for use (ifu) it indicates: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. dilators, catheters and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. the device was in use for treatment. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
38043 | 7604883 | 7604883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7604883&pc=PBL | 06/06/2018 | 06/11/2018 | 06/11/2018 | 06/15/2018 | 06/11/2018 | 06/11/2018 |  | 06/15/2018 |  |  |  |  | 06/15/2018 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | PBL |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | Yes | [5 empty] | No | Physician |  |  | User Facility report | [14 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The purge pressure on the impella became elevated (>1000), tubing was checked for kinks, none were found. the impella rep was called and we were advised to change the impella cassette. impella cassette was changed with no resolution of the alarm. the impella rep was called a second time, and he advised us to turn the p level down to zero for 10 seconds. this was done with the supervision of the attending physician. the purge pressure remained high after this intervention as well. it was attempted to aspirate from the purge line, and still the purge pressure remained elevated. the decision was made to exchange the impella in the cath lab.
38044 | 1220648-2018-00065 | 7606673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7606673&pc=OZD | 05/16/2018 | 05/16/2018 |  | 06/15/2018 |  |  | 05/16/2018 | 06/15/2018 |  |  | 01/31/2020 |  | 06/15/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1329565 |  | 0048-0020 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Patient Problem/Medical Problem | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A patient admitted with acute mi and cardiogenic shock had angioplasty and intraaortic balloon pump (iabp) placed. as the patient decompensated the iabp was replaced by the impella. after five days of support the patient had a cardiac bypass (cabg) surgery done with the impella left in the left ventricle in surgical mode. during the cabg the team observed air within the left ventricle and the patient had a neurological event. the cause of the stroke was not known. two days after the cabg the patient had his impella explanted. the patient had been successfully hemodynamically supported for 7 days. | [5 empty] | The pump product and data logs were returned for analysis. the manufacturing history of the pump was analyzed and there have been no other complaints made against the lot of impella pumps. the root cause of the air in the patient's left ventricle, and associated stroke, was unable to be determined. whether the surgical team introduced the air, or not, due to user error is unknown and can not be confirmed. no corrective action is warranted at this time. the failure mode will be monitored and trended.
38045 | 1035166-2018-00055 | 7607737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7607737&pc=DYB | 05/06/2018 | 10/16/2018 |  | 06/15/2018 |  |  | 05/17/2018 | 06/15/2018 |  |  | 07/01/2020 |  | 06/15/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED INTRODUCER KIT FOR IMPELLA 23F |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | C1-13437 |  | 0052-3006 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN | HEALTH PR |  |  |  | 60 YR |  |  | 90 |  |  | Death | Yes | No | Leak/Splash | [6 empty] | Death | Ischemia | Myocardial Infarction | Cardiogenic Shock | Blood Loss | Multiple Organ Dysfunction Syndrome | [7 empty] | Death | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | The patient was presented with acute mi and cardiogenic shock. the impella was attempted to be placed on (b)(6) 2018; insertion at the right axillary for placement of the impella 5.0. at placement the team observed a bleed due to problems with the valve of the introducer (aix kit); the valve was leaking. the device appeared fine but started to leak in the cross opening of the valve. the leak was detected as soon as the first guidewire and diagnostic catheters were inserted. one 5.0 impella was inserted through the valve, but both 6 and 8 french introducers were used and passed through the valve to try to contain the bleeding. the patient loss 500cc of blood. the device was not tested outside of the patient, prior to use in the surgery. the impella 5.0 serial number (b)(4) was explanted sunday (b)(6). due to ischemia of the leg. the patient died of multiple organ failure according to treating physician, the impella supported the patient and was not was a contributing factor to the patients poor outcome. specific date of death unknown, however patient expired between (b)(6) 2018. | [5 empty] | Country of incident occurrence - (b)(6) . | Although requested specific date of death, it was not provided. patient expired between (b)(6) 2018. country of incident occurrence - (b)(6). the device was used in treatment. the device was discarded, therefore will not be returned for analysis; as a result, the clinical observation cannot be confirmed. a review of the device history record (dhr) was conducted and there was no manufacturing rejects or anomalies recorded in the dhr affecting released product. a complaint review found no additional complaints. per quality assurance procedure (abiomed introducer sheath) in-process & final inspection: device is tested one hundred percent for cap and seal inspection, a leak test is perform, also an occlusion test is perform. sideport with valve assembly samples are randomly selected for visual inspection: with the naked eye, inspect sideport assembly for excessive glue. smeared glue to the sideport assembly is not acceptable. verify the sideport tube is properly bonded to hub and stopcock valve by gently tugging on both joints. the instructions for use, abiomed adelante s2 - hemostatic peel away & sideport. under precaution: aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. aspirate all air from the sheath valve assembly by using a syringe connected to the sideport. flush the sheath with 5 cc of saline immediately before peeling the sheath away in order to minimize back bleeding. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
38046 | MW5077873 | 7605184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7605184&pc=KFM | 06/05/2018 | 06/11/2018 |  | 06/14/2018 |  |  |  | 06/14/2018 |  |  |  |  | 06/14/2018 |  | PARTIAL CIRCULATORY SUPPORT |  |  | ABIOMED IMPELLA HEART PUMP |  |  | KFM |  |  | IMPELLA CP0048-0003 |  |  | ABIOMED |  |  | US |  |  |  | # SAA |  | Yes | [5 empty] | No | No Information |  |  | Voluntary report | [8 empty] | 79 YR |  |  | 88 |  |  | Injury | No | No | Detachment Of Device Component | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Pigtail catheter separated from body of device during removal. diagnosis or reason for use: status post cardiac arrest.
38047 | 1035166-2018-00052 | 7584425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7584425&pc=DYB | 03/29/2018 | 10/12/2018 |  | 06/08/2018 |  |  | 05/11/2018 | 06/08/2018 |  |  | 06/01/2019 |  | 06/08/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 23F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-12603 |  | 0052-3021 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | Yes | Leak/Splash | [6 empty] | Cardiomyopathy | Left Ventricular Failure | Blood Loss | [10 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | It was reported that on (b)(6) 2018 during a medical procedure, the patient was slowly decompensating, with right heart failure while on echo. an impella rp was placed for support for cardiomyopathy. access to rfv; some difficulty placing swan initially into lpa branch. rp advanced over wire, but would not make turn into pa, despite rotations with push-pull technique. an amplatz 0.035 wire placed as buddy wire, still not successful in turning into pa. attempted to advance for at least 35 minutes. had issues with bleeding from the peel-away sheath valve; patient required administration of two units of blood, d/t pressures dropping, gtts (drips) increased. decision made to exchange sheath for another rp peel-away sheath. rp re-advanced through new sheath but still would not make turn. after another 20 minutes the second peel-away began to exhibit bleeding back through the valve, so decision made to abort rp and place patient on tandem heart via right ij. bleeding was resolved by removing sheath, suture placed and manual pressure held. patient was reported to be critical, tandem heart placed and patient was successfully supported with the tandem heart. | [5 empty] | Conclusion not yet available as the device is undergoing an evaluation but has not yet been completed. | Two 23f abiomed introducer sheaths from lot c1-12603 were returned from the customer without the dilators. there were no other accessories. the device history record was reviewed to confirm that the device passed all applicable in process and final inspections. blood was found on and inside both sheaths. both sheaths were kinked and both sheaths had torn hemostasis valves. upon evaluation of the returned devices under a 10x microscope, it was found that the 1st sheath was kinked at approximately 16cm from the distal tip and the hemostasis valve was torn. the 2nd sheath was kinked at approximately 17.5cm from the distal tip and the hemostasis valve was torn. it was found that the insertion points into the hemostasis valves were identified to be considerably off center of the cross cuts, creating a tear in both valves. returned device analysis reveals two 23f abiomed introducer sheaths that are within manufacturing specifications. according to the complaint report, the patient experienced bleeding from the sheath. both sheaths were found to have a torn hemostasis valve which would account for the blood loss from the sheaths. the valves were torn as a result of off center insertion due to use error/mishandling in the field. the introducer sheaths passed all in-process and qa final inspection steps before shipping to the customer including visual, dimensional and leak testing. per qa procedure, abiomed introducer sheath in-process and final inspections: perform pressure and vacuum leak test per procedure, latest revision. this procedure is performed by qa 100%. as per instructions for use: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
38048 | 1220648-2018-00051 | 7579697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7579697&pc=OZD | 05/08/2018 | 05/08/2018 |  | 06/07/2018 |  |  | 05/08/2018 | 06/07/2018 |  |  | 09/30/2019 |  | 06/07/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1310896 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Fracture | [6 empty] | Blood Loss | [12 empty] | Life Threatening | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that the physician placed an impella cp in a (b)(6) year old female patient. the insertion and placement were reported as fluid and flawless. the repositioning sheath was inserted and sutured. upon getting ready to transport the patient to the floor, the placement signal became erratic. blood was observed to be leaking from the red pump housing. a normal saline pressure bag was attached to red port with the roller clamp open which made the bleed out of the housing worse. the pressure bag was removed and the red side port was capped, and the red housing was kept above the patient's heart. upon arrival to the floor, it was noticed that the bleeding had slowed to a near stop. patient blood loss was estimated to have been between 50 to 75cc's, and the patient as administered 1 unit of packed red blood cells. it was decided to continue supporting the patient with the impella cp, and the patient was successfully supported for 49.42 hours. there were no additional adverse effects on the patient during patient support. | [5 empty] | The impella cp was returned for evaluation, along with the automatic impella controller's sole data logs. an analysis of the data logs confirmed that the placement signal was not reliable and that "placement signal lumen blocked" alarms occurred during the event. purge pressure and purge flow appeared normal throughout case the evaluation of the returned impella cp revealed that the placement signal was not functioning properly. during the evaluation the red pump plug was cracked open revealing a significant amount of biomaterial. there was no corrosion found to the internal mechanisms. a twin pin tear down was performed which revealed blood in the catheter sleeve and a crack in the adhesive. in conclusion, the root cause of the bleeding from the red pump plug and placement signal unreliable alarms was a twin pin fracture. [(b)(4)].
38049 | 1220648-2018-00052 | 7574086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7574086&pc=OZD | 05/08/2018 | 05/08/2018 |  | 06/06/2018 |  |  | 05/08/2018 | 06/06/2018 |  |  | 12/31/2019 |  | 06/06/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1321944 |  | 0048-0026 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A patient was admitted for coronary artery bypass surgery. two grafts were placed. two days later the patient decompensated with vt storm and had an impella placed for hemodynamic support. a femoral-t-femoral external bypass was placed to perfuse the patient's lower limbs due to known history of peripheral vascular disease. after two days of support the impella was explanted due to limb ischemia. after explant the surgical team repaired the limb. | [5 empty] | The pump was not returned for analysis. the data logs do not pertain to the failure mode. imaging was requested in order to observe the patient's level of previously known peripheral disease, but imaging was not shared. should more information come to light regarding the cause of the limb ischemia, a supplemental will be filed. (b)(4).
38050 | 1220648-2018-00057 | 7574249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7574249&pc=OZD | 05/15/2018 | 05/18/2018 |  | 06/06/2018 |  |  | 05/18/2018 | 06/06/2018 |  |  | 01/31/2020 |  | 06/06/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1323719 |  | 005048-JP |  | No answer provided | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 85 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella 2.5 was placed for support for ami/cgs and provided support for over five days. the clinical team reported pump positioning issues during the case. the team aso reported the patient experienced renal failure while on support, requiring dialysis treatment. the clinical team noted signs of hemolysis until the pump was repositioned, however the suspected hemolysis was never confirmed using a quantitative test. the patient was reported to be hemodynamically stable following removal of the impella 2.5. | [5 empty] | Neither the impella 2.5 nor the console data logs were returned for evaluation. the cause of the renal failure and suspected hemolysis is unable to be determined because the product could not be inspected. there were no nonconformances found during a review of the device history. complaints of this type will continue to be monitored for trends.
38051 | 1220648-2018-00050 | 7574573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7574573&pc=OZD | 04/18/2018 | 05/16/2018 |  | 06/06/2018 |  |  | 05/16/2018 | 06/06/2018 |  |  | 12/31/2019 |  | 06/06/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1321320 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 76 YR | [5 empty] | Malfunction | No | Yes | Device-Device Incompatibility | [6 empty] | No Known Impact Or Consequence To Patient | No Code Available | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A patient with acute mi and cardiogenic shock. the team was to complete an elective tavr (percutaneous aortic valve replacement) within the cardiac catheterization lab. during the tavr procedure the patient arrested and the team performed cpr and placed an impella cp for support. the patient had suction alarms with low pump flows, and so the cp pump was explanted and sent for complaint analysis. | [5 empty] | The cp pump and data logs were returned for analysis. no imaging was shared despite requests made. the pump was examined and found to have a fractured impeller blades. the root cause of the suction and low flows was the fractured blades. the broken blades were highly likely due to an interaction with another device. this can not be confirmed as imaging was not available for review. (b)(4).
38052 | 1220648-2018-00048 | 7569791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7569791&pc=OZD | 05/07/2018 | 05/09/2018 |  | 06/05/2018 |  |  | 05/09/2018 | 06/05/2018 |  |  | 01/31/2020 |  | 06/05/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1335495 |  | 005042 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 94 YR | [5 empty] | Malfunction | No | Yes | Kinked | Difficult to Remove | Activation, Positioning or Separation Problem |  |  |  |  | No Consequences Or Impact To Patient | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella 2.5 was placed via the right femoral to support an atherectomy patient in the cath lab. the team lost left ventricle (lv) position and so re-advanced the wire and repositioned the pump within the lv. upon the team's attempted removal of the guidewire, the pump kinked back upon itself and the wire was unable to be removed. the team therefore attempted to remove the pump and wire in combination, but needed surgical removal techniques to be employed, as contralateral leg snaring of the product was not successful for pump and wire removal. a vessel repair/closure was done to the surgical site and the arteriotomy was closed. | [5 empty] | The returned system (of pump, wire, and introducer) was still intact. the guidewire was still inside the pump, and was in a loop within itself. the guidewire and pump were also stuck inside the introducer. the root cause of the pump's inability to cross the aortic valve was unable to be determined. the root cause of the inability to remove the guidewire, along with the pigtail kink, was the loop created by the wire. the root cause of the wire core fracture was most likely the snare used to attempt to remove the wire. there were no other complaints against this lot of pumps or guidewires.
38053 | 7600521 | 7600521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7600521&pc=PBL | 02/10/2018 | 06/01/2018 | 06/01/2018 | 06/05/2018 | 06/01/2018 | 02/10/2018 |  | 06/05/2018 |  |  |  |  | 06/05/2018 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | PBL | [5 empty] | ABIOMED |  |  | US | [5 empty] | Device was returned to manufacturer | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 81 YR |  |  | 62 |  |  | Death | No | Yes | Pumping Stopped | Defective Component | Device Displays Incorrect Message |  |  |  |  | Cardiac Arrest | [12 empty] | Death | [5 empty] | IMPELLA CATHETER | IMPELLA CONTROLLER: SERIAL # (B)(4) USED WITH | [21 empty] | Initial submission |  |  |  |  | 02/10/2018 | Y | HOSPITAL |  |  | Impella alarmed with screen showing "impella stopped / retrograde flow" p-level dropped to 0. no kinks or disconnects noted. despite changing white cables and contacting impella rep for guidance, pt deteriorated and expired despite all resuscitative efforts. critically ill pt in icu on impella support. on (b)(6) 2018, impella alarms sounded and screen showed "impella stopped, retrograde flow." impella initially at p-level 6 but dropped to p-level 0 with impella flow + 0.00l/min. attempts were made to dial up the p-level back to 6 but device continued to drop to 0. device assessed for disconnects and kinks. impella rep contacted, white cables changed. pt became hemodynamically unstable, cpr initiated. pt expired despite all resuscitative efforts. device was immediately turned over to the mfr's rep for investigation. recently received report from a biomed dated (b)(6) 2018 indicating that a coil short in the impella catheter was likely the source of the product malfunction.
38054 | 1220648-2018-00043 | 7553054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7553054&pc=OZD | 04/30/2018 | 04/30/2018 |  | 05/30/2018 |  |  | 04/30/2018 | 05/30/2018 |  |  | 02/28/2019 |  | 05/30/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  |  |  | 1283704 |  | 0048-0004 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 45 YR | [5 empty] | Injury | Yes | Yes | Appropriate Term/Code Not Available | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | Complainant reported that the physician was treating a (b)(6) male patient admitted in acute myocardial infarction and in cardiogenic shock. an impella cp was put in place to unload the left ventricle (lv) while on ecmo. during patient support bleeding was noted to be coming from the catheter and the insertion site. the bleeding was described as a very low drip. the decision was made to continue supporting the patient with the impella cp and to just continue to monitor the leak. the patient was successfully supported 19.60 hours. the patient was reported to require replacement blood products as a result of the bleeding, although the amount of replacement blood administered to the patient was unable to be determined. the patient support was reported as successful. | [5 empty] | The device was returned for evaluation. the evaluation of the returned impella cp revealed that the cause of the patient bleeding was that there was a crack in the twin pin assembly at the placement lumen. a device history record review and lot analysis determined that all applicable systems are adequate and remain effective. no corrective action is recommended at this time because the failure mode was determined to be low risk index based on very low occurrence and moderate severity. failures of this type will continue to be monitored for trends. (b)(4). | Since the initial filing the discovery was made that the serial number was listed incorrectly.
38055 | 7544290 | 7544290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7544290&pc=PBL | 05/08/2018 | 05/14/2018 | 05/14/2018 | 05/25/2018 | 05/14/2018 | 05/14/2018 |  | 05/25/2018 |  |  |  |  | 05/25/2018 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PBL |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  | 140515 |  |  |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | User Facility report | [14 empty] | Malfunction | No | Yes | Incorrect Or Inadequate Test Results | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | After placing the impella catheter, blood was backing out of the red connection, making the arterial waveform and pressure inaccurate. the pump continued to perform correctly so the impella stayed in the patient for 48 hours. no harm to patient upon removal. manufacturer response for impella catheter, (brand not provided) (per site reporter): device was returned to rep for manufacturer evaluation.
38056 | 1220648-2018-00047 | 7536760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7536760&pc=OZD | 05/03/2018 | 05/03/2018 |  | 05/23/2018 |  |  | 05/03/2018 | 05/23/2018 |  |  | 12/31/2019 |  | 05/23/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1325093 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Malposition of Device | Positioning Problem | [5 empty] | Low Blood Pressure/ Hypotension | Perforation | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A patient was admitted in shock and as she decompensated, the team placed two impella, for both right and left heart support. the impella cp was placed to the left ventricle, and the rp to the right heart. during the rp placement the patient coded and cpr was administered. due to concern of a pericardial effusion, the patient was sent to the operating room. while in the or the team observed that the cp had perforated the heart wall of the left ventricle. the perforation was repaired and the bleed stopped. | [5 empty] | The team returned the data logs and the pump product for investigation. the logs revealed the support/case was four hours in length. the data logs did reveal position and suction alarms. the data logs, along with the clinical account, point to the administered cpr being the root cause of the ventricular perforation. the returned pump was largely intact, but with a slight deformation/bend in the pigtail, indicating interaction with, most likely, the left ventricular wall. despite requests, the imaging of the procedure was not shared for analysis. the root cause of the ventricular perforation was most likely the administered cpr, but cannot be definitively determined. no corrective action is recommended at this time as the root cause was unable to be definitively determined. internal reference: (b)(4).
38057 | 1220648-2018-00042 | 7536809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7536809&pc=OZD | 04/30/2018 | 04/03/2018 |  | 05/23/2018 |  |  | 04/30/2018 | 05/23/2018 |  |  | 02/29/2020 |  | 05/23/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1331968 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Malposition of Device | Positioning Problem | Device Handling Problem |  |  |  |  | Perforation | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | The patient presented in cardiogenic shock to a hospital in (b)(6). the impella cp was placed for support, but shortly after placement an effusion was noted. the patient had the perforation of the lv repaired and the impella was explanted. the team placed ecmo, but the patient expired days later with multiple comorbidities. in conversation with the physician, the consensus was the .018 guidewire caused the perforation. the left ventricle wall was noted to be necrotic due to underlying cardiac co-morbidities and previous myocardial infarctions. | [5 empty] | The device was not returned for analysis. the data logs were returned. the logs concur with the clinical report of early suction and poor placement signal that later remedies with the repair of the ventricular wall. upon analysis of the pump lot, no other complaints have been made against the lot of pumps. the most likely cause of the perforation is the pump or guidewire puncturing the necrotic tissue. the definitive cause was unable to be determined as the device was not returned for analysis. internal reference: (b)(4). | Since the initial filing of the medwatch, the european facility returned the pump product for analysis. the pump was tested. inspection revealed damage to the pigtail, which had been severed at the curled region. the tip of the pigtail was not included for analysis. the fracture of the pigtail was re-created in testing and the fracture was most consistent with fracture created by shears. the most likely cause of the perforation is the pump or guidewire puncturing the necrotic tissue.
38058 | 1220648-2018-00044 | 7537803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7537803&pc=OZD | 04/27/2018 | 04/28/2018 |  | 05/23/2018 |  |  | 04/28/2018 | 05/23/2018 |  |  | 02/29/2020 |  | 05/23/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1332857 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Malfunction | No | Yes | Fracture | [6 empty] | Patient Problem/Medical Problem | [41 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella cp was inserted for support for ami/cgs. when the peel away sheath was removed and the repositioning sheath advanced, a piece of the easy guide lumen was observed in the patient's vasculature. the piece was easily removed and the repositioning sheath was advanced into position without difficulty. the pump provided support to the patient for approximately three days. | [5 empty] | The returned red lumen was confirmed to be fractured and the failure mode was confirmed. the length of returned lumen, approximately 14cm in length, would align with the pump housing of the impella cp if the lumen was placed correctly. however a full analysis can not be completed because the pump was not returned.
38059 | 1220648-2018-00040 | 7539048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7539048&pc=PYX | 04/27/2018 | 04/27/2018 |  | 05/23/2018 |  |  | 04/27/2018 | 05/23/2018 |  |  | 08/31/2020 |  | 05/23/2018 |  | TEMPORARY NON ROLLER RIGHT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1302551 |  | 004334 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Perforation | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that the physician was treating a (b)(6) year old male post heart transplant patient who was having difficulty being weaned off the cardiopulmonary bypass machine. the patient was placed on va ecmo and pa pressure started increasing. the heart team evaluated the patient and made decision to place an impella rp and to remove the ecmo. the patient was taken to the cvor and the impella rp was placed via the right femoral artery using the 23fr sheath. the cardiothoracic surgeon de-cannulated the ecmo. the patient was reported to be bleeding following the removal of the peel away sheath, after one hour of manual compression the impella site continued to bleed; therefore the decision was made to cut down on the impella site and explorer the right femoral artery. the exploration revealed a dissection to the right femoral artery. the artery was successfully repaired. the patient remained stable and hemostasis was achieved with the surgical repair of the artery. | [5 empty] | The facility did not return the impella rp product for investigation. the data logs were not returned for analysis. the root cause of the tear to the femoral artery was unable to be determined. (b)(4).
38060 | 7532873 | 7532873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7532873&pc=OJE | 05/10/2018 | 05/16/2018 | 05/16/2018 | 05/22/2018 | 05/16/2018 | 05/16/2018 |  | 05/22/2018 |  |  | 11/30/2019 |  | 05/22/2018 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA |  |  | OJE |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  |  |  | 004334 |  | No answer provided | [5 empty] | Invalid/Incomplete | No Information |  |  | User Facility report | [8 empty] | 68 YR | [5 empty] | Malfunction | No | Yes | Device Operates Differently Than Expected | [6 empty] | Thrombus | [18 empty] | NOT BY REORT | [22 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The impella device "sucked up" a clot. "initially tolerated coming off bypass but then abruptly deteriorated secondary to right ventricular assist device (rvad) dysfunction and rv failure. ino and isoproterenol were started with minimal improvement. the rp rvad was removed and the impeller was found to be obstructed by thrombus.
38061 | 7525250 | 7525250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7525250&pc=KFM | 05/10/2018 | 05/16/2018 | 05/15/2018 | 05/18/2018 | 05/15/2018 | 05/15/2018 |  | 05/18/2018 |  |  |  |  | 05/18/2018 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA |  |  | KFM | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No answer provided | [5 empty] | Invalid/Incomplete | No Information |  |  | User Facility report | [8 empty] | 51 YR | [5 empty] | Malfunction | No | Yes | Device Displays Incorrect Message | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL | SN154672 |  | Impella alarmed "impella stopped retrograde flow." attempted to change to different controller per troubleshooting instructions on impella controller, same alarm on new controller, the icu team removed the impella device and the arterial sheath. the impella rep was notified and at the bedside post removal.
38062 | 1220648-2018-00045 | 7523373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7523373&pc=OZD | 04/20/2018 | 05/14/2018 |  | 05/17/2018 |  |  | 04/20/2018 | 05/17/2018 |  |  | 12/31/2019 |  | 05/17/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1325136 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 71 YR | [5 empty] | Malfunction | No | Yes | Device Operates Differently Than Expected | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Other | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella cp was inserted to support a patient for post-cartiotomy cardiogenic shock, following open heart surgery. after approximately one hour, pump flows dropped suddenly. the position of the pump was verified using fluoro and echo. the pump was exchanged for a second impella cp without issue. | [5 empty] | The impela cp was returned for evaluation into the low pump flows that were reported. the pump was unable to be started in the lab. a tear down of the pump revealed broken blades on the impeller. the root cause of the low flow is broken blades. the cause of the broken blades is unknown at this time. if abiomed learns of additional information a follow up report will be sent.
38063 | 1220648-2018-00035 | 7524042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7524042&pc=OZD | 04/20/2018 | 04/20/2018 |  | 05/17/2018 |  |  | 04/20/2018 | 05/17/2018 |  |  | 01/31/2020 |  | 05/17/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1327020 |  | 0048-0026 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Appropriate Term/Code Not Available | [5 empty] | Thrombus | Injury | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On (b)(6) 2018 and impella cp was placed in a (b)(6) year old male patient. the patient remained in stable condtion was successfully supported for 44 hours. on (b)(6) 2018 the patient continued to be in stable condition, and the impella cp was providing successful hemodynamic support. on this same date, and after 44 hours of patient's support the device was weaned from the patient and was explanted. following removal of the impella cp the patient had a cerebral vascular incident. the patient was sent to interventional radiology where the clot was removed. it was reported that upon removal of the device biomaterial was seen attached to the pump's outlet cage. the physician reported that the patient's aorta had some atherosclerotic plaque, and that the impella's cook sheath may have scraped the aorta collecting some of the plaque along the way when it was removed with the pump. during patient support, heparin was given systematically and in the impella's purge system. the act was 150 at the time the impella cp was removed. only the last act reading at the time of the removal of the pump was shared with the manufacturer. | [5 empty] | Neither the impella cp nor the automatic impella console logs were returned for evaluation and analysis. the device had been held in the facility's risk management department, and then the impella cp and the cook sheath were reported to have been accidently discarded. a photograph of the biomaterial on the outlet cage was submitted to the manufacturer. the root cause of this event was unable to be determined because the device was not returned for evaluation. (b)(4).
38064 | 1220648-2018-00036 | 7518513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7518513&pc=OZD | 04/19/2018 | 04/19/2018 |  | 05/16/2018 |  |  | 04/19/2018 | 05/16/2018 |  |  | 01/31/2019 |  | 05/16/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1279798 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Appropriate Term/Code Not Available | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that a (b)(6) male patient was scheduled to undergo a cabg procedure. the physician placed an impella cp to provide the patient with hemodynamic support during the cabg procedure. following the initiation of pump support the patient's creatine levels increased up to over 2.0, and the patient was reported to have exhibited hemolysis symptoms. in an effort to alleviate these symptoms the physician repositioned the cannula and administered volume to the patient, but the hemolysis did not resolve. during this event there were no plasma free hg drawn to confirm or rule out hemolysis, and no renal lab values were taken. the pump was weaned from the patient and removed due to the inability to resolve the hemolysis symptoms. it was reported that following the removal of the impella cp the patient's creatine levels increased to >5.0. it was reported that patient dialysis was no provided and that the patient's hemolysis symptoms eventually resolved after the pump was explanted. | [5 empty] | The impella cp was returned for evaluation. in addition, the automatic impella console's logs were returned for analysis. the analysis of the logs did not reveal any performance abnormalities. the evaluation of the returned impella cp was performed. the evaluation determined there was no damage or defects present on the impeller or the pump housing. hemolysis testing of the device revealed that it met all the required values. a lot analysis was performed and it was determined that there had been no additional complaint associated with any of the impella cp pumps built with under this lot number. the root cause of the hemolysis was unable to be determined, as there were no relevant issues found during the evaluation of the returned impella cp. (b)(4).
38065 | 1220648-2018-00041 | 7496264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7496264&pc=OZD | 04/23/2018 | 04/23/2018 |  | 05/08/2018 |  |  | 04/23/2018 | 05/08/2018 |  |  | 10/31/2019 |  | 05/08/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1319167 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 52 YR | [5 empty] | Injury | Yes | No | Kinked | Difficult to Remove | Device Issue |  |  |  |  | Perforation of Vessels | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A patient admitted with cardiogenic shock and acute mi had an impella cp placed with difficulty via the right femoral artery during cpr. due to the pump kinking, when the guidewire was removed, the pump was explanted to resolve the kink. upon reinsertion the right femoral artery was observed to be dissected. the patient had the dissection tamponade by the placement of a 14fr introducer and the case and cpr continued, until his death from cardiogenic shock and failed cpr and defibrillations. | [5 empty] | The impella pump was returned for analysis. the data logs were not retrieved. upon examination of the pump there was damage noted as the catheter was stretched and broken. biomaterial was observed within the catheter's cannula. the pump housing was noted to be deformed and damaged. the root cause of the femoral artery dissection was most likely the excessive force used to remove the guidewire and pump. the root cause of the necessity for this excessive force is unable to be determined. no corrective action is warranted, the failure mode will be monitored and trended. internal reference (b)(4).
38066 | 1220648-2018-00030 | 7487724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7487724&pc=OZD | 02/21/2018 | 02/21/2018 |  | 05/04/2018 |  |  | 02/21/2018 | 05/04/2018 |  |  | 12/22/2017 |  | 05/04/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1316051 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 51 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that a (b)(6) year old female patient with peripheral artery disease (multi-vessel vessel coronary artery disease with high grade stenosis) had an impella cp placed via the right leg's femoral artery. this patient had previously undergone a lower left extremity amputation. the impella cp was placed to rest the heart. the peel-away sheath was left in place to allow the cardiovasular surgeon to eventually remove the device without having to manage the access site. following bypass surgery, the patient developed a cold leg, resulting in vascular surgery, during which the leg was amputated. the patient expired 3 days later on (b)(6) 2018. the physician reported that the patient outcome was unrelated to any issue with the impella rp, as the patient would not have made it to the or without the support of the pump. | [5 empty] | The device was returned to the manufacturer for evaluation. the evaluation of the impella rp revealed a split in the outer sheath of the repositioning unit/sheath, as well as biomaterial around the impeller of the pump. in conclusion, the peel away sheath was left inside of the patient's femoral artery during surgery. the patient had severe pvd and an amputation on the other leg. following the surgery there was a delay in restoring pump flow due to complications requiring additional surgical intervention. the ischemic leg may have been the result of leaving the peel-away sheath inside of the patient's artery during surgery, despite the patient being high risk for limb ischemia, along with the extended loss of hemodynamic support following the bypass operation; however, this cannot be definitively concluded. (b)(4).
38067 | 1220648-2018-00031 | 7472345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7472345&pc=OZD | 02/12/2018 | 04/03/2018 |  | 04/30/2018 |  |  | 04/03/2018 | 04/30/2018 |  |  | 02/28/2019 |  | 04/30/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1282175 |  | 0048-0002 |  | No answer provided | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Death | Yes | No | Positioning Problem | [6 empty] | Injury | [12 empty] | Hospitalization | Required Intervention | Disability | [26 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella cp was placed for support into a (b)(6) male for ami/cgs. pump flows were low and position alarms were triggered immediately after the pump was started. echo imaging revealed papillary muscle head and tendinous chords were drawn into the pump inlet. the clinical team repeatedly attempted to reposition the device. the patient's condition was deteriorating therefore the pump was removed. the patient expired due to cardiogenic shock. | [5 empty] | The products and clinical images were requested from the field. a follow up report will be submitted if they are received and additional information is discovered.
38068 | 1220648-2018-00032 | 7472473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7472473&pc=OZD | 04/04/2018 | 04/04/2018 |  | 04/30/2018 |  |  | 04/04/2018 | 04/30/2018 |  |  | 09/30/2019 |  | 04/30/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1313154 |  | 005066-JP |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Injury | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella 5.0 was placed into a (b)(6) patient for support for post-pci, severe myocardial regurgitation and cardiogenic shock on (b)(6). on the 6th day of support, the patient received scheduled mvr/cabg/pv isolation. although the patient had mr , mr deterioration and deviation of anterior leaflet was observed by echo during surgery. a chordae rupture in center of the anterior leaflet was observed during the surgery. the physician noted the patient's chordae tendineae was fragile. | [5 empty] | The pump and case images were returned for evaluation. there were no abnormalities found with the pump during post-case testing. tee imaging revealed an anterior leaflet prolapse of the mitral valve, which was caused by a rupture of the chordae in the center of the leaflet. a review of the case revealed the implanting physician stated the "patient's chordae tendinae was fragile, and the rupture of the chordae was observed during exchanging of the pump." a review of the case by an abiomed physician revealed the pump was in an improper position due to patient condition. the patient had a sigmoid septum and an obstructive left ventricular outflow tract, which caused the pump to be malpositioned. the root cause of the mitral valve damage is related to patient condition.
38069 | 1220648-2018-00027 | 7468267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7468267&pc=OJE | 03/29/2018 | 03/29/2018 |  | 04/27/2018 |  |  | 03/29/2018 | 04/27/2018 |  |  | 03/31/2018 |  | 04/27/2018 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA RP |  |  | OJE |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1250295 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | Yes | Difficult To Position | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | The complainant reported that (b)(6) year old female patient was in a slowly decompensating condition, and was in right hear failure. a swan was placed in the patient via the right femoral vein, but there was some initial difficulty experienced by the physician when placing the swan into the left pulmonary artery (lpa). the physician then attempted to place an impella rp. when advancing the pump over the wire the device would not make the turn into the pulmonary artery (pa). an 0.035 amplatz wire was placed as a buddy wire, but this was not successful in making the turn into the pa. during the 35 minute attempt to place the impella rp, patient bleeding occurred from the peel-away sheath valve. two units of replacement blood was administered to the patient as a result of the pressures dropping. the sheath was exchanged for another rp peel-way sheath, but the same issue occurred, and the pump would not make the turn into the pa. it was decided to abort the impella rp placement, as a result of the patient's cardiac anatomy. the physician then placed a tandem heart in the patient via the right internal jugular. | [5 empty] | The impella rp and peel-away sheaths have not yet been returned for evaluation. at this time there is no definitive information indicating that the inability to place the impella rp was the result of any issue with the device. (b)(4). | The impella rp and the 23fr oscor introducer sheaths were returned for evaluation. the returned oscor introducer sheaths were analyzed. upon close examination of the valves from a top side perspective, a slight cut (tear) on one half of the valve hub was seen in both of the introducers. the cause of bleeding was damage most likely incurred as a result of off-centered insertion of the dilator through the oscor sheath. this may have led to bleeding from the introducer valve. internal reference: spr 18-0367
38070 | 1220648-2018-00033 | 7450414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7450414&pc=OZD | 03/29/2018 | 03/29/2018 |  | 04/23/2018 |  |  | 03/29/2018 | 04/23/2018 |  |  | 07/31/2019 |  | 04/23/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1303861 |  | 0048-0026 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Difficult or Delayed Positioning | Activation, Positioning or Separation Problem | Human-Device Interface Problem |  |  |  |  | Low Blood Pressure/ Hypotension | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An (b)(6) male patient in cardiogenic shock had an impella cp placed for hemodynamic support. due to the patient size the sheath could not be advanced into position to achieve hemostasis. the tissue track was too deep due to his body habitus. the patient lost blood from the groin access site and needed vascular intervention to achieve hemostasis. a cut down was done and sutures were placed. one unit of blood product was infused and the support continued for the angioplasty procedure and transfer to the icu. | [5 empty] | The product was discarded at the hospital and was unable to be investigated. the lot of impella pump has no other complaints leveled against it. most likely the cause of the bleeding was due to the deep vasculature in the obese patient. the ifu states to observe for (b)(6), and this patient was noted to be obese. his body habitus did not allow the repositioning sheath to traverse the length of the tissue tract completely, to achieve hemostasis and tamponade the artery. in addition the ifu states the following: "when securing the repositioning sheath, vascular closure may be difficult in obese patients with extensive adipose tissue." (b)(4).
38071 | 1035166-2018-00032 | 7439593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7439593&pc=DYB | 03/20/2018 | 09/17/2018 |  | 04/18/2018 |  |  | 03/21/2018 | 04/18/2018 |  |  | 01/12/2020 |  | 04/18/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-13895 |  | 0052-3025 |  | No answer provided | [5 empty] | Invalid/Incomplete | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION | [10 empty] | Malfunction | Yes | Yes | Break | Component Missing | [5 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup | Followup |  |  |  | N | Invalid/Incomplete |  |  | Additional information was received that included the following: the short introducer was used with a long dilator, and when the impella was introduced, there was bleeding associated with the orange hub cap breaking off the introducer as the valve was disrupted. this bleeding was controlled. the amount of blood was negligible and did not require infusion of blood replacement products. | The following information was inadvertently ommitted from initial report: it was reported according to circ support in or, no short 14f introducer included, only two long introducer dialators. | It was reported that an incident occurred during a surgical procedure for placement of an impella cp on (b)(6) 2018. it was during use, while accessing the right femoral artery in the groin area that the issue was observed. once the impella cp was inserted, the 14f introducer peel-away sheath broke away at the orange hub on one side, disrupting the double valve. the patient experienced blood loss at groin until attending physician stepped in and took control. a second 14f introducer was then used and the procedure was completed. although it is reported that no harm was experienced by the patient, the patient status and blood loss amount is unknown at this time. patient involved is a cardiac patient on ecmo. |  |  |  | After review of the record data and additional information received, the category was changed from serious injury to reportable malfunction. the category was changed from serious injury to reportable malfunction based on review of reported and additional information received that indicated the amount of blood loss was negligible and did not require infusion of blood replacement products. there was no report of concern for the patient related to this event. oscor inc., is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc., which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor, inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor, inc., or its employees, that the report constitutes an admission that the device, oscor, inc., or its employees caused or contributed to the event described in this report. | The device was used in treatment, and was not returned for analysis. the investigation will focus on a review of product documentation. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. complaint review of reported lot identified one additional unrelated report. based upon the investigation, two (2) separate corrective actions have been opened to investigate these issues. components are packaged per abiomed 14f introducer kit packaging inspection for impella, 13cm and 25cm as follows: verify contents of packaging according to the bom. as per instructions for use (ifu): device description - the introducer set is composed of sheath, dilator(s) and guidewire. although the ifu does not list specific sheath quantities, the label does note the kit should contain one 14f 13cm and one 14f 25cm in length sheath. per qa procedure adelante s2s introducer sheath in-process and final inspection: "with naked eye at a distance of 12" to 18", verify the sheath hub and split caps are free of excess silicone oil applied between the seals during assembly. verify split caps are properly placed and secured onto sheath hub and free of cracks/damages or excessive adhesive." additionally, destructive testing is performed per sampling plan. inspector is to "insert dilator into sheath, lock dilator hub lock nut. hold dilator hub with one hand and one of the sheath hub wing with the other hand. bend the dilator hub to the side until dilator hub lock nut is disconnected from the sheath hub. split caps should remain in place. there should be no damage and fracture to split cap and tabs." the instructions for use (ifu) instructs the user: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | A preliminary review showed there were no manufacturing rejects or anomalies of this event type recorded in the device history record. conclusion not yet available as the investigation is on-going.
38072 | 1220648-2018-00028 | 7440418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7440418&pc=PYX | 03/19/2018 | 03/29/2018 |  | 04/18/2018 |  |  | 03/29/2018 | 04/18/2018 |  |  | 11/30/2019 |  | 04/18/2018 |  | TEMPORARY NON ROLLER RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1314718 |  | 004334 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Appropriate Term/Code Not Available | [6 empty] | Perforation of Vessels | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An (b)(6) male was admitted with cardiogenic shock and had an aortic valve repair (avr) and right coronary artery repair. post operatively an impella rp was placed. the placement of the rp was difficult due to tortuous pulmonary arteries. after four days of support the patient was deteriorating and he was treated for a hemothorax. the next day he was taken to the operating room where they were unable to determine the source of the bleeding. as his condition was not improving, the patient had his impella explanted. for a cannulation of ecmo. he required blood products (approximately 50 units), the abdomen was distended, and the medical team observed a tear to the inferior vena cava. the cause of the tear is not known. the family chose to withdraw care, and he expired. | [5 empty] | The facility did not return the impella rp product for investigation. the data logs were returned for analysis. the root cause of the tear to the ivc was unable to be determined. (b)(4).
38073 | 1220648-2018-00039 | 7441254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7441254&pc=OZD | 03/20/2018 | 03/20/2018 |  | 04/18/2018 |  |  | 03/20/2018 | 04/18/2018 |  |  | 11/30/2019 |  | 04/18/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1324054 |  | 0048-0002 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The physician reported that the patient experienced bleeding issues at the groin after insertion. the physician managed the bleeding with protamin and pressure dressing. | [5 empty] | The device was discarded by the hospital. no device evaluation can be completed.
38074 | 1035166-2018-00028 | 7424676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7424676&pc=DYB | 03/14/2018 | 10/09/2018 |  | 04/12/2018 |  |  | 03/15/2018 | 04/12/2018 |  |  | 12/01/2020 |  | 04/12/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-13865 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | Invalid/Incomplete | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN | HEALTH PR |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Peeled/Delaminated | [6 empty] | Blood Loss | [12 empty] | Other | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | It was reported that on (b)(6) 2018 during a medical procedure, the physician noticed the peel was detached from the dilator tip as it was being pushed forward, causing severe bleeding from the artery. the issue was first noted during the 14fr peel away introducer insertion via the right femoral artery for placement of the impella cp. it is noted a second introducer was used after issue was observed. manual compressing and a second oscor sheath was placed to stop the bleeding. the impella support was discontinued 2 days later. the patient had been admitted for a high risk pci and it is reported that the patient is stable post surgery. patient is reported to be a (b)(6) year old middle (b)(6) male. no weight is known. | [5 empty] | Report source: country of incident occurrence - (b)(6). a preliminary review showed there were no manufacturing rejects or anomalies of this event type recorded in the device history record. conclusion not yet available as the investigation is on-going | The introducer sheath passed all in-process and qa final inspection steps before shipping to the customer including visual, dimensional and leak testing. no manufacturing defects were found. the device history record was reviewed to confirm that the device passed all applicable in process and final inspections. based on the investigation, capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. one 14f, 13cm abiomed introducer sheath was returned from the customer without the dilator. there were no other accessories. blood was found on the interior and exterior of the sheath. upon evaluation of the returned sheath under a 10x microscope, it was found that there was a split in the sheath tube starting approximately 10 cm from the distal end continuing through the distal tip. there was also extensive damage found around the circumference of the distal tip edge. there was a split in the sheath tube starting approximately 10 cm from the distal end and continuing through the distal tip. there was also extensive damage found around the circumference of the distal tip edge. potential contributing factors to the damage is if the device was advanced through an area of resistance (as stated in the precautions of the ifu), or if the device was subjected to unusual stresses (as stated in the handling and storage section of the ifu). per procedure abiomed introducer sheath in-process and final inspection, the devices in this lot were inspected as follows: 9.2.2 visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. slight discoloration from tipping process is acceptable. verify proper tip shape according to drawing. per instructions for use : never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
38075 | 1035166-2018-00030 | 7425373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7425373&pc=DYB | 03/16/2018 | 09/07/2018 |  | 04/12/2018 |  |  | 03/19/2018 | 04/12/2018 |  |  | 05/01/2020 |  | 04/12/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 23F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC |  |  | US |  | C1-13315 |  | 0052-3021 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 63 YR | [5 empty] | Malfunction | No | Yes | Fracture | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | It was reported that on (b)(6) 2018, during an acute myocardial infarction/cardiogenic shock (ami/cgs) procedure, the 23fr introducer sheath fractured inside the vessel. it was during insertion of the impella rp pump that it was noted the 23f introducer sheath had fractured while inside the vessel. the fractured sheath was removed, a new 23f sheath was obtained and the procedure was completed, the impella rp was delivered. this (b)(6) male patient has history of runs of ventricular tachycardia and the pulmonary artery pulsatility index was trending down. it was reported that there was no injury to the patient nor any extended hospitalization. no further patient or incident information was provided. | [5 empty] | The device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. one 23 fr abiomed introducer sheath was received in a used condition with the dilator. no other accessories were received with this device. blood was found on the exterior and interior of the sheath and side port tubing. upon evaluation of the returned product, it was found that there were two splits in the sheath tube; 1.5cm and 3.5cm from the distal end. a kink was found in the sheath approximately 14cm from distal tip. there was also damage found around the circumference of the distal tip edge. the hemostasis valve was damaged/torn. the device did not fracture into separate pieces. the returned device had two splits and one kink in the sheath, a damaged/dented tip, and a torn/damaged hemostasis valve. returned device analysis reveals one 23f abiomed introducer sheath that is within manufacturing specifications. no manufacturing defects were found. potential contributing factors to the damages is if the device was advanced through an area of resistance (as stated in the precautions of the ifu), or if the device was subjected to unusual stresses (as stated in the handling and storage section of the ifu). as per procedure (abiomed introducer sheath in-process and final inspection), the devices in this lot were inspected as follows: 9.2.2 visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. per section 12.2, visual inspection, with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. perform leak test as per procedure. per ifu : never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Device has been requested for retrieval but not yet received. a preliminary review showed there were no manufacturing rejects or anomalies of this event type recorded in the device history record. conclusion is not yet available as the investigation is on-going.
38076 | 1220648-2018-00020 | 7421463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7421463&pc=OJE | 03/15/2018 | 03/16/2018 |  | 04/11/2018 |  |  |  | 04/11/2018 |  |  | 10/31/2019 |  | 04/11/2018 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA RP |  |  | OJE |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1311325 |  | 004334 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 44 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Death | Life Threatening | Required Intervention | Disability | [25 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella rp was placed for support for right ventricular failure, post-pulmonary embolism and pulmonary hypertension on (b)(6). while the clinical team was removing the 23 x 30 peel away sheath, a mattress suture was placed; however it was not sufficient to maintain hemostasis. a second mattress was placed and femostop was applied. upon arrival to cvicu, there was still continuous bleeding from the right femoral vein. the patient was brought to the operating room for repair of the vasculature and removal of the impella rp. the patient was then returned to the cvicu where she continued to bleed, requiring the administration of a blood transfusion. the patient went into pulseless electrical activity (pea) and arrested. cpr was performed however the patient expired. | [5 empty] | There is no known malfunction that we are aware of, however neither the pump nor the introducer was returned for evaluation.
38077 | 1220648-2018-00026 | 7415628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7415628&pc=OZD | 03/16/2018 | 03/16/2018 |  | 04/10/2018 |  |  | 03/16/2018 | 04/10/2018 |  |  | 02/28/2019 |  | 04/10/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1282414 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | Yes | Kinked | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that following stent placement in a 7(b)(6) female patient an impella cp was placed in the left femoral artery without the aid of fluoro imaging. the pump was placed deep into the left ventricle when the cannula kinked and the pigtail of the impella cp appeared to be entrapped in the pump's inflow cage. the physician was unable to free the pigtail. the patient was taken to ct surgery where the device was removed from the left ventricle via the aid of a snare. another impella cp was obtained and successfully placed in the patient. the patient was supported for 3.25 hours when the patient's family decided to withdraw care, and patient expired. | [5 empty] | The device was not returned for evaluation. the console logs were returned and the analysis of the logs revealed that the pump did not appear to be well positioned throughout the case. the pump was never started, and was plugged in for roughly one hour. the submitted photo of the device along with the clinical account of the event did confirm that the cannula did kink. although it is highly unlikely for the pigtail of the impella cp to enter the inflow cage, it is possible. further analysis cannot be completed, however, as no product was returned from for evaluation, and therefore the root cause of the pigtail kinking could not be determined, nor can the failure mode be confirmed without either the product or imaging. a manufacturing review found no other complaints for this failure mode made against this lot of impella pumps. the root cause of the kinking was unable to be determined as the pump was not returned for investigation. failures of this type will be monitored and trended. internal reference: (b)(4).
38078 | 7406441 | 7406441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7406441&pc=KFM | 02/21/2018 | 03/30/2018 | 03/30/2018 | 04/09/2018 | 03/30/2018 | 03/30/2018 |  | 04/09/2018 |  |  | 10/31/2019 |  | 04/09/2018 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | KFM | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003 |  | Yes | [5 empty] | No | Physician |  |  | User Facility report | [8 empty] | 83 YR | [5 empty] | Malfunction | No | Yes | Difficult to Insert | Malposition of Device | [5 empty] | Low Blood Pressure/ Hypotension | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Heart catheterization was performed, no ischemic vessels noted. considering patients continued hypotension, impella cp support was initiated. the first and second placement attempts were unsuccessful as the device did not sit properly within the patients anatomy (left ventricle). after removing a second time, the device was tested and flushed. it did not perform as expected. the device d/c'd from the field and removed in place of a second device. the device then was advanced and performed as expected.
38079 | 1220648-2018-00029 | 7402564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7402564&pc=OZD | 03/09/2018 | 03/09/2018 |  | 04/06/2018 |  |  | 03/09/2018 | 04/06/2018 |  |  | 04/30/2019 |  | 04/06/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1300129 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The customer reported that the patient was experiencing hemolysis and bleeding while being supported by an impella cp. during transport, significant bleeding occurred at the groin. femstop was placed to attempt to resolve the bleeding. the patient received 8 units of blood and there was no lasting harm. the physician escalated to an impella 5.0 to resolve hemolysis. | [5 empty] | The pump was not returned for evaluation. there is no known malfunction that we are aware of. if the manufacturer becomes aware of any additional information, a follow-up will be submitted.
38080 | 1220648-2018-00034 | 7402781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7402781&pc=OZD | 08/19/2017 | 03/13/2018 |  | 04/06/2018 |  |  |  | 04/06/2018 |  |  |  |  | 04/06/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | UNK |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report | [8 empty] | 70 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On (b)(6) 2018 a cardiol res and elmer press inc.¿ journal article was found. the date of the event was not reported and abiomed was unable to find a corresponding event that had been reported for any customer. abiomed has used the manuscript submission date of (b)(6) 2017 as the date of the event. a (b)(6) year-old male with known coronary artery disease and stable angina was admitted to our facility with a prolonged chest pain episode at rest, associated with dyspnea. an impella cp was successfully placed. following pci, the patient suffered acute profound hypotension necessitating intravenous fluids and vasopressor support with epinephrine. the impella device was removed and selective left common iliac angiogram from the right femoral access revealed a vascular injury site with an indication of a retroperitoneal hematoma. an 8.0 × 59 mm icast covered stent was deployed to the left external iliac artery with successful sealing of the perforation. the impella device site was closed with two perclose devices. computed tomography (ct) of the abdomen revealed large retroperitoneal bleeding along the plane of left external iliac artery. the patient required 4 units of packed red blood cell transfusion. the patient was discharged from the hospital 10 days later. | [5 empty] | On march 13, 2018 a cardiol res and elmer press inc¿ journal article was found. the date of the event was not reported and abiomed was unable to find a corresponding event that had been reported for any customer. abiomed has used the manuscript submission date of august 19, 2017 as the date of the event. no device was returned for evaluation, and abiomed was unable to obtain any additional information regarding this event. (b)(4).
38081 | 1220648-2018-00022 | 7403703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7403703&pc=OZD | 03/11/2018 | 03/11/2018 |  | 04/06/2018 |  |  | 03/11/2018 | 04/06/2018 |  |  | 12/13/2019 |  | 04/06/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1325093 |  | 0048-0026 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | Device Operates Differently Than Expected | Human-Device Interface Problem | [5 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) year old male with post cardiotomy cardiogenic shock was being transferred to the tertiary medical center while on support with an impella cp and the patient's limb went cold without distal pulses. the impella cp had been in for approximately 20 hours when it was explanted due to limb ischemia and a fasciotomy was performed and blood products were infused. | [5 empty] | The impella cp product was discarded at the user facility and is not available for analysis. requests for further clinical detail and imaging have not been met. should further detail become available a supplemental medwatch will be completed. internal reference (b)(4).
38082 | 1035166-2018-00027 | 7404344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7404344&pc=DYB | 01/12/2018 | 09/07/2018 |  | 04/06/2018 |  |  | 03/15/2018 | 04/06/2018 |  |  | 10/01/2020 |  | 04/06/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-13745 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION | [10 empty] | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | It was reported that on (b)(6) 2018 during a medical procedure, the physician noticed a leak at the valve of the oscor 14f x13, lot # c1-13745. the issue was noted during product insertion at the left femoral artery for placement of the impella cp. when the physician noted a leak, the introducer was replaced and the case continued. patient is reported to be stable after procedure. the patient had been admitted for a high risk pci with history of multivessel coronary artery disease and turned down for surgery by cardiothoracic team. | [5 empty] | A preliminary review showed there were no manufacturing rejects or anomalies of this event type recorded in the device history record. conclusion not yet available as the investigation is on-going. | The device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. one 14f abiomed introducer sheath was returned from the end user without the dilator. the end user returned the introducer to the customer with the two orange split cacps and hemostasis valve attached. the customer removed the two orange split caps and hemostasis valve prior to returning to oscor. oscor received the device in the following manner: one 14f, 13 cm long introducer sheath was returned. there were two orange split caps and a hemostasis valve returned, detached from the sheath's hub. there were no other accessories. visible blood was found on and inside the sheath. the hub was checked for a crack and the valve was checked for damage under a 10x microscope. no crack was found on or inside the hub. the valve looked normal and no damage or fatigue was found. it was not possible to perform a leak test on the product as the valve and caps were already detached. no crack was found on or inside the hub. the valve looked normal and no damage or fatigue was found. per ifu when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. infusion through the sideport can be done only after all air is removed from the unit. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
38083 | 1220648-2018-00023 | 7400364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7400364&pc=OZD | 03/22/2018 | 03/22/2018 |  | 04/05/2018 |  |  | 03/22/2018 | 04/05/2018 |  |  | 11/30/2019 |  | 04/05/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1318488 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 69 YR | [5 empty] | Malfunction | Yes | No | Device Issue | Material Deformation | Positioning Problem |  |  |  |  | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) year old male was admitted in cardiogenic shock and had an impella cp placed for hemodynamic support. while on impella support he had a coronary angioplasty and then when the team was preparing the patient for transport to the icu, they noticed they were unable to position the pump's repositioning sheath due to a split. the team had to reverse the anticoagulation, monitor the patient for his liter of blood lost, and place a femostop to the groin site. | [5 empty] | The impella pump was returned for analysis. the wound closure tip was observed to be damaged. this confirms the reported failure and blood loss. in addition scratch mark were noted at the tip. a capa has been opened to address this failure. the failure mode will be monitored and trended. (b)(4).
38084 | 1220648-2018-00018 | 7395628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7395628&pc=OZD | 03/13/2018 | 03/13/2018 |  | 04/04/2018 |  |  | 03/13/2018 | 04/04/2018 |  |  | 12/31/2019 |  | 04/04/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1322666 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella cp was placed for support for hrpci into a (b)(6) old female patient on (b)(6). after the stent was placed echo revealed bleeding in the pericardium. the patient was brought to the or where 400 units of blood was removed, and a perforation of the left ventricle was discovered. the origin of the perforation is unknown. | [5 empty] | The impella cp was returned for evaluation. there were no abnormalities found, and no kinks on the cannula or pigtail were detected. according to the clinical case the physician noted the patient had fragile tissue at the apex, and the patient had an apex infarction in the past. the root cause of the ventricle perforation is likely related to patient condition.
38085 | 1035166-2018-00022 | 7393771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7393771&pc=DYB | 03/02/2018 | 10/05/2018 |  | 04/03/2018 |  |  | 03/05/2018 | 04/03/2018 |  |  |  |  | 04/03/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  |  |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 64 YR | [5 empty] | Death | Yes | Yes | Break | [6 empty] | Cardiomyopathy | Death | No Known Impact Or Consequence To Patient | [10 empty] | Death | [28 empty] | Initial submission | Followup | Followup | Followup |  |  | N | Invalid/Incomplete |  |  | The customer provided the additional information: patient is a (b)(6) old male. within representative notes on (b)(6) 2018 the patient expired but unsure what day, as it occurred over the weekend. there was no extended hospitalization. the patient had a previous catheterization and the team knew the vasculature was tortuous. imaging was requested by the customer but none has been/or will be shared. it was reported there was profuse bleeding; the amount of blood loss is unknown however the patient received four (4) units. upon insertion the bleeding was seen from the valve and immediately the introducer was replaced with the 14 x13 oscor. it was reported the orange portion of the hub came off. the product issue was resolved by using a second oscor introducer. the additional devices were used during the procedure a second introducer, a 10 mm balloon, and a covered stent. the customer was not given any details of the manufacturer, model or serial numbers, or lot numbers of the balloon or stent for the additional devices. the lot number of the introducer is unknown used in this procedure, however the customer was able to provide the model number 0052-3025. | The customer reported: what is the date of event? (b)(6) 2018. what is the name of the facility where event occurred? (b)(6). what is the outcome of the patient? (e.g. patient stable, patient alive, no injury, serious injury, death) stable, on impella support at this time. what procedure was the product used for? insertion at the right femoral artery. what was the condition of the patient when presented to the physician? the patient has cardiomyopathy and "tortuous iliacs". was there an issue with the product ? if yes , then provide details of the issue and answer the questions a) and b). yes, the long 14fr x 25 cm oscor introducer had the orange portion of the hub come off. when was the issue identified? (before , during or after the procedure) in insertion there was profuse bleeding and the physician noticed the orange hub was off. what was the user (physician/ doctor/ nurse) doing when the issue occurred? inserting at rfa. please provide the contact information of the physician who used the product to obtain additional clinical information of the event in question. n/a what is the model number and lot no. of the device used? unknown at this time. was there any additional devices used during the procedure? 3 sheaths in total, all oscor. what is the name, model no. and lot no. for any additional device used during the procedure? unknown lot at this time. in which sequence the devices were used if multiple devices were involved? 14x25 cm oscor, bleed/remove, 14x13fr oscor for hemostasis, noted external iliac artery perforation, 10mm balloon, covered stent, new 14x25 cm oscor, impella 2.5. are the devices available to return to oscor for evaluation? product hold at hospital at this time. |  |  |  |  | As per additional information received march 2018 was reported date of death, actual day of death is unknown. device evaluation anticipated, but not yet begun. conclusion not yet available as the investigation is on-going. a follow-up will be submitted as soon as the investigation is complete. | Was intentionally left blank. information is currently unknown. conclusion not yet available, investigation on-going. | A review of device history record (dhr) could not be performed as the customer did not provide a lot number, however, inspection procedures require any oscor product to pass all in-process and qa final inspection before shipping to the customer. one 14f, 25cm long abiomed introducer sheath was returned from the customer with the dilator. there were two orange split caps returned, detached from the sheath's hub. there were no other accessories. visible blood was found inside the sheath. the returned product was found to have adhesive on hub side a along with orange material from the split cap. hub side b and both orange split caps showed no signs of visible adhesive residue. the returned device's orange split caps were separated from the device. the device had a small amount of glue on one side of the hub and no evidence on the other side of the hub or on either of the orange split caps. the lot number was not provided so we could not determine if the lot had any manufacturing rejects. all lots are visually inspected for properly placed and secure split caps and each lot performed destructive testing per a sampling plan to ensure the split cap remains fixed to the sheath. based on the returned device analysis, it appears that the lack of glue on the hub/orange cap could have contributed to this failure. no manufacturing defects were found. based upon the investigation, a corrective and preventive action has been initiated to address this issue. oscor will continue to monitor this event type and risk. the event will be re-evaluated if additional information becomes available. per qa adelante s2s introducer sheath in-process and final inspection procedure: "with naked eye at a distance of 12" to 18", verify the sheath hub and split caps are free of excess silicone oil applied between the seals during assembly. verify split caps are properly placed and secured onto sheath hub and free of cracks/damages or excessive adhesive." additionally, destructive testing is performed per sampling plan. inspector is to "insert dilator into sheath, lock dilator hub lock nut. hold dilator hub with one hand and one of the sheath hub wing with the other hand. bend the dilator hub to the side until dilator hub lock nut is disconnected from the sheath hub. split caps should remain in place. there should be no damage and fracture to split cap and tabs." the instructions for use (ifu) informs the user: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. | Per instructions for use (ifu): once assembly is fully introduced into the vascular system, separate the dilator cap from the sheath valve housing by rotating the dilator cap off the hub.
38086 | 1220648-2018-00025 | 7388219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7388219&pc=OZD | 03/06/2018 | 03/06/2018 |  | 04/02/2018 |  |  | 03/06/2018 | 04/02/2018 |  |  | 11/30/2018 |  | 04/02/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1267104 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | A patient had an impella cp placed for support for cardiogenic shock post cabg. within 6 hours of impella placement, at the right femoral artery, the limb showed signs of ischemia. as the leg was cold and lacked pedal pulses, the impella was explanted and ecmo was placed instead. | [5 empty] | The product and data logs have not yet been returned from (b)(6) facility for evaluation. the clinical account is still being further clarified. upon completion of investigation, a final medwatch will be completed. (b)(4). | Since the filing of the initial medwatch, the investigation has completed. the device was not returned for analysis, nor has the data logs. the device history analysis revealed no other similar failure mode complaints from the lot of pumps. the root cause of the limb ischemia was unable to be determined. (b)(4).
38087 | 1220648-2018-00019 | 7370378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7370378&pc=OZD |  | 03/15/2018 |  | 03/26/2018 |  |  | 03/15/2018 | 03/26/2018 |  |  | 11/30/2019 |  | 03/26/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1315475 |  | 0048-0026 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Injury | Patient Problem/Medical Problem | [10 empty] | Required Intervention | Disability | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella cp was placed via right femoral artery for support in a (b)(6) female patient for ami/cgs on (b)(6). the pump supported the patient through the intervention. after the intervention the patient had experienced lack of pedal pulses in both sides, and ischemia in the right leg. the patient was weaned from support and the pump was explanted on (b)(6). it was reported to the abiomed field team that the patient's right leg had been amputated. | [5 empty] | The device was not returned for evaluation and there is no indication that the device contributed to the adverse event.
38088 | MW5076088 | 7369253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7369253&pc=KFM | 03/16/2018 | 03/22/2018 |  | 03/23/2018 |  |  |  | 03/23/2018 |  |  | 07/31/2019 |  | 03/23/2018 |  | IMPELLA PURGE CASSETTE/TUBING |  |  | ABIOMED IMPELLA 5.0 |  |  | KFM |  |  | IMPELLA 5.0 |  |  | ABIOMED |  |  | US |  |  |  | 221791 |  | No answer provided | [5 empty] | No | Health Professional |  |  | Voluntary report | [14 empty] | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Information | [41 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | The purge cassette tubing was loaded into the impella driver. there was leaking noted after the disc snapped into place at the time of priming tubing. the disc was removed and fluid was seen draining alongside opening. the disc was examined and noted to be delaminated and leaking on the posterior portion.
38089 | 1220648-2018-00015 | 7363553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7363553&pc=OZD | 02/27/2018 | 02/28/2018 |  | 03/22/2018 |  |  | 02/28/2018 | 03/22/2018 |  |  | 09/30/2019 |  | 03/22/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1305355 |  | 005060 |  | Yes | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella 5.0 was placed with difficulty for support for cardiomyopathy on (b)(6). according to the clinical account, advancement of the pump into the aorta was not possible despite several attempts. the patient became unstable during this process and was bleeding from the axillary site. the decision was made to place the device via open chest through the aorta and support was initiated. the patient received blood replacement products and stabilized. | [5 empty] | The impella 5.0 was returned for investigation, however the introducer that was used durng the axillary attempt was not returned. a review of the clinical case revealed that axillary insertion was difficult therefore the pump was inserted via open chest. the patient had bleeding issues requiring blood and plasma replacement products, and the bleeding subsided on the day of implant. evaluation of the pump did not reveal any abnormalities. the root cause of the access site bleeding is unable to be determined as the introducer could not be inspected. a review of the device history did not reveal any nonconformaties.
38090 | 7356683 | 7356683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7356683&pc=PBL | 03/08/2018 | 03/15/2018 | 03/15/2018 | 03/21/2018 | 03/15/2018 | 03/15/2018 |  | 03/21/2018 |  |  | 12/31/2019 |  | 03/21/2018 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | PBL |  |  | 0048-0003.0814 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003 |  | Yes | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 71 YR |  |  | 109 |  |  | Malfunction | No | Yes | Device Operates Differently Than Expected | Infusion or Flow Problem | [5 empty] | Death | [12 empty] | Other | [28 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The impella cp catheter failed while implanted in patient. the flow started out between 3.2 and 3.4 and then dropped to 1.2. the catheter was removed, flushed out, and re-implanted, but flow never went above 1.2. it was eventually replaced with a 2.5 l pump that was not sufficient to support the patient. a right axillary cutdown was performed and 5l pump was place. although unrelated to the pump failure, the patient expired in the operating room. per site reporter: (b)(4) reps were present during surgery. the device will be returned for failure analysis.
38091 | 1220648-2018-00016 | 7353343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7353343&pc=OZD | 02/19/2018 | 02/20/2018 |  | 03/20/2018 |  |  | 02/20/2018 | 03/20/2018 |  |  | 04/30/2019 |  | 03/20/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT HEART PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1295141 |  | 005042 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 62 YR | [5 empty] | Injury | Yes | No | Difficult To Position | [6 empty] | Perforation | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | A patient with diagnosed multivessel coronary artery disease was having an impella placed for hemodynamic support during a st elevation mi. the physician made between 3 or 4 attempts to place the 13fr introducer into the right femoral artery, due to disease in the common femoral artery. during these attempts, the physician observed a dissection of the femoral and aborted impella placement. the dissection was treated with balloon angioplasty. | [5 empty] | The investigation is on-going at this time. a supplemental medwatch will be submitted at conclusion of investigation. (b)(4). | Since the initial medwatch was filed, the investigation has completed. the guidewire and impella pump were not returned for analysis. the manufacturing review found no other complaints for the lot of pumps. there were no inspection failures for the lot of guidewires. the root cause of the cfa dissection could not be determined as the pump and wire were not returned. the most likely root cause of the dissection was the occlusion in the cfa resulting in difficult placement. (b)(4).
38092 | 1220648-2018-00017 | 7353369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7353369&pc=OZD | 03/02/2018 | 03/02/2018 |  | 03/20/2018 |  |  | 03/02/2018 | 03/20/2018 |  |  | 05/31/2019 |  | 03/20/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1297263 |  | 005042 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR | [5 empty] | Malfunction | Yes |  | Material Separation | Device Operates Differently Than Expected | [5 empty] | Perforation | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | A cardiomyopathy patient, with known tortuous iliac arteries, was taken to the cath lab for placement of an impella cp. the team chose to use a long 14fr x25 cm introducer sheath due to the tortuous iliac. the introducer had a blood leak shortly after placement, and was replaced. the hub had come off of the introducer and caused a blood loss that necessitated 4 units of blood replacement. after placement of the second introducer, the team noted a perforation had occurred and placed a balloon and covered stent to the area. after the blood loss and perforation were treated, the placement of the impella cp was successful and the cardiomyopathy patient was supported. | [5 empty] | The investigation is on-going at this time. a supplemental medwatch will be filed at conclusion of the investigation. (b)(4). | Since the original medwatch was submitted, the investigation has completed. the introducer was returned for analysis. the orange cap, which was stated to have fallen off of the introducer when the dilator was removed, was compared against known good samples and identified as having less adhesive remaining. this suggests insufficient adhesive was applied. the root cause of the orange hub failure was insufficient application of adhesive. the origin of the introducer is not known, as it is not kitted with the impella 2.5 pump. as such, lot analysis cannot be conducted. the failure mode will be monitored and trended. internal refernce (b)(4).
38093 | 1220648-2018-00014 | 7348269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7348269&pc=OZD | 02/23/2018 | 02/24/2018 |  | 03/19/2018 |  |  | 02/24/2018 | 03/19/2018 |  |  | 09/30/2019 |  | 03/19/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1312629 |  | 005048-JP |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | No | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | Support with an impella 2.5 was initiated for acute myocardial infarction and cardiogenic shock on (b)(6). the patient's urine was reportedly red after he was checked into the intensive care unit. the pump was repositioned and the red urine resolved. the following day the clinical team suspected the patient had low blood volume therefore blood replacement products were administered. the patient's red urine returned at this time and was unable to be resolved despite proper pump positioning. the pump was removed on (b)(6) as the patient had stabilized. | [5 empty] | The impella 2.5 and aic console data logs were returned and inspected. analysis of the device determined there was no damage or defects found that could cause hemolysis. additionally, the pump passed hemolysis testing, therefore there was no evidence the hemolysis was caused by the device.
38094 | 1220648-2018-00021 | 7349476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7349476&pc=OZD | 02/19/2018 | 02/22/2018 |  | 03/19/2018 |  |  | 02/20/2018 | 03/19/2018 |  |  | 09/30/2019 |  | 03/19/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1311597 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | DISTRI |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | Kinked | Difficult To Position | [5 empty] | Perforation | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | The complainant reported that the clinicians were attempting to place an impella cp in a (b)(6) year old female patient with a very small aortic arch. the clinicians successfully placed the impella cp, but upon the attempted removal of the 0.018 guide-wire the cannula kinked. the physicians then tried to reposition the device but was unsuccessful. the pump was removed from the patient, and reinsertion was attempted but the same issue occurred. soon after a tamponade was identified which required emergency surgery. when the patient's chest was opened, a left ventricle perforation was found with massive bleeding which required recirculation of blood via cardiopulmonary bypass. the perforation was successfully repaired and hemostasis was achieved, and a coronary artery bypass graft was successfully placed. following this event, the patient was reported as doing well. the physician reported that he believes that patient size and the patient's small aortic arch contributed to the event. | [5 empty] | The device has been returned to the manufacturer; however the evaluation has not yet been completed. upon the completion of the evaluation of the impella cp a supplemental mdr will be submitted. (b)(4). | The impella cp and the guide-wire were returned for evaluation. an image returned from the case confirms that a kink formed in the cannula close to the pump housing. a visual evaluation was performed on the returned pump and guide-wire. during that evaluation a kink was found on the guide-wire, and was located approximately 25cm from the guide-wire tip. in addition, evidence of bending and kinking was found on the pigtail. the bent cannula, pigtail, and guide-wire all point to the pump and guide-wire having been inserted too deeply into the patient's left ventricle. given the fact that a tamponade formed in that same timeframe, it is most likely that either the impella or the guide-wire caused the tamponade; however, without additional evidence indicating which device injured the patient's heart or additional clinical details, the root cause of this event is unable to be definitively determined. a device history review was performed which revealed that there have been no issues or rework relevant to this failure mode documented in the impella cp's device history record. in addition a complaint review was performed, which found that there had been no complaints filed for other pumps from this same lot. internal reference: spr (b)(4). please note that this supplemental mdr 1220648-2018-00021-01 is the correct supplemental mdr reporting the evaluation results from this event. this report was originally submitted on 05/01/2018 under the incorrect supplemental mdr number of 1220648-2018-00011-01.
38095 | 1220648-2018-00012 | 7329022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7329022&pc=OZD | 02/07/2018 | 02/07/2018 |  | 03/09/2018 |  |  | 02/07/2018 | 03/09/2018 |  |  | 09/30/2018 |  | 03/09/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1310666 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 68 YR | [5 empty] | Malfunction | No | Yes | Device Issue | [6 empty] | No Known Impact Or Consequence To Patient | No Code Available | [40 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that a (b)(6) man was supported with an impella cp for 46.49 hours. after approximately 24 hours of patient support suction alarms occurred. patient support continued for another 22.49 hours. the device was removed from the patient. the patient was alert and oriented. there were no adverse patient effects as a result of the reported suction. the device was returned and the evaluation of the pump revealed large biomaterial in the pump housing and revealed impeller damage consisting of a detached impeller blade. biomaterial ingestion is known to cause suction; however, the biomaterial needed to be removed to complete the device evaluation and the pump could not be run with it in place. although the pump was not run with the biomaterial present during testing, it can be confidently stated that there is a chance this is what caused the suction issues reported in the case. | [5 empty] | The impella cp and data logs were returned for analysis. the pump evaluation revealed large biomaterial in the pump housing, as well as impeller damage consisting of a detached impeller blade. as stated in the event description the device was returned and the evaluation of the pump revealed large biomaterial in the pump housing and revealed impeller damage consisting of a detached impeller blade. biomaterial ingestion is known to cause suction; however, the biomaterial needed to be removed to complete the device evaluation and the pump could not be run with it in place. although the pump was not run with the biomaterial present during testing, it can be confidently stated that there is a chance this is what caused the suction issues reported in the case. the root cause of the breakage was unable to be definitively determined. the failure mode will be trended and monitored. (b)(4).
38096 | 1220648-2018-00009 | 7306953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7306953&pc=OZD | 02/04/2018 | 02/04/2018 |  | 03/01/2018 |  |  | 02/04/2018 | 03/01/2018 |  |  | 08/31/2019 |  | 03/01/2018 |  | TEMPORARY, NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | 00813502010879 IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1306511 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 83 YR | [5 empty] | Injury | Yes | No | Material Separation | Use of Device Problem | [5 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A patient with cardiomyopathy was admitted to the hospital, and taken to the cardiac cath lab for implantation of the impella cp, for hemodynamic support. upon insertion of the impella there was bleeding from the sheath hub. to control the bleed the patient received 2 units of blood and had a manual hold to the access site. the bleed was controlled and the patient was successfully supported. | [5 empty] | The impella cp was returned for analysis. the data logs were not analysed as they do not pertain to the failure mode. upon inspection of the impella there was an observation made of insufficient glue applied at the repositioning sheath. the lack of glue resulted in bond failure and allowed the sheath to move during the case support, and this resulted in the bleeding. when the manufacturing review was done no other complaints have been made against this lot of pump. the failure mode will be monitored and trended. (b)(4)
38097 | 1220648-2018-00013 | 7304742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7304742&pc=OZD | 01/30/2018 | 01/30/2018 |  | 02/28/2018 |  |  | 01/30/2018 | 02/28/2018 |  |  | 09/30/2019 |  | 02/28/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, PRODUCT |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1312624 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 47 YR | [5 empty] | Injury | Yes | Yes | Break | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that the patient was taken to the coronary catheterization lab for removal with bilateral access, catheter cross-over and balloon tapenade for 20 minutes. in addition 10 to 15 minutes of manual compression was performed. following removal of the impella cp it was found that the patient's femoral artery was dissected. the patient was taken to surgery where the arterial repair was successfully performed. the patient's 19 hours of impella cp support was reported as successful and the patient was reported to be in stable condition. | [5 empty] | The device was nor returned for evaluation as it was discarded by the user following patient support. there are no definitive information indicating that the use of the impella was the cause of the femoral dissection. the clinicians reported that they believed that per-close challenges indicated potential arterial calcification and mild peripheral vascular disease. the root cause of vascular damage could not be determined because the pump was not returned to abiomed for investigation. (b)(4).
38098 | 1220648-2018-00008 | 7288543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7288543&pc=OZD | 02/01/2018 | 02/02/2018 |  | 02/22/2018 |  |  | 02/02/2018 | 02/22/2018 |  |  | 11/30/2019 |  | 02/22/2018 |  | TEMPORARY NON-ROLLER LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1320825 |  | 0048-0020 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Kinked | Patient-Device Incompatibility | Human-Device Interface Problem | Torn Material |  |  |  | Bradycardia | Low Blood Pressure/ Hypotension | Blood Loss | [10 empty] | Hospitalization | Required Intervention |  |  |  |  | PERCLOSE | PROTAMINE | [21 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) male with known 3 vessel coronary artery disease and peripheral vascular disease, presenting with bilaterally occluded iliacs, was taken to the cardiac catheterization lab for insertion of the impella cp and angioplasty of 3 coronary arteries. the impella cp was placed through a 14fr introducer that kinked at insertion into the axillary artery. the cp supported the patient successfully for the angioplasty procedure and was then explanted. after the introducer was explanted a perclose was placed and protamine was given to reverse the anticoagulation. the patient had a re-bleed at the axillary site and had a vagal episode in the recovery room. the team gave the patient two units of blood and iv fluids and he was reported as stable and doing well. | [5 empty] | The 14 fr x 25 cm introducer was returned for device analysis. the impella cp and data logs were not returned for this investigation. the introducer was measured and found to be within specifications. the introducer was found to have 2 kinks, as the clinical report had mentioned. in addition there was damage noted at the tip of the introducer as well. the root cause of the vascular damage and associated bleed, was a tear at the tip of the introducer sheath. no corrective action is warranted at this time, as the failure mode was determined to be low risk. the failure mode will be monitored and trended. (b)(4).
38099 | 1220648-2018-00010 | 7285565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7285565&pc=OJE | 01/26/2018 | 01/26/2018 |  | 02/21/2018 |  |  | 01/26/2018 | 02/21/2018 |  |  | 06/30/2019 |  | 02/21/2018 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA RP |  |  | OJE |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1296280 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | Yes | Material Integrity Problem | Positioning Problem | [5 empty] | Blood Loss | Patient Problem/Medical Problem | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | The patient was brought to the cardiac care lab, where a swan-ganz was placed in the rca. a balloon tipped pulmonary artery catheter was placed in the left pulmonary artery. the physician was unable to advance the impella rp and lost wire placement. the pump was backed out, and a portion of the sheath was noted in the cannula of the rp and was exiting through the cannula cage. a new impella rp was obtained and the sheath was exchanged with a new 23fr peel-away sheath. approximately 20 cc's of blood was lost at the femoral vein access site, but did not require any replacement blood products to be administered. the patient was supported with the replacement impella rp for 7.92 hours, when he began to decompensate and was severely anemic and acidotic. the family decided to withdraw care. | [5 empty] | The impella rp was returned for evaluation, although the introducer used during this event was discarded by the user. the evaluation of the returned pump has not yet been completed. (b)(4). | The impella rp was returned for evaluation. the 23fr peel-away sheath used during this event was not returned for evaluation, as it was discarded by the staff following the procedure. during the evaluation no remains or material from the introducer sheath were detected in the cannula, and/or around the inflow and outflow cages. the largest outside diameter (od) of the pump was measured and it was within specification. the rp impella was then inserted through and retracted out of a 23fr introducer sheath using a guide-wire without any problems. the device history record was reviewed and revealed that there are no other complaints relevant for this failure mode and no other complaint opened for another pump in this lot. the lot for the 23fr introducer sheath kit (lot no c1-13045, pn: 0052-3021, (b)(4)) had passed incoming inspection. in conclusion, because the introducer sheath was not returned for evaluation with the pump, the root cause of the impella rp pump damaging the introducer sheath was unable to be determined. no corrective action is recommended at this time because the failure mode was determined to be low risk. failures of this type will continue to be monitored and trended. (b)(4).
38100 | 1220648-2018-00005 | 7282770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7282770&pc=OZD | 01/26/2018 | 01/26/2018 |  | 02/20/2018 |  |  | 01/26/2018 | 02/20/2018 |  |  | 07/31/2019 |  | 02/20/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1299265 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Device Issue | [6 empty] | Low Blood Pressure/ Hypotension | Perforation | Blood Loss | [10 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) male was admitted with influenza. in his workup multivessel coronary artery disease was diagnosed. he was sent for bypass surgery with impella cp being utilized for hemodynamic support. after the bypass grafts were placed, and the patient was taken off-pump, the impella groin access site was seen to have pulsatile blood flow. due to the blood loss, the impella was explanted, and the patient had a surgical intervention and repair of the torn common femoral artery. a new impella cp was placed for continued support at the left femoral artery, via a newly placed 6cm hemashield platinum graft, and his hemodynamic support continued until support was withdrawn at the request of the family. | [5 empty] | The product was returned for analysis. the data logs are not relevant to the failure mode, and were not gathered for analysis. there was damge noted on the 15fr portion of the pump, at the wound closure. the timing of when the damage occurred is unknown. the damage most likely contributed to the bleeding observed during the case support. the root cause of the bleeding is unable to be determined. the failure will be monitored and trended. (b)(4).
38101 | 1220648-2018-00007 | 7282801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7282801&pc=OZD | 12/26/2017 | 01/24/2018 |  | 02/20/2018 |  |  | 01/24/2018 | 02/20/2018 |  |  | 08/31/2019 |  | 02/20/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1307590 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 42 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | [6 empty] | Hematoma | Low Blood Pressure/ Hypotension | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) male was admitted with influenza. upon admission to the hospital he had a cardiac catheterization which revealed multivessel coronary artery disease, had an impella cp placed for support, and an angioplasty was performed. during the overnight hours in the icu, the patient developed a hematoma at the groin site and he became hypotensive. volume, in the form of iv fluids, and blood was given to the patient and the team explanted the impella cp with the help of the surgical team. the patient had a new impella 5.0 placed axillary and his hemodynamic support continued. | [5 empty] | The root cause of the patient's hematoma could not be determined, as the product and data logs were not returned. a manufacturing review found no other complaints for this failure mode made against the lot of impella pumps. failures of this type will be monitored and trended. (b)(4).
38102 | 1220648-2018-00004 | 7253045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7253045&pc=OZD | 01/18/2018 | 01/18/2018 |  | 02/08/2018 |  |  | 01/19/2018 | 02/08/2018 |  |  | 09/30/2019 |  | 02/08/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1310279 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | DISTRI |  |  |  |  | 58 YR | [5 empty] | Malfunction | No | Yes | Human-Device Interface Problem | [6 empty] | Perforation of Vessels | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | An impella cp was to be utilized for support on a high risk pci patient. the 14fr introducer was placed via the right femoral artery. there was difficulty with the placement of the cp through the introducer, and the team noted a vessel damage had occurred. the team chose to explant the cp and perform femoral artery repair with two covered stents. the team proceeded to perform the pci procedure without any impella hemodynamic support. the patient was noted to be stable at case end. | [5 empty] | The investigation is on-going at this time. a supplemental medwatch will be completed at conclusion of the investigation and analysis. (b)(4). | Since the initial medwatch, the product has arrived and the investigation has completed. the introducer was returned and a tip split was confirmed. the damage to the tip was consistent with a difficult insertion. the dilator was slightly bent as well. the data logs were not analyzed as they did not pertain to the failure mode. the manufacturing review found no other complaints leveled against this lot of pump/introducer. no corrective action is warranted. the failure was determined to be low risk. the failure mode will be monitored and trended. internal reference (b)(4).
38103 | 1220648-2018-00003 | 7238301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7238301&pc=PYX | 01/12/2018 | 01/12/2018 |  | 02/02/2018 |  |  | 01/12/2018 | 02/02/2018 |  |  | 10/31/2018 |  | 02/02/2018 |  | TEMPORARY RIGHT HEART SUPPOR BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1269486 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 81 YR | [5 empty] | Malfunction | No | Yes | Kinked | Device Operates Differently Than Expected | [5 empty] | Hematoma | Blood Loss | Patient Problem/Medical Problem | [10 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | A (b)(6) male was admitted for a bypass surgery (cabg) in which five grafts were to be placed. after the cabg an impella rp was placed for patient support. while placing the rp the team kinked two 23fr x 30 cm introducers. the third introducer allowed for successful placement. the rp was being used in addition to an already placed iabp, which was also placed at the right femoral/groin access site. the patient developed a hematoma at the right groin on the following day due to bleeding. the team administered an estimated 20 units of blood product, and three additional purse string sutures were placed at the rp access site in the operating room, to control the access site bleeding. | [5 empty] | For the failure analysis both the introducer product and data logs were returned. the investigation is on-going at this time. a supplemental medwatch will be submitted at conclusion of the investigation. (b)(4) | Since the first filing of the initial medwatch, the investigation has completed. two of the three utilized introducers were returned. both were kinked at the same point along the length of the sheath. without any imaging or evidence of the kinking occurring within the patient's anatomy, the root cause of the kinking could not be determined. the kinks were speculated to result from a hematoma that formed due to the iabp. no corrective action is warranted. the failure was determined to have low risk index based upon very low occurrence and moderate severity. the root cause could not be determined. the failure will be monitored and trended. internal reference (b)(4).
38104 | 1220648-2018-00006 | 7234380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7234380&pc=OZD | 01/24/2018 | 01/26/2018 |  | 02/01/2018 |  |  | 01/26/2018 | 02/01/2018 |  |  | 09/30/2019 |  | 02/01/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1311823 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 43 YR | [5 empty] | Injury | Yes | No | Physical Property Issue | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | An impella cp was implanted on (b)(6). there was difficulty inserting the repositioning sheath completely in, and the physician attempted to reposition it. the patient received two units of blood due to bleeding from the insertion site during the case. upon removal of the impella cp, the repositioning sheath appeared to be damaged. | [5 empty] | The products have not been returned from the user facility. (b)(4). | The impella cp was returned for investigation. the wound closure tip was found to be damaged with two tears noted. a closer inspection revealed there were visible scrape marks on the wound closure, leading directly into the cracks at the tip. the root cause of the access bleeding was the fractured wound closure tip. the root cause of the wound closure damage was unable to be determined. (b)(4).
38105 | 1220648-2018-00001 | 7227258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7227258&pc=OZD | 01/03/2018 | 01/03/2018 |  | 01/30/2018 |  |  | 01/03/2018 | 01/30/2018 |  |  | 08/31/2019 |  | 01/30/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1304914 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 56 YR | [5 empty] | Malfunction | Yes | No | Device Operates Differently Than Expected | Torn Material | [5 empty] | Hematoma | Low Blood Pressure/ Hypotension | Blood Loss | [10 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) male transported himself to the hospital emergency room and promptly coded, went into ventricular tachycardia, and after successful resuscitation, was taken to the cardiac cath lab. an angioplasty was performed with an impella cp being utilized for hemodynamic support. after the angioplasty the patient was sent to the icu with the cp for continued support and monitoring. while in the icu the pump position was checked via echo and repositioned. after repositioning, the team noticed a groin access site hematoma was developing and began to hold pressure at the site. the patient's heart rate increased and blood pressure was decreasing at this time. when the bleeding was unable to be controlled, and with the patient condition deteriorating, the team chose to send the patient to the operating room (or) for evaluation of the groin site. while in the operating room (or), the cardiothoracic surgeon performed a cutdown to the site and the vessel was repaired. the team did not infuse any blood products and noted the patient to be stable after vessel repair and impella cp explant. upon review of the cp post explant the team noted that the repositioning sheath of the cp had a tear of unknown etiology. | [5 empty] | The impella cp product was returned for analysis. the data logs are not applicable to the failure mode and were therefore not a part of the event analysis. upon review of the returned product, there were two tears discovered along the tip of the wound closure. the lot of wound closure had no other complaints leveled against it. the root cause of the bleeding was the tear at the wound closure. there is a capa for this failure mode. the failure will be monitored and trended. (b)(4).
38106 | 1220648-2018-00113 | 7225907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7225907&pc=OZD | 08/18/2018 | 08/18/2018 |  | 01/29/2018 |  |  | 08/18/2018 | 01/29/2018 |  |  | 05/31/2020 |  | 01/29/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP, PRODUCT |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1349680 |  | 0048-0024 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2018 a (b)(6) male with cardiomyopathy was placed on impella cp support via a wireless right axillary artery insertion. the pump was placed and started without issue. the introducer was peeled away and the repositioning sheath was advanced. oozing was observed at the access site, and the act was noted to be 252. no heparin was administered at this time as the physician was waiting for the act to drop to 160-180 before administering heparin in the purge solution. on (b)(6) 2018, after 7.92 hours of patient support the physician noted that the flow to the patient's hand was diminished. the axillary impella cp was then explanted and another impella cp was placed without issue through the patient's left femoral artery. the patient support was reported as successful, with no lasting adverse effect to the patient's limb. | [5 empty] | The device was returned to the manufacturer for evaluation. the guide-wire repositioning unit's outer diameter was measured and found to be within specification. there are no other complaints associated with pumps from this lot. both guide-wire repositioning unit lots associated with this pump lot had passed all inspections. there are no other complaints associated with units from either guide-wire repositioning unit lot relevant to the failure mode. the clinical account suggests that the patient had a very high act at the beginning of the case, and that heparin was not administered in the purge solution while the impella cp was in place. the impella ifu recommends the usage of heparin in the purge solution for proper anticoagulation. the root cause of the limb ischemia was most likely improper coagulation, as heparin was not utilized in the purge system due to the high act at the beginning of the case. | The pump was returned on december 27, 2017 for evaluation for vascular complication requiring a thrombectomy procedure. a review of the clinical case revealed the physician had noted that "all the clot appeared to be well organized and therefore had been in the body for a while." the device was returned in two pieces from the field. the field team stated the repositioning sheath was removed from the body intact, therefore the fracture occurred after explant. the repositioning sheath appeared to be otherwise intact. the root cause of the vascular complication is unable to be determined. a review of the device history did not reveal nonconformities. (b)(4) follow up.
38107 | 1035166-2018-00006 | 7211751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7211751&pc=DYB | 01/19/2018 | 06/29/2018 |  | 01/23/2018 |  |  | 01/22/2018 | 01/23/2018 |  |  | 11/01/2020 |  | 01/23/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14 F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC |  |  | US |  | C1-13834 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION | [10 empty] | Death | Yes | No | Break | [6 empty] | Death | Cardiogenic Shock | [11 empty] | Death | [28 empty] | Initial submission | Followup | Followup |  |  |  | N | Invalid/Incomplete |  |  | 1. what is the date of event? (b)(6) 2018. 2. what is the name of the facility where event occurred? (b)(6). 3. what is the outcome of the patient? (e.g. patient stable, patient alive, no injury, serious injury, death) death. 4. what procedure was the product used for? insertion via the left femoral artery of the impella cp. 5. what was the condition of the patient when presented to the physician? patient was in cardiogenic shock being supported by va ecmo. 6. was there an issue with the product ? if yes , then provide details of the issue and answer the questions a) and b). yes. a. when was the issue identified? (before , during or after the procedure) during implant of the impella cp the team noticed the 14fr peel away sheath was broken/cut sideways. b. what was the user (physician/ doctor/ nurse) doing when the issue occurred? pending more details. 7. please provide the contact information of the physician who used the product to obtain additional clinical information of the event in question. n/a. 8. what is the model number and lot no. of the device used? c1-13834. 9. was there any additional devices used during the procedure? when the first oscor was seen to be broken, a second oscor 14fr was used. 10. what is the name, model no. and lot no. for any additional device used during the procedure? n/a. 11. in which sequence the devices were used if multiple devices were involved? pending more details. 12. are the devices available to return to oscor for evaluation? not yet received from rep in the eu. | Additional information was received on 12 april 2018 that indicated the following: as for the cause of death, the patient prognosis was poor and when he was unable to be resuscitated, he did expire. the reporter provided that since the european representative was not on site for the case, it is only known that 1 dilation was done prior to insertion of the introducer. it was during use that the bleed was observed and the introducer was replaced. although requested repeatedly, no additional information has been received as of date of this report. |  |  |  |  | Conclusion not yet available-evaluation in progress. this initial mdr is being submitted to meet our requirements of reporting. a follow-up will be submitted as soon as the investigation is complete. | A review of device history record (dhr) was conducted and there were no manufacturing rejects or anomalies recorded in the dhr affecting released product. the introducer and dilators passed all in-process and qa final inspection steps before shipping to the customer including visual, dimensional, functional and leak testing. a complaint review of the reported lot found no additional reports involving this lot number. per the abiomed introducer sheath in-process and final inspection procedure, it states the following regarding visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. slight discoloration from tipping process is acceptable. verify proper tip shape according to drawing. in the instructions for use (ifu) it indicates : never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. one 14f abiomed introducer sheath was returned from the customer. blood was found on and inside the sheath and side port. upon evaluation of the returned product, it was found that the introducer most probably encountered an obstruction, either anatomy or calcification during insertion, resulting in an irregular force; likely exceeding the design of the introducer, that caused a portion of the sheath tip to tear pre-maturely and deform inwards. once the introducer tip was damaged, it would be susceptible to pre-mature tearing along score lines once the dilator is removed and/or indwelling device is delivered/withdrawn from introducer. under the microscope, the introducer sheath provided was observed to be torn axially from the tip in two different sections axially down the score-lines. the returned introducer provided was visually inspected and met all requirements except for three split sections that initiated from the tip of the sheath along the score lines. in addition to the 135 mm tear, there is a smaller tear 6 mm on the opposite side of the sheath. using micro vu equipment, the sheath with missing portion was shown to have peeled away. there are obvious signs of stress on the tip and along the adjacent portions of the missing peeled back section. sheath void was shown at 25x magnification with signs of stress at tip and along outer score lines. sheath tip portion shows the tip was bent inward at an extreme angle. returned device analysis reveals one 14f abiomed introducer sheath that is within manufacturing specifications and requirements. oscor engineering reviewed the introducer tipping process and identified the following process/product controls in-place to ensure with adequate confidence and reliability that the tip design requirements are met for each lot of introducers produced. based on these process and product controls, oscor engineering reviewed the dhrs for the tipped introducer lots that went into the finished affected product lot and there were no failures reported for the requirements listed in tipping process controls. therefore, the product met the requirements prior to leaving oscor and there is no indication of risk within the process or design. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk.
38108 | 1035166-2018-00007 | 7213306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7213306&pc=DYB | 12/08/2017 | 05/31/2018 |  | 01/23/2018 |  |  | 12/28/2017 | 01/23/2018 |  |  | 12/01/2019 |  | 01/23/2018 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-12952 |  | 0052-3021 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 85 YR | [5 empty] | Injury | Yes | No | Kinked | Split | [5 empty] | Perforation | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | The customer reported this introducer was being used for insertion of the impella rp at the femoral vein. the patient was presented to the cath lab for rp placement post coronary artery bypass graft (cabg). when attempting to place the 23fr introducer for rp implant, the introducer was seen to be compromised at insertion and kinked/split. there was an iliac vein perforation, which did require stenting and between 4-6 units of blood product replacement. the devices involved with this procedure were used in the following sequence: 23fr introducer, followed by impella rp and a stent of unknown manufacturer. the outcome of the patient is stable. | Additional information was received indicating the following: an lmpella rp was prepared to support an 85 year old male patient in cgs post-cabg. the physician had difficulty manipulating the rp so that the outlet cage sat in the pa. during this manipulation, the 23 fr oscor introducer sheath split and the 11 fr rp catheter exited the introducer. an iliac vein perforation was discovered. the pump and introducer were removed and the perforation was repaired with covered stents. damage was found at the outlet of the rp. placement was aborted. the representative reported that physician said the patient's iliac vein was "very tortuous." |  |  |  |  | Conclusion not yet available, evaluation in process. this initial mdr is being submitted to meet our requirements of reporting. a follow-up will be submitted as soon as the investigation is complete. | There were no manufacturing rejects or anomalies of this event type recorded in the device history records. the introducer sheath passed all in-process and qa final inspection steps before shipping to the customer; including visual, dimensional and leak testing. a review of complaints against this lot number identified no additional reports. per the abiomed introducer sheath in-process and final inspection procedure it is mentioned in the visual inspection section: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. slight discoloration from tipping process is acceptable. verify proper tip shape according to drawing. per the instructions for use (ifu) it advises: never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. one 23f abiomed introducer sheath was returned from the customer. there were no other accessories. blood was found on the exterior and interior of the sheath and side port. upon evaluation of the returned product, it was found that there was a 7cm split in the sheath tube, starting at approximately 9cm from the distal end. there was also a kink in the sheath tube approximately 15cm from the distal end. as mentioned in the complaint, the patient's iliac vein was "very tortuous." putting any excessive stress on the sheath may cause it to kink, deform or break. it is believed this sheath was subjected to stresses beyond its capabilities during use out in the field. returned device analysis reveals one 23f abiomed introducer sheath that is within manufacturing specifications. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc., is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc., which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor, inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor, inc., or its employees, that the report constitutes an admission that the device, oscor, inc., or its employees caused or contributed to the event described in this report.
38109 | 1035166-2018-00004 | 7205341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7205341&pc=DYB | 01/12/2018 | 06/20/2018 |  | 01/19/2018 |  |  | 01/16/2018 | 01/19/2018 |  |  | 09/01/2019 |  | 01/19/2018 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | C1-12794 |  | 0052-3021 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 81 YR | [5 empty] | Death | Yes | No | Deformation Due to Compressive Stress | [6 empty] | Death | Blood Loss | [11 empty] | Death | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | Additional information received on 20june2018 that indicates the family withdrew medical care on (b)(6) 2018 but, although requested, an exact date of death is unknown. | 1. what is the date of event? (b)(6) 2018. 2. what is the name of the facility where event occurred? (b)(6). 3. what is the outcome of the patient? (e.g. patient stable, patient alive, no injury, serious injury, death) patient is stable. 4. what procedure was the product used for? insertion for impella rp to rfv. 5. what was the condition of the patient when presented to the physician? right heart failure post cabg and iabp. 6. was there an issue with the product ? if yes , then provide details of the issue and answer the questions a) and b). yes. a. when was the issue identified? (before , during or after the procedure) during insertion attempt the introducer kinked. b. what was the user (physician/ doctor/ nurse) doing when the issue occurred? gaining access. 7. please provide the contact information of the physician who used the product to obtain additional clinical information of the event in question. n/a 8. what is the model number and lot no. of the device used? c1-12794. 9. was there any additional devices used during the procedure? 2 introducers kinked, both the 23fr oscor. 10. what is the name, model no. and lot no. for any additional device used during the procedure? n/a. 11. in which sequence the devices were used if multiple devices were involved? pending more information from field rep. 12. are the devices available to return to oscor for evaluation? not yet retrieved from field rep. additional information received on 01/17/2018 if there was patient impact then please provide details. · was there any adverse effect to the patient? the artery dissected at the iliac venous junction requiring 3 purse string sutures to control bleeding. it also appeared to the surgeon that this may be the same place where the first two sheaths kinked. · was there any extended hospitalization? the family withdrew care and the patient expired, no extended hospitalization. how was the product issue resolved? the patient had hypercoagulopathy and was bleeding and being given blood products even before impella insertion. when he went back to or the following morning i was told he had been given 20+ units. · used another product? used the 23 fr oscor sheath, in total 3 were used. · used same product? same. · used product of another brand? n/a do you have any images available ? not at this time, though they have been asked for. was the patient vessel calcified ? unknown at this time. he was a (b)(6) male. the cause of death was the family withdrawing care. |  |  |  |  | Additional information has been requested, however not received as of date of this report. date of death of the patient is still unknown at time of this report. a review of device history record (dhr) was conducted and there were no manufacturing rejects or anomalies recorded in the dhr affecting released product. the released device passed all in-process and qa final inspection steps before shipping to the customer. a complaint review of the reported lot number identified two additional reports for unrelated events. according to the abiomed introducer sheath in-process and final inspection procedure, qa performs the following inspection 100 percent: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded foreign material, or any other damages. the instructions for use (ifu) provide these precautions: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. infusion through the sideport can be done only after all air is removed from the unit. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. the device was not returned for analysis, therefore, the clinical observation could not be confirmed. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. | Conclusion not yet available-evaluation in progress. this initial mdr is being submitted to meet our requirements of reporting. a follow-up will be submitted as soon as the investigation is complete.
38110 | 1220648-2017-00131 | 7205713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7205713&pc=OZD | 12/28/2017 | 12/28/2017 |  | 01/19/2018 |  |  | 12/28/2017 | 01/19/2018 |  |  | 10/31/2019 |  | 01/19/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP | [7 empty] | 1314907 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Malposition of Device | [6 empty] | Perforation | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) female was admitted with cardiogenic shock and acute mi. upon admission to the cardiac cath lab she had an angioplasty (pci) of the left anterior descending (lad) coronary artery. after the lad pci was successfully completed the team chose to place an impella cp for hemodynamic support. shortly after the cp was placed there was a "wrong position" alarm and the physician chose to adjust the position within the left ventricle (lv) under fluoroscopic guidance. while moving the cp deeper into the lv there was a feeling of resistance and so the physician stopped advancing the cp. the position alarm resolved at that time, though now the console alarmed for suction. again under imaging guidance, this time echo rather than fluoroscopy, the cp was repositioned and at this time the team noted the lv to be depressed and have an effusion. the team urgently began to treat her for the effusion with a code being called and a pericardiocentesis being performed. the pericardiocentesis withdrew blood from the cavity. the team did infuse blood products back to the patient and took her to the or to repair the lv perforation. the perforation was noted to be on the posterior portion of the heart. it was repaired surgically, the cp was explanted, and the patient was reported to be stable on medical therapy. no other hemodynamic support device was placed. | [5 empty] | The product and data logs have not yet been returned for analysis. requests have been made for more clinical details and imaging have not been met to date. should more information be shared, or product returned, a detailed analysis and subsequent medwatch will be filed. (b)(4)
38111 | 1220648-2017-00130 | 7205840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7205840&pc=OZD | 12/22/2017 | 12/26/2017 |  | 01/19/2018 |  |  | 12/26/2018 | 01/19/2018 |  |  | 04/30/2018 |  | 01/19/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1239018 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Positioning Problem | [6 empty] | Hemolysis | Patient Problem/Medical Problem | [11 empty] | Hospitalization | Required Intervention | Disability | [26 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) year old female patient presented with a history of prior cabg and mitral valve replacement, acute myocardial infarction and cardiogenic shock on (b)(6). an impella cp was implanted successfully on (b)(6) following removal of an iabp. the patient was experiencing leg ischemia on (b)(6). the clinical staff noted the patient's posterior tibial pulses had been dopler audible at baseline, however on the 21st tibial pulses became dopler absent distal to the impella site. ultrasound revealed an occlusive thrombus, however the decision was made not to do intervention at that time due to the grave prognosis. the patient was reportedly very unstable and was shocked several times on the (b)(6), and the family elected to withdraw care. | [5 empty] | The product has not been released by the user facility. if the product is received by the manufacturer for evaluation, a follow up report will be submitted. (b)(4).
38112 | 1220648-2017-00129 | 7205913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7205913&pc=OZD | 12/19/2017 | 12/19/2017 |  | 01/19/2018 |  |  | 12/19/2017 | 01/19/2018 |  |  | 12/31/2018 |  | 01/19/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1273798 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | Yes | Crack | [6 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On (b)(6) 2017 an impella cp was inserted in the left groin of a (b)(6) male patient. the pump was placed without issue. the patient was then removed from the cath lab and was taken to the operating room for by-pass surgery. at this time it was noticed that the pump flows and pressures were dropping. a hematoma was found in the patient's groin. manual pressure was held. it was then observed that blood was actually coming from the repositioning sheath, which was found to be cracked. a cut-down procedure was performed and proximal and distal purse string were applied to the femoral artery. the repositioning sheath was removed from the artery, while the impella cp remained in the patient. a three vessel cabg was successfully performed. the impella cp was sutured to the patient's groin, with a plan to leave the pump in the patient's for a couple of days. pulses were palpable in the affective groin, and the site was reported to look much better with no hematoma, and the area was soft but was much bruised. the impella was removed from the patient on (b)(6) 2017, after 54.98 hours of successful patient support. the physician reported that he was very happy with the patient's recovery and how well he did with the impella cp support. | [5 empty] | The device was returned to the manufacturer for evaluation. the evaluation confirmed that both the desivalve, repo 9f (part number 0046-0552) and tapered wound closure 15f (part number 0046-0517) were found to be damaged. both parts were found to have multiple cracks at their tips and were scratched in the areas behind cracks. this damage is consistent with damage that usually caused by calcified vasculature; however, according to the complainant no calcifications were visualized during this case. a manufacturing review was done, which revealed that there were no material review reports or reworks that apply to this failure mode. in addition, a review of the device history record was performed and it did not reveal any complaints against this lot of impella cp pumps. in conclusion the root cause of the access site hematoma was a damaged repositioning unit. the damage to the repositioning unit has been confirmed, but the root cause was unable to be determined. no corrective action is recommended because the root cause was unable to be determined. (b)(4).
38113 | 1220648-2017-00128 | 7200599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7200599&pc=OZD | 12/20/2017 | 12/20/2017 |  | 01/17/2018 |  |  |  | 01/17/2018 |  |  | 09/30/2019 |  | 01/17/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1310928 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 53 YR | [5 empty] | Malfunction | Yes | Yes | Difficult to Remove | Device Operational Issue | [5 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | An impella 5.0 was implanted into a 53 year-old female patient with cardiomyopathy and elevated filling pressures. the pump was implanted on (b)(6) and provided good support for six days before the pump stopped. the clinical staff reported that on (b)(6), the pump stopped at the same time the patient was being turned in bed. the pump was pulled back into the axillary artery and the patient went to the operating room for placement of a lvad. after the lvad was placed the physician attempted to remove the impella, however during removal the pump separated at the outlet cage. the physician performed a sternotomy to remove the remaining portion of the device. the patient remained stable throughout the procedure without any further complications. | [5 empty] | The product has not been released by the user facility. if the product is received by the manufacturer for evaluation, a follow up report will be submitted. (b)(4). | The impella 5.0 was received on (b)(6) and was evaluated with the console data logs from the case. the data logs confirmed the clinical account, whereby the pump operated for several days before the pump stopped. the pump was investigated and was unable to be started in the lab. a retain pump was used to reproduce the log signature as seen in the case. this was accomplished by disconnecting the motor drive lines to simulate broken lines. a broken drive line would cause a pump to stop, and would cause a log signature similar to what was seen in the logs from the case. the pump was also inspected to determine the cause of the cannula separation that occurred when the pump was removed during the case. an adhesive failure was discovered at the point where the cannula is bonded to the pump housing, which caused the separation. a review of the device history did not reveal any nonconformities.
38114 | 1220648-2017-00126 | 7185613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7185613&pc=PYX | 12/13/2017 | 12/13/2017 |  | 01/12/2018 |  |  | 12/13/2017 | 01/12/2018 |  |  | 09/30/2019 |  | 01/12/2018 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1309261 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 71 YR | [5 empty] | Injury | Yes | Yes | Difficult To Position | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that a (b)(6) year old male patient presented to the hospital a couple of days previous to (b)(6) 2017 after having "c/o indigestion". the patient was having an acute myocardial infarction (ami) with right ventricular failure. the physician scheduled the patient for an impella rp placement procedure. the patient was brought to the cardiac catheterization lab on (b)(6) 2017 where the impella rp was placed via the right femoral artery (rfa) using a swan catheter, and then exchanged the swan for an .027 wire deep into the left pulmonary artery. the pump was tracked over wire and flows reached 4.2 lpm; however upon watching the patient under fluoroscope in the anteroposterior (ap) view, the device placement was in question. a right anterior oblique, radiographic projection was done, which revealed that the posterior anterior view was due to the patient's anatomy, and not because the device was mal-positioned. the physician then attempted to place the device again but a second ap view revealed the same issue. it was then decided to leave the device in the patient and remove the peel-way sheath. the peel-way of the sheath went very well until the mattress suture was deployed. a nurse held the held manual pressure for 10 minutes, but noted gushing of blood as the mattress sutures were not supporting hemostasis. the patient was reported to have lost approximately 2 ml of blood during this event, and an unknown amount of replacement blood products were administered to the patient. another physician attempted to place multiple sutures with deep tissue bites, but the patient's girth and size precluded all attempts to control the bleeding. at this point the repositioning sheath had been manipulated by the suturing attempts and was pulled back. the first physician then attempted to reposition the pump in the pa but was unsuccessful, and the device hinged in the right atrium and came out of the pa. without sheath access and the wire the doctor was unable to reposition the device. the impella rp was removed from the patient and a .035 stiff wire was from another impella rp kit was used to retain access for new sheath placement. this was done by placing the wire through the inlet cage and removing device over the wire. a new 23fr peel-away sheath was placed in the right femoral artery, and another impella rp was placed in the patient. the replacement device successfully supported the patient for 57.62 hours, but the patient was reported to have expired on (b)(6) 2017. | [5 empty] | The impella rp was returned for evaluation; however the introducer and guidewire used in this event were discarded following use. the console data logs were returned, and the analysis of the logs revealed 6 brief continous suction alarms that were triggered after the pump was repositioned. the visual inspection of the impella rp identified evidence of kinking and twisting of the canula. these deformation were likely caused by the repeated attempts to positon the pump due to the patient's unusual anatomy. without the guidewire and introducer used during this event, a more comprehensive evaluation was unable to be performed. a review of the device history record and the lot analysis indicated all applicable systems were deemed adequate and remain effective. in conclusion, the root cause of the positioning issues was the patient's cardiac anatomy. (b)(4).
38115 | 1220648-2017-00136 | 7182095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7182095&pc=OZD | 12/14/2017 | 12/14/2017 |  | 01/11/2018 |  |  | 12/14/2017 | 01/11/2018 |  |  | 09/30/2019 |  | 01/11/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1310563 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) male with a history of multivessel coronary artery disease was admitted for revascularization via an angioplasty of the chronic totally occluded right coronary artery (rca). the patient had a successful procedure, but when being moved from the cardiac catheterization lab, the patient went into ventricular fibrillation and subsequently coded due to a thrombosed rca. the patient was again intervened upon and the rca was re-opened. due to his instability, the impella cp was placed for hemodynamic support, a temporary pacemaker was placed, and a hypothermia protocol was initiated. once the impella cp was placed, the patient was transferred to the icu for monitoring. during his two days in the icu, the impella pump position was confirmed and the patient treated with cardiac iv drips and evaluated by vascular surgery. vascular was consulted due to a cool limb in the leg the impella was placed. the surgeon chose not to treat the leg, rather to allow the impella to remain within the left leg for hemodynamic support until the patient could be weaned from the impella hemodynamic support. the following day the patient's limb was improving but it was still decided to wean the patient from his impella cp support. the weaning process took place over 6 hours. after the weaning completed, the patient had his impella explanted. shortly after explant, the patient began to decompensate and did expire. the cause of the patient's outcome is not known at this time. the patient was in need of further coronary revascularization in the form of coronary artery bypass surgery. | [5 empty] | To date there are no results available. no product has been returned, nor have requests for clinical patient information been satisfied. the data logs have recently arrived for analysis, but the investigation is on-going. when results have been found and/or additional information is discovered, a supplemental report will be filed. internal reference 17-1219
38116 | 1220648-2017-00134 | 7182115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7182115&pc=OZD | 12/13/2017 | 12/13/2017 |  | 01/11/2018 |  |  | 12/13/2017 | 01/11/2018 |  |  | 09/30/2019 |  | 01/11/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1310773 |  | 005082 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 73 YR | [5 empty] | Injury | Yes | Yes | Material Separation | [6 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | Complainant reported that on (b)(6) 2017 following 5 days of successful impella ld support for a (b)(6) year old male patient the pump was removed from the patient. during removal of the pump it fractured where the motor connects to the cannula. the canula and inflow cage became detached and remained in the patient. the vascular team removed the detached portions of the device from the patient via the use of a snare. no other patient intervention was required, and were no adverse effects to the patient as a result of the reported issue. the patient was reported to be in stable condition following the removal of the device. | [5 empty] | The failure of the cannula detaching from the pump housing was unable to be confirmed because neither product nor console data logs were returned; no visual evidence was provided by the user facility of the failure for investigation. a review of the device history and lo history records did not reveal any nonconformities. if additional information is obtained a supplemental report will be submitted. (b)(4).
38117 | MW5074526 | 7181679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7181679&pc=OZD | 12/20/2017 | 01/09/2018 |  | 01/10/2018 |  |  |  | 01/10/2018 |  |  |  |  | 01/10/2018 |  | IMPELLA 5.0 |  |  | IMPELLA 5.0 |  |  | OZD | [5 empty] | ABIOMED |  |  | US | [5 empty] | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 53 YR |  |  | 98 |  |  | Injury | No | Yes | Detachment Of Device Component | Device Dislodged or Dislocated | [5 empty] | Aortic Insufficiency | Failure of Implant | Foreign Body In Patient | [10 empty] | Life Threatening | Required Intervention | [27 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | The device completely stopped functioning while still in the left ventricle. due to this malfunction the cardiologist withdrew the device outside the heart to prevent left ventricular thrombus 10cm from 46 to 36cm from the visible marker on the right chest. later, the patient was taken to the operating room for insertion of a left ventricular assist device, during the removal of the failed device it was noticed that the distal portion of the pump was not attached and this was seen on the tee in ascending aorta and later retrieved. additionally, there was moderate-severe aortic valve insufficiency after the removal. retained device was later found in the aortic arch and retrieved.
38118 | MW5074532 | 7181755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7181755&pc=KFM | 01/05/2018 | 01/09/2018 |  | 01/10/2018 |  |  |  | 01/10/2018 |  |  |  |  | 01/10/2018 |  | IMPELLA HEART PUMP |  |  | IMPELLA HEART PUMP |  |  | KFM |  |  | 0048-003 |  |  | ABIOMED |  |  | US | [5 empty] | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 57 YR | [5 empty] | Injury | No | Yes | Device Displays Incorrect Message | [6 empty] | Cardiogenic Shock | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Impella heart pump was inserted on (b)(6) 2018 into a pt suffering from shock after a stemi. the device began alarming (b)(6) 2018, cardiologist and rep from abiomed attempted to troubleshoot the alarms in the cvicu without success. the pt was transported to the cardiac cath lab for possible repositioning and or reinsertion. abiomed rep remained assisted troubleshooting in cardiac cath lab. a new device was inserted. the original device was removed in fluoro guidance and the pt was transported back in cvicu. the abiomed rep retained the original device for reimbursement and investigation purposes. dates of use: (b)(6) 2018. diagnosis or reason for use: cardiogenic shock s/p stemi.
38119 | 1220648-2017-00127 | 7172448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7172448&pc=OZD | 12/03/2017 | 12/21/2017 |  | 01/08/2018 |  |  | 12/21/2017 | 01/08/2018 |  |  | 08/31/2019 |  | 01/08/2018 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1296055 |  | 0048-3074 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Malfunction | No | Yes | Device Operational Issue | [6 empty] | No Known Impact Or Consequence To Patient | No Code Available | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) female with a history of coronary artery disease, treated by stenting to the left main, left anterior descending, and left circumflex, 9 months prior was admitted to an outlying facility and was then transferred to a tertiary hospital for care. upon admission to the tertiary hospital and being sent to the cardiac catheterization suite, the impella cp was placed for hemodynamic support via the right femoral artery. within the first minutes of support, the aic was alarming for suction and the team troubleshot by removing the impella cp to look at the pump for the presence of any biomaterial and to flush it accordingly. the pump was then re-inserted and flows were seen to be less than optimal. the pump was again removed, and at this point the field representative arrived to assist the hospital team. the team chose to remove the cp and advance a new pump to position in the left ventricle. the second impella cp successfully supported the patient for the angioplasty, was weaned, and explanted. there was no lasting harm or injury from any delay in successful support initiation. | [5 empty] | Both the impella cp pump and data logs were returned for analysis. the data logs confirmed the clinical report of flows initially being appropriate, but then suction alarms were triggered after 8 minutes of support. the suction alarms were unable to be resolved. the pump was analyzed and broken down within the testing environment. upon microscopic examination there was an observation made of a broken impeller blade. the root cause of the low flows and suction was most likely the result of a broken impeller blade. this failure mode will be monitored and trended. in addition, an incident investigation was initiated to address failures of the type. (b)(4).
38120 | 1035166-2018-00003 | 7172452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7172452&pc=DYB | 12/14/2017 | 06/15/2018 |  | 01/08/2018 |  |  | 12/15/2017 | 01/08/2018 |  |  | 09/01/2020 |  | 01/08/2018 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC |  |  | US |  | C1-13677 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION | [10 empty] | Injury | Yes | No | Backflow | [6 empty] | Hematoma | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | The customer wanted to notify the manufacturer of the information they have regarding this event; hope to have more details from the field representative soon. date of event (b)(6) 2017. what is the outcome of the patient? patient stable. what procedure was the product used for? insertion for the impella cp; for a high risk pci. what was the condition of the patient when presented to the physician? patient presented for hrpci due to surgical turndown. patient history of hypertension, diabetes, shortness of breath. when was the issue identified? during procedure. hematoma developed. what was the user (physician/ doctor/ nurse) doing when the issue occurred? utilizing the impella cp for hemodynamic support. was there any additional devices used during the procedure? none. in which sequence the devices were used if multiple devices were involved? pending information from rep. are the devices available to returned for evaluation? not yet returned from field. | Additional information received includes the following: device was used in a procedure for insertion of the impella cp on a patient with high risk percutaneous coronary intervention. the reported issue of blood loss/ leak is from the seam of the peel away sheath. the issue occurred during the procedure when the user was utilizing the impella cp for hemodynamic support. the team infused 1 unit of packed red blood cells due to blood loss estimated at 150cc. the patient developed a hematoma at the groin site, which the physician attributed to the leak at the peel away sheath. the patient is reported to be stable. |  |  |  |  | Conclusion not yet available, evaluation in process. this initial mdr is being submitted to meet our requirements of reporting. a follow-up will be submitted as soon as the investigation is complete. | The following sections were updated in follow-up #1: sex, date of report, event, concomitant medical products, date rec¿d by mfr, pma/510k, if follow-up, what type and device evaluated by mfr. a review of device history record (dhr) was conducted and there were no manufacturing rejects or anomalies recorded in the dhr affecting released product. there was one (1) reject out of (b)(4) units found for the reported lot that failed the leak test, however that specific piece/ device was not shipped. the released introducer sheath and dilator passed all in-process and qa final inspection steps before shipping to the customer including visual, dimensional and leak testing. a complaint review of the reported lot identified one additional report but for an unrelated event. per the vacuum leak test qa procedure, it states the following: perform pressure and vacuum leak test per established procedure using the latest revision. this procedure is performed by qa 100 percent. in the instructions for use (ifu) it indicates: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. one 14f abiomed introducer sheath was returned from the customer. there were no other accessories. traces of blood were found on and inside the sheath. upon evaluation of the returned product, the sheath including the hemostasis valve looked normal and there were no splits found down the sheath tube. no leakage was found when manually leak tested at the site where the sideport tubing and stopcock attaches to the hub. the abiomed introducer was then tested using the iso 11070:2014 liquid leakage test. the device was connected at the distal tip and pressurized to 38kpa and then examined for liquid leakage. the part passed the leak test, and a cut off dilator with a hub lock was connected to put pressure on the seal. (if this is not done, the seal will bow outward from the pressure and eventually leak from the slit in the seal due to the bowing.) the part was ramped up to 300kpa for 30 seconds and examined for leakage with the following results: no leakage observed at 38kpa, no leakage observed at 300kpa. returned device analysis reveals one 14f abiomed introducer sheath that is within manufacturing specifications. the device passed leak testing was performed per applicable iso standard. based on the analysis, the complaint cannot be confirmed. no manufacturing defects were found. a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc., is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc., which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor, inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor, inc., or its employees, that the report constitutes an admission that the device, oscor, inc., or its employees caused or contributed to the event described in this report.
38121 | 1220648-2017-00121 | 7163853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7163853&pc=OJE | 12/08/2017 | 12/08/2017 |  | 01/04/2018 |  |  | 12/08/2017 | 01/04/2018 |  |  | 03/31/2019 |  | 01/04/2018 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA RP |  |  | OJE |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1282657 |  | 0046-0019 |  | Device was returned to manufacturer | [5 empty] | No | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 85 YR | [5 empty] | Injury | Yes | No | Material Perforation | [6 empty] | Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An (b)(6) year old male patient presented to the cardiac cath lab post-cabg and with decreasing cardiac index. the decision was made to implant an impella rp for support. the patient's iliac vein was noted to be very tortuous according to the implanting physician. as the physician attempted to manipulate the pump into place, the peel away introducer was damaged, and the catheter portion of pump exited the side of the damaged introducer and perforated the patient's iliac vein. the patient's vein was surgically repaired with covered stents and blood replacement products were given. the decision was made to abort the placement of the impella rp as the patient's condition had improved to the point where there was no immediate need for other support. | [5 empty] | The impella rp was returned for investigation. the struts of the outflow cage were found to have failed at the solder joint between the strut and ball of the outflow cage. according to the physician, the patient's iliac vein was "very tortuous." the damage to the struts on the outflow cage may have occurred while the pump was passing through the introducer, which had kinked due to the iliac tortuosity. the root cause of the pump damage is likely delivery through the kinked introducer as a result of the iliac tortuosity, however this cannot be confirmed. therefore, the vascular perforation was likely caused by the damaged struts of the outflow cage, however this cannot be confirmed. a review of the device history did not reveal any abnormalities. (b)(4).
38122 | MW5074247 | 7147681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7147681&pc=KFM | 12/13/2017 | 12/22/2017 |  | 12/26/2017 |  |  |  | 12/26/2017 |  |  | 10/31/2019 |  | 12/26/2017 |  | IMPELLA 5.0 HEART PUMP |  |  | IMPELLA 5.0 |  |  | KFM |  |  | 005062 |  |  | ABIOMED |  |  | US |  | 005062 |  | 005062 |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 75 YR | [5 empty] | Malfunction | No | Yes | Crack | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Abiomet impella 5.0 pump purge tubing was discovered to be cracked. needle inserted into purge set in sicu as per instruction of company rep. pump unit identified as sn (b)(4). purge filter released to company rep by sicu nurse; (b)(6) 2017: pt improving, impella discontinued in main operating room. impella 5.0 set, ref: (b)(4).
38123 | 1220648-2017-00117 | 7140206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7140206&pc=OZD | 11/22/2017 | 11/23/2017 |  | 12/21/2017 |  |  | 11/23/2017 | 12/21/2017 |  |  | 06/30/2019 |  | 12/21/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1294891 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that a (b)(6) male patient was being treated at the facility as a result of having experienced 2 successive days of internal cardiac defibrillator (icd) shocks of increasing frequency due to a ventricular tachycardia heart rhythm. the patient was reported to have a history of multiple comorbidities including a coronary bypass graft performed 22 years ago, hypertension, hypersensitivity lung disease, chronic kidney disease, diabetes mellitus, chronic myofascial pain, cardiomyopathy, and an implanted biventricular pacemaker and cardioverter-defibrillator. upon admission to the hospital the patient had an ejection fraction of 15%, and his act was over 300. on (b)(6) 2017 a ventricular tachycardia ablation was performed in the ep lab. after the physician burned a couple of sites during the procedure the patient went into a pulseless electrical activity (pta) and arrested. cpr was then initiated. the impella cp was then placed in the patient via the left femoral artery (lfa) without difficulty. during the impella support the patient would drift into period of no ejection, then no contractility whatsoever, requiring further compressions. the decision was then made to escalate to a v-a ecmo the impella peel-away sheath was removed and the repositioning sheath was advanced. during this time a small hematoma developed immediately. manual pressure was held, and the site was then soft and looked good. when the ecmo was placed, and after 10,000 units of a heparin bolus was administered, the hematoma rapidly expanded with continued bleeding around the repositioning sheath. manual pressure continued to be held, and vascular surgery was called. the ecmo was completed and the impella purge pressure was reduced for left ventricle decompression, exploration by the vascular surgeon revealed bleeding all around the repositioning sheath. the impella cp was then removed from the patient after 4.85 hours of hemodynamic support. the vessel was repaired and a chimney graft was placed. another impella cp was placed for left ventricle venting. patient blood loss was estimated to be 2,000 cc's, although the patient's act was over 300/second. the patient was administered 6 units of replacement packed red blood cells (prbc.) the patient was eventually transferred to another facility for further treatment. there was no indication from the complainant of any additional adverse effects to the patient as a result of the reported issue. | [5 empty] | The impella cp and other disposables were discarded following removal from the patient; consequently an evaluation of the pump at issue could not be performed. a console data log analysis was not included in this report because data logs are not relevant to this failure mode. without the device or repositioning sheath returned for investigation the cause of the femoral artery damage could not be determined. a manufacturing review was performed for the impella cp used during this event, which revealed that there were no material review reports or reworks that apply to this failure mode. in addition, a review of the device history record was performed and it did not reveal any complaints against this lot of impella cp pumps. the repositioning unit was from lot number 1282116, and its tapered wound closure was from lot number 1701271. all units selected from this lot passed incoming inspection. there were no complaints for failures of other repositioning units from this lot. the root cause of the hematoma was unable to be determined as the product was not returned from the field.. no corrective action is recommended as the failure mode was determined to have a low risk index based on moderate severity and very low occurrence and the root cause was not determined. failures of this type will continue to be monitored for trends. failures of this type will continue to be monitored for trends. internal reference: (b)(4).
38124 | 1220648-2017-00118 | 7129778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7129778&pc=OJE | 11/27/2017 | 11/28/2017 |  | 12/19/2017 |  |  | 11/28/2017 | 12/19/2017 |  |  | 10/31/2018 |  | 12/19/2017 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA RP |  |  | OJE |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1273471 |  | 0046-0011 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | DISTRI |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | Yes | Difficult to Remove | Material Separation | [5 empty] | Cardiogenic Shock | Device Embedded In Tissue or Plaque | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | Italian complainant reported that on (b)(6) 2017 an impella rp was placed in a (b)(6) female patient. the patient was reported to have been successfully supported, per the case information relayed through european distributor. on (b)(6) 2017, day 9 of patient support, the patient was weaned from the rp device. the physician tried to remove the device, but damaged the catheter in his attempt and fractured it into two pieces during removal. as a result, the detached portion of the cannula and pump remained in the patient's iliac vein, which necessitated surgical removal of the portion of the cannula. the detached portion of the device was successfully removed from the iliac vein. the patient was reported to be in cardiogenic shock during the surgical removal of the device. an ecmo was placed in the patient for hemodynamic and respiratory support subsequent to impella rp explant. the patient was reported to be in critical condition. there was no additional information available regarding the patient's condition. | [5 empty] | The impella rp was returned for evaluation, with the cannula detached from the motor housing. during the evaluation the silicone was found over the delo bead that was lined up with the differential pressure sensor, and the hydex ring was deformed. while the particular circumstances of this failure is unknown, silicone contamination and deformation of the hydex ring is known to reduce the strength of the pump housing/cannula bond. this could predispose it to failure under the stress applied during pump explant. in this instance is not clear whether the hydex ring deformation occurred prior to the intervention to extract the cannula from the patient's iliac vein. (b)(4) was created to address rp cannula detachments. in conclusion the root cause of the cannula detachment was unable to be determined, but silicone contamination may have contributed to the failure. (b)(4). | The impella cp pump product and data logs were returned for analysis, however the investigation is on-going at this time. a supplemental medwatch will be completed upon investigation completion. (b)(4).
38125 | 1035166-2017-00112 | 7131732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7131732&pc=DYB | 11/29/2017 | 05/25/2018 |  | 12/19/2017 |  |  | 11/30/2017 | 12/19/2017 |  |  | 08/01/2020 |  | 12/19/2017 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-13629 |  | 0052-3025 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 91 YR | [5 empty] | Injury | Yes | No | Backflow | Kinked | [5 empty] | Hematoma | Hemostasis | Blood Loss | [10 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | There was a complaint made yesterday (b)(6) 2017, regarding an oscor introducer/impella support case. what is the date of event? (b)(6) 2017. what is the name of the facility where event occurred? (b)(6) hospital, (b)(6). what is the outcome of the patient? (e.g. patient stable, patient alive, no injury, serious injury, death) patient is stable. what procedure was the product used for? implant of the impella cp what was the condition of the patient when presented to the physician? the patient was admtited for a high risk percutaneous coronary intervention (pci) for a chronic total occlusion of the right coronary artery (rca). was there an issue with the product ? if yes , then provide details of the issue and answer the questions a) and b). there was a bend/kink at the tip of the 14x25 introducer. they then placed the 14x13 introducer and there was a constant ooze at the diagphragm. awaiting more case particulars about the amount of blood lost from the ooze and the resulting hematoma. when was the issue identified? (before, during or after the procedure) during. what was the user (physician/ doctor/ nurse) doing when the issue occurred? placement of the 2 introducers. please provide the contact information of the physician who used the product to obtain additional clinical information of the event in question. n/a what is the model number and lot no. of the device used? both the 14x25 and 14x13 were used (b)(4). was there any additional devices used during the procedure? when the case was completed the arterial closure device failed, and so the groin access site had a competitor's 14x30 introducer placed until hemostasis could be achieved. in which sequence the devices were used if multiple devices were involved? serial dialtions, 14x25, 14x30, impella cp, arterial closure device, a competitor's 14x30. are the devices available to return to the manufacturer for evaluation? product has been requested from the field but, not yet returned. received additional information on (b)(6) 2017: what is the age, sex, and initials of the patient? the patient is a (b)(6) male. did the patient experience any extended hospitalization? if yes, please provide details. will enquire regarding the patient discharge etc. what date was the pump explanted? (b)(6) 2017. customer still inquiring about amount of blood loss | [5 empty] | A review of device history record (dhr) was conducted and there were no manufacturing rejects or anomalies recorded in the dhr affecting released product. the sheaths released to inventory passed all in-process and qa final inspection steps before shipping to the customer, customer including visual, dimensional and leak testing. as per the procedure, adelante s2s introducer sheath in-process and final inspection, it instructs "with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages." as the procedure introducer set, silicone seal slitting states, "the operator visually inspects all the slit seals for defects 100 percent with a microscope. also, a leak test is performed on 100 percent of the introducer sheaths by qa per the leak test - introducer with seal- electronic tester procedure. two 14f abiomed introducer sheaths, one 13 cm and one 25 cm in length were returned from the customer along with one 33 cm long dilator. there were no other accessories. blood was found on all components. upon evaluation of the 25 cm long sheath, there was damage found around the circumference of the distal tip edge that looked like a bend/kink as stated in the complaint. the distal tip of the 33 cm dilator is curved as well. putting any excessive stress on the sheath may cause it to kink, deform or break. it is believed the sheath and dilator were subjected to stresses beyond their capabilities during use out in the field. upon evaluation of the 13 cm long sheath, it was found that the insertion point into the hemostasis valve was identified to be considerably off center of the helical cuts, creating a hole in the valve. this caused the blood to ooze out of the hemostasis valve. it is stated in the instructions for use (ifu) never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. it also indicates "when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane." returned device analysis reveals two 14f introducer sheaths and one dilator that are within manufacturing specifications. the reason for return cannot confirmed. no manufacturing defects were found. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, there was no new failure mode identified and the risk remains acceptable for both events reported. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. | Conclusion not yet available, evaluation in process. this initial mdr is being submitted to meet our requirements of reporting. a follow-up will be submitted as soon as the investigation is complete.
38126 | 1220648-2017-00120 | 7107989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7107989&pc=OZD | 11/20/2017 | 11/20/2017 |  | 12/12/2017 |  |  | 11/20/2017 | 12/12/2017 |  |  | 06/30/2019 |  | 12/12/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1298129 |  | 0048-0026 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 84 YR | [5 empty] | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that an (b)(6) male patient was brought to the ccl on (b)(6) 2017 for an emergent placement of an impella cp. the patient was reported to be suffering from severe takotsubo, mild coronary artery disease and severe mitral regurgitation. the impella cp was placed at the left groin, without issue. during placement the physician employed a long peel-away 14french oscor introducer sheath to bypass the diseased iliac. following insertion, the physician attempted to remove the peel-way introducer sheath from the groin, but reported that he was unable to remove the sheath completely from the common femoral artery (cfa) before cracking the sheath because it was situated up against the pump plug. as a result the physician was forced to start cracking the peel-away sheath while a portion of the sheath still resided inside the patient's vasculature. this resulted in patient bleeding, with the patient's blood loss being estimated to be approximately 1 unit (500 ml). it was determined that no replacement blood products were necessary to administer to the patient as a consequence of the blood loss. the patient was successfully supported for 3.76 days with the impella cp. the physician reported that the patient's condition had improved drastically, and was weaned from the impella cp. the impella cp was explanted from the patient without issue and was reported to be doing well. | [5 empty] | The abiomed impella cp and the oscor 14 french introducer were discarded by the physician and staff following patient use; consequently no device evaluation was able to be performed. . the console data logs were not relevant to this failure mode. the device history record was reviewed. the introducer used during this event was part of batch c1-13578. this batch passed incoming inspection as inspection lot 10000079342. the review revealed one additional complaint had been submitted for a broken hub cap on a 14x25 introducer sheath from the same batch. no other complaints have been filed for other pumps from this lot. a risk assessment was performed, and due to very low occurrence and moderate severity, this failure mode has been determined to be low risk (occurrences = 8; usage = 12463; rate = .064%) in conclusion, and from the clinical description of the event provided by the complainant the root cause of the patient injury is believed to be the result of the peeling away of the sheath while it was still inside of the patient's vasculature. the reason the sheath was peeled away before completely removing it from the patient's vasculature is because the top of the hub was pressed against the pump plug and the sheath could not be retracted any more. corrective action: no corrective action is recommended because this failure mode has been determined to be low risk index due to very low occurrence and moderate severity. the failure will continue to be monitored and trended. (b)(4).
38127 | 1220648-2017-00114 | 7108981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7108981&pc=OZD | 10/11/2017 | 11/14/2017 |  | 12/12/2017 |  |  | 11/14/2017 | 12/12/2017 |  |  | 07/31/2019 |  | 12/12/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1302006 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Malfunction | No | Yes | Device-Device Incompatibility | [6 empty] | No Known Impact Or Consequence To Patient | No Code Available | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) year old male was admitted for a tavr (transcatheter aortic valve replacement) on (b)(6) 2017. the cardiac team successfully deployed the new valve and subsequently the patient arrested, was coded for 30 minutes, and was resuscitated. due to his cardiac instability and need of hemodynamic support, the impella cp was placed via left femoral artery access. the impella cp had good pump flows, but prior to leaving the operating room, he began to have suction alarms so the team attempted to reposition the cp under echo guidance. the repositioning did not remedy the flow and suction alarms, and so the team replaced the cp with another impella pump. the second pump was an impella 2.5 due to inventory at the facility, rather than another cp pump. the second pump supported the patient and allowed for him to be taken out of the operating room and be admitted to the icu. the patient had ventricular tachycardia and needed cpr as he coded again, within the first 3 hours of admission to the icu. the family did choose to withdraw care at this time. | [5 empty] | Both the impella cp pump and data logs were returned for analysis. the data logs confirmed the clinical report of flows initially being appropriate, but then suction alarms were triggered. the suction alarms were prolonged and unable to be resolved. the logs further displayed a pump stop due to high motor current within the first seven minutes of support. the device had been in and out of proper lv position during this time. the pump stop and drop in motor speed are indicative of contact with the pump's impeller. the pump was analyzed and broken down within the testing environment. the visual inspection did demonstrate a kink within the pump catheter. when the pump was run there were suction alarms triggered, as there had been during patient use. upon microscopic examination there was an observation made of a broken impeller blade. the root cause of the broken blade was most likely the result of an interaction with the newly placed tavr. this failure mode will be monitored and trended. in addition, an incident investigation was initiated to address failures of the type. (b)(4).
38128 | 1035166-2017-00110 | 7105422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7105422&pc=DYB | 11/20/2017 | 05/24/2018 |  | 12/11/2017 |  |  | 11/21/2017 | 12/11/2017 |  |  | 07/01/2020 |  | 12/11/2017 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-13578 |  | 0052-3025 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 84 YR | [5 empty] | Injury | Yes | No | Difficult to Remove | [6 empty] | Cardiomyopathy | Mitral Regurgitation | Blood Loss | [10 empty] | Hospitalization | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | Additional information received indicates that the blood lost was approximately 500 cc. information received states the physician was forced to peel away the sheath, with part of the sheath still within the vasculature, due to the introducer being up against the impella pump plug. as per the reporter, the introducer used in this case was the 14f, 25 cm introducer (as opposed to the 14f 13 cm option). | The customer wanted to notify the manufacturer of a recent complaint that was relayed to them by a field representative on 11/20/2017: what is the outcome of the patient? stable. what procedure was the product used for? insertion/access for an impella cp at the left femoral artery. what was the condition of the patient when presented to the physician? the patient had known cardiomyopathy, severe takotsubo, cad, and mr was there an issue with the product unknown at this time, awaiting more details from field rep. when was the issue identified? during procedure. what was the user doing when the issue occurred? after insertion, the md was peeling away and lost 1 unit of blood, as there was an inability to remove peel away completely from femoral before cracking the sheath. was there any additional devices used during the procedure? no. in which sequence the devices were used if multiple devices were involved? awaiting more details from rep in field regarding serial dilations and other case particulars. are the devices available to return for evaluation? noted to be discarded. additional information received on 11/27/2018: the updates i have for you are as follows: the patient is listed as "doing well" as of (b)(6) 2017. the impella has been explanted. patient is a (b)(6) male. the patient had the impella cp placed for hemodynamic support of a cardiomyopathy patient. the bleeding did not extend the hospitalization, nor were blood products given. regarding the inability to remove the peel away, the rep has yet to respond with much more details beyond what i initially had for you. the representative was not present at the insertion to witness the issue. the rep has been able to tell us that the patient was not obese and that the serial dilations were done. the blood loss is estimated at 1 unit. yes, the introducer was discarded. |  |  |  |  | A review of device history record (dhr) was conducted and there were no manufacturing rejects or anomalies of this type recorded in the dhr. the device passed all in-process and qa final inspection steps before shipping to the customer including visual, dimensional and leak testing. a complaint review of the reported lot identified one additional report for an unrelated event. the procedure, adelante s2s introducer sheath in-process and final inspections, ensures sheaths are inspected for visual, dimensional, and peel functional compliance. the destructive testing sampling plan instructs to: manually break the sheath and hub and verify that the seals split easily without extreme elongation. also verify that the split cap and seals remain secure and do not break free or loosen from the sheath hub. it also instructs to "manually peel the sheath and verify the sheath peels easily along the sheath body and tip". please note that this product is supplied bulk non-sterile without an ifu (instructions for use) to the customer for further processing. since the device was discarded, it will not be returned for analysis and as a result, the allegations against this introducer cannot be confirmed. no subsequent device or clinical adverse events have been reported. based on the investigation a capa is not required. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. | Conclusion not yet available, evaluation in process. this initial mdr is being submitted to meet our requirements of reporting. a follow-up will be submitted as soon as the investigation is complete.
38129 | 1220648-2017-00112 | 7100745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7100745&pc=OZD | 11/12/2017 | 11/12/2017 |  | 12/08/2017 |  |  | 11/12/2017 | 12/08/2017 |  |  | 06/30/2019 |  | 12/08/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1303706 |  | 005042 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | Device Operates Differently Than Expected | Positioning Problem |  |  |  |  | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) male was admitted for a high risk angioplasty, with an impella 2.5 placed at the left femoral artery for hemodynamic support. the plan of care was to perform atherectomy and drug eluting stenting to the left main, left circumflex, and left anterior descending arteries. the patient's access site of the left femoral artery was observed to be bleeding, after approximately two hours of support due to patient movement on the procedure table. the movement had caused the impella's sheath to back out of the arteriotomy. the team was unable to advance it further into proper position. cardiac drugs and blood products were infused as his condition deteriorated. the fellow also began to hold pressure at the access site to remedy and reduce the blood loss and hematoma development. vascular surgery was consulted and the team did contralateral ballooning from the right femoral artery to close the bleed access site. the impella 2.5 was explanted and a 12fr sheath was placed in the left femoral artery. | [5 empty] | The data logs and clinical account were analyzed. no product was returned for analysis. there was a manufacturing review done which reveals no other complaints of this failure mode against this lot of pumps. the root cause of the bleeding could not be determined as no product was returned. no corrective action will be taken as the root cause was unable to be determined. (b)(4).
38130 | 1220648-2017-00113 | 7096372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7096372&pc=OZD | 11/03/2017 | 11/08/2017 |  | 12/07/2017 |  |  | 11/08/2017 | 12/07/2017 |  |  | 09/30/2018 |  | 12/07/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1261021 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombus | Vascular System (Circulation), Impaired | Vascular Dissection | [10 empty] | Death | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) female patient presented with a large anterior, lateral myocardial infarction and an occlusion of the ostial left anterior descending artery. the patient was being supported by an intra-aortic balloon pump during the intervention, and reportedly decompensated after the reperfusion. an impella cp was implanted on (b)(6) and supported the patient for the next three days, and the patient was reported to be doing well on the evening of (b)(6). the physician attempted a quick wean from impella support, decreasing by 1 performance level every 10 minutes, and the patient did not tolerate this. the following morning on (b)(6), the physician began to wean the patient from impella support again. the patient was brought to the cardiac catheterization lab for removal of the impella cp. during the procedure the patient complained that her legs were going to sleep. an arteriogram revealed the left circumflex and iliac arteries were open, however there was no blood flow. an angiogram then revealed a massive clot or dissection in the abdominal aorta, beginning in the renals and extending to the aortoiliac bifurcation. vascular surgery was consulted and the patient was brought to the operating room with the intention of performing either a thrombectomy or repair of an ascending aortic aneurism (aaa). the percutaneous thrombectomy was performed with angiojet in the operating room. renal and inferior mesenteric artery stents were placed, and the patient was sent back to the unit, however pulses were lost in the patient's left leg. the patient was brought back to the operating room for an open thrombectomy. the procedure was unsuccessful and the patient expired. | [5 empty] | The impella cp was not returned for investigation, despite multiple requests for product return. if additional information is obtained by the manufacturer, a follow up report will be submitted. (b)(4).
38131 | 1220648-2017-00108 | 7088735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7088735&pc=OZD | 11/07/2017 | 11/07/2017 |  | 12/06/2017 |  |  | 11/07/2017 | 12/06/2017 |  |  | 07/31/2019 |  | 12/06/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1301345 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) female patient presented with myocardial infarction, ischemic mitral valve disease as well as diseased distal vessels. the patient was being supported by an intra-aortic balloon pump (iabp) when she arrived at the hospital. the iabp was removed for escalation of care to an impella cp, and the patient was reported to have decompensated following the removal of the iabp. the impella cp was implanted for support on (b)(6) through the left femoral artery. the pump supported the patient with the intention to allow the left ventricle to rest so the patient could undergo mitral valve repair. the patient experienced oozing from the access site on the evening the pump was implanted, and the introducer was reported to have come out of position. the introducer was readvanced and gauze was applied, and the access site oozing was remedied. blood replacement products were ordered, however none were given at this time, and the pump continued to support the patient. the patient experienced signs of hematuria on (b)(6), and the position of the pump was managed in order to prevent hemolysis from developing. the pump was repositioned on (b)(6), after which the patient experienced bleeding from the access site. the pump was inserted further into the sheath hub to control the bleeding, and the patient received blood replacement products at this time. it was discovered the patient was experiencing a gi bleed on (b)(6), and she received additional blood replacement products. later in the day on (b)(6) the pump stopped, and was unable to be restarted. the patient was brought to the cardiovascular operating room for removal of the impella. she was decompensating at this time, and the decision was made not to start cpr. the patient expired due to comorbidities and a new onset of sepsis. | [5 empty] | The impella cp was returned for inspection. the repositioning sheath was undamaged and there were no defects found on the catheter. the clinical case was also reviewed. the access site oozing was observed during the times when the repositioning unit was found to have backed out of the arteriotomy. each time the repositioning sheath was re-advanced, the oozing would resolve. according to the clinical representative the patient's bmi was (b)(6), and the patient weighed (b)(6) pounds. the patient received 5 units of blood and two packs of platelets during the case, and the majority of the blood loss was attributed to a gastrointestinal bleed. the estimated access site ooze rate was less than 15 cc/min. the root cause of the oozing at the access site was most likely the repositioning unit backing out of the arteriotomy. it is possible the repositioning sheath backed out of the arteriotomy as a result of the patient's size, however this cannot be confirmed. a review of the device history did not reveal any nonconformities. (b)(4).
38132 | 1220648-2017-00109 | 7083377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7083377&pc=OZD | 11/07/2017 | 11/07/2017 |  | 12/05/2017 |  |  | 11/07/2017 | 12/05/2017 |  |  | 06/30/2019 |  | 12/05/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1305208 |  | 0048-0020 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 46 YR | [5 empty] | Injury | Yes | Yes | Use of Device Problem | Patient-Device Incompatibility | [5 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) female was admitted with a non-stemi myocardial infarction. she had a known history of a previous coronary bypass surgery (cabg) and aneurysm (aaa) repair. upon admission to the cardiac catheterization lab she was seen to have a thrombotic burden in the distal left anterior descending coronary artery. for hemodynamic support during this high risk procedure, of a balloon angioplasty and infusion of the thrombolytic tpa, the physician placed an impella cp via the left femoral artery. the patient had pain in her groin access site and experienced hypotension while in the cath lab post-procedure holding area. the patient was diagnosed with a retroperitoneal bleed. the impella was explanted and the patient transferred to surgery. the patient was taken to the operating department for a vascular repair and infusion of blood products. the patient did later expire due to complications. | [5 empty] | To date the data logs and imaging from the case support have not yet been returned for analysis. the product itself was discarded by the user at the time of patient explant. a supplemental medwatch will be completed if any additional information is received by the manufacturer. (b)(4).
38133 | 1220648-2017-00116 | 7083382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7083382&pc=OZD | 11/14/2017 | 11/28/2017 |  | 12/05/2017 |  |  | 11/28/2017 | 12/05/2017 |  |  | 08/31/2017 |  | 12/05/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1293083 |  | 0048-3096 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 80 YR | [5 empty] | Malfunction | No | Yes | Device-Device Incompatibility | [6 empty] | Thrombosis | No Known Impact Or Consequence To Patient | [11 empty] | Hospitalization | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An (b)(6) male was admitted with and st elevation myocardial infarction, and upon evaluation in the cardiac cath lab was seen to have no coronary artery disease, but rather abnormal movement in his previously placed tavr corevalve. the team conducted an echo and believed thrombus to be present on the valve. the patient was known to have a coagulation disorder. the team chose to place an impella cp pump for support and left ventricle unloading. they also began high dose anticoagulation to treat the clot burden. if this failed, they would escalate to another valve-in-valve tavr procedure. on the second day of support, the team noted continuous suction alarms with low flows on the impella console. conversation was made as to the plan of care, whether to replace the impella cp or add volume. the decision was made to exchange to a new impella. a second cp was placed, and the first was retained for analysis. | [5 empty] | The impella cp product, data logs, clinical account, and imaging were evaluated and analyzed. the impella cp pump was returned in good condition. when analyzed visually and microscopically the impeller blades were seen to be damaged. when the pump was run in the lab, the flows were out of specification, as they were noted to be in the clinical account. the data logs further supported the noted drop in flows and suction alarms. in addition the logs showed an out of position alarm. the data and device analysis indicate that the impella cp had interacted with the valve. the low flow was the result of the damage observed to the impeller blades. this failure mode will be monitored and trended. in addition, an incident investigation was initiated to address the failure mode. the failure mode was determined to be low risk index based upon very low occurrence and moderate severity. internal reference (b)(4).
38134 | 1220648-2017-00110 | 7063003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7063003&pc=OZD | 10/30/2017 | 11/03/2017 |  | 11/28/2017 |  |  | 11/03/2017 | 11/28/2017 |  |  | 07/31/2019 |  | 11/28/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1301783 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH | [10 empty] | Injury | Yes | No | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A male patient presented with st-elevated myocardial infarction and with a depressed ejection fraction. an impella cp was implanted through the left femoral artery on (b)(6) 2017. the impella cp came out of position of the patient's left ventricle soon after implantation, and the physician was unable to recross the aortic valve. the patient was also experiencing blood loss from the access site which prompted the clinical team to perform type and cross for the patient in preparation of administering blood replacement products, however no blood products were ever given. according to the clinical staff the introducer was peeled away before the introducer was completely removed from the vasculature, and the repositioning sheath was inserted in order to control the bleeding. the patient's bleeding did not subside, and the impella cp was removed. the patient was placed on a vassopressor for support and was reported to be stable following the removal of the impella cp. | [5 empty] | The impella cp was returned from the facility, however the introducer was not returned for investigation. a review of the clinical case revealed the introducer was prematurely split before it was removed from the femoral artery. the impella cp came out of position from the patient's left ventricle during this exchanging of the introducer for the repositioning sheath, and the catheter kinked when the physician attempted to recross the aortic valve. a review of the device history did not reveal any nonconformities. the root cause of the bleeding from the access site is most likely the splitting of the introducer prematurely. (b)(4).
38135 | 1220648-2017-00111 | 7060828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7060828&pc=OZD | 11/05/2017 | 11/05/2017 |  | 11/27/2017 |  |  | 11/05/2017 | 11/27/2017 |  |  | 03/31/2017 |  | 11/27/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1291379 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR | [5 empty] | Malfunction | No | Yes | Kinked | Torn Material | [5 empty] | No Code Available | [41 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | Complainant reported that during the week of (b)(6) 2017 a (b)(6) female patient had undergone an aortic valve replacement (avr). the patient was reported to be obese with a large panis. following the avr the patient decompensated, and as a result an emergent 3 vessel coronary artery bypass graft (cabg) was performed. after the cabg was performed the patient, in a semi-stable condition, was brought to the icu with an intra-aortic balloon pump (iabp) that had been supporting the patient. the iabp was weaned from the patient and removed. on (b)(6) 2017 the patient decompensated. the physician then decided to place an impella cp for support. the patient coded as she was being prepped. cpr was started and continued throughout the attempted placement of the impella cp. the 14fr introducer sheath was placed in the patient's right femoral artery that had also been accessed with the iabp. visualization was reported to be poor during the placement of the pigtail, and it would not advance to the left ventricle (lv). a fluoroscopy was performed which revealed that the wire was looped within the artery. the physician then attempted to gain access of the left femoral artery, but was unable to gain access due to active cpr and low pulsility. the venous was accessed only. without the return of spontaneous circulation (rosc), the procedure was aborted and the patient expired. upon removal of the 14fr introducer sheath from the right femoral artery it was found to be kinked and that the peel-away portion of the sheath had cracked. the physician reported that the patient outcome was not as a result of any issue with the impella cp. the physician categorized this case as a "salvage situation", and that access attempts were ceased. | [5 empty] | The impella cp was discarded following removal from the patient; consequently an evaluation of the pump at issue could not be performed. the 14french introducer is being held by the hospital's adminstration department. a console data log analysis was not included in this report because data logs are not relevant to this failure mode. the root cause of this event could not be determined. no corrective action has been recommended because the evaluation has not yet been performed. internal reference: (b)(4).
38136 | 7057417 | 7057417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7057417&pc=PBL | 10/03/2017 | 10/10/2017 | 10/10/2017 | 11/24/2017 | 10/10/2017 | 10/10/2017 |  | 11/24/2017 |  |  |  |  | 11/24/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003 |  | Yes | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 56 YR | [5 empty] | Malfunction | No | Yes | Loss of Power | Device Operates Differently Than Expected | [5 empty] | No Information | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL | 453054401-00480003 |  | Device stopped due to fail safe perimeters.
38137 | 7057575 | 7057575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7057575&pc=KFM | 10/02/2017 | 10/09/2017 | 10/09/2017 | 11/24/2017 | 10/09/2017 | 10/09/2017 |  | 11/24/2017 |  |  | 03/31/2019 |  | 11/24/2017 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA |  |  | KFM |  |  | 005062 |  |  | ABIOMED, INC. |  |  | US |  | 127985 |  |  |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | User Facility report | [8 empty] | 76 YR | [5 empty] | Injury | Yes |  | Material Fragmentation | [6 empty] | Device Embedded In Tissue or Plaque | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella 5.0 was removed from trans caval sheath. trans caval 24f cook sheath was left in place. patients hemodynamic remains unchanged immediately after removal of impella 5.0. note: inlet cage and pigtail of impella 5.0 was stuck on inner side of cook sheath (at silicon hub), and was profusely bleeding --after pulling impella, inlet cage and pigtail broke off and remain in the hub.
38138 | 1220648-2017-00106 | 7053007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7053007&pc=OZD | 11/02/2017 | 11/02/2017 |  | 11/22/2017 |  |  | 11/02/2017 | 11/22/2017 |  |  | 08/31/2019 |  | 11/22/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1305450 |  | 0048-0026 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | CONSUMER | HEALTH PROFESSIONAL |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | [6 empty] | Patient Problem/Medical Problem | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | An impella cp was implanted into a (b)(6) female patient for support for an acute myocardial infarction and cardiogenic shock. this patient has a history hypertension, hyperlipidemia, peripheral artery disease, diminished pulses in the left leg, and transient ischemic attack. the patient was coding when presenting with total occlusion of the right coronary and circumflex arteries, 70% occlusion of the mid-left anterior descending (mlad) and left proximal anterior descending (plad). an intra-aortic balloon pump (iabp) was inserted while the patient was at (b)(6), and the patient was transferred to (b)(6) for implantation of the impella cp. after the iabp was removed the impella cp was implanted through the right femoral artery on (b)(6) and was performed independent of abiomed clinical representative support. the abiomed clinical representative reported that the physician did not perform assessment of the groin access site prior to insertion of the impella cp. in addition, it is not known whether the patient's activated clotting time (act) was therapeutic at the time of implant. the first case notes report that echocardiogram revealed thrombus in the left ventricle. on (b)(6) the patient was sent to the cardiac catheterization lab with the plan to have a stent placed in the left anterior descending artery, and with the plan to explant the impella cp. the patient's (act) was not at a therapeutic level, therefore the pump was left in place and systemic heparin was deceased. two hours later the right coronary and left marginal arteries were stented, however the circumflex artery was not able to be repaired. a percutaneous transluminal coronary angioplasty (ptca) of the circumflex artery was performed and the clinical team planned to do a contralateral balloon tamponade for explant. the left femoral artery was unable to be used for this part of the procedure due to a dissection, therefore the right radial artery ( rra ) was used. upon removal of the impella cp a thrombus was noted on the inlet and outlet cages, with a large thrombus remaining in the artery. angiojet was used to remove the thrombus, however the clinical team was unable to remove all of the thrombus this way, and the patient was brought to surgery for a thrombectomy. the patient was reported to be in stable condition following the procedure. | [5 empty] | The impella cp and clinical case were evaluated. according to the clinical description, the clinical staff could not confirm the patient's acts were at 250 at the time the pump was implanted, which is the level that is recommended in the impella catheter ifu. upon removal of the device two days later, thrombus was noted on the inflow and outflow cages. in addition, thrombus remained in the common femoral artery requiring surgical removal and the patient was scheduled for a thrombectomy later that day to rectify the issue. when the pump was inspected by the manufacturer, biomaterial was found in the pump housing. there were no deformities on the pump noted. the root cause of the patient's thrombus is unable to be determined, as there were no features on the pump that would have led to the development of the thrombus that was found on the device. a review of the device history did not reveal any nonconformities. (b)(4). | The products were returned on november 15, 2017 however the investigation is still underway. a followup report will be submitted with the results of the investigation. (b)(4).
38139 | 1220648-2017-00097 | 7046271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7046271&pc=OZD | 10/21/2017 | 10/21/2017 |  | 11/20/2017 |  |  | 10/21/2017 | 11/20/2017 |  |  | 05/31/2019 |  | 11/20/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1290478 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Human-Device Interface Problem | [6 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On (b)(6), a (b)(6) year old male was admitted with an acute myocardial infarction, having several rounds of cpr and shocks prior to admission to the cardiac cath lab. while the field representative was on route for clinical support, the impella cp, placed via the right femoral artery, was exhibiting low pump flows. per recommendation, the team adjusted and repositioned the pump position within the left ventricle multiple times. to troubleshoot further, the impella console was exchanged. the angioplasty procedure continued with the hemodynamic support of the impella cp, and the physician made the decision to leave the cp in for support within the icu. the team prepped the patient for transfer to the icu, by removal of the 14fr introducer and advancement of the impella cp's repositioning sheath. the team at this point noticed the patient was developing a hematoma and began to hold pressure. the impella was explanted and the team attempted to place a 14 fr x 13cm cook introducer at the arteriotomy to lessen the blood loss. the cook introducer was then removed and the groin closure device, the 8fr angioseal system, was to be placed. the angioseal placement was not successful, and so the 14fr x 30cm introducer was left at the groin arteriotomy site with a femostop placed on top to stop the groin bleed. in addition, to treat the blood loss at least 3 units of blood replacement product was infused. | [5 empty] | The impella cp pump product and data logs were returned for analysis. the manufacturing review found no other complaints leveled against this lot of impella cp pumps. upon analysis of the pump no defects were found. the introducers were not returned for analysis. the cause of the hematoma was not able to be identified. given that the root cause was not determined, there will be no corrective action recommended. the failure mode will be monitored and trended. (b)(4).
38140 | 1220648-2017-00107 | 7046303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7046303&pc=PYX | 10/03/2017 | 11/07/2017 |  | 11/20/2017 |  |  | 11/07/2017 | 11/20/2017 |  |  | 05/31/2019 |  | 11/20/2017 |  | RIGHT VENTRICLULAR BYPASS ASSIST DEVICE |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1289373 |  | 0046-0019 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 44 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | EKOS THROMBOLYSIS SYSTEM | [22 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) female patient was admitted with right heart failure post pulmonary embolism. an impella rp was placed on (b)(6), in conjunction with the ekos thrombolysis system. on the following day, (b)(6), the abiomed field representative was notified that the rp was explanted due to bleeding. the representative was told there was bleeding around the sheath at the access site. the hospital did not retain the pump or sheath for analysis. there was no need for intervention in the form of blood product replacement or vascular repair. the patient was noted to be stable post rp explant. | [5 empty] | Neither the product nor data logs were returned for analysis. the manufacturing review found no other complaints against this lot of impella rp pumps. without product or data logs, no root cause of the bleed could be determined. the failure mode will be trended and monitored. (b)(4).
38141 | 1220648-2017-00105 | 7046334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7046334&pc=OZD | 10/27/2017 | 10/27/2017 |  | 11/20/2017 |  |  | 10/27/2017 | 11/20/2017 |  |  | 08/31/2019 |  | 11/20/2017 |  | TEMPORARY LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1304520 |  | 0048-0026 |  | Yes | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | No | Device Operates Differently Than Expected | Device-Device Incompatibility | [5 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) female was admitted with an acute myocardial infarction (st elevation mi) and multiple cardiac comorbidities. she was taken to the cardiac cath lab and had an intra-aortic balloon pump (iabp) placed while cardiothoracic surgery was consulted. when the surgical team turned her down for surgery, she was readmitted to the cath lab and had the iabp replaced to escalated care, in the form of an impella cp. the impella cp was placed via the same groin site as the previous iabp after the "pre-close" technique of insertion of two perclose devices at the access site in the right femoral artery. once the impella support was initiated the patient had her multivessel angioplasty (pci) performed, via the opposing leg access, at the left femoral artery. the patient was transferred to the icu where a hematoma began to develop at the pci site, the left femoral artery. the hematoma prompted the team to infuse blood products and hold pressure. a blood loss /ooze began overnight at the right femoral artery, which lead the team to place a femostop over the impella site, against the recommendations of the field representative. the patient continued to have blood loss from the two groin sites, both the impella access site and the angioplasty groin site. in total 11 units of blood were infused and eventually vascular surgery was consulted to treat the angioplasty groin access site. the impella was weaned and explanted after over 25 hours of support. the patient is noted to be stable. | [5 empty] | The impella cp was returned for analysis. the pump investigations found no defects. the cause of the bleeding is unknown because the pump was found to be within specifications. the manufacturing review found no other complaints leveled against this lot of pumps. the root cause of the bleeding is unknown, and therefore no corrective action is warranted. the failure mode will be monitored and trended. (b)(4).
38142 | 1220648-2017-00100 | 7040667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7040667&pc=OZD | 10/23/2017 | 10/23/2017 |  | 11/17/2017 |  |  | 10/23/2017 | 11/17/2017 |  |  | 07/31/2017 |  | 11/17/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1296291 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 27 YR | [5 empty] | Injury | Yes | Yes | Material Integrity Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On (b)(6) 2017 a (b)(6) female patient presented to the emergency room with extreme shortness of breath. the patient was in postpartum cardiomyopathy. the patient's infant was (b)(6). in addition, the patient had a history of lupus. the patient decompensated when laying down and exhibited a post-intubation pulmonary embolism. cooling protocol with an ecmo was initiated, but the patient's stats dropped to the 40's. an impella cp was added to the ecmo to vent the left ventricle. the cp placement was performed without incidence. on (b)(6) 2017 patient bleeding occurred emanating from the hub of the impella cp. it was then determined that there was a compromised hemostasis valve in the hub of the pump which resulted in the patient bleeding. a surgical team member successfully sealed the valve with surgiseal and patient support successfully continued. patient bleeding from the hub of the impella ceased. on (b)(6) 2017 bleeding was observed around the insertion site of the repositioning sheath, and was unrelated to the earlier bleeding issue. one unit of replacement blood was administered to the patient, hemostasis was regained, and the patient was in stable condition. on (b)(6) 2017 after 8.4 days of impella cp support the pump was removed from the patient. in addition, the ecmo was also removed. | [5 empty] | A definitive root cause of this event could not be determined as the impella cp was discarded by the hospital staff following removal from the patient; consequently, no device analysis could be performed. the console logs were also not returned for analysis, although console data log analysis would not be applicable to this failure mode. the manufacturer will continue to investigate all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. (b)(4).
38143 | 1220648-2017-00103 | 7036943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7036943&pc=OZD | 10/18/2017 | 10/18/2017 |  | 11/16/2017 |  |  | 10/18/2017 | 11/16/2017 |  |  | 05/31/2019 |  | 11/16/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1292315 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | On (b)(6) 2017 a (b)(6) old female patient was admitted to the hospital with severe left main and multi-vessel disease. on (b)(6), 2017 the physician was planning on taking the patient to surgery, but the patient decompensated. she was then intubated and taken to the cath lab for impella cp placement. the pump and swan were placed through the same access site in the patient's groin. the physician planned to let the heart rest while hemodynamically supporting the patient with the impella cp, and to re-evaluate the patient over the next couple of days. following impella placement quite a bit of oozing was seen from the access site. manual pressure was held for 30 minutes, and then a pressure dressing was applied. the patient was also more heavily sedated as the oozing from the access site increased when she was agitated. the patient was successfully supported for 11.32 hours, but then in the late evening of (b)(6) 2017 the pump stopped. trouble shooting was unable to identify the problem and the pump could not be restarted. another pump was obtained and during the insertion of a replacement impella lp2.5 the patient's blood pressure dropped and she became maxed out on drips, and coded. the patient was defibrillated 3 times and was given cpr. following the placement of the impella lp2.5 the patient was administered 2 units of replacement blood products. the patient was then stable and was removed to the icu. the patient was successfully supported with the impella lp2.5 for 2.27 hours. it was reported that the patient coded several time during the night and eventually expired. the physicians reported that they felt that the patient would not have survived the bypass surgery, had she recovered this episode. there was no indication from the physician that the impella cp stopping caused or contributed to the patient outcome. | [5 empty] | The device continues to be held with the user facility, although the automatic impella controller data logs were returned for analysis. the analysis of the data logs did confirm multiple pump stop alarms for reverse flow during troubleshooting, but they give no indication as to what caused the pump stop. review of effectiveness of applicable systems was not completed as the root cause could not be determined. in conclusion, the pump stop was confirmed. the root cause of the pump stop ws not determined as the impella cp was not returned for evaluation. no corrective action was recommended at this time because as previously stated the root cause was unable to be determined. failures of this type will continue to be monitored and trended. internal reference (b)(4). | A definitive root cause of this event could not be determined as the impella cp is currently being held at the user facility; consequently, no device analysis could be performed. the console logs were also not returned for analysis. the manufacturer will continue to investigate all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. internal reference:(b)(4).
38144 | 1220648-2017-00104 | 7034817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7034817&pc=OZD | 10/16/2017 | 10/16/2017 |  | 11/15/2017 |  |  | 10/16/2017 | 11/15/2017 |  |  | 04/30/2019 |  | 11/15/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1288374 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2017 a (b)(6) male patient was admitted to the hospital with an anterior st elevation myocardial infarction (stemi.) the patient was hemodynamically unstable, and an impella cp was placed via the right femoral artery (rfa), prior to having a successful percutaneous coronary intervention (pci) performed on the patient's right coronary artery (rca). the impella was successfully placed and the patient was transported to another unit. during transport the patient became agitated and was moving excessively. this movement caused a large amount of arterial bleeding from the rfa insertion site. manual pressure was held, but the bleeding continued. the impella cp was removed from the patient and an impella lp 2.5 pump was placed. the patient was then taken emergently to the operating room for successful repair of the rfa. the patient was administered 12 units of replacement blood products. there were no additional reports of any adverse effects to the patient. | [5 empty] | The impella cp was discarded following removal from the patient; consequently an evaluation of the pump at issue could not be performed. a console data log analysis was not included in this report because data logs are not relevant to this failure mode. a device history review was performed and revealed that there were no issues for any pump in this lot set. in addition, there have been no other complaints relevant to this failure mode for pumps in this log set. the root cause of this event could not be determined. no corrective action has been recommended because the failure mode has a low risk index. this failure mode will continue to be monitored. (b)(4).
38145 | 1220648-2017-00101 | 7028673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7028673&pc=OZD | 10/15/2017 | 11/21/2017 |  | 11/14/2017 |  |  | 10/15/2017 | 11/14/2017 |  |  | 10/31/2018 |  | 11/14/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1266569 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 62 YR | [5 empty] | Malfunction | No | Yes | Kinked | Device Issue | Positioning Problem |  |  |  |  | No Consequences Or Impact To Patient | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | On (b)(6) 2017 a (b)(6) male was admitted with acute myocardial infarction and cardiogenic shock. he was in need of hemodynamic support for a staged cardiac procedure. the plan was to first complete a coronary artery intervention/pci, followed by a valvuloplasty for aortic valve stenosis. the pci was completed and then the patient decompensated during the valvuloplasty, and so the impella cp was placed. upon insertion of the impella pump, via the right femoral artery, the cp pump was seen to kink and twist back upon itself. the team needed to intervene in the form of a snare to remove and explant the pump. a new impella was then placed. the second impella was a 2.5 rather than another cp. the second pump supported the patient successfully for three days. | [5 empty] | Since filing the initial medwatch for on (b)(6) 2017, the investigation has been completed. the pump, data logs, and imaging were evaluated and analyzed. the manufacturing review found no other complaints for this failure mode for the lot of impella cp pumps. the pump was analyzed and found to have significant and numerous kinks on the catheter shaft. in addition, there was kinking observed on the .018 guidewire. imaging further confirmed the failure mode and kinking on the catheter. the root cause of the kink, which necessitated the use of a snare to remove the pump, was unable to be determined. of note, in the clinical case review, there is mention of the patient having aortic valve stenosis (as). within the ifu of the impella cp there is a cautionary statement of: "patients with aortic stenosis or other abnormal aortic valve performance may be compromised by the use of the impella catheter. " and aortic stenosis is further classified as contraindicated in: "aortic valve stenosis/calcification (equivalent to an orifice of 0.6 cm2 or less)." the fact that the patient had known as could increase the difficulty of crossing the valve with the pump. no corrective action is warranted at this time, however the failure mode will be trended and monitored. (b)(4). | The impella cp pump product and data logs were returned for analysis, however the investigation is on-going at this time. a supplemental medwatch will be completed upon investigation completion. (b)(4).
38146 | 1035166-2017-00098 | 7026390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7026390&pc=DYB | 11/01/2017 | 06/07/2018 |  | 11/13/2017 |  |  | 11/10/2017 | 11/13/2017 |  |  | 08/01/2020 |  | 11/13/2017 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3025 |  |  | OSCOR INC. |  |  | US |  | C1-13589 |  | 0052-3025 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 46 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Insufficient Information | [5 empty] | Aneurysm | Death | Exsanguination | Hemorrhage/Blood Loss/Bleeding | Myocardial Infarction | Blood Loss | [7 empty] | Death | [5 empty] | IMPELLA CP | [22 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | Additional information was received that included: the root cause of the bleed was never determined. the physicians indicated that the "patient had a retroperitoneal bleed at impella site and was not overtly bleeding out from site. they also reported regarding the event and patient death: "might be from the serial dilations and believes it may have been from the vessel and lack of recoil around arteriotomy." the reporter also stated that there was no specific device failure noted. reporter was able to confirm that the device in question is not available for return and analysis. although requested, additional information specific to exact cause and details surrounding the death of the patient is still unknown at time of this report. | The customer reported, today i received details of an adverse event regarding the combo kit oscor introducer packaged with the impella cp: 1. what is the date of event? (b)(6) 2017. 2. what is the name of the facility where event occurred? (b)(6). 3. what is the outcome of the patient? (e.g. patient stable, patient alive, no injury, serious injury, death) death. 4. what procedure was the product used for? a high risk pci, access for insertion of the impella cp. 5. what was the condition of the patient when presented to the physician? the patient was a 46yo female presented with nstemi, history of cabg and aaa repair. 6. was there an issue with the product ? if yes , then provide details of the issue and answer the questions a) and b). yes. a. when was the issue identified? (before , during or after the procedure) post procedure. b. what was the user (physician/ doctor/ nurse) doing when the issue occurred? the impella cp was explanted and patient was in hospital when the retroperitoneal bleed was discovered. 7. please provide the contact information of the physician who used the product to obtain additional clinical information of the event in question. n/a. 8. what is the model number and lot no. of the device used? oscor introducer c1-13589. 9. was there any additional devices used during the procedure? n/a. 10. what is the name, model no. and lot no. for any additional device used during the procedure? n/a. 11. in which sequence the devices were used if multiple devices were involved? serial dilations were done for 14x13 oscor insertion. peeled away, and then the impella cp was placed. the bleed was discovered post procedure. vascular repair was performed and 2-3 units of blood were given. the patient did expire. 12. are the devices available to return to oscor for evaluation? the field rep has been unable to locate any of the dissposable product to date. |  |  |  |  | A review of device history record (dhr) was conducted and there were no manufacturing rejects or anomalies recorded in the dhr affecting released product. the released device passed all in-process and qa final inspection steps including visual, dimensional and mechanical testing before shipping to the customer. a complaint review of the reported lot found no additional reports involving this lot number. the device was in use for treatment. the device is not available for return to oscor for further analysis and there was no allegation brought against the device reported. no investigation was performed in this case as the event did not involve a reported failure. no further follow-up is required. the clinical team did not attribute the death to the device or any device failure. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type." | Conclusion not yet available, evaluation in process. this initial mdr is being submitted to meet our requirements of reporting. a follow-up will be submitted as soon as the investigation is complete.
38147 | MW5073201 | 7013607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7013607&pc=PBL | 10/30/2017 | 11/06/2017 |  | 11/09/2017 |  |  |  | 11/09/2017 | 11/09/2017 |  |  |  | 11/09/2017 |  | TEMPORARY CARDIAC SUPPORT | TEMPORARY CARDIAC SUPPORT |  | IMPELLA CP | IMPELLA CP |  | PBL | PBL |  | 0048-0003 |  |  | ABIOMED INC. | ABIOMED INC. |  | US | US | 138963 |  |  |  | Yes | No answer provided |  |  |  |  | No | Health Professional | No Information |  | Voluntary report | [8 empty] | 54 YR |  |  | 91 |  |  | Malfunction | No | Yes | Communication or Transmission Problem | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Right fem anterial access was obtained and an impella cp was advanced into the lv. despite appropriate connections, the impella device was not recognized by the console. the console was restarted and connecting cables exchanged, but the console was still not recognizing the impella. that impella was removed and a second impella cp was rapidly exchanged and successfully placed into the lv. vf was noted on the monitor and required defibrillation x3 followed by initiation of an amiodarone infusion. minimal flow was generated by the impella. decision was made to escalate hemodynamic support to va ecmo. the impella cp was removed from the r-fem.
38148 | 1220648-2017-00094 | 7006583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7006583&pc=OJE | 10/13/2017 | 10/16/2017 |  | 11/07/2017 |  |  | 10/16/2017 | 11/07/2017 |  |  | 01/31/2019 |  | 11/07/2017 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA RP |  |  | OJE |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1276209 |  | 004334 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Difficult to Advance | Positioning Problem | [5 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella rp was implanted into a (b)(6) female patient for support for an off-pump coronary artery bypass. the patient presented with an ejection fraction range of 35-40%, with a friable (thin and weak) right atrium. after the bypass, the clinical team was unable to float the swan to the left pulmonary artery using a pigtail catheter, thefore the catheter was exchanged for a multipurpose catheter. the multipurpose catheter was removed and the impella rp was inserted. insertion of the impella rp was difficult and the physician noted the guide wire was difficult to advance, and he was unsure whether the guide wire was in the right place. the impella rp was unable to advance into the left pulmonary artery despite several attempts. during these attempts the patient went into atrial fibrillation, and the patient's right atrium and right ventrical were perforated. the right heart was repaired with sutures and began to pump again. the patient was reported to be in stable condition and there was no indication of further adverse effect to the patient following the surgical repair. | [5 empty] | The products involved in this event were discarded by the user. a review of the clinical case revealed the implanting physician reported the patient's heart was in very poor condition and was "falling apart...like paper tearing." after the patient's heart was repaired the right heart began pumping again and the patient was stable. an abiomed physician also reviewed the case and believes the perforation was due to the continued manipulation of the guidewires and guiding catheters. the root cause of the right heart perforations during manipulation of the various guidewires and catheters was most likely due to the patient's friable cardiac tissue. 17-0985.
38149 | 1220648-2017-00099 | 7007351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7007351&pc=OZD | 10/13/2017 | 10/13/2017 |  | 11/07/2017 |  |  | 10/13/2017 | 11/07/2017 |  |  |  |  | 11/07/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD | [5 empty] | ABIOMED |  |  | US | [5 empty] | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | LITERATURE | [5 empty] | 33 YR | [5 empty] | Injury | Yes | No | Malposition of Device | Device Operates Differently Than Expected | [5 empty] | Perforation | [12 empty] | Hospitalization | Life Threatening | Required Intervention | [26 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On (b)(6), while conducting a literature review, a publication in cci was discovered by abiomed in which a patient expired after suffering from a left ventricular (lv) perforation following use of an impella 2.5. the patient was a (b)(6) female who expired post surgical intervention of the perforation. she had neurological comorbidities and apparent viral myocarditis. the myocarditis was noted in the publication as "it is likely that the friability of the tissue secondary to presumed viral myocarditis predisposed her to left ventricular perforation". from the publication, alone, the following details were gathered. there was no documentation of the adverse event from the field, nor complaint made to an abiomed field member. an impella 2.5 was placed to support a (b)(6) female patient in biventricular failure secondary to suspected viral myocarditis. after 24 hours on support, the patient was transferred to a larger hospital for treatment of the presumed myocarditis via air transport. the patient's knee was immobilized prior to transport, and she was hypotensive and had cold extremities upon arrival. the patient's condition worsened despite medical management; it was unclear whether this was due to right ventricular failure or issues with impella support. the placement signal was noted to be displaying a ventricular waveform. a tee showed pericardial effusion, and the pigtail of the pump through the wall of the lv. the pump was "never noticeably dislodged" at the groin. pericardiocentesis was performed bedside to remove blood from the patient's chest, but there was "no meaningful improvement." the patient was then brought emergently to the or for surgery and possible ecmo placement. the inflow cage was not visible in the heart under tee, and the "entire device" appeared to be deeper than 4-4.5 cm below the aortic valve. the patient was then put on cardiopulmonary bypass. the pump had visibly perforated the lv wall. it was then pulled back into the aorta and the perforation was sutured shut. the physician noted that the patient's lv wall was friable. biventricular centrimag devices were the placed, and the patient was returned to the icu. after weeks of treatment, care was later withdrawn due to neurological comorbidities. once the publication was found all due diligence was made to ascertain more details with extensive internal searches through databases and historical case documentation. no lot number, expiration dates, or serial number have been discovered for the medwatch report. | [5 empty] | Neither the device logs nor pump product were returned from the field. no investigation is possible to confirm root cause of the lv perforation. the lv perforation was most likely due to the patient's friable wall tissue secondary to viral myocarditis. no corrective action is recommended as the failure mode was determined to have low risk index due to major severity and very low occurrence. the failure mode will be monitored for trends. (b)(4).
38150 | 1220648-2017-00092 | 7003893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7003893&pc=OZD | 10/07/2017 | 10/07/2017 |  | 11/06/2017 |  |  | 10/07/2017 | 11/06/2017 |  |  | 07/31/2019 |  | 11/06/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1295955 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR |  |  | 112 |  |  | Injury | Yes | No | Restricted Flow rate | Difficult To Position | [5 empty] | Hemolysis | Patient Problem/Medical Problem | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella cp was inserted into a (b)(6) male patient for support for acute myocardial infarction and cardiogenic shock following a vt ablation procedure. the patient decompensated after the procedure. the patient had a prior major myocardial infarction when he was (b)(6), ischemic cardiomyopathy, hypertension, paroxysmal atrial fibrillation, chronic kidney disease asthma, ventricular tachycardia and an ejection fraction of 15%. this patient was transferred from (b)(6) hospital to (b)(6) for the vt ablation procedure, and needed to be shocked several times while at (b)(6). the impella was implanted through the left femoral artery without issue on the evening of (b)(6) 2017,however the patient experienced hemolysis within one hour of placement of the pump. the following morning the patient was still experiencing hemolysis and was not producing adequate urine. the patient developed a sizeable hematoma in the groin. echo revealed the pump required repositioning which was done several times by the physician, however the pump continued to have positioning issues and was providing restricted flow. the decision was made to explant the pump. the patient was successfully weaned from impella support and the pump was explanted on (b)(6). an update from the clinical staff revealed the patient remains intubated and on dialysis as of (b)(6). | [5 empty] | The impella cp was discarded by the user facility therefore the product could not be inspected, and the root cause is unable to be determined . if abiomed receives additional information from the field a supplemental report will be submitted. (b)(4).
38151 | 1220648-2017-00091 | 7003923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=7003923&pc=OZD | 10/10/2017 | 10/10/2017 |  | 11/06/2017 |  |  | 10/10/2017 | 11/06/2017 |  |  | 02/28/2019 |  | 11/06/2017 |  | PUMP, BLOOD, CARDIOPLULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1280511 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 50 YR | [5 empty] | Injury | Yes | No | Difficult to Remove | Retraction Problem | Human-Device Interface Problem |  |  |  |  | Patient Problem/Medical Problem | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | HEMASHIELD GRAFT | [22 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) male was transferred to (b)(6) medical center from another hospital for hemodynamic monitoring and escalated care. on the (b)(6) an impella 5.0 was placed via the right axillary artery. of note, the graft was said to be sutured at 90' and tunneled under the skin. the pump was successfully supporting the patient for 16 days, and then the weaning process began with plan for surgical removal. on the (b)(6) the pump was being removed and an iabp was to be placed. the iabp was not enough hemodynamic support and so the 5.0 remained for support. on the (b)(6) the physician decided to add ecmo support for ongoing respiratory needs. the patient was again taken to the or for surgical explant of the 5.0 on (b)(6). while attempting to explant the pump, the team met resistance and could only torque and remove the pump partially. the pump was cut and the parts of the pump were removed through the anastomosis. due to the explant complications, there was over a liter of blood lost and 3 liters were infused back to the patient, in addition to platelets and ffp. the axillary artery was repaired in the operating room, it had been damaged but, not dissected. the family is determining the care plan due to questionable neurologic status, and the patient's final outcome is not known at the writing of this medwatch report. | [5 empty] | The product and data logs were returned for analysis. in addition, the hemashield graft was returned at the conclusion of the 18 day hemodynamic support. the product lot was investigated and no other complaints have been made against this lot of 5.0 pumps. when the hemashield graft was analyzed there was a large thrombus burden observed. the thrombus likely contributed to the adverse event and pump damage. the pump's steering catheter was found to be damaged upon analysis. the damage did not affect support as the pump was being explanted. the impella got stuck in the graft anastomosis but, the team was able to partially torque it, though to fully remove it the pump had to be cut. the large clot found in the graft likely contributed to the damage. due to the low risk index, from a minor severity and very low occurrence, there will be no capa but, the failure mode will be monitored and trended. internal reference (b)(4).
38152 | 1220648-2017-00088 | 6997195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6997195&pc=PYX | 10/06/2017 | 10/06/2010 |  | 11/02/2017 |  |  | 10/06/2017 | 11/02/2017 |  |  | 10/31/2017 |  | 11/02/2017 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1216991 |  | 004334 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | Yes | Crack | Difficult to Insert | [5 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2017 a (b)(6) male patient had undergone a cervical disk repair. during the following overnight the patient suffered a myocardial infarction (mi) and heart rhythm changes. it was reported that the patient had previously undergone coronary artery bypass graft procedures. at approximately noon on (b)(6) 2017 the patient was transported to the cath lab. the patient was found to be in biventricular failure. the physician decided that the patient required right and left heart hemodynamic support. an impella cp was placed for left side heart support. the physician then attempted to place an impella rp for right side heart support. the rp impella insertion was reported as very difficult. during the insertion the 23fr peel away cracked near hemostatic valve while inserting the dilator. further examination of the sheath found that it was cracked and partially split on proximal portion near flush side arm. the physician was unsure if the crack happened prior to or when the dilator was placed within it. a large kelly clamp was placed around the sheath to keep together. the impella rp was then successfully placed. hemodynamic support was provided and resulted in the patient's right ventricle being no longer extended. at least 1 unit of blood product was given in the ccl. it was reported that the patient was a bleeder due to anticoagulation from the spinal surgery that had been performed the previous day. the right coronary artery (rca) graft occlusion was successfully repaired. the patient was then successfully supported for 5.09 hours, with no additional impella rp issues reported. the patient was reported to be in stable condition following the impella rp support. | [5 empty] | The impella rp and repositioning sheath were discarded by the user following patient use, so an evaluation of the device and event were unable to be performed. the console data log analysis is not applicable to this failure mode. there were no complaints for any other pump in this pump set lot. there were five other sprs associated with this 23 fr rp introducer lot. three of the complaints were directly related and displayed the same failure mode. this failure mode has a medium risk assessment based on low occurrence (0.5%, 2 occurrences out of 393 uses from 2016-10-16 to 2017-10-16) and major severity. the root cause of the broken sheath was unable to be determined because the product was not returned. no corrective action is recommended because the root cause was unable to be determined. (b)(4).
38153 | 1220648-2017-00096 | 6997463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6997463&pc=OZD | 10/04/2017 | 10/04/2017 |  | 11/02/2017 |  |  | 10/04/2017 | 11/02/2017 |  |  | 06/30/2019 |  | 11/02/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1299130 |  | 0048-0020 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 59 YR |  |  | 82 |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | Complainant reported that on (b)(6) 2017 a (b)(6) female patient arrived at the hospital by emergency medical services. the patient presented with a slow ventricular tachycardia (vt) heart rhythm. the patient quickly decompensated. the physicians tried to cardiovert the patient twice without success. the patient required cpr and intubation. an impella cp was placed and the patient's 100% occluded left anterior descending artery (lad) was opened. the patient quickly stabilized. the patient was transferred to another facility. during transfer the impella cp was damaged, causing the pump to stop. on (b)(6), 2017 the first impella cp was removed from the patient and a replacement impella cp was inserted in the patient without issue. later that day the nurse reported that the patient's insertion side leg was cold, pulseless and mottled. the patient was administered higher doses of pressors and the device purge pressure was lowered. the clinical team then decided to continue supporting the patient with this pump due to patient being hemodynamically unstable. the pump supported the patient for 4 more days. on (b)(6) 2017, after successfully supporting the patient for a total of 5.79 days, the device was removed from the patient as a result of the patient's leg becoming ischemic. during the device explant a surgical closure was performed at the insertion site. the complainant reported that there was no lasting harm to the patient and the patient was in stable condition and doing well. | [5 empty] | The impella cp and cook sheath were both discarded by the user following patient use, so an evaluation of the device and event were unable to be performed. the console data log analysis is not applicable to this failure mode. a review of the device history record found no other issues related to this failure mode in this lot, in addition there were no other complaints reported related to this failure mode. this failure mode is determined to be low risk index based on moderate severity and very low occurrence (occurrences: 7, usages: 13741, rate: 0.05%) the root cause of leg ischemia was unable to be determined because the product was not returned. no corrective action is recommended at this time because the root cause was unable to be determined. failures of this type will continue to be monitored and trended. (b)(4).
38154 | 1220648-2017-00084 | 6997464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6997464&pc=OZD | 09/30/2017 | 10/03/2017 |  | 11/02/2017 |  |  | 10/03/2017 | 11/02/2017 |  |  | 07/31/2019 |  | 11/02/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1301783 |  | 0048-0026 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 82 MO | [5 empty] | Malfunction | Yes | Yes | Difficult To Position | Difficult to Remove | Malposition of Device |  |  |  |  | Bradycardia | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that an 82 year old female patient presented to an outside facility with an st elevation myocardial infarction. the patient was transferred to the hospital and was taken to the cath lab for a percutaneous coronary intervention to be performed. while in the cath lab the patient became bradycardic requiring a temporary pacer to be placed. because the patient was also hypotensive an intra-aortic balloon pump (iabp) was also placed. during the performance of the pci the patient went into a ventricular fibrillation (vfib) arrest which required defibrillation, cpr and increased pressor support. following a return of spontaneous circulation (rosc) the physician removed the iabp and placed an impella cp. the pump was placed without difficulty. the pci was successfully completed with no impella cp alarms or issues occurring. the patient was then transferred to the icu, with an arterial line placed for monitoring, along with a swan. an echo was performed to confirm correct pump placement, but it revealed that the impella cp was 1.5 cm from the inlet to the arteriovenous (av). due to the patient's size the physician advanced the pump under echo. movement was noted, but loss of motor current mc and reduced flows was observed. the physician then attempted to move the impella back but was met by abnormal resistance, and patient related constrictive pericarditis (cp) and jaw pain. the impella was advanced but at this point there was a loss of all motor current, placement signal and impella flow. at this point what looked like the pigtail on top of the impella catheter was seen on the echo. resulting in chest x-rays being taken. chest x-ray showed that the impella catheter was kinked. the junction of the inlet cage and catheter, with the pigtail was laying against the catheter. multiple attempts to reposition the catheter failed. the patient was transferred to the cath lab, and it was decided to attempt to remove the impella through the opposite femoral site. the physician believed that the impella was caught in a papillary muscle or mitral apparatus. he gained access with a 16 fr. sheath and snared the impella pigtail and cut the impella removing the pigtail first. the remaining portion of the impella was snared and pulled into the bifurcation of the iliacs. during this procedure the patient had multiple episodes of profound hypotension and upon further investigation it was found that the impella catheter had become entangles with the mitral valve and could not be freed. the patient was then taken to surgery. during surgery it was found that the catheter had become caught between two leaflets. there was no damage to the mitral valve identified. the impella was successfully removed from the patient. the patient survived the procedure. | [5 empty] | The impella cp was discarded by the user following patient use, so an evaluation of the device and event was unable to be performed. abiomed will continue to make good faith attempts to retrieve data and information, and will submit a follow up report if they become available. (b)(4).
38155 | 6993143 | 6993143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6993143&pc=OZD | 09/29/2017 | 10/30/2017 | 10/30/2017 | 11/01/2017 | 10/30/2017 | 10/30/2017 |  | 11/01/2017 |  |  |  |  | 11/01/2017 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA AUTOMATED CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0000-US |  | Device was returned to manufacturer | [5 empty] | No | NURSE |  |  | User Facility report | [8 empty] | 65 YR | [5 empty] | Malfunction |  | Yes | Loss of Power | Device Operational Issue | [5 empty] | Myocardial Infarction | [18 empty] | NO OTHER THERAPIES | VENTILATOR, STENTS WERE PLACED IN CATH LAB, IV INF | [21 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Arrived as an st-elevation myocardial infarction (stemi), papillary muscle rupture, quickly went into respiratory distress and intubated. frequent episodes of afib. taken to catheter lab. stented three arteries and placed the impella device. inferior and post. myocardial infarction following des. transferred to icu. approximately three hours later, the impella just stopped. troubleshooting occurred; even changed the console. interventional cardiologist arrived to place intra-aortic balloon pump and removal the impella device. later that day patient transferred for mitral valve replacement on [date redacted]. discharged home ten days later with home health.
38156 | 1220648-2017-00089 | 6990219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6990219&pc=OZD | 10/02/2017 | 10/02/2017 |  | 10/31/2017 |  |  | 10/02/2017 | 10/31/2017 |  |  | 03/31/2019 |  | 10/31/2017 |  | PUMP,BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1287177 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | Yes | Break | Difficult to Remove | Physical Resistance | Device-Device Incompatibility |  |  |  | Patient Problem/Medical Problem | Device Embedded In Tissue or Plaque | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | A (b)(6) male was admitted to an outlying facility and received an iabp for hemodynamic support. the patient was transferred to the tertiary facility with an ef of 25% and the decision was made to place an impella 5.0 via the transcaval approach on (b)(6) . the transcaval access was performed via the cook 24fr x 65 cm sheath. the iabp was then explanted. the patient began to experience signs and symptoms of hemolysis in the form of hemolyzed labs on the following day of support, the (b)(6) . the pump position was assessed by echo and not repositioned. the patient was monitored and the renal function was observed during his ongoing dialysis. the team began to question if the hemolysis was due to rhabdomyolysis on (b)(6) , rather than solely the use of the impella 5.0. as the patient's deterioration continued, the family chose to withdraw care. upon removal of the transcaval sheath and explant of the 5.0 pump, the pump met resistance coming out of the sheath. they continued to pull on the device and the pump came out of the sheath but, the cannula had become detached from the inlet cage. the inlet cage and pigtail remained inside the sheath while the rest of the pump was explanted. as the care was withdrawn, the team did not snare or remove the remaining portion or exchange the sheath. the patient was allowed to expire with product retained. | [5 empty] | On december 12, 2017 a user initiated medwatch was received, (b)(4). no new clinical details were provided since manufacturer filed medwatch 1220648-2017-00089. internal reference (b)(4). | The root cause of the cannula detachment could not be determined as not all product was returned. the portion of the pump explanted through the cook sheath was returned, but not the sheath itself, nor the pump portion that remained within the sheath. the pump ring and inlet cage of the pump were not able to be analyzed. the cannula was analyzed and found to be stretched when compared to a sample 5.0 pump. the outflow cage was twisted and the bond between the pump housing and motor housing was damaged. no data logs were returned for analysis. a manufacturing review found no other complaints against this lot of impella 5.0 pumps. the failure will continue to be monitored and trended. (b)(4).
38157 | 1220648-2017-00095 | 6990235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6990235&pc=OZD | 08/29/2017 | 10/18/2017 |  | 10/31/2017 |  |  | 10/18/2017 | 10/31/2017 |  |  | 04/30/2019 |  | 10/31/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1288788 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 42 YR | [5 empty] | Malfunction | Yes | No | Kinked | Difficult To Position | [5 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) year old male presented with cardiac arrest and was taken to the cardiac catheterization lab for care. the patient had a 14fr x13cm sheath placed for insertion of an impella cp pump. this was done independent of any abiomed field representative support. the abiomed representative was later told that the insertion of the impella cp was not successful, as the cp pump had kinked back upon itself after insertion. the kink was noted to have occurred in the left ventricle near the pump motor housing. the team had attempted to straighten the pump, but the flows were low, and the pump placement was aborted and the pump explanted. as the 14fr introducer had been peeled away, the team had to perform a surgical removal of the impella cp, and a surgical groin wound closure. | [5 empty] | On (b)(4), a user medwatch 3500a was received. this medwatch form did contain new information regarding the patient event, inclusive of the fact that the pump placement took place at the left femoral artery. the hospital did not return for analysis the impella cp, 14 fr introducer, or data logs for analysis. because the product and data could not be inspected, the root cause of the cannula kink was unable to be determined. a manufacturing review was done, which revealed no other complaints leveled against this lot of pumps. the failure mode was determined to be a low risk index based on very low occurrence and moderate severity. the failure mode will be trended and tracked. internal reference (b)(4).
38158 | 1220648-2017-00102 | 6990532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6990532&pc=OZD | 10/04/2017 | 10/06/2017 |  | 10/31/2017 |  |  | 10/06/2017 | 10/31/2017 |  |  | 08/31/2019 |  | 10/31/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1298639 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 47 YR | [5 empty] | Malfunction | Yes | No | Entrapment of Device | Positioning Problem | [5 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) year old male patient presented with an ejection fraction of 15 and was anuric. an impella cp was implanted for support for hrpci on (b)(6) 2017. on (b)(6), fluoroscopy imaging revealed the pump came out of position out of the left ventricle, requiring repositioning back across the aortic valve. the imaging also revealed the impella pump was completely in the aorta and the pigtail had become hooked on the catheter. the patient was brought to the cardiac catheterization lab in order for the device to be snared from the prolapsed position. once the pigtail was released, the pump straightened out and was advanced across the aortic valve and into position in the left ventricle. the patient was returned to the unit and support continued. the patient continued to decompensate over the next several days, and was placed on dialysis on (b)(6). the patient developed a fever on (b)(6) and the family made the decision to withdraw care. the impella did not cause or contribute to the patient outcome. | [5 empty] | The impella cp was discarded by the user facility. the aic console data logs involved in this event have not been returned for evaluation by the user. abiomed will continue to make good faith attempts to retrieve data and information, and will submit a follow up report if they become available. (b)(4).
38159 | MW5073006 | 6986486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6986486&pc=DSQ | 10/18/2017 | 10/26/2017 |  | 10/30/2017 |  |  |  | 10/30/2017 |  |  | 10/18/2017 |  | 10/30/2017 |  | LEFT VENTRICLE ASSIST DEVICE |  |  | IMPELLA 5.0 |  |  | DSQ |  |  | IMPELLA 5.0 |  |  | ABIOMED |  |  | US |  | 134310 |  | 134310 |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 58 YR | [5 empty] | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Information | [41 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | A biomed impella 5.0 console demonstrated a "high purge pressure" alarm. alarm resolved shortly following the introduction of heparinized solution.
38160 | 1220648-2017-00083 | 6987941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6987941&pc=OZD | 10/02/2017 | 10/02/2017 |  | 10/30/2017 |  |  | 10/02/2017 | 10/30/2017 |  |  | 12/31/2018 |  | 10/30/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1276506 |  | 5042 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR |  |  | 88 |  |  | Malfunction | Yes | No | Difficult to Remove | Physical Resistance | [5 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) male patient presented to the hospital with chest pain. this patient was scheduled to have a protected percutaneous coronary intervention 5 days later. an impella 2.5 was inserted through the right femoral artery for support for hrcpi. upon removal of the 0.018 wire during implantation of the pump, resistance was felt and the wire was unable to be removed. the impella 2.5 was pulled out of the left ventricle inadvertently, as the wire was lodged in the outlet cage, and the catheter prolapsed on itself. the device became lodged in the common femoral artery requiring surgical removal of the device. the arteriotomy was surgically closed and the patient was reported to be stable at this time. | [5 empty] | The impella 2.5, 0.018" wire and clinical case were evaluated. both the impella 2.5 and guidewire were returned with kinks. a sample impella 2.5 was used to test the ease of insertion and removal of the returned guidewire. the returned guidewire was difficult to remove from the returned pump and the sample pump. engineering is unable to discern if the guidewire was the root cause of the issue as it was returned from the field with damage. the root cause of the guidewire interaction with the impella 2.5 was unable to be determined. a review of the device history did not reveal any nonconformities.
38161 | 1220648-2017-00082 | 6987950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6987950&pc=OZD | 01/18/2016 | 10/02/2017 |  | 10/30/2017 |  |  | 10/02/2017 | 10/30/2017 |  |  | 09/30/2017 |  | 10/30/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1217335 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 68 YR | [5 empty] | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Consequences Or Impact To Patient | No Known Impact Or Consequence To Patient | [40 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2017 a (b)(6) male patient was struggling following a transcatheter aortic valve replacement (tavr) that has been performed in (b)(6) 2015 (exact date unknown), and was admitted to the hospital with a low ejection fraction (ef.) the patient began to decompensate, so an impella cp (serial (b)(4)) was inserted for hemodynamic support. a pump stop occurred early in the case, but the pump restarted and did not stop again. doctors were happy with support and planned to leave the pump in until the patient recovered. patient condition improved and pump explant was scheduled for monday (b)(6). on the fourth day of support pump flow decreased and suction alarms occurred. the first impella cp was explanted, and a new impella cp (serial (b)(4)) was inserted. insertion for the second impella was normal. a new sheath was used with a new guide-wire. seven hours after insertion a pump stop occurred secondary to a coughing episode. pump was restarted in the patient. the pump produced low flow that continued to decrease throughout the case. eventually, the pump was explanted and the patient was scheduled to be switched to a tandem heart. patient expired in the ccl before surgery could be started. the physician reported that he did not feel that the patient outcome was caused or contributed to by any issue with the impella cp. the patient support was considered a "hail mary" attempt to save this patient. this mdr is reporting the issue with the second impella cp used during this event. the first pump (serial (b)(4)) has been reported on mdr 1220648-2017-00081. the physician reported that he did not feel that the patient outcome was caused or contributed to by any issue with the impella cp. the patient support was considered a "hail mary" attempt to save this patient. on october 2, 2017 this complaint was re-opened to review this case and the failure investigation. at that time it was found that the failure analysis had determined that a close inspection of the impeller blades found damage to both blades which caused the reported suction alarms. this review determined that this event is an mdr reportable event. | [5 empty] | The console data logs were returned for this investigation. the log analysis showed the impella cp stopped early in the case, and was restarted, but low flow then occurred. later in the case a spike in motor current occurred and suction alarms sounded. the motor was set to p-6 with flow continuing to decrease until the end of the case. the impella cp was returned for inspection. the pump was in good condition with no obvious defects or signs of misuse. an initial inspection did not find any biomaterial in the cannula or inflow cage. the cannula was cut open to look for signs of biomaterial on the impeller. no biomaterial was found attached to the impeller. in order to reproduce the failure, the pump was run in a beaker of water. the pump produced low flow and triggered suction alarms when run. an inspection of the outflow cage found that sections were missing from both impeller blades, additionally, damage to the inflow cage was found. a portion of inflow cage window was missing and bent toward the impeller. there appeared to be some contact between the damaged portion and the impeller. one of the struts was also damaged and had a dent along one side. additional information received from the physicians confirmed that the doctors used medtronic's corevalve for the tavr procedure that had been performed the previous december. the drops in motor speed suggest contact of the impeller with an object. because fluoroscopic images of the pump were not returned a definitive root cause could not be determined. the manufacturing review revealed no other complaints against this lot of impella cp pumps. there is no corrective action because a definitive root cause could not be determined. this failure mode will continue to be monitored. (b)(4).
38162 | 1220648-2017-00081 | 6987959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6987959&pc=OZD | 01/17/2016 | 10/02/2017 |  | 10/30/2017 |  |  | 10/02/2017 | 10/30/2017 |  |  | 08/31/2017 |  | 10/30/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1212781 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 68 YR | [5 empty] | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Consequences Or Impact To Patient | No Known Impact Or Consequence To Patient | No Code Available | [39 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2017 a (b)(6) male patient was struggling following a transcatheter aortic valve replacement (tavr) that has been performed in (b)(6) 2015 (exact date unknown), and was a admitted to the hospital with a low ejection fraction (ef.) the patient began to decompensate, so an impella cp (serial (b)(4)) was inserted for hemodynamic support. a pump stop occurred early in the case, but the pump restarted and did not stop again. doctors were happy with support and planned to leave the pump in until the patient recovered. patient condition improved and pump explant was scheduled for monday (b)(6). on the fourth day of support pump flow decreased and suction alarms occurred. the first impella cp was explanted, and a new impella cp (serial (b)(4)) was inserted. insertion for the second impella was normal. a new sheath was used with a new guidewire. seven hours after insertion a pump stop occurred secondary to a coughing episode. pump was restarted in the patient. the pump produced low flow that continued to decrease throughout the case. eventually, the pump was explanted and the patient was scheduled to be switched to a tandem heart. patient expired in the ccl before surgery could be started. the physician reported that he did not feel that the patient outcome was caused or contributed to by any issue with the impella cp. the patient support was considered a "hail mary" attempt to save this patient. this mdr is reporting the issue with the first impella cp used during this event. the second pump (serial (b)(4)) has been reported on mdr 1220648-2017-00082. the physician reported that he did not feel that the patient outcome was caused or contributed to by any issue with the impella cp. the patient support was considered a "hail mary" attempt to save this patient. on october 2, 2017 this complaint was re-opened to review this case and the failure investigation. at that time it was found that the failure analysis had determined that a close inspection of the impeller blades found damage to both blades which caused the reported suction alarms. ths review determined that this event is an mdr reportable event. | [5 empty] | The console data logs were returned for this investigation. the log analysis confirmed the clinical account of the event. a pump stop was observed in the beginning of the case that was quickly resolved. pump support resumed with normal flow. later in the case several motor current spikes preceded a drop in pump flow with suction alarms. mean flow dropped below 1 l/min. flow was low for the rest of the case. the impella cp was returned for inspection. the pump was in good condition with no obvious defects or signs of misuse. biomaterial was found around the outflow cage on the pump. the cannula was cut open to look for biomaterial attached to the tip of the impeller. minimal biomaterial was found and it was unlikely that it was causing the low flow. when the pump was run, all biomaterial was expelled from the pump housing; however, the pump produced low flow and triggered suction alarms. inspection of the impeller found one blade was fractured and the other was damaged. without an intact impeller the motor current will be lower than normal and trigger suction detection by the minimum motor current algorithm. additional information received from the physicians confirmed that the doctors used medtronic's corevalve for the tavr procedure that had been performed the previous december. the root cause of the suction alarms was fractured impeller blades. because the fluoroscopic image of the pump was not returned, a definitive root cause of the fracture could not be determined. the manufacturing review revealed no other complaints against this lot of cp pumps. there is no corrective action because a definitive root cause could not be determined. this failure mode will continue to be monitored. (b)(4).
38163 | 1220648-2017-00098 | 6989261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6989261&pc=OZD | 10/03/2017 | 10/03/2017 |  | 10/30/2017 |  |  | 10/03/2017 | 10/30/2017 |  |  | 07/31/2019 |  | 10/30/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1301925 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 56 YR | [5 empty] | Malfunction | Yes | No | Pumping Stopped | Device Issue | [5 empty] | Cardiac Arrest | Death | [11 empty] | Death | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | A user medwatch report received by abiomed. there is corrected information included. the original age listed in initial report is (b)(6), however the corrected patient age in this report is (b)(6). | (b)(6) year old male patient presented to the er with a st-elevation myocardial infarction (stemi), coded, was intubated, and was brought to the cardiac catheterization laboratory (ccl). a temporary pacer was placed as well as an impella cp as a last ditch effort. the impella cp was inserted through the right femoral artery and support was initiated. left heart catheterization revealed 100% occlusion of the left anterior descending artery and 90% occlusion of the circumflex artery. intervention was performed and the patient was sent to icu, where the impella catheter system was switched to standard configuration. the following day, the impella cp purge pressure and motor current increased and during a call with the abiomed clinical support center it was discovered that 5% dextrose in lactated ringer's solution was used as the purge solution initially during the case instead of 5% dextrose in water. lactated ringer's solution contains additives such as sodium, calcium, lactate, potassium and chloride that are damaging to the pump. the impella cp ifu recommends using dextrose dw solution in the purge system, and warns the user against using saline in the purge system. the purge solution was exchanged from 5% dextrose in lactated ringer's solution to 5% dextrose solution, however the purge pressure and motor current continued to rise. the patient was continuing to decompensate while the clinical team made plans to exchange the pump. the pump stopped and the patient expired while the pump exchange was being prepared. |  |  |  |  | The user medwatch report received by abiomed included the patient age, which has been added to this follow up report. the user medwatch report listed the device serial number as (b)(4), however the serial number is (b)(4). 17-0929 follow up :user mdr from fda mw # (b)(4). | The returned impella cp, aic console data and case data was evaluated. the console data revealed the purge pressure begins to rise to above 1000 mmhg 36 hours into the case (about 2x the normal value). the console data logs show the case began at 11 pm on october 1st; and show the system was changed to standard configuration at 11:45 pm. according to the salesforce database, at 9:18 am the following morning the abiomed clinical support center (csc) received a call regarding high purge pressure, at which point the csc representative realized that 5% dextrose in lactated ringer's solution was in the purge system. the representative instructed the clinical staff to change the purge fluid immediately. during the time between 9:30 am and 12:00 pm, there were three purge fluid changes and one purge cassette change, however the purge pressure remained high. the data logs also reveal the pump stopped due to high motor current at 1:30 pm. post-case inspection of the pump revealed an accumulation of some type of gelatinous amber material in the gap between the impeller and sleeve bearing of the motor, which is where the purge fluid exits. the root cause of the high purge pressure is the buildup found in the motor. the pump most likely stopped due to the material that was accumulating in the gap between the impeller and sleeve bearing, causing higher friction on the impeller shaft until the pump exceeded its current rating and shut off. the most likely cause of the buildup is use of the 5% dextrose in lactated ringer's solution in the purge system. this solution contains multiple additives that likely can cause a reaction with the pump leading to the accumulation that was found in the motor. the impella cp ifu recommends using dextrose in water as a purge solution. (b)(4).
38164 | 1220648-2017-00090 | 6984062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6984062&pc=OZD | 09/29/2017 | 09/29/2017 |  | 10/27/2017 |  |  | 09/29/2017 | 10/27/2017 |  |  | 04/30/2019 |  | 10/27/2017 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1290987 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombus | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | Complainant reported that on a (b)(6) 2017 a (b)(6) female patient who was in cardiogenic shock (cgs) after suffering with an acute myocardial infarction (ami) had undergone an aortic valve replacement (avr). during the avr procedure the patient had required multiple defibrillations for ventricular tachycardia (vt) heart rhythm. following the avr procedure the patient was taken to the cath lab to place an impella cp and stent in the left main artery. the patient was then transferred to the icu. . on the following day, (b)(6) 2017 due to oozing at the impella access site a femstop had been applied, but was not hemostatic and the dressing was saturated. it was noted that the dressing had obstructed the view of the centimeter marker (cm) and the tuohy borst valve lock; consequently, the site was re-dressed and the femstop re-applied resulting in clear hemostasis. two units of replacement blood products were administered to the patient. on (b)(6) 2017 the patient was reported to be doing "great", and was being weaned from the impella cp. after 43 hours of successful patient support the impella cp was removed from the patient on (b)(6) 2017. upon removal a thrombus was noted on the impella pigtail, and another piece was found on the inflow cage. the complainant reported that there was no harm to the patient as a result of the clot formation. the patient's impella cp support was reported as successful, and that the patient was "doing fine." | [5 empty] | The impella cp and the automatic impella controller data logs were returned to abiomed for evaluation. a visual examination of the impella cp was performed. during the evaluation material was confirmed to be inside the inflow cage, but the material reported to be on the pigtail had already been removed at the user facility. there were no defects in the pigtail or the inflow cage; in addition, there was no material found on the impeller or inside the pump's outflow cage. the device history record review did not reveal any other impella cp complaints for pumps with this lot number. the aic data logs reflected that the pump was operating properly during the case and the logs confirmed the clinical narrative of weaning the patient off of the pump. there were no low flows and no suction alarms during the case. a definitive root cause of the thrombus on the pump could not be determined as the pump was running properly, without low flows, and contained no defects that would have caused thrombus to occur. no corrective action is recommended at this time because the failure mode was determined to have a low risk index based on minor severity and very low occurrence. failures of this type will continue to be monitored and trended. (b)(4).
38165 | 1220648-2017-00093 | 6984066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6984066&pc=OZD | 09/29/2017 | 09/29/2017 |  | 10/27/2017 |  |  | 09/29/2017 | 10/27/2017 |  |  | 04/30/2019 |  | 10/27/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 128955 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | Yes | Appropriate Term/Code Not Available | [6 empty] | Mitral Regurgitation | Rupture | [40 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2017 an emergent case of a 65 year old male patient who had a heart attack and who was in cardiogenic shock was admitted to the hospital. the patient was also found to be suffering from mitral regurgitation (mr), mitral valve disease (mvd) and a left ventricular end-diastolic pressure of 40. the physician placed an impella cp prior to performing a percutaneous coronary intervention (pci.) the physician reported that the impella placement went perfectly and the impella worked failure free during the pci. the physician performed a complete revascularization, of the circumflex artery, left anterior descending artery and the right coronary artery. following the completion of the procedure the oscor sheath was peeled away and the repositioning sheath was advanced and secured. the patient was successfully supported for 6.65 hours. post impella cp explant the patient was noted to have a ruptured mitral valve papillary muscle. it was reported that prior to the impella cp insertion, the patient was noted to have mitral regurgitation and a dilated left atrium (la.) the physicians stated that they were unsure of whether or not the papillary damage was new or if the damage was old as the la was dilated. they also stated that the patient's mr could have been chronic. patient support was reported as successful, and the patient was reported to be in stable condition following patient support. additional event and patient information regarding repair of the ruptured mitral valve papillary muscle was requested, but was unable to be obtained. | [5 empty] | The impella cp and the automatic impella controller data logs were returned for evaluation. pump was returned in good condition. visual inspection of the returned impella cp found no defects on the pigtail, inflow cage, cannula, pump housing or motor housing that might contribute to the damage of the mitral valve. the returned pump passed the ring gage test. the console data logs were not reviewed for this event as they did not pertain to the failure mode. a manufacturing review was done and no other complaints have been made against this lot of impella cp pumps that pertain to this failure mode. . in conclusion, no defects that might contribute to the damage of the mitral valve muscle were found on this pump; consequently, the root cause of the ruptured mitral valve pap muscle was unable to be determined. no corrective action is recommended at this time because the failure mode was determined to have a low risk index based on moderate severity and very low occurrence. failures of this type will continue to be monitored and trended. (b)(4).
38166 | 1220648-2017-00086 | 6964456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6964456&pc=OZD | 10/02/2017 | 10/02/2017 |  | 10/20/2017 |  |  | 10/02/2017 | 10/20/2017 |  |  | 05/31/2019 |  | 10/20/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1300179 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Positioning Problem | [6 empty] | Perforation of Vessels | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On the (b)(6) 2017, the cardiac catheterization lab notified the abiomed representative that an impella cp was being urgently placed for a (b)(6) year old male admitted with cardiogenic shock and acute myocardial infarction. the patient had multivessel coronary artery disease and was declined by the cardiothoracic team. the hospital intubated the patient as he was decompensating, and began the placement of the 14fr sheath for the impella cp. prior to inserting the sheath, the team evaluated the vessel size and believed it to be of a diameter to accommodate the 14fr sheath. while the team was placing the guidewire for the impella cp through the sheath the team felt resistance. an angiogram of the access site showed the iliac artery had dissected. the patient began to further decompensate, was coded, and expired. the team never placed the impella cp for hemodynamic support. in post procedure discussions the team stated that the patient's death was not due to the dissected artery, but rather to the rapid decompensation with cardiogenic shock. | [5 empty] | The hospital did not return the 14fr sheath for evaluation. the data logs were not collected, as the impella pump was never utilized for support. the manufacturing review of the lot of 14fr sheath found no other failures from any units in the lot. the root cause of the vascular injury was unable to be determined, as the product was not returned. no corrective action is recommended as the failure mode has a low risk index. the failure mode will be trended and monitored. internal reference (b)(4).
38167 | 6965040 | 6965040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6965040&pc=OZD | 06/05/2017 | 09/26/2017 | 09/26/2017 | 10/20/2017 | 09/26/2017 | 09/26/2017 |  | 10/20/2017 |  |  | 03/21/2021 |  | 10/20/2017 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 16313 |  |  |  | Yes | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 64 YR | [5 empty] | Malfunction |  | Yes | Device Inoperable | [6 empty] | No Information | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL | 1067416796 |  | Surgery for coronary artery bypass graft (cabg) with mitral valve repair (mvr) was performed successfully. when surgeon tried to restart impella, the device would not work.
38168 | 1220648-2017-00076 | 6962433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6962433&pc=OZD | 09/17/2017 | 09/19/2017 |  | 10/19/2017 |  |  | 09/19/2017 | 10/19/2017 |  |  | 02/28/2019 |  | 10/19/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1282508 |  | 0048-0002 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | Yes | Difficult to Insert | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2017 an impella cp was placed in a (b)(6) female patient who had an acute myocardial infarction (ami) and who was in cardiogenic shock (cgs). the impella pump was placed in the patient following the performance of a percutaneous transluminal coronary angioplasty (pci.) the impella cp placement was reported to have been "very very" difficult, and took 4 hours to place the pump. the patient was supported with the impella cp for approximately 24 hours, but during that time patient bleeding occurred, and it was decided to replace the pump, with another impella. the replacement impella cp was successfully placed in the patient, but bleeding continued. the bleeding was thought to be coming from the back of the repositioning sheath, but closer observation by the clinicians revealed that damage had occurred to the hemostatic valve/ ring inside the repositioning sheath. the clinicians reported that it was assumed that the damage to the valve was due to the repeated in and out movements and manipulations of the impella catheter during insertion. an estimated 2 liters of patient blood was lost, and 7 bags of replacement packed red blood cells were administered to the patient. the patient expired on (b)(6) 2017 as a result of multi-organ failure. there was no indication from the physician that the patient outcome was caused or contributed by any issue with the impella cp. | [5 empty] | The impella cp was not returned for evaluation, and the console data logs are irrelevant to this case. a review of the clinical case was done. the repositioning unit (p/n 002164) was part of lot batch 1264867, and was manufactured at abiomed, inc. in (b)(4). there are no complaints for other repositioning units in this lot. the valve subassembly of the repositioning unit (p/n 0046-0552) was part of lot batch 0122015 and was inspected in the repositioning unit (p/n 002164) and was part of batch 1264867, which was also made in abiomed, inc. in (b)(4). there are no complaints for other repositioning units in this lot. the valve subassembly of the repositioning unit (p/n 0046-0552) was part of batch 0122015 and was inspected at abiomed, inc. in (b)(4) on 10/28/2016. the lot passed inspection and there are no complaints for other valve subassemblies in this lot. this failure mode was determined to be low risk due to very low occurrence ((b)(4)) and moderate severity. the root cause of the blood leaking from the repositioning unit into the sterile sleeve was unable to be determined because the product was not returned. no corrective action is recommended at this time because the failure mode was determined to have a low risk index based on very low occurrence and moderate severity. failures of this type will continue to be monitored and trended. (b)(4). | This supplemental mdr is being submitted to correct the second paragraph of the event description contained in initial medwatch report 1220648-2017-00076. the second paragraph should have statedthe following: after transfer of the patient to the icu, patient bleeding was observed: blood was leaking into the steril repositioning sleeve on the catheter. the bleeding was thought to be coming from the back of the repositioning sheath, but closer observation by the clinicians revealed that damage had occurred to the hemostatic valve/ ring inside the repositioning sheath. the clinicians reported that it was assumed that the damage to the valve was due to the repeated in and out movements and manipulations of the impella catheter during insertion. the clinicians considered replacing the pump; however, considering the very difficult initial placement procedure, they decided to continue with the pump with the bleeding repositioning sheath. the patient was supported with the impella cp heartpump for over 24 hours.
38169 | 6947533 | 6947533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6947533&pc=KFM | 10/02/2017 | 10/03/2017 | 10/03/2017 | 10/13/2017 | 10/03/2017 | 10/03/2017 |  | 10/13/2017 |  |  |  |  | 10/13/2017 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | KFM |  |  | IMPELLA 2.5 AIC PUMP SET |  |  | ABIOMED, INC. |  |  | US |  |  |  | 005042 |  | Yes | [5 empty] | No | Physician |  |  | User Facility report | [8 empty] | 75 YR |  |  | 88 |  |  | Malfunction | No | Yes | Bent | Collapse | Sticking | Dent in Material |  |  |  | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella was inserted via femoral artery to left ventricle. after placement, tried to pull out the guide wire but it was stuck in the outlet of the catheter and it prolapsed onto itself. pulled back the entire device into smaller vessels but the device bent onto itself more, causing two big dents in the wire and the catheter was bent. this caused the wire to be stuck in the motor housing per the representative. this opted for an open surgical removal from a surgeon. manufacturer response for impella catheter 2.5 (per site reporter): once device is released by risk and pathology, then the manufacturer would like to pick up the device for their analysis.
38170 | 1220648-2017-00077 | 6947727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6947727&pc=OZD | 08/17/2017 | 09/21/2017 |  | 10/13/2017 |  |  | 09/21/2017 | 10/13/2017 |  |  | 05/31/2019 |  | 10/13/2017 |  | IMPELLA CP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1294473 |  | 0048-3058 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 83 YR | [5 empty] | Malfunction | Yes | No | Device-Device Incompatibility | [6 empty] | No Known Impact Or Consequence To Patient | No Code Available | [11 empty] | Hospitalization | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An (b)(6) male was admitted in cardiogenic shock after a transcatheter aortic valve replacement (tavr) was done at the cleveland clinic. the tavr was a corevalve evolut r. the impella cp was placed via the tavr 20 fr sheath at the left femoral. within 10 minutes of support, the pump's flow dropped and suction alarms were noted. the team evaluated the pump position and ruled out the presence of clot in the left ventricle by echo. the pump was replaced by another impella cp. the second and third impella cp, will be discussed in subsequent mdrs ( medwatch 1220648-2017-00078 and 1220648-2017-00087) while on support with the third impella cp the flows were again decreased, and so the team chose to instead place a tandem heart and ecmo for oxygenation needs. the patient did expire after all of the above mentioned support was attempted. | [5 empty] | The impella cp pump and data logs were returned for analysis. the data logs show less flow early in the case, which means that the impella and tavr corevalve interacted after initial placement when the impella shifted out of position. in the data logs there was an alarm for "impella position wrong". upon examination of the pump, the returned pump has broken impeller blades. the blade fracture was the root cause of the low pump flows. the manufacturing review revealed no other complaints against this lot of cp pumps. no corrective action is warranted at this time because the failure mode was low occurrence. the occurrence rate was noted to be 0.01%. the failure mode will be monitored and trended.
38171 | 1220648-2017-00078 | 6947731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6947731&pc=OZD | 08/16/2017 | 09/21/2017 |  | 10/13/2017 |  |  | 09/21/2017 | 10/13/2017 |  |  | 05/31/2019 |  | 10/13/2017 |  | IMPELLA CP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1292087 |  | 0048-3058 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 83 YR | [5 empty] | Malfunction | Yes | No | Device-Device Incompatibility | [6 empty] | No Known Impact Or Consequence To Patient | No Code Available | [11 empty] | Hospitalization | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An (b)(6) male was admitted in cardiogenic shock after a transcatheter aortic valve replacement (tavr) was done at the cleveland clinic. the impella cp (medwatch 1220648-2017-00077) was placed via the tavr 20 fr sheath at the left femoral. within 10 minutes of support, the pump's flow dropped and suction alarms were noted. the team evaluated the pump position and ruled out the presence of clot in the left ventricle by echo. the pump was replaced by another impella cp. this second impella cp, also exhibited low pump flows. once again, the team evaluated the pump position and lack of left ventricular clot by echo before explanting and replacing the impella cp with a third impella cp (medwatch 1220648-2017-00087) . while on support with the third impella cp the flows were again decreased, and so the team chose to instead place a tandem heart and ecmo for oxygenation needs. the patient did expire after all of the above mentioned support was attempted. | [5 empty] | The impella cp pump and data logs were returned for analysis. the data logs show less flow early in the case, which means that the impella and tavr interacted after initial placement when the impella shifted out of position. in the data logs there was an alarm for "impella position wrong". upon examination of the pump, the returned pump has broken impeller blades. the blade fracture was the root cause of the low pump flows. the manufacturing review revealed no other complaints against this lot of cp pumps. no corrective action is warranted at this time because the failure mode was low occurrence. the occurrence rate was noted to be 0.01%. the failure mode will be monitored and trended.
38172 | 1220648-2017-00087 | 6947732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6947732&pc=OZD | 08/16/2017 | 09/21/2017 |  | 10/13/2017 |  |  | 09/21/2017 | 10/13/2017 |  |  | 05/31/2019 |  | 10/13/2017 |  | IMPELLA CP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1290935 |  | 0048-3058 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 83 YR | [5 empty] | Malfunction | Yes | No | Device-Device Incompatibility | [6 empty] | No Known Impact Or Consequence To Patient | No Code Available | [11 empty] | Hospitalization | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | (b)(6) year old male was admitted in cardiogenic shock after a transcatheter aortic valve replacement (tavr) was done at the cleveland clinic. the impella cp was placed via the tavr 20 fr sheath at the left femoral (medwatch 1220648-2017-00077). within 10 minutes of support, the pump's flow dropped and suction alarms were noted. the team evaluated the pump position and ruled out the presence of clot in the left ventricle by echo. the pump was replaced by another impella cp (medwatch 1220648-2017-00078). a second impella cp, also exhibited low pump flows. once again, the team evaluated the pump position and lack of left ventricular clot by echo before explanting and replacing the impella cp with a third impella cp. while on support with the third impella cp the flows were again decreased, and so the team chose to instead place a tandem heart and ecmo for oxygenation needs. this third impella was discarded subsequent to explant, and was not able to be retrieved for this medwatch report's investigation. the patient did expire after all of the above mentioned support was attempted. | [5 empty] | The impella data logs were returned for analysis, although the pump was not. the data logs show low flow early in the case, which means that the impella and tavr interacted after initial placement when the impella shifted out of position. the manufacturing review revealed no other complaints of the same failure mode against this lot of cp pumps. the root cause of the low flows was most likely fractured impeller blades. no corrective action is warranted at this time because the failure mode was low occurrence. the occurrence rate was noted to be (b)(4). the failure mode will be monitored and trended. (b)(4).
38173 | 1035166-2017-00087 | 6948640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6948640&pc=DYB | 10/12/2017 | 03/08/2018 |  | 10/13/2017 |  |  | 10/12/2017 | 10/13/2017 |  |  | 05/01/2020 |  | 10/13/2017 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3014 |  |  | OSCOR INC |  |  | US |  | C1-13369 |  | 0052-3014 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 64 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Myocardial Infarction | [12 empty] | Death | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | Case information: (b)(6) year old male with stemi brought to the hospital. femoral angiogram performed prior to sheath insertion showing sufficient vessel size. the 14 fr oscor introducer sheath was inserted in the rfa. iliac dissection was noted post sheath insertion with angiogram. impella insertion was aborted. the patient did not survive. the sheath was taken away with patient and could not be retrieved. we will not be receiving this product back, so this will be all the information we will be getting regarding the case. | [5 empty] | Conclusion not yet available-evaluation in progress. this initial mdr is being submitted to meet our requirements of reporting. a follow-up will be submitted as soon as the investigation is complete. | The device will not be returned for analysis. the clinical team did not attribute the death to the device. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity and no new failure mode identified. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type. quality assurance procedure (abiomed introducer sheath) in-process & final inspection: device is tested one hundred percent for cap and seal inspection, a leak test is perform according to procedure, also an occlusion test is perform. sideport with valve assembly samples are randomly selected for visual inspection: with the naked eye, inspect sideport assembly for excessive glue. smeared glue to the sideport assembly is not acceptable. verify the sideport tube is properly bonded to hub and stopcock valve by gently tugging on both joints. note: if one part fails, one hundred percent inspection of the lot for bond integrity of sideport tube to hub and stopcock valve is required. instructions for use (ifu), abiomed adelante s2 - hemoslaticp peel away & sideport. under precaution: aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. aspirate all air from the sheath valve assembly by using a syringe connected to the sideport. flush the sheath with 5 cc of saline immediately before peeling the sheath away in order to minimize back bleeding.
38174 | 1220648-2017-00074 | 6941218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6941218&pc=OZD | 09/11/2017 | 09/12/2017 |  | 10/11/2017 |  |  |  | 10/11/2017 |  |  | 01/31/2019 |  | 10/11/2017 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1276840 |  | 005062 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 52 YR | [5 empty] | Death | Yes | No | Positioning Problem | [6 empty] | Perforation | [12 empty] | Death | Required Intervention |  |  |  |  | CARDIOPULMONARY BYPASS MACHINE | [22 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On (b)(6) 2017 a (b)(6) male patient presented with chronic heart failure, shortness of breath, coronary artery disease, severe mitral valve regurgitation (mr) and tricuspid regurgitation (tr). the patient's ejection fraction was estimated to be 25%.the physician believed the patient should be transferred to a transplant center, however the patient wished to stay where he was. the patient underwent a coronary artery bypass grafting (cabg) procedure, as well as mitral and tricuspid ring repair while on cardiopulmonary bypass support. after the procedure the impella 5.0 was implanted without difficulty and allowed complete separation from the bypass machine. the patient was transferred to the unit in stable condition. the impella 5.0 provided support for the next five days. the clinical team was happy with the stability of the patient. on (b)(6) the patient was beginning to wean off of the device. the motor current became flat at the time the patient was turned on his side, and a "position unknown" alarm was triggered. a chest x-ray revealed the impella had retracted out of the left ventricle. the patient was hemodynamically stable at this time, and remained stable for the next several hours with the device completely out of position in the descending aorta. the patient was brought to the or for the removal of the pump and exchanging of the endotracheal (et) tube. the patient became unstable when the anesthesiologist attempted to exchange the et tube intraoperatively. when the et tube was exchanged the patient's blood pressure dropped, and cpr was initiated. transesophogeal echocardiogram (tee) revealed massive aortic insufficiency and a tear in the aortic valve leaflet. the patient was unable to maintain a blood pressure and expired in the or. according to the case the physician did not feel the impella contributed to the valve problem. | [5 empty] | The pump was not returned from the field for evaluation. several attempts were made to retrieve the pump. the root cause of the aortic valve tear is unable to be determined because the product was not returned. the root cause of the pump retracting from the patient's left ventricle is most likely due to patient movement. a review of the device and lot histories did not reveal any nonconformities. (b)(4).
38175 | 6935619 | 6935619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6935619&pc=OJE | 08/29/2017 | 09/06/2017 | 09/06/2017 | 10/10/2017 | 09/06/2017 | 09/06/2017 |  | 10/10/2017 |  |  | 04/30/2019 |  | 10/10/2017 |  | IMPELLA LEFT VENTRICULAR ASSIST DEVICE |  |  | IMPELLA |  |  | OJE |  |  | 129265 |  |  | ABIOMED |  |  | US |  |  |  | 00480003 |  | Yes | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 42 YR |  |  | 58 |  |  | Malfunction | No | Yes | Device Dislodged or Dislocated | [6 empty] | Foreign Body In Patient | [12 empty] | Other | [28 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The patient required placement of an impella ventricular assist device via left common femoral introducer sheath. the device became entrapped on attempt at removal.
38176 | 1220648-2017-00085 | 6926310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6926310&pc=OZD | 09/09/2017 | 09/10/2017 |  | 10/07/2017 |  |  | 09/10/2017 | 10/07/2017 |  |  | 07/31/2019 |  | 10/07/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1295775 |  | 0048-0004 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Device Operates Differently Than Expected | Positioning Problem | [5 empty] | Hemolysis | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | (b)(6) year old male, with known renal dysfunction, was admitted in cardiogenic shock from a late presentation anterior myocardial infarction, to a hospital in (b)(6) on (b)(6) 2017. the impella cp was placed for support and the patient was sent to the icu for continued monitoring. upon admission to the icu, the echo was performed, as recommended to assess pump position, and the pump was pulled back 1 cm in relation to the aortic valve. the urine was dark. and later that evening, hemodialysis was begun. during support on the following day ((b)(6) 2017), there was an observation made that the groin access site at the left femoral artery continued to ooze blood. there were no surgical interventions or measures taken to remedy the ooze, rather simply a re-do of the suture. the pump was again repositioned in the lab under fluoroscopy. the team at this time noted that the urine color was clearing. to remedy the signs of hemolysis in the patient, dialysis was begun and the patient received 3 units of blood products while urine output was minimal. the patient's support with the impella cp continued in the icu with labs hemolyzing and repeated echo done to assess pump position. on the (b)(6), the patient's condition began to further deteriorate with arrhythmias which required shocks. the ventricular tachycardia continued and on the (b)(6), the patient expired when the cardiac code failed. | [5 empty] | The impella pump and data logs were returned and analyzed since the initial medwatch was filed. the data logs show multiple "impella wrong position" alarms which are occurring at the same time as the field representative was told there was hemolysis by the staff nurse. the pump was run through two separate hemolysis tests in-house at abiomed. the pump was then broken down and analyzed microscopically. an anomaly was found along most of the circumference of the pump housing. there was a misalignment of the cannula and pump housing that caused a step feature to be exposed. this resulted in an edge of the outflow cage to be in the path of the blood flow and contribute to hemolysis. the pump lot was evaluated and found no other complaints leveled against it. the failure mode was determined to be low risk and there will be no corrective action for the failure, though it will be trended and monitored. internal reference (b)(4). | The pump was returned for analysis to abiomed, and the analysis and hemolysis testing is ongoing. the data logs have yet to be returned for analysis. a supplemental medwatch will be completed upon conclusion of the investigation. (b)(4).
38177 | 1220648-2017-00079 | 6913446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6913446&pc=OZD | 09/14/2017 | 09/14/2017 |  | 10/04/2017 |  |  | 09/14/2017 | 10/04/2017 |  |  | 06/30/2019 |  | 10/04/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1292373 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 48 YR | [5 empty] | Malfunction | Yes | No | Device Operates Differently Than Expected | [6 empty] | Hemolysis | Hematuria | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On (b)(6) an impella cp was placed for support on a (b)(6) male with ischemic cardiomyopathy during a scheduled coronary angioplasty (pci) and defibrillator (icd) placement. unfortunately, due to insurance coverage issues, the pci was aborted and postponed. the cp was left on for support with the plan to attempt the pci the following day. upon admission to the icu for monitoring the patient's urine color was changed, and tinged red, which the hospital team attributed to hemolysis. the cp was repositioned multiple times in the following days in an attempt to alleviate the hemolysis. the patient was monitored over the weekend and labs were hemolyzing. the patient did go into total heart block and on the monday and was taken to the electrophysiology (ep) lab for an emergent icd placement. the impella cp was supporting the patient during the ep study and icd placement, but did exhibit suction and wrong position alarms. the cp position was again checked by echo. positioning can result in hemolysis if there is suction or interaction with the aortic valve. the patient did have the cp explanted on (b)(6), after the icd placement, and was placed on dialysis. in total the cp supported the patient for over 57 hours. | [5 empty] | The root cause of the hemolysis could not be determined, as neither the product nor data logs were returned for analysis. the analysis of the manufacturing record reveal no other complaints leveled against this lot of cp pumps. the abiomed team did conduct post event re-training, at the hospital, as a result of the reported hemolysis. there will be no corrective action because the root cause was unable to determined. the failure mode will be trended and monitored. (b)(4).
38178 | 1220648-2017-00070 | 6907829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6907829&pc=OZD | 09/08/2017 | 09/11/2017 |  | 10/02/2017 |  |  | 09/11/2017 | 10/02/2017 |  |  | 09/30/2018 |  | 10/02/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GT |  | 1261680 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 49 YR | [5 empty] | Injury | Yes | No | Missing Value Reason | [6 empty] | Hematoma | Injury | Blood Loss | [10 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | Complainant reported that on (b)(6) 2017 a (b)(6) male, who was a noncompliant diabetes mellitus patient, was admitted to the hospital with reports of worsening chest pains for the last 2 days. an echo was performed which revealed that the patient had a large ventricular septal defect. the patient was brought to the icu where he decompensated requiring cpr and advanced airway placement. the patient was then transferred to the ccl for impella cp placement due to refractory cardiogenic shock (cgs.) the impella cp catheter was placed in the left femoral artery (lfa) without issue. shortly after arriving in the icu the patient became agitated and attempted to sit up in bed, bending both legs. the patient became adequately sedated with the administration of medication. shortly thereafter, the impella site was notably saturated with blood. manual compression was applied by the nurse. a large developing hematoma soon became apparent around insertion site, which extended down to the patient's scrotum. bleeding around the insertion site gradually worsened despite applied manual compression, initially noted as a steady trickle of blood and ultimately progressing to a pulsatile spray. the physician was notified and applied a vicryl suture to the site in an attempt to better control bleeding; however the bleeding continued, and the patient was administered 4 units of packed red blood cells. vascular surgery was then consulted and it was decided that the best course of action was to take the patient to the or for explant of the impella cp and to perform a surgical repair of site, and to then place another impella cp to the contralateral leg. due to concern over adequacy of vessel size secondary to vasoconstriction from high dose vasopressors, the physician chose to place the smaller diameter impella lp2.5 catheter. in an attempt to minimize the patient's time off of hemodynamic support, access to the contralateral leg with 13fr peel away sheath was made and the pigtail catheter was inserted into lv, and made ready for delivery of the 0.18 wire. the 0.18 wire was then delivered into lv once the cp catheter pulled into the left iliac. the impella lp2.5 was inserted and patient support re-initiated, initially delivering max flows at 2.5l/min. moments later all pulsatility was lost and the patient was noted to be in asystole. cpr initiated and several rounds were administered. patient rhythm was eventually reestablished for a brief period. impella pump position was reconfirmed under c-arm and device explant/vascular repair was momentarily resumed. the patient ultimately re-entered lethal arrhythmia requiring cpr; however the patient was unable to be resuscitated and was pronounced in the or. the physician reported that the issues with the impella cp did not cause or contribute to the patient outcome. | [5 empty] | The impella cp pump was not returned for evaluation, as it was discarded subsequent to patient use; consequently, no device analysis could be performed and the root cause of the complaint bleeding and the occurrence of the hematoma were unable to be determined. no corrective action is recommended at this time because the failure mode was determined to have a low risk index based on moderate severity and very low occurrence. failures of this type will continue to be monitored and trended. (b)(4).
38179 | 1220648-2017-00075 | 6880327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6880327&pc=OZD | 09/05/2017 | 09/05/2017 |  | 09/20/2017 |  |  | 09/05/2017 | 09/20/2017 |  |  | 06/30/2019 |  | 09/20/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 1294452 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 53 YR | [5 empty] | Injury | Yes |  | Device Operates Differently Than Expected | [6 empty] | Hemolysis | Thrombus | Hematuria | Blood Loss | [9 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) active male was admitted to a hospital with complaints of shortness of breath (sob) from recent work outdoors. upon admission, the patient's troponin was elevated and he was intubated for respiratory support. the cardiac catheterization revealed multivessel coronary artery disease and an ejection fraction of only 20%. the team chose to place an impella cp for support, via a cook introducer, in the left femoral artery. the patient was then transferred to the icu for monitoring while awaiting transfer to the tertiary medical center for cardiothoracic surgery to perform bypass grafts, cabg. the patient was successfully transferred to the tertiary center and a groin access site ooze was observed at the left femoral artery (lfa), as well as a change in urine color. the team called for an echo to be performed to rule out hemolysis, which was an assumed cause of the pink tinged urine. labs were also drawn that could determine the diagnosis of hemolysis. the team began to treat the groin bleed with manual pressure, while awaiting transfer to the operating department. while the patient was in the operating department, the team gave one unit of blood and the lfa was closed by a suture, to alleviate continued blood loss. despite the blood loss, the limb was intact and there was no vascular damage needing further vascular intervention. the patient at this time also had the care escalated to the impella 5.0 via the right axillary artery, and the impella cp was explanted. when the cp was explanted thrombus was noted, but this did not cause any harm or injury and could have been due to changes in the anticoagulation regimen throughout his two days of cp support. the patient is successfully still supported to date with the second impella, the 5.0, while awaiting a durable vad to be placed. the vad placement is delayed due to his colon cancer and planned resection surgery. | [5 empty] | Neither the data logs nor the product in question were returned for investigation. without the data logs or the device, a complete investigation was not able to be completed. the causes of the hematuria and access site bleed were unable to be determined. no corrective action is recommended at this time, and the failure mode will be monitored and trended. internal reference (b)(4).
38180 | 6894724 | 6894724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6894724&pc=KFM | 08/23/2017 | 09/07/2017 | 09/07/2017 | 09/19/2017 |  |  |  | 09/19/2017 |  |  |  |  | 09/19/2017 |  | IMPELLA 2.5 |  |  | IMPELLA 2.5 |  |  | KFM | [5 empty] | ABIOMED |  |  | US |  | 130683 |  |  |  | No | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 86 YR |  |  | 82 |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 08/22/2017 | Y | HOSPITAL |  |  | Pt presented for elective percutaneous coronary intervention to left main coronary artery which was severely calcified and significantly stenosed. the revascularization plan included use of an impella device given pt's age and comorbidities. pt was considered high risk for surgical revascularization and had declined surgery. after placement of impella device, pt became diaphoretic, pale, developed tachycardia and became unresponsive. cpr initiated. pt became hemodynamically unstable in hypovolemic/cardiogenic shock with pericardial effusion and tamponade from perforation of ventricle. massive transfusion protocol initiated. transported to operating room suite for repair of perforation. post procedure, pt transferred to icu. subsequently, developed hypotension and pea arrest. resuscitation efforts unsuccessful. pt expired.
38181 | 6868311 | 6868311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6868311&pc=PBL | 09/06/2017 | 09/07/2017 | 09/07/2017 | 09/15/2017 | 09/07/2017 | 09/07/2017 |  | 09/15/2017 |  |  |  |  | 09/15/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Yes | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 55 YR | [5 empty] | Malfunction |  | Yes | No Display/Image | Device Emits Odor | Smoking | No Flow |  |  |  | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Patient has an impella, and around 0830 the writer was in room checking the device. while i was writing down numbers to transfer into the document flow sheet, the flows dropped to 0 lpm, but the device was still on. i started troubleshooting the connections/device when i noticed it started to smell like burning and some smoke was seen coming from the impella console and the screen had gone black and would not restart. i unplugged the device and replugged it into a new outlet. the charge registered nurse and critical care anesthesia attending were then in the room and attempted to again unplug and replug it back in and restart the console with no success. the smoke had stopped but the smell of burning continued. once the rep had reached the room the console restarted, but it was continuing to smell like burning so the console was changed out with another one with successful resumption of flow. the patient remained stable throughout and the device was off for about 5 minutes. perfusion was paged and at bedside. the dr. was also paged and at bedside - decision made prior to resumption of flow to take the patient to the operating room (or) early for his coronary artery bypass grafting (was already scheduled for the same day at noon). once flow was re-established the surgical team decided to move forward with going to the operating room and the patient was transported to the operating room at approximately 1000.
38182 | 1220648-2017-00072 | 6870166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6870166&pc=OZD | 09/09/2017 | 09/10/2017 |  | 09/15/2017 |  |  | 09/10/2017 | 09/15/2017 |  |  | 01/31/2019 |  | 09/15/2017 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS NON-ROLLER TYPE |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1276840 |  | 005062 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | [6 empty] | Mitral Regurgitation | [12 empty] | Hospitalization | Required Intervention | Disability | [26 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | (b)(6) year old male was admitted with a stemi and placed upon impella 5.0 support as decisions were made for angioplasty versus cabg. he was noted to have an ejection fraction of only 5-10%. the impella 5.0 was inserted through an axillary artery cutdown with hemashield 10mm graft placement. the insertion was successful and support initiated on (b)(6) 2017. the pump supported the patient successfully for 9 days and then the patient was transferred to a tertiary medical center for further evaluation and possible escalation of care to vad or transplant. during these days, the physician and team were able to attempt intervention upon the native coronary arteries and place stents within the left anterior descending and left main arteries. while admitted to the second hospital, the patient's weaning off the impella 5.0 began and the support team observed mitral valve regurgitation (mr). the mr prompted the team to replace the valve while also undergoing a single bypass (cabg) graft placement. the physician noted that the posterior cord was torn and prolapsed. in total the impella 5.0 supported the patient from (b)(6). the patient recovered and is stable. | [5 empty] | The pump was discarded within the operating department. the data logs were returned for analysis. the logs show no suction alarms were triggered during support and there is no indication from the placement signal or motor current that suction occurred during support. logs did not indicate anything abnormal during support. the manufacturing review did not reveal any other complaints against this lot of aachen, germany manufactured pumps. the product was not returned and could not therefore be inspected to determine the cause of the torn chordae. in addition, no echo imaging was shared for analysis. the complaint is a medium risk index based upon low occurrence and major severity. no corrective action is warranted, as the root cause was unable to be determined. the failure mode will be monitored and trended. (b)(4).
38183 | 1220648-2017-00067 | 6866544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6866544&pc=OZD | 08/15/2017 | 08/15/2017 |  | 09/14/2017 |  |  | 08/15/2017 | 09/14/2017 |  |  | 04/30/2019 |  | 09/14/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1285513 |  | 0048-0020 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 71 YR | [5 empty] | Malfunction | Yes | No | Difficult To Position | Device Issue | Patient-Device Incompatibility |  |  |  |  | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On (b)(6) 2017 a (b)(6) year old male was taken to the emergency room after experiencing chest pain during a cat scan. he had a cardiac arrest in the er and was taken to the cardiac cath lab for care. the impella cp was placed for support due to the need for support during angioplasty to the left main, left anterior descending, and left circumflex coronary arteries. that evening, he was admitted to the icu for monitoring. upon admission to the icu the patient's urine output began to change color and an echo was obtained to evaluate position in the left ventricle (lv) of the impella cp in relation to the valve. the standard blood draws were also hemolyzing at the time. the team did not take the advice of the abiomed representative, and re-adjust position of the impella. the following morning, the (b)(6), the position of the impella cp was adjusted from 1.4cm to 3.4cm. according the ifu of the impella cp: "positioning the inlet area of the catheter 3.5 cm below the aortic valve annulus and in the middle of the ventricular chamber, free from the mitral valve chordae." the patient was infused with 3 units of red blood cells after repositioning. the patient's hematocrit and hemoglobin were 8.9 and 24.2. throughout the days of the (b)(6), the urine output and color were monitored. the color was still indicative of hemolysis. the patient was also placed on dialysis with his creatinine level being 6 mg/dl. throughout the day of (b)(6) the team made adjustments to the pump's p level in attempted troubleshooting for the signs of hemolysis, although this did little to remedy the urine color change. only with repositioning of the pump did the urine output improve. on the (b)(6) the urine color was clearing and impella weaning began. on the night of the (b)(6) the impella was explanted and the patient recovered without harm or injury. | [5 empty] | The impella cp pump and data logs were returned for analysis. the data log analysis confirmed the clinical report of the pump position being adjusted in relation to the valve. as the ifu states hemolysis is a "known potential adverse event" and furthermore: "evaluate the position of the impella® catheter if the automated impella® controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it." the p-level changes, observed within the data logs, also coincide with the clinical report. the p-levels were decreased in an attempt to resolve hemolysis. later the p-level was increased, while the pump was in proper position, and the hemolysis issues worsened, according to the clinical account, within hours. the p-level was again decreased and hemolysis issues reportedly resolved. the cp pump itself was analyzed, and nothing was identified on the pump that could have caused or contributed to the hemolysis. the pump was run through standard hemolysis testing procedures and passed. the reported failure was not reproduced. the lot of cp pumps was investigated in a manufacturing review. no other pumps have complaints leveled from the field for hemolysis. there is no corrective action as the root cause of the hemolysis could not be determined. the failure will continue to be monitored and trended. internal reference (b)(4).
38184 | 1220648-2017-00071 | 6856311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6856311&pc=OJE | 08/10/2017 | 09/21/2017 |  | 09/11/2017 |  |  | 08/11/2017 | 09/11/2017 |  |  | 05/31/2019 |  | 09/11/2017 |  | PUMP, BLOOD,CARDIOPULMONARY BYPASS NON-ROLLER TYPE |  |  | IMPELLA RP |  |  | OJE |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1289242 |  | 0046-0019 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Malfunction | Yes | No | Device Issue | Patient-Device Incompatibility | [5 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | On the (b)(6), a (b)(6) year old male was admitted with back and shoulder pain. the patient was taken to the cardiac catheterization lab and a ventricular septal defect (vsd) was observed. the vsd was thought to be caused by a myocardial infarction. an impella cp was placed for support. four days later, on the (b)(6), due to right ventricular failure, and impella rp was placed for right heart support. the patient returned to the icu and was monitored for right heart filling and given fluids. the patient's condition decompensated on the night of the (b)(6) and the team drew labs to confirm a diagnosis of hemolysis. after twenty two hours of impella rp support the pump, the rp was explanted and va- extracorporeal membrane oxygenation (ecmo) was placed due to concerns of hemolysis. on the (b)(6), the team cared for the hemolysis and proceeded with support. the patient's support continued with the va- ecmo and impella cp for ten more days. during these days of support, the patient experienced multiorgan failure, was placed on dialysis, received blood product replacement, and eventually care was withdrawn. the patient expired on the (b)(6) 2017. | [5 empty] | After the initial medwatch report was filed and submitted on september 11, 2017 the product and data logs were returned for analysis. the impella rp pump was inspected and run through abiomed in-house hemolysis testing. the pump passed hemolysis testing. the root cause of the hemolysis could not be determined. the data logs were analyzed and the logs show that the pump performed without issue. no corrective action is recommended as the failure mode was unable to be determined. the failure mode will be trended and monitored. internal reference (b)(4). | To date the impella rp pump product and data logs have not been returned for investigative analysis. the return process from the field has begun. pending return of the product and data logs, the only review possible has been that of the manufacturing records. there are no other complaints made against this lot of impella rp pumps for the failure mode of hemolysis. the root cause of the hemolysis can not be determined at this time. the manufacturer will continue to investigate all reasonable obtainable source information, and will file a supplemental medwatch report if they become available. (b)(4).
38185 | 1220648-2017-00069 | 6851558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6851558&pc=OZD | 02/23/2017 | 02/23/2017 |  | 09/08/2017 |  |  | 02/23/2017 | 09/08/2017 |  |  | 12/31/2018 |  | 09/08/2017 |  | TEMPORARY CARDIAC BLOOD SUPPORT PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1270545 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 89 YR | [5 empty] | Injury | Yes | No | Human-Device Interface Problem | [6 empty] | Hypovolemia | Blood Loss | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On the (b)(6) a (b)(6) year old female was admitted after suffering angina while on a treadmill. she coded in the field on route to the hospital. upon admission she was hypotensive and bradycardic. the physician placed an impella cp for support while cardiothoracic surgery was consulted for a totally occluded right coronary artery (rca) and left anterior descending artery (lad). on the evening of the (b)(6) the patient had a noted bleed from the access site in the left femoral artery (lfa). the lfa was compressed with a manual hold and the impella cp was propped up at the access site to relieve the profuse bleeding. the patient developed a hematoma at the site due to the bleed. the patient coded due to hypovolemia and had a vascular surgery consult. the team returned to the cardiac cath lab to re-evaluate her coronary artery disease. at this time a diagnosis of cardiogenic shock was made and the patient had two units of blood product infused. the patient was stabilized in the following days, and the impella cp was explanted on the (b)(6). the groin access site bleed had resolved and she was noted to be doing well. | [5 empty] | The adverse event was discovered though internal case review on september 7, 2017. due to an oversight, the event had not been reported previously. the event had occurred in (b)(6) of 2017. due to the delay in reporting, there was no product to evaluate nor, any data logs available for analysis. the root cause of the access site bleeding was unable to be determined as the product was not returned from the field. no other complaints of adverse events have been made against this lot of impella pump or introducer. no corrective action is recommended as the root cause was unable to be determined as the product was not returned. (b)(4).
38186 | 1220648-2017-00066 | 6853340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6853340&pc=OZD | 08/22/2017 | 08/22/2017 |  | 09/08/2017 |  |  | 08/22/2017 | 09/08/2017 |  |  | 04/30/2019 |  | 09/08/2017 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1295302 |  | 005042 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 86 YR |  |  | 82 |  |  | Injury | Yes | No | Human-Device Interface Problem | [6 empty] | Low Blood Pressure/ Hypotension | Perforation | Blood Loss | [10 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | On (b)(6) 2017 a (b)(6) female was admitted for a high risk coronary artery angioplasty. for support, the team was to place an impella 2.5 via the left femoral artery. the physician placed the impella via the abiomed provided .018 guidewire. the wire was shaped into a pigtail-like configuration and advanced to the left ventricle; a ventricle noted to be small. when the physician started the impella 2.5 within the left ventricle the patient became hypotensive and had lower than expected flows. the team then performed an echo and diagnosed a perforation of the ventricle. the patient was taken to the operating department for a patch of the perforation. the perforation was seen to measure just less than 1 centimeter. the physician felt that the .018 guidewire had prolapsed and caused the perforation. the patient did not receive any other forms of cardiac support and did subsequently expire. | [5 empty] | A user initiated medwatch (b)(4) was received on october 23, 2017 with more patient details than when manufacturer originally submitted. new details do not change original manufacturer medwatch conclusions written in 1220648-2017-00066. internal reference (b)(4). | The facility did not return the product or console data logs for analysis. the adverse event of a left ventricle perforation was confirmed soley by clinical account. the physician stated he thought the perforation occurred due to the .018 guidewire prolapse. the abiomed representative was unable to return the wire for analysis of a prolapse or kink. there were no other complaints of an adverse event related to the lot of impella 2.5 pumps. the failure mode of left ventricle perforation is of a very low risk index due to very low occurrence. the root cause of the perforation could not be determined due to product not being returned for analysis. internal reference (b)(4).
38187 | 1220648-2017-00064 | 6841943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6841943&pc=OZD | 08/15/2017 | 08/15/2017 |  | 09/05/2017 |  |  | 08/15/2017 | 09/05/2017 |  |  | 12/31/2018 |  | 09/05/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1273798 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 53 YR | [5 empty] | Malfunction | No | Yes | Kinked | Device-Device Incompatibility | Human-Device Interface Problem | Material Deformation |  |  |  | No Code Available | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | COOK SHEATH 14FR X 30CM | PERCLOSE | [21 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) male was admitted with multivessel coronary artery disease, diabetes, hypertension, and mild mitral valve disease. the plan was to place an impella cp for support during multivessel revascularization in the cardiac catheterization lab. while in the cath lab, a swan was placed to monitor right heart pressures, and the cp was placed via a cook 14fr x 30cm (abiomed approved) sheath in the left femoral artery. the left femoral was known to have tortuosity, and the cook sheath would provide adequate support. to assist with the groin access site closure the team placed 2 perclose in the "pre-closing" method to the left femoral artery. during the catheterization procedure, the patient's pressures were labile and there was intermittent hypertension. therefore, the cp pump weaning process was begun in anticipation of explant. during pump explant, the cp was unable to be eased out of the cook sheath's valve. the physician then applied force to pull the pump back. during this pull through the sheath, the cannula of the pump was damaged with the internal nitinol coil observed outside of the sheath and the remaining portion of the pump caught within the sheath. the team then used fluoroscopy to confirm the remaining portion of the pump was intact and within the cook sheath. the sheath was clamped and cut so that the remaining portion of the pump could be removed. as the physician tried to maintain vascular access, the remaining portion of the cook sheath migrated into the femoral artery. a vascular surgeon was able to emergently remove the remaining portion of the cook sheath. the sheath was snared out and removed via the contralateral leg. the team then completed the perclose procedure of the left femoral and the patient was sent to the icu for a successful recovery. | [5 empty] | The cp pump was returned for investigation. the cannula of the pump was torn and still within the a portion of the cook sheath. the sheath was cut in two placed and had kinks throughout its length. there were numerous kinks observed. the kinks distal to where the sheath was cut by the physician were likely caused during the clamping and subsequent snaring of the sheath by the vascular surgeon. the damage to the cannula was caused when the physician applied excessive force to the pump when it became caught within the sheath during explantation. the pump was unable to pass through the sheath due to the observed kinks in the sheath. these kinks were most likely caused by the vascular tortuousity. this known tortuosity had prompted the team to utilize the cook sheath. the data logs did not pertain to the investigation, and therefore were not analyzed. the root cause of the damge to the cannula was the excessive force applied. failures of this type will be monitored and trended. (b)(4).
38188 | 1220648-2017-00063 | 6842452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6842452&pc=OZD | 08/15/2017 | 08/15/2017 |  | 09/05/2017 |  |  | 08/15/2017 | 09/05/2017 |  |  | 03/31/2019 |  | 09/05/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP,, |  |  | IMPELLA CP PUMP SET, US |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1282366 |  | 0048-0020 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Missing Value Reason | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | On (b)(6) 2017 a (b)(6) male patient was undergoing a diagnostic catheterization in the facility's cath lab when a wire dissected the left anterior descending artery and the circumflex artery. the patient was intubated and an impella cp was placed in the patient's left femoral artery. the team repaired the patient's dissected arteries. the clinicians felt it best for the patient to leave the pump in the patient to allow the heart to rest. the physician removed the 14fr peel-away sheath, and advanced the repositioning sheath into the patient. the physician verified placement, locked the tohey valve, and lastly sutured the catheter in place. lastly he placed a swan-ganz in the patient's right femoral vein. upon completion, the doctor noticed a "large" hematoma at the impella's left femoral site. the team applied pressure, and re-accessed the right femoral artery to perform an arteriogram to assess the artery and to locate the source of the bleed. no source of bleeding was identified, so team believed bleeding to be coming from impella sheath. the doctor cut the sutures and attempted to push the impella's repositioning sheath further into the arteriotomy (the last visible ridge on the suture pad was the second one). this approach did not provide hemostasis, and the team had to continue holding manual pressure on the left femoral artery. vascular surgery was consulted and the surgeon repaired the artery by performing a femoral artery cut down, and then placed a purse-string suture with a rummel tourniquet. the patient was then transferred to the icu for further management. on (b)(6) 2017 the patient was weaned from the impella cp, and the pump was explanted after 30 hours of successful hemodynamic support. vascular surgery assisted with the explant using a purse-string technique over the left femoral artery site. upon removal the patient was reported in stable condition. | [5 empty] | The manufacturer will continue to investigate all reasonable and obtainable sources of information. (b)(4). | The impella cp was returned for evaluation. the automatic impella controller's data logs were not returned for analysis. a visual examination of the returned pump was performed, which determined that there were no obvious defects or signs of misuse. a review of the clinical narrative suggested that the hematoma originated at the impella inserion site. the repositioning unit was also evaluated; there were no obvious defects or signs of misuse to the device relevant to the failure. the sheath's outside diameter was evaluated and found it to be within specification. the root cause was of the reported issue was unable to be determined. a review of the device history record showed no other devices from this lot with this or any other failure mode. in conclusion, the root cause of the hematoma was unable to be determined. no corrective action is recommended because the root cause was unable to be determined. the failure will continue to be monitored and trended. (b)(4).
38189 | 1220648-2017-00062 | 6842939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6842939&pc=OZD | 08/09/2017 | 08/09/2017 |  | 09/05/2017 |  |  | 08/09/2017 | 09/05/2017 |  |  | 02/28/2019 |  | 09/05/2017 |  | IMPELLA 5.0 AIC PUMP SET |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1283643 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | Device Operates Differently Than Expected | [5 empty] | Hemolysis | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | (b)(6) year old female was admitted with cardiomyopathy and an ef of 10% for an lvad workup. an intraaortic balloon pump (iabp) was placed initially for support, with a plan to remove the iabp when the anticoagulation was regulated and the ptt was lowered. on (b)(6) 2017 the impella 5.0 was placed via the right axillary artery surgical cutdown and medications were titrated down accordingly. the patient postoperatively was moving excessively in the bed and needed the team to evacuate a hematoma in the right axillary. while in the or the urine appeared to be tinged, indicative of hemolysis. the pump was repositioned and the urine became clearer. three units of blood replacement products (prbc's) were infused to the patient. the following day, (b)(6), the urine was concentrated in appearance and color, which was a presumed sign of renal failure and/or hemolysis, and the pump was again repositioned. the pump was visualized as being very close to the patient's mitral valve. early in the morning, the iabp was explanted and the patient was extubated. later in the evening, the urine was darker in color, the bilirubin and ldh levels were rising, and so the team chose to remove the 5.0 pump due to concerns of hemolysis. | [5 empty] | The device and data logs were both returned for investigation. no echo or fluoroscopy imaging of the device, and any repositioning of it were returned for analysis. the data logs did not reveal any positioning alarms during the time of the reported hemolysis. the pump was stated to be very close to the mitral valve, but no imaging of this was shared for analysis. furthermore, no suction events or low flows were seen in the data logs at the time of the reported hemolysis. the 5.0 pump itself was analyzed and run through hemolysis testing in-house at abiomed. no pump damage was noted that could have caused or contributed to hemolysis. the pump was run through hemolysis testing and passed, according to abiomed specific values. while reviewing the lot of pumps, no other pumps have been cited in complaints for hemolysis, or any other adverse events. the root cause of the hemolysis could not be determined. no corrective action is recommended as the pump tested within specification for hemolysis performance testing. (b)(4).
38190 | 1220648-2017-00065 | 6814300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6814300&pc=OZD | 07/07/2017 | 07/26/2017 |  | 08/23/2017 |  |  | 07/26/2017 | 08/23/2017 |  |  | 11/30/2018 |  | 08/23/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET, EU |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1262469 |  | 0048-0002 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | Patient Problem/Medical Problem | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | On (b)(6) 2017 an impella cp was placed in a (b)(6) female patient. eleven hours later the clinicians reported purge management issues following the repositioning of the pump. an x-ray and an echo were performed that revealed that the impella cannula was completely wrapped around the mitral cords causing suction alarms. due to the more than 8 hours of the cannula being severely malpositioned poor pump flow continued. as a result the patient's kidneys suffered from hemolysis, and not enough perfusion causing very low urine output and dark red urine. the patient's haptoglobin decreased and ldh increased. the clinicians tried to reposition the pump at the bedside under echo for almost 2 hours. the patient's respiratory status worsened and the decision was made to go back to cath lab for better visual access. in the cath lab the impella was pulled back and the clinicians could see the pigtail unwrapping and extending; however, the pigtail was wrapped around the mitral cords. repositioning under echo finally resulted in the pump's pigtail being freed from the mitral cords. once the pigtail was free, and the cannula was successfully repositioned, good pump flow was restored, and the patient's respiratory status quickly improved. the patient was transferred to a larger icu for better observation and dialysis was started. blood products were transfused to the patient, although it was uncertain how much of the administration of blood products was a result of the hemolysis, as the patient also had a bleed from a gastric ulcer. following this event and on the same day training was given to the medical and nursing team, as they were not familiar with the device. the patient was supported by the impella cp for 5.88 days, without further issue. the device was removed from the patient on (b)(6) 2017, as it was reported that the impella cp's hemodynamic support was insufficient for the patient's needs. the clinicians then started an ecmo to improve renal perfusion and give better circulatory support. the patient had developed arrhythmias such as ventricular fibrillation, and was defibrillated several times. the patient continued to worsen and expired a few days later. the clinicians reported that the patient outcome was not attributed to any issue that occurred with the impella cp. | [5 empty] | The impella cp was not returned for evaluation, as it was discarded by the user subsequent to the event. the automatic impella controller's console logs were returned for analysis. the analysis of the logs determined the following: the long time data log from the field revealed that there were 159 impella position wrong alarms. in addition, during the first 11hours of the case 31 suction alarms were recorded were also recorded. positioning issues like this are known to contribute to hemolysis. the pump remained poorly positioned for 8 hrs. an x-ray and an echo were performed that revealed that the impella was sitting too deep inside the ventricle and was caught in the mitral chords. also, the returned data logs returned contained numerous positioning and suction alarms, including an alarm for the outlet that was under the valve. this positioning is also known to contribute to hemolysis. images from the echo and fluoroscopy were returned. the fluoroscopic image clearly showed the pump was sitting too deep in the ventricle as indicated by the image bent of the pigtail. a review of the device history record showed no other devices from this lot relevant to the failure mode. in conclusion, the root cause of the hemolysis was most likely caused by improper positioning of the pump. no corrective action is recommended at this time because the failure mode was determined to have a low risk index based on very low occurrence and moderate severity. failures of this type will continue to be monitored and trended. internal reference: (b)(4). | The impella cp has not yet been returned for evaluation. without an evaluation of the device used in this case, the root cause of the event is unable to be determined. no corrective action is recommended as the root cause was unable to be determined. the manufacturer will continue to investigate all reasonable and obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. (b)(4).
38191 | 1220648-2017-00058 | 6802491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6802491&pc=OZD | 07/18/2017 | 07/18/2017 |  | 08/17/2017 |  |  | 07/18/2017 | 08/17/2017 |  |  | 03/31/2019 |  | 08/17/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP, |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, US |  |  | ABIOMED, INC |  |  | US |  | 128748 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 71 YR |  |  | 100 |  |  | Injury | Yes | Yes | Difficult To Position | Material Separation | [5 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On (b)(6) "20107" a (b)(6) year old male patient presented in the facility's emergency room with a st-elevation myocardial infarction (stemi). a percutaneous coronary intervention (pci) procedure was started and during the case the patient became unstable, at which time the decision was made to place and impella cp. a14fr oscor introducer sheath was then inserted into the right femoral artery (rfa) in preparation for the placement of the impella cp. while gaining access via the 14fr oscor introducer, a guidewire was introduced several times, but would not advance. fluoroscopy showed the guidewire to be outside of the sheath. the physician reported that the introducer sheath was difficult to place. a clinician reported that the insertion dilator "backed out" of the sheath while it was being advanced. the introducer sheath was then removed, and was found to have split in three different locations. the patient remained in stable condition. the patients blood pressure slowly began to drop again, so the decision was made to place the impella cp in the lfa. the pump was then placed without issue. a large softball size hematoma developed at the rfa access site. the hematoma was reported to be dry and soft. manual pressure was held for approximately one hour due to the administration of 10,000 units of heparin, aggrastat, and reopro. the hematoma was successfully expressed out. there was slight ooze at the impella groin site, but this was minimal and resolved easily with the administration of 10-200mm hemoglobin (hg) on the femstop. there was no hg drop, no blood loss needing any intervention with blood products. the patient was supported with the impella cp for 28.57 hours. the patient expired on (b)(6) 2017. the physician reported that she did not feel that the sheath problem contributed to the patients overall critical status or outcome. | [5 empty] | Neither the impella cp nor the oscor 14 french introducer (lot number 1275384) was returned to the manufacturer for evaluation. the impella cp was discarded and the oscor 14 french introducer is being held "for the foreseeable future" by the facility's risk management department. the device history record for the oscor 14 french introducer was reviewed. all introducers, selected from lot number 1275384 per aql sampling, passed incoming inspection. the automatic impella console data logs were not reviewed for this case, as the failure occurred during insertion of the device. based on additional information obtained from the physician, the introducer dilator was possibly not advancing with the sheath. advancing the sheath over the guidewire, without the dilator through its tip, may have caused the guidewire to cause the sheath to split from the tip along its score lines; however, this cannot be definitively determined as the 14 french oscor introducer sheath was not returned for evaluation. the failure mode was determined to be medium risk due to moderate severity and low occurrence. in conclusion, the root cause of the tip split was unable to be determined as the introducer was not returned for evaluation by the customer. no corrective action is recommended as the root cause was unable to be determined. internal reference: (b)(4).
38192 | 1220648-2017-00059 | 6799197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6799197&pc=OZD | 07/20/2017 | 07/20/2017 |  | 08/16/2017 |  |  | 07/20/2017 | 08/16/2017 |  |  | 01/31/2019 |  | 08/16/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET, US |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1279798 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 78 MO | [5 empty] | Injury | Yes | No | Improper or Incorrect Procedure or Method | Adverse Event Without Identified Device or Use Problem | [5 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On (b)(6) 2017 a (b)(6) year old female patient presented at the facility with an st elevation myocardial infarction (stemi). an impella cp was placed via the right femoral artery on this same date. the patient arrived in the cath lab where the removal of the peel-away sheath was completed and the repositioning sheath was advanced and sutured in place. the patient was found to have severe multi-vessel coronary artery disease. a drug eluding stent was placed in the left anterior descending artery (lad). the physician repositioned the pump under echo and advanced device 1cm. the impella was found to measure at approximately 3.5cm from the av. on (b)(6) 2017 it was observed that the device was anchored down with occlusive dressings and a knee immobilizer causing the catheter to be positioned non-coaxial to the artery. this resulted in bleeding at site. the catheter was then positioned so to allow the catheter and sheath to be repositioned with a more natural curve to the catheter when secured in place. the bleeding then slowed down but did not completely stop. the vascular surgeon was consulted, and sutures were placed around the repositioning sheath. bleeding improved but the patient's hemoglobin fell from 11 to 7. in response the physician ordered the heparin and integeilin be discontinued and the purge solution changed to d5 only. patient bleeding subsided. a total of 3 units of replacement packed red blood cells (prbc) were administered to the patient. on (b)(6) 2017 the patient was weaned, and the device was successfully removed after 46.50 hours of patient support. surgical cut-down with closure of the artery was performed during explant of impella. the patient was reported to be in stable condition. | [5 empty] | The impella cp was returned intact for evaluation. in preparation of the evaluation of the pump the clinical information supplied by the facility was reviewed. a visual evaluation of the pump was performed; this evaluation was unable to identify any defects on the pump. the clinical information supplied by the staff reported that a combination of manual adjustment were performed on the repositioning sheath, and that in addition, the heparin dosage infused in the purge system had been lowered during patient support. this lowering of the heparin dosage infused in the purge system, and the removal of systemic heparin resulted in the patient bleeding resolving. a device history review was performed. there were no issues identified during the manufacture of this device, and there were no other complaints reported for other pumps in this same lot. in conclusion, the patient bleeding decreased when the repositioning sheath was manually adjusted, and the bleeding was completely stopped with the adjustment of the heparin dosage. the root cause of the patient bleeding has been determined to be user related, as it was caused by an incorrect dosage of heparin being administered to the patient. therefore, the patient bleeding was as a result of improper anticoagulation therapy. the instructions for use (ifu) of this device advises the user of the following: do not add any alternative anticoagulant (such as a direct thrombin inhibitor) to the purge fluid. the impella® catheter has not been tested with any alternative anticoagulants in the purge solution. dextrose solution (typically 5% dextrose in water with 50 iu/ml of heparin) is used as the purge fluid through the impella® catheter. a 500 cc bag of dextrose solution for purge solution (20% recommended; 5% to 20% acceptable) with 50 iu heparin/ml. as a corrective action retraining of the facility's staff was recommended, and training was performed on july 21, 2017, and is ongoing. (b)(4).
38193 | 1220648-2017-00060 | 6794746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6794746&pc=OZD | 07/16/2017 | 07/16/2017 |  | 08/15/2017 |  |  | 07/16/2017 | 08/15/2017 |  |  | 05/31/2019 |  | 08/15/2017 |  | IMPELLA CP PUMP SET |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1289906 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 51 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | Compatibility Problem | [5 empty] | Ischemia | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On (b)(6) 2017 a (b)(6) diabetic female presented to the hospital in cardiogenic shock. the cardiac catheterization revealed an occluded left anterior descending artery (lad) and pulseless electrical activity (pea). to treat the patient a hypothermia protocol was begun as well as placement of an intra-aortic balloon pump (iabp). on (b)(6) the patient decompensated and care was escalated, with the removal of the iabp and placement of an impella cp, via the right femoral artery. an external fem-fem bypass was performed due to poor perfusion in the right leg two days later, on (b)(6). during the support with the cp the patient was diagnosed with heparin induced thrombocytopenia (hit). the hit diagnosis caused the patient's anticoagulation regimen to change. heparin was removed from the purge solution to the cp as well as systemically, and the direct thrombin inhibitor, angiomax, was begun systemically on (b)(6). the purge flow was seen to be decreasing, which is indicative of clot. on (b)(6) the cp was explanted due to the concern of possible clot formation, and a second cp was placed via the right femoral artery. the hospital team was re-educated by abiomed personnel of the value of a purge fluid anticoagulation protocol. the patient was taken to surgery, on (b)(6), for an endarterectomy of the right leg due to limb ischemia. the vascular surgeon was able to restore flow to the limb. pedal pulses returned and there was no lasting harm or injury. the physician was unaware of why the thrombus had developed in the right leg. another impella cp was placed in the contralateral leg, via the left femoral artery, on (b)(6). the patient was decompensating and was paralyzed medically. in addition, she was being treated for pulmonary edema and her urine production was decreasing. the team was infusing four different pressor drugs to elevate her arterial blood pressure and cardiac output. on (b)(6) the patient coded and did expire, as the family chose to withdraw care, as she was not a candidate for escalation to an lvad. in total the patient had support with an impella cp for 13 days. | [5 empty] | The impella pumps were all discarded at the user facility. no product was returned for analysis. the only review possible was of the clinical report. the console data logs were not relevant to this failure mode, of limb ischemia. the limb ischemia was caused by a clot developing in the right femoral artery, which was used as vascular access for the first two impella cp pumps. the flow was so poor in the patient's right leg after the placement of the first cp, that an external fem-fem bypass was performed to maintain perfusion. the clot may have developed due to the restricted flow, but the root cause can not be determined as no product was available for evaluation. the patient did have anticoagulation issues due to the hit diagnosis. anticoagulation and the heparin concentration are noted in the ifu as the following: "the impella catheter is designed to be operated with a purge solution that contains heparin. operation of the system without heparin in the purge solution has not been tested. in the event that a patient is intolerant to heparin, due to heparin-induced thrombocytopenia (hit) or bleeding, physicians should use their clinical judgment to assess the risks versus benefits of operating the impella system without heparin. if it is in the best interest of the patient to operate the system without heparin, the dextrose solution is still required, and physicians should consider systemic delivery of an alternative anticoagulant. the impella catheter has not been tested with any alternative anticoagulants in the purge solution. use of alternative anticoagulants may reduce the longevity or performance of the impella catheter." the physician's decision to operate without heparin, and with angiomax instead is at his discretion. the abiomed field personnel did have conversations regarding this fact. no corrective action is recommended at this time as the risk index in this failure mode is low due to moderate severity and very low occurrence. the root cause of the leg ischemia was the development of clot in the right femoral artery. the root cause of the clot was unable to be determined. (b)(4).
38194 | 1035166-2017-00067 | 6761331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6761331&pc=DYB | 01/06/2017 | 02/09/2018 |  | 08/02/2017 |  |  | 06/03/2017 | 08/02/2017 |  |  | 11/01/2019 |  | 08/02/2017 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC |  |  | US |  | C1-12829 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN | HEALTH PR |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Unintended Movement | [5 empty] | Angina | Calcium Deposits/Calcification | Blood Loss | [10 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | The complainant reported that an (b)(6) female patient with 3-vessel disease was scheduled to have a subacute pci to be performed on (b)(6) 2017 with the aid of an impella cp. good local arterial pulsation was found in both groins. a standard 6f femoral sheath was inserted into the right femoral artery and iliacography showed both iliac arteries as good caliber. a standard 6fr femoral sheath was inserted into the left femoral artery. 5000 iu of heparin was administered intra-arterially. a long standard j-wire was advanced to the aortic arch. the sheath was then removed from the left groin and the impella cp arterial dilators were used with no resistance. the impella peel-away introducer was inserted into the artery. the wire moved freely, but following the removal of the introducer dilator there was no arterial flow in the sheath. the patient immediately deteriorated into shock due to severe bleeding from the left femoral artery, and was quickly intubated and ventilated. the complainant reported that during the insertion of the introducer it accidently peeled and damaged the vessel wall, requiring surgical repair. during the surgical procedure the patient was administered 6 units of blood. it was estimated that the patient lost a total of 2000 ml of blood. the punctured profunda artery and a traumatic lesion in the superficial femoral artery required repair. both lesions successfully surgically repaired. additional information was obtained from the complainant, which reported that the patient had very heavily calcified and tortuous vessels and the profunda artery was situated very deep and to the lateral. due to this some degree of angulation of the sheath occurred during entry into the artery. thus, the introducer dilator followed the bending and entered the artery, while the peel-away portion of the sheath was peeled off on the edge of the heavily calcified artery; consequently, the peel-away sheath ended up on the outside of the artery. this finding was confirmed visually during the surgical repair. the pci was aborted, and was scheduled for the following week. the impella cp supported the patient without issue, and the patient was hemodynamically doing fine. the device successfully supported the patient for approximately 25 days. | [5 empty] | One 14f abiomed introducer sheath with dilator was returned from the customer. there were no other accessories. blood was found on and inside the sheath and dilator. the introducer and dilator were both returned assembled together. the dilator provided has a bend of approximately 13 degrees. the bent section starts at approximately 45mm from the distal tip. additionally, the o.d. and the i.d at the distal tip is distorted from the typical round dimension. other than aforementioned issues the dilator passes visual inspection. the introducer provided passes visual inspection except for split at the tip that extents 9.8cm down the sheath. additionally there is an indentation 2cm from distal tip 1cm wide. this area most likely buckled from the angulation force of the dilator. dilator dimensions were measured all within spec. sheath dimensions were measured all within spec. upon evaluation of the returned product, it was found that the dilator most probably exerted extreme diagonal force, (determined by the bent tip on dilator) to the introducer and the indentation of the sheath. the diagonal force translated stresses to the tip of the introducer which caused it to split along one of the score lines. in addition the notation from the (senior interventional cardiologist) states that "artery was heavily calcified". the calcified artery may have helped cause the tip of the introducer to tear because the abrasive surface, further illustrated by the distorted dilator tip. the angular force applied during the procedure likely exceeded the design of the introducer. based on the investigation, a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. in addition, no new failure mode was identified and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type. per procedure (abiomed introducer sheath in-process and final inspection) 9.2.2 visual inspection: using 10x magnification, verify the tip is round, no flash protruding greater than 0.4 mm (0.016") and no flash with width (around circumference) greater than 0.7 mm (0.028"), no splitting, cracks or other damages. slight discoloration from tipping process is acceptable. verify proper tip shape according to drawing. insert a tipped dilator of appropriate size through the tipped sheath and check if dilator inserts without excessive force or obstruction. with the naked eye at a distance of 12" to 18", inspect where the sheath and dilator meet and check the gap is minimal and a smooth transition is present. per instructions for use (ifu): no representation or warranty is made that an oscor inc. product will not fail. oscor inc. disclaims responsibility for any medical complications, including death, resulting from the use of its products. except as expressly provided by this limited warranty, oscor inc. is not responsible for any direct, incidental, or consequential damages based on any defect, failure, or malfunction of its products whether the claim is based on warranty, contract, tort, or otherwise. some states do not allow the exclusion or limitation of consequential damages and so the above limitation or exclusion may not apply to you. | Conclusion not yet available, evaluation in process. this initial mdr is being submitted to meet our requirements of reporting. a follow-up will be submitted as soon as the investigation is complete.
38195 | 1220648-2017-00056 | 6749977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6749977&pc=OZD | 07/02/2017 | 07/03/2017 |  | 07/28/2017 |  |  | 07/03/2017 | 07/28/2017 |  |  | 03/31/2019 |  | 07/28/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1285687 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | No | Malposition of Device | [6 empty] | Cardiac Tamponade | Cardiac Perforation | [11 empty] | Hospitalization | Life Threatening | Required Intervention | [26 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella cp was implanted into a (b)(6) male patient who had suffered an acute myocardial infarction and had an ejection fraction of 10%. heart catheterization revealed proximal left anterior descending and chronic total occluded right coronary artery disease. a stent was placed, and then the impella cp was implanted through the left femoral artery. shortly after arriving to icu the patient became tachycardic with heart rate dropping to 40 beats per minute. an echocardiogram revealed cardiac tamponade and a pericardiocentesis was performed . the physician believes that while loading the impella cp onto the 0.018" guide wire, the tip of the wire may have caused a perforation at the apex of the left ventricle. a pericardiocentesis patch and two coronary bypass grafts were placed . the impella cp was explanted in the operating room, and the patient was reported to be in stable condition at that time. | [5 empty] | The products involved in the event were discarded by the user after the event. (b)(4).
38196 | 1220648-2017-00054 | 6748356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6748356&pc=OZD | 06/28/2017 | 06/29/2017 |  | 07/27/2017 |  |  | 06/29/2017 | 07/27/2017 |  |  | 01/31/2019 |  | 07/27/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1280049 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Death | Yes | No | Patient-Device Incompatibility | Activation, Positioning or Separation Problem | [5 empty] | Hematoma | Blood Loss | [11 empty] | Death | Hospitalization | Life Threatening | [26 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6)obese male patient weighing (b)(6) presented to the cardiac cath lab as an out-patient for a diagnostic catheterization, which revealed multivessel disease in the left main, left anterior descending and left circumflex arteries. an intra-aortic balloon pump was inserted for support, however the patient lost pulse. the decision was made to escalate care from a balloon pump to impella. the balloon pump was removed and exchanged for an impella cp. the impella cp was implanted through the same left femoral artery access site. the occluded coronary vessels were stented while the patient was supported by impella. after the intervention, the 14 fr cook peel-away introducer was removed,and the clinical team attempted to insert the repositioning sheath. however the repositioning sheath was unable to be positioned due to the patient's large size. extensive bleeding at the insertion site occurred as well as the formation of a hematoma. the exact amount of blood loss is unknown. manual pressure was held to control the bleeding. the patient's condition did not improve and he became hypertensive. the patient was unable to achieve return of spontaneous circulation (rosc), and expired in the cardiac cath lab. | [5 empty] | According to the impella cp instructions for use (ifu), " when securing the repositioning sheath, vascular closure may be difficult in obese patients with extensive adipose tissue." the impella cp and aic console data were returned for evaluation, however the console data is not relevant to this issue. the pump was returned with a tear in the wound closure of the repositioning sheath, and a portion of the wound closure was cracked at the tip. additionally, a portion of the wound closure had peeled away from the catheter and contained rough edges. the root cause of the reported hematoma and bleeding from the insertion site is a combination of the patient's size and the damaged repositioning unit. the damage to the device is likely due to patient condition. a review of the device history record revealed there were no other issues reported for any other device with this lot number. the manufacturer will continue to investigate all reasonable and obtainable sources of information and will provide results and conclusions if additional information is received. (b)(4).
38197 | 1220648-2017-00057 | 6743411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6743411&pc=OZD | 06/30/2017 | 06/30/2017 |  | 07/26/2017 |  |  | 06/30/2017 | 07/26/2017 |  |  | 03/31/2019 |  | 07/26/2017 |  | IMPELLA 5.0 |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED INC. |  |  | US |  | 1285998 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 51 YR | [5 empty] | Injury | No | No | Kinked | Patient-Device Incompatibility | Difficult to Advance |  |  |  |  | Ischemia | Thrombosis | Blood Loss | [10 empty] | Hospitalization | Required Intervention |  |  |  |  | 10MM VASCULAR GRAFT | VA- ECMO | [21 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) male was taken to the cardiac cath lab for a stemi. during the cath procedure, the patient arrested and cpr measures were taken for approximately 1 hour. the team placed a va ecmo and two days later the impella 5.0 for additional support. the 5.0 was being positioned through a 10 mm graft in the right axillary artery when there was difficulty with insertion. the right axillary artery was presumed to be smaller than 7 mm, though exact measurements were not taken and shared with abiomed. after an hour of attempted placement, the team decided to place the 5.0 instead via the femoral artery, but were unable to do so due to the small artery diameter. the impella placement was aborted and the patient remained supported with the va ecmo. the right arm did experience ischemia as well as bleeding from the anastomosis site of the 10 mm graft. the cardiologist believed that he felt clot in the arm. the anticoagulation maintenance of the patient was difficult throughout the icu stay. the patient had to have approximately 15,000 units of heparin to achieve an act of 250. the impella 5.0 was never successfully placed. the arm ischemia was treated with a covered stent. the patient did regain flow in the arm and was supported with an impella cp inserted into the left arm, via the left axillary artery. the team attempted to wean off the ecmo, however this was not possible due to patient condition. there was an internal bleed of unknown location that required 8 units of blood product replacement. on the second day of support with the impella cp, the family chose to withdraw care, and the patient expired. | [5 empty] | The impella 5.0 pump was returned for investigation in good condition. the data logs did not relate to this failure mode. there were no reworks or other complaints for the lot of pumps.. upon inspection, a kink was found on the 5.0 pump located at the 15 cm mark and the cannula was damaged near the pump housing. no other defects were found on the rest of the pump. when the outer diameter of the pump was measured, the size was below the upper limit of specification. the size of the pump did not contribute to the pump kinking. the root cause of the kinking during insertion was the excessive force used to advance through the anastamosis. no corrective action is recommended for this root cause as it is low risk, based upon very low occurrence and moderate severity. these types of kink failures will continue to be monitored and trended. in regards to the limb ischemia and question of thrombosis, there was difficulty in the icu management of the act (activated clotting time). according to the facility, and excessive amount of heparin was given; approximately 15,000 units to reach the act of 250 for insertion. per the abiomed instructions for use: "administer heparin and achieve act of at least 250 seconds." "maintaining act after insertion of the catheter (and until explant), act should be maintained at 160 to 180 seconds." the physician stated that he felt there was clot in the right arm of the patient, who had a small axillary artery. the root cause of the arm ischemia and/or presence of clot was not ever able to be determined as the failure mode could not be reproduced. no corrective action is recommended because the root cause is unable to be determined. the failure mode is low risk based on very low occurrence and moderate severity. (b)(4).
38198 | 1220648-2017-00061 | 6745298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6745298&pc=OZD | 10/20/2016 | 06/26/2017 |  | 07/26/2017 |  |  | 06/26/2017 | 07/26/2017 |  |  |  |  | 07/26/2017 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA LP2.5 AIC PUMP SET, US |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC |  |  | US |  | UNK |  | 005042 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | HEALTH PROFESSIONAL | [5 empty] | 30 YR | [5 empty] | Injury | Yes | Yes | Appropriate Term/Code Not Available | [6 empty] | Aortic Insufficiency | Cusp Tear | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On june 26, 2017 during the 2017 annual conference of (b)(6) a poster was exhibited that reported an adverse event that had occurred with an impella lp2.5 pump on (b)(6) 2016. this patient support and adverse event with this impella lp2.5 has not been reported to abiomed. continuing attempts to obtain additional device, event and patient information regarding this issue has been unsuccessful. the following is the report that was displayed on the poster and seen at the asio conference: a (b)(6) female patient presented with post-partum cardiomyopathy presented with critical cardiogenic shock after sustaining 2 episodes of cardiac arrest. transesophageal echocardiogram (tee) demonstrated dilated lv with ef of 15-20%. femoral va-ecls was initiated and the impella 2.5 was inserted via the right axillary artery in the icu. the impella required repositioning 3 times under bedside tte. the impella was turned down to p prior to manipulation. post- operative day 9 durable axial flow lvad was implanted. impella was turned off, and explanted. immediately following lvad implantation tee revealed new moderate to severe aortic insufficiency (ai). heart was arrested, and upon direct and upon direct inspection, the non-coronary cusp had sustained a tear from node of arantius to the nadir of the leaflet. left and right leaflets appeared normal. bio-prosthetic avr was performed. the patient recovered and was discharged with lvad 52 days after presentation. | [5 empty] | This adverse event was not reported to abiomed, inc. at the time it occurred. the impella 2.5 pump was not returned for evaluation, as it was discarded subsequent to patient use; consequently, no device analysis could be performed and the root cause of the tissue tear in the non-coronary cusp of the av was unable to be determined. as stated in describe event or problem" numerous and continued attempts are being made to obtain additional device, event and patient information regarding this issue. at this time abiomed has been unable to obtain a definitive date of the event, the serial number and lot number of the device; consequently, the device manufacture and expiration dates were unable to be determined. in addition, the implant and explant dates are unknown. the manufacturer will continue to monitor and investigate all reasonable and obtainable sources of information and will provide results and conclusions in a supplement medwatch report, if and when any additional information is received. no corrective action is recommended at this time because the failure is determined to be low risk. the failure will continue to be monitored and tracked. (b)(4).
38199 | 1220648-2017-00055 | 6726330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6726330&pc=OZD | 10/20/2016 | 06/29/2017 |  | 07/19/2017 |  |  | 06/26/2017 | 07/19/2017 |  |  | 06/30/2018 |  | 07/19/2017 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE, |  |  | IMPELLA LP2.5 AIC PUMP SET, US |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1248774 |  | 005042 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 20 YR | [5 empty] | Injury | Yes | Yes | Appropriate Term/Code Not Available | [6 empty] | Aortic Insufficiency | Cusp Tear | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | On june 26, 2017 during the (b)(6) a poster was exhibited that reported an adverse event that had occurred with an impella lp2.5 pump on (b)(6) 2016. although patient support with the device had been documented, the adverse event had never previously been reported to abiomed. the following is the report that was displayed on the poster and seen at the asio conference: a (b)(6) female presented with critical cardiogenic shock and cardiac arrest after a recent illness from viral cardiomyopathy. a transesophageal echocardiography (tte) demonstrated an ejection fraction (ef) of 20-25% with severe global hypo-kinesis and dilation. a femoral va-ecls was initiated and the impella 2.5 placed via the left femoral artery. the impella required repositioning 3 times with tte guidance. the impella was turned down to p0 prior to manipulation. a pod 15, durable centrifugal flow lvad was implanted and the impella was withdrawn. following lvad initiation, an intraoperative tee revealed new moderate-severe aortic insufficiency (ai) and showed a floating object in the non-coronary cusp of the aortic valve (av). the patient's heart arrested and upon visual inspection, a tissue tear was noted in the non-coronary cusp of av. av repair was performed. tee showed trace to mild ai after the leaflet repair. patient recovered and was discharged with lvad 57 days after presentation. the patient use of this device was reported to abiomed on october 5, 2016, with updates through october 20, 2017. there were no indications from the hospital staff, at anytime during or after patient support, to an abiomed employee of the above detailed issue; consequently, the issue was not reported in our standard case report. the last information received by abiomed (prior to observing the poster at the (b)(6) conference on june 29, 2017) was on october 20, 2016, at which time the following was entry was added to our standard report: 10/20 pm update: spoke with rn- impella explanted today. patient doing ok. patient is on heartware. | [5 empty] | This adverse event was not reported to abiomed, inc. at the time it occurred. the impella 2.5 pump was not returned for evaluation, as it was discarded subsequent to patient use; consequently, no device analysis could be performed and the root cause of the tissue tear in the non-coronary cusp of the av was unable to be determined. the manufacturer will continue to monitor and investigate all reasonable and obtainable sources of information and will provide results and conclusions in a supplement medwatch report, if any additional information is received. no corrective action is recommended at this time because the failure is determined to be low risk. the failure will continue to be monitored and tracked. (b)(4).
38200 | 1220648-2017-00053 | 6721032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6721032&pc=OZD | 06/26/2017 | 06/26/2017 |  | 07/18/2017 |  |  | 06/26/2017 | 07/18/2017 |  |  | 02/28/2019 |  | 07/18/2017 |  | IMPELLA 5.0 |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED INC. |  |  | US |  | 1269191 |  | 002162 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 47 YR | [5 empty] | Injury | Yes | No | Kinked | Difficult to Remove | Patient-Device Incompatibility |  |  |  |  | Perforation | Thrombosis | Blood Loss | [10 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | A (b)(6) female with cardiomyopathy, of an unknown origin, had her care escalated from an impella cp to a 5.0 while the team consulted with the transplant team. after 20 days of successful support, with no alarms and the ability to ambulate, the 5.0 pump began to be weaned. the pump wean began as the team was planning to place a durable vad. while attempting to explant the pump there was difficulty pulling the pump back. the pump was stuck at the junction of the carotid, subclavian, and vertebral arteries. the pump's position and integrity were seen under fluoroscopy. in the fluoro imaging a left ventricular clot was observed. at this time the team attempted intervention in the form of balloon dilations to the artery to ease the pump explant. this did not work and so a sternotomy was performed for an open explant of the pump, at the suggestion and with consultation from abiomed's medical team. the hospital team was not able to explant via the leg due to a limb ischemia from another device in the femoral. the sternotomy was successful for pump explantation. the sternotomy allowed for successful explant of the impella 5.0. when the team observed the vessel it was noted to be soft and non-calcified, though a dissection was seen at the graft just proximal to the vertebral artery. thrombus burden was also noted at the site of the dissection. most likely the thrombus occurred at the location of the innominate bifurcation during the implantation of the pump. the thrombus was removed and the patient recovered without any support devices while awaiting vad. | [5 empty] | The 5.0 pump was returned for analysis. the pump had been cut at the 9fr catheter. the engineer did discover a kink at the proximal end of the cannula. this kink is most likely what prevented the pump from being explanted successfully. the root cause of the kink was unable to be determined. no ancillary products to the pump, such as the introducer or access kit, were returned. upon a review of the manufacturing lot of the pump, no other complaints have been filed nor were there any reworks or mrrs of the lot of 5.0 pumps. the data logs were not evaluated as they would not be applicable to this failure. the clinical context at the time of the kink occurring is not known. no imaging has been shared for evaluation and investigation. there is no mention of the cannula kinking during the case. most likely, the thrombus was a result of a dissection that occurred at the same location during the pump's implant. the pump itself did not contribute to the issue directly. these points were clarified in conversation between the hospital and a senior medical director at abiomed. the conclusions drawn were the following: the thrombus at the innominate bifurcation was preventing a successful explant of the device. the thrombus was present at this location prior to the explant of the pump. the dissection caused this thrombus to form. the dissection likely occurred during the implantation of the pump. no corrective actions will be opened for this event. based on the investigation, the pump was found to have not contributed to the failure directly, therefore no corrective action is recommended at this time. the failure will be monitored and tracked. should further information regarding this event be gathered, a supplemental medwatch will be filed. (b)(4).
38201 | 1035166-2017-00059 | 6722947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6722947&pc=DYB | 06/20/2017 | 03/30/2018 |  | 07/18/2017 |  |  |  | 07/18/2017 |  |  | 02/01/2020 |  | 07/18/2017 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3014 |  |  | OSCOR INC |  |  | US |  | C1-13113 |  | 0052-3014 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Backflow | Adverse Event Without Identified Device or Use Problem | [5 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | Additional information was received from the reporter that includes: on the (b)(6) 2017 a 58 year old obese male was admitted to the hospital and found to have an acute in-stent thrombosis in the coronary artery due to non- compliance with medications. the patient had an intervention (pci) in the cardiac catheterization lab to open up the left anterior descending artery. the pci was successfully completed and the cardiac team made the decision to place an impella cp for support while in recovery in the intensive care unit. of note listed in the instructions for use (ifu) for the cp is a statement regarding "evaluation prior to inserting the impella catheter: observations: obesity". a review of the clinical information revealed the patient to be in excess of 300 pounds. the physician began to prep the patient for the impella insertion by removing the 6fr sheath and introducing the 14fr sheath to the right femoral artery, using a .035 x145 cm wire. per the (abiomed) ifu the wire provided in the access kit is listed as a : ".035 inch stiff access guide wire". the patient was obese and so the access site was assessed. the femoral angiogram revealed the artery to be 9mm in diameter. when the team attempted to place the 14fr sheath into the artery there was an observation made of the "odd angulation" of the insertion and associated bleeding. due to the bleeding the impella insertion was aborted. the 14fr sheath was immediately replaced by a larger 16fr dryseal sheath to occlude the artery and decrease the blood loss. the patient was then transferred to the operating department where there was a vascular repair. the anterior stick and posterior laceration were surgically repaired. the physician did inform the representative that this stick was improper technique, rather than product malfunction. the patient is recovering from the massive acute mi and had a permanent pacemaker placed. the impella cp and 14fr sheath were not returned for evaluation. the cp was never inserted into the patient and the 14fr sheath was discarded subsequent to the event. | The customer reported a (b)(6) male patient had an acute in-stent thrombosis. the patient was brought to the catheter lab, to which his vessel was opened. when oscor sheath was inserted it was immediately noted that sheath was at an "odd," angle and that there was significant bleeding. the case was aborted, and the oscor sheath was removed and replaced with a different sheath. the new sheath was placed and occluded the femoral and iliac vessel. as a result, the patient was taken to the o.r. for vascular repair of vessel. an anterior stick and posterior laceration were noted and repaired. the device is not available for analysis. |  |  |  |  | The 14fr sheath was not returned for evaluation as it was discarded subsequent to the event. consequently, no device analysis was performed. in addition, no data logs were returned as the console and cp were never utilized for support. a review of the clinical information revealed the patient to be in excess of 300 pounds. due to the patient's large body habitus the angulation of access was atypical. the physician stated the combination of the patient's large habitus and his technique caused the bleeding and vascular injury. according to abiomed introducer sheath in-process and final inspection procedure sampling per ansi z 1.4, gen level i, normal aql 0.4, random selection: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded foreign material, or any other damages. the instructions for use (ifu) provide these precautions: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. infusion through the sideport can be done only after all air is removed from the unit. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. though the device was not returned for analysis, extreme angles of introducer insertion are known to facilitate this type of failure. there were no manufacturing rejects or anomalies of this event type recorded in the device history record. a review of complaints against this lot number identified no additional similar events. the device was not returned for analysis. as a result, the allegations against this device cannot be confirmed. a capa is not required as findings did not identify a design, labeling or manufacturing non-conformity. a risk analysis is not required as there was no failure reported against the device. the event will be re-evaluated if additional information becomes available. oscor inc., is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc., which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor, inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor, inc., or its employees, that the report constitutes an admission that the device, oscor, inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. this initial mdr is being submitted to meet our requirements of reporting. a follow-up will be submitted as soon as the investigation is complete. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda.
38202 | 1220648-2017-00052 | 6723323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6723323&pc=OZD | 06/20/2017 | 06/22/2017 |  | 07/18/2017 |  |  | 06/22/2017 | 07/18/2017 |  |  | 08/31/2018 |  | 07/18/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1257333 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 40 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Injury | Vascular Dissection | [10 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | N | Invalid/Incomplete |  |  | An impella cp was implanted into a (b)(6) caucasian male for support during a revascularization procedure, following an acute myocardial infarction (ami). the patient's act was reported to be 371 when the test was performed in the cardiac catheterization lab (ccl). the proper position of the pump was verified using fluoroscopy, however the patient was bradycardic with a heart rate of approximately 45-50 bpm. the pump was inserted after the coronary arteries were assessed, with the proximal left anterior descending (lad) noted to be occluded. the clinical staff reported the insertion of the impella cp went well, however the patient experienced bleeding from the insertion site. the clinical staff held pressure on the insertion site to control the bleeding from the time the peel-away introducer was removed, until the repositioning sheath was inserted. a large hematoma also formed at the insertion site. in response to the hematoma a balloon was used over the iliac bifurcation to tamponade the insertion site. one drug eluding stent (des) was placed and the protected coronary intervention (pci) was performed successfully. bleeding around the sheath continued and the patient was taken to surgery to assess the source. a small tear was found in the artery above the arteriotomy and the patient was weaned from impella support. the insertion site was surgically closed. the impella cp provided support to the patient for approximately two hours. there was no indication of any further adverse effect to the patient following the surgical removal of the device. | [5 empty] | The impella cp was not returned for evaluation, and the console data logs are irrelevant to this case. a review of the clinical case was done. an impella cp pump was implanted to support a (b)(6) male patient who had suffered an acute myocardial infarction. the impella insertion went well however bleeding occurred and a hematoma formed at the insertion site. the physician confirmed that they had pulled the peal-away sheath out of the patient before pealing it away. vascular surgery was consulted and the patient was moved to the operating room (or) to stop the bleeding. the patient was stable in the or and was found to have a small tear in the artery above the arteriotomy. the pump then was explanted and the access site was closed without further incident. the impella cp was discarded upon explant. the root cause of the arterial tear is unable to be determined because the product could not be evaluated. if additional information is obtained a supplemental report will be submitted. a review of the device and lot histories did not reveal any nonconformities. (b)(4).
38203 | 1035166-2017-00056 | 6699576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6699576&pc=DYB | 05/11/2017 | 11/09/2017 |  | 07/10/2017 |  |  | 06/12/2017 | 07/10/2017 |  |  | 08/01/2019 |  | 07/10/2017 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3001 |  |  | OSCOR INC. |  |  | US |  | C1-12705 |  | 0052-3001 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | FOREIGN | HEALTH PR | [9 empty] | Injury | Yes | No | Backflow | [6 empty] | Blood Loss | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | During the implantation of an abiomed impella pump on (b)(6) 2017, which procedure kit contains and requires the use of an 23f peel away introducer. after the 23f introducer was inserted into the patient body, the dilator was removed, and the hemostatic valve housing started to leak immediately. the leak was estimated to be about 15ml/minute, which resulted in a total blood loss before, during and after the pump implant procedure of the patient from 300ml to 500ml before. this resulted that the patient had to be infused with several bags of plasma blood transfusion. the abiomed impella pump was implanted successfully and the 23f introducer removed by means of peeling as per procedure (peel away). the patient had the abiomed pump devices implanted until (b)(6). the heart and of the patient worked normally after the procedure. | [5 empty] | Conclusion not yet available, evaluation in process. | The device was not returned for evaluation, as it was discarded. as a result, the allegations against this introducer (leaky valve) cannot be confirmed. a review of the device history record (dhr) identified no rejects during manufacture of this lot. a review of complaints against this lot number identified no other reported complaints. since the device was not returned, a root cause of the failure could not be determined, however, the potential cause of this failure may be: user does not insert dilator into the center of seal, improper insertion of dilator, and valve tear/damage. the potential effect(s) of this failure may be: device damage, tearing of seal, leaks-device replacement, and procedure delay. the potential cause(s) of this failure may be: device not used by adequately trained personnel and/or does not follow instructions for use (ifu). a review of the risk for this failure mode identified the risk to be medium. based on this investigation a capa is not required. oscor will continue to monitor this device for complaint trends and risk. per quality assurance (qa) leak test - introducer with seal-electronic tester procedure, a leak test is performed on 100% of the introducer sheaths. the abiomed adelante s2 hemostatic peelaway and sideport, instructions for use (ifu) provide these precautions: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. infusion through the sideport can be done only after all air is removed from the unit. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only.
38204 | 1220648-2017-00047 | 6680405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6680405&pc=OZD | 06/01/2017 | 06/01/2017 |  | 06/30/2017 |  |  | 06/01/2017 | 06/30/2017 |  |  | 01/31/2019 |  | 06/30/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET, EU |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1279436 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | DISTRI |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Bent | Material Separation | Patient-Device Incompatibility |  |  |  |  | Perforation | Injury | Blood Loss | [10 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that an (b)(6) female patient with 3-vessel disease was scheduled to have to have a subacute pci to be performed on (b)(6) 2017 with the aid of an impella cp. good local arterial pulsation was found in both groins. a standard 6f femoral sheath was inserted into the right femoral artery and iliacography showed both iliac arteries as good caliber. a standard 6fr femoral sheath was inserted into the left femoral artery. the 5000 iu of heparin was administered intra-arterially. a long standard j-wire was advanced to the aortic arch. the sheath was then removed from the left groin and the impella cp arterial dilators were used with no resistance. the impella peel-away introducer was inserted into the artery. the wire moved freely, but following the removal of the introducer there was no arterial flow in the sheath. the patient immediately deteriorated into shock due to severe bleeding from the left femoral artery, and was quickly intubated and ventilated. the complainant reported that during the insertion of the introducer it accidently peeled and damaged the vessel wall, requiring surgical repair. during the surgical procedure the patient was administered 6 units of blood. it was estimated that the patient lost a total of 2000 ml of blood. the traumatic lesion in the superficial femoral artery required repair. both lesions successfully surgically repaired. additional information was obtained from the complainant, which reported that the patient had very heavily calcified and tortuous vessels and the profunda artery was situated very deep and to the lateral. due to this some degree of angulation of the sheath occurred during entry into the artery. thus, the introducer followed the bending and entered the artery, while the peel-away portion of the sheath was peeled off on the edge of the heavily calcified artery; consequently, the peel-away sheath ended up on the outside of the artery. this finding was confirmed visually during the surgical repair. the pci was aborted, and was scheduled for the following week. the impella cp continued to support the patient without issue, and the patient was hemodynamically doing fine. the device successfully supported the patient for approximately 25 days. | [5 empty] | The device was returned for evaluation. the inspection of the introducer revealed that the sheath was split and the dilator was confirmed to be bent. the splitting was confirmed to occur along one of the score-lines. the score-line was not a "peel away" score line. a device history record was reviewed and revealed that the 14fr introducer had passed all incoming inspection criteria. the lot analysis review did not reveal any relevant failures. the clinical narrative of this event revealed that this patient had heavy calcification in the area where the introducer was being inserted. furthermore the artery was relatively deep. both of these points were speculated by the physician as the likely cause of the failure. the root cause of the sheath splitting was the difficult insertion angle caused by the patient's calcified vasculature and the deep nature of the vessel relative to the insertion site. a capa has been initiated to address 14fr introducer sheath failures of this kind. (b)(4).
38205 | 1220648-2017-00048 | 6680584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6680584&pc=OZD | 04/13/2017 | 06/01/2017 |  | 06/30/2017 |  |  | 06/01/2017 | 06/30/2017 |  |  | 11/30/2018 |  | 06/30/2017 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPEALL 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1271090 |  | 005062 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | [6 empty] | Hematoma | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The impella 5.0 was implanted through right axillary site on (b)(6) 2017 to support a (b)(6) male patient who was awaiting a heart and kidney transplant. the sterile anchoring ring that secures the sterile sheath was observed to be broken on (b)(6), however the clinical staff reported there was nothing else out of the ordinary found with the pump. the nurse used tegaderm to secure the component. the pump was reported to be successfully supporting the patient on the (b)(6), with a note made stating the patient's ptt was reported to be 44 seconds. on (b)(6) a hematoma was discovered at the insertion site and the hematoma was drained. we are reporting this event for the sake of consistency and are awaiting additional information from the field regarding the treatment of the hematoma. on (b)(6) the purge line was damaged, therefore the pump operated without purge for approximately one day before it was exchanged for a second impella 5.0. the patient was successfully weaned from impella support on (b)(6) 2017 and did not require any further mechanical circulatory support. there was no indication of any further adverse effect to the patient following the surgical removal of the device. | [5 empty] | Neither the pump nor the console data logs were returned from the field for evaluation. several attempts were made to retrieve the patient, console data and case information. multiple attempts have been unsuccessful, however if any additional information is obtained a supplemental report will be submitted.
38206 | 1220648-2017-00049 | 6680636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6680636&pc=OZD | 05/30/2017 | 06/01/2017 |  | 06/30/2017 |  |  | 06/14/2017 | 06/30/2017 |  |  | 04/30/2018 |  | 06/30/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | 0048-0002 |  |  | ABIOMED, INC. |  |  | US |  | 1239296 |  | IMPELLA CP |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Malposition of Device | [6 empty] | Mitral Regurgitation | Tissue Damage | Rupture | [10 empty] | Required Intervention | [5 empty] | ECMO | [22 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | An impella cp was implanted into a (b)(6) male patient who suffered an acute myocardial infarction, for support during a protective coronary intervention (pci). the patient had a history of 80 % stenosis of the right coronary artery (rca) , 100% stenosis of the left circumflex (lc), ventricular fibrillation, pulmonary edema and was administered cpr. the patient had a cardiac infarction with percutaneous transluminal coronary angioplasty (ptca) and stenting procedure and was placed on ecmo on (b)(6)2017. the impella cp was implanted for left ventricle (lv) decompression on (b)(6) 2017. ecmo was explanted on (b)(6) 2017 and impella cp was explanted on (b)(6) 2017. the physician observed a structure on the inlet area of the impella catheter upon explant. after explantation the patient developed mitral regurgitation. the physician reported the impella was close to the mitral valve area during the case, however there were no alarms and the flow rate was normal. the team did not feel that intervention or correction was necessary. the ifu recommends "the impella inlet area of the catheter should be 3.5 cm from the aortic valve annulus and well away from papillary muscle and subannular structures." the patient later developed sepsis and expired. | [5 empty] | The last of the console data logs, product and clinical images were received by the manufacturer on june 21, 2017, therefore the investigation is underway. a follow up report with the results of the investigation will be submitted following the completion of the evaluation. (b)(4). | The last of the console data logs and products were received by the manufacturer on june 21, 2017. this investigation included a review of the clinical description, returned product, tee clinical images, tissue analysis and console data logs. according to the clinical description, an impella cp was being explanted and biomaterial structure was noted to be found in the inlet cage of the pump. the patient then began to deteriorate and it was discovered they were suffering from mitral regurgitation. the plan was to perform a mitral clip but the patient expired before the procedure could be completed. the biomaterial structure found in the inlet cage was sent for testing by the hospital. the biomaterial structure was found to not be mitral structures. a review of the tee clinical images of the heart that were taken after the pump was explanted revealed the mitral valve was damaged. the original report stated "the physician reported the impella catheter was close to the mitral valve, however there were no alarms and the flow rate was normal. the team did not feel that intervention or correction was necessary." however, a review of the console data logs revealed the pump was out of position triggering "impella position in ventricle" and "position unknown" alarms on the fifth day of support, and prior to explant. the pump was returned for evaluation. the pump's inlet cage struts were free from any protrusions, burrs and sharp edges. without having clinical imaging of the device taken while inside the ventricle, it is not possible to determine the extent of the interaction between the pump inlet cage and the mitral structures. there were no defects found on the returned product components that would be likely to contribute to this issue during normal pump performance. the root cause of the mitral valve damage was unable to be determined because the extent of the interaction between the pump and mitral structures could not be determined from the returned imaging that was returned for evaluation. a review of the complaint system did not reveal any other complaints related to this reported issue for other pumps in this lot. a review of the device history did not reveal any mrrs or reworks related to the failure mode. (b)(4).
38207 | 1220648-2017-00051 | 6668763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6668763&pc=OZD | 05/29/2017 | 05/29/2017 |  | 06/26/2017 |  |  | 05/29/2017 | 06/26/2017 |  |  | 01/31/2019 |  | 06/26/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1280007 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Malfunction | No | Yes | Device Stops Intermittently | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that the physician placed an impella cp in a (b)(6) male patient on (b)(6) 2017. after the pump placed across patient's valve the guide wire removed and pump flow initiated, but the pump stopped after signaling an "impella stopped motor current high" alarm. the pump was successfully restarted only to ramp up and repeat alarm and then stopped 3 more times. despite the "performance level" the patient's hemodynamic status and ejection fraction were optimum. the decision was made to maintain the impella cp's performance level 2 as the motor current, if, purge pressure, purge flow and mean atrial pressure (map) were adequate to maintain a map >80mmhg. the patient was transferred to the icu, where the abiomed clinical educator (ce) observed that the automatic impella controller screen indicated that an impella lp2.5 was supporting the patient despite verification that an impella cp device was in place. the ce reported her observation to the physician and staff, and instructed the staff and physician that the motor current and pump flows were adequate, safe and would be used in place of plevels. she also warned them to be prepared for pump stop if the motor current spiked. the option of replacing the console was discussed along with the possibility that this could result in an inability to re-start pump. the option of replacing the pump with an impella lp2.5 pump was also discussed. the physician reported to the cm that he was comfortable with leaving the current pump in place and to continue running the pump at current level, as the patient had already been weaned completely off of pressors. the patient was reported to have been supporting the patient adequately without suction or hemolysis. this impella pump was verified as an impella cp, although it was registering on the aic, as well as running as an impella 2.5 pump. mc and flows are as documented in case and the pump was supporting the patient adequately without suction or hemolysis. patient support with this impella continued and it was reported to be supporting the patient "very well" for 30 hours. the patient was successfully weaned and the device was explanted without issue on (b)(6) 2017. the patient was reported to have been successfully supported and was in stable condition at the conclusion of the impella support. | [5 empty] | The impella cp pump was returned for investigation. the evaluation confirmed that this impella cp was recognized by the aic as an impella 2.5. during the evaluation the pump was run in a blood loop and was functional up to p6 before stopping for high motor current. in addition, the true flow and displayed flows were recorded at 60 mmhg of backpressure. the pump was unable to start on auto when initiating support. the device could only support the patient while operating at p2 without having a pump stop. the automatic impella controller data logs were also returned. the analysis of the data logs revealed that the pump stopped several times when started and was only able to sustain support while operating at p2, without triggering a high motor current pump stop. data logs confirm the reported issue and device analysis shows that the pump was programmed as an impella 2.5 during production rather than an impella cp. the rpms assigned to a 2.5 differ from the rpms of a cp at equivalent performance levels a review of the complaint database confirmed that there were no other pumps in this lot that apply to this failure mode. a device history record review was also performed and did not reveal any other impella pumps from this lot number having this issue. a review of the programming steps from the production lot show that the first programming step was performed correctly, assigning the eeprom the impella cp file structure, however, the second programming step was performed incorrectly assigning the eeprom the 2.5 file structure in conclusion the root cause of the pump stops and restricted performance levels was improper programming of the eeprom during production. as a corrective action a (b)(4) has been opened to address this failure mode. the failure will continue to be monitored and trended. internal reference (b)(4).
38208 | 1220648-2017-00044 | 6665631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6665631&pc=OZD | 05/22/2017 | 05/22/2017 |  | 06/23/2017 |  |  | 05/22/2017 | 06/23/2017 |  |  | 10/13/2018 |  | 06/23/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1265993 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | Yes | Crack | [6 empty] | Injury | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2017 a patient was to undergo an elective coronary artery bypass graft (cabg) for a critical proximal left ascending artery (lad) lesion and moderate (50-60%) left main (lm) lesion amenable to bypass. upon induction the patient went into ventricular fibrillation arrest. cpr was performed and the patient was resuscitated. the physicians had attempted to place an intra-aortic balloon pump (iabp) during the time of the patient crashing, but due to uncertain access without pulsatility during the code, the decision was made to open the patient's chest and proceed with the placement of the iabp, and then place the impella cp. the patient's chest was opened, but the physician was unable to locate the target for the circumflex artery (cx) or the coronary arteries. a saphenous vein graft (svg) to the lad was placed. about 3 hours into the case the patient experienced significant blood loss requiring massive packed red blood cells (prbc) and fresh frozen plasma (ffp) and platelets. the patient's hgb was 7 and hct was at 22. it was decided to wait for four more units of blood prior to the patient coming off iabp, and the impella cp being placed. during the wire exchange ventricular fibrillation ensued, requiring cardiac resuscitation via defibrillation. normal sinus rhythm resumed. consequently the impella cp was placed via the right femoral artery prior to the blood products being administered. following the impella cp's placement the additional 4 units of blood were administered to the patient, and the iabp was removed. impella support was started after the oscor sheath was peeled away and the repositioning sheath was inserted without issue. due to what appeared to be a significant ooze from insertion the site, and given the massive amounts of blood products that had been given to the patient as well as the hgb and hct levels the surgeons immediately decided to open the insertion site and place a purse-string suture at the arteriotomy for controlling any further bleeding. during this procedure the physician discovered a puncture at the arteriotomy site that was located high and above the inguinal ligament. the surgeon then placed an additional incision to gain access to the puncture site; however, it appeared there was a significant amount of pulsatile blood exiting the repositioning sheath. the sheath was pulled back out of the arteriotomy and a purse-string suture was placed around the impella catheter's 9fr shaft, adequately providing hemostasis. upon an initial and brief inspection the surgeons attempted to identify the problem and identified a small "slit" in the distal end of the repositioning sheath, which was allowing blood to pulse out during initial placement. the patient was obese and it was difficult to delineate where exactly the "split" was located on the repositioning sheath. the decision was made to simply secure the sheath partway out of the tissue track and maintain hemostasis with purse-string suture around impella shaft, and this appeared to control any blood loss. during an assessment of the patient, the right extremity appeared slightly less perfused and colder than left extremity. the vascular surgical team evaluated the patient and it was decided to re-open rfa impella site and assess. a fogerty catheter was placed and a very small clot was removed from the superficial femoral artery (sfa). it was noted that the patient was administered protamine prior to coming off iabp and it was decided not to add heparin into the purge solution due to blood loss and bleeding concerns. the 8k of heparin had been given by iv during the vascular intervention. the wound closed and patient was re-packed and closed w/staples. the patient was then sent to the icu in stable condition. on (b)(6) 2017 the patient was successfully weaned down from the impella cp. the following day, (b)(6) 2017, the impella cp was removed from the patient. additional information obtained from the staff indicated that this patient had a coagulopathy of some type that required him to be on coumadin for life. it was reported that the patient bleeding was likely the result of an assumed high inr for the coumadin. | [5 empty] | The impella cp and introducer were not returned for evaluation, as they are being held at the customer facility; consequently, the cause of this event could not be determined. in addition, the data logs were not reviewed as they did not pertain to the failure mode of this event. the manufacturer will continue to investigate all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report, if additional information is received. internal reference (b)(4). device being held at user facility.
38209 | 6646353 | 6646353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6646353&pc=KFM | 05/23/2017 | 06/07/2017 | 06/07/2017 | 06/16/2017 | 06/07/2017 | 06/07/2017 |  | 06/16/2017 |  |  |  |  | 06/16/2017 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | KFM |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  | 128704 |  |  |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | User Facility report | [14 empty] | Malfunction | No | Yes | Device Stops Intermittently | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella was running fine once started on auto flow setting. after a few minutes it was noted that the impella had switched itself off somehow and had stopped running. flow rate said off; it was set back to auto and it turned back on and began running again. impella rep was called; switched the device and console.
38210 | 1220648-2017-00043 | 6642128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6642128&pc=OZD | 05/15/2017 | 05/15/2017 |  | 06/14/2017 |  |  | 05/15/2017 | 06/14/2017 |  |  | 07/31/2018 |  | 06/14/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP,, |  |  | IMPELLA CP PUMP SET, US |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1254064 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 44 YR | [5 empty] | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Injury | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | On (b)(6) 2017 a (b)(6) year old female patient was admitted to the facility to have an elective procedure performed. the patient was reported to have had a history of carcinoma, a resent c-spine decompression, right coronary artery, acute myocardial infarction and a middle cerebral artery stroke in (b)(6)2017. the patient's ejection fraction was at 20%. during the elective percutaneous cardiac intervention (pci) the patient had 3 drug eluting stents placed at the left ascending artery. two hours following the pci the patient became hypotensive, and had an arterial blood pressure (abp) of 50/30, with anterior st elevations; consequently the patient was brought back to the cath lab were an impella cp was placed. the pump was placed via the left femoral artery. an angiogram revealed a large thrombus in the proximal left anterior descending artery (plad) which was aspirated. an intravascular ultrasound was performed which showed good stent apposition. the impella's peel-away sheath was removed and the repositioning sheath was advanced. sometime later that same day heavy oozing from the impella insertion site was observed. heparin and integrilin were administered intravenously. the activated clotting time (act) was over 320 seconds. manual pressure was held at the insertion site. during this time the impella cp purge pressure was running very high. when the patient was moved to the unit site bleeding increased. the staff continued to maintain pressure at the pump insertion site. the impella cp's purge pressure continued to rise. the physician checked for kinks in the cannula and the entry angle for the repositioning sheath, but none were seen. d5w with heparin was requested from the hospital pharmacy, but there was a reported long delay in the staff receiving it from the pharmacy. while waiting the purge pressure rose to over 1300 hg. the purge system was changed, but there was no improvement. when the purge solution arrived, it was changed, and the tubing was flushed so heparin would enter the motor housing more quickly. within a short time the purge pressure began to decrease. the patient then developed hemoptysis, and had coffee-grounds emesis. the patient became hypotensive; consequently, volume was given, and 4 units of packed red blood cells were administered. the heparin and integrilin were discontinued and the patient was intubated. vascular surgery was consulted and a small cut-down was made at the lateral circumflex femoral artery and sutures were placed around the vessel and the repositioning sheath. there was no retroperitoneal bleed, although there was still oozing. at this point the impella cp exhibited rising motor current. a thrombus was suspected, and the staff was advised that the pump might stop. a plan was then made to return the patient to cath lab when possible. prior to the return to the cath lab an echo was then performed which showed that the left and right ventricles were barely contracting. the impella cp flows were confirmed to still be above 2.6. a pump stop then occurred, but it did restart. the team suspected disseminated intravascular coagulation (dis) or some other coagulopathic process. the patient was then returned to lab where the staff prepped for an ecmo. approximately 10 minutes before ecmo cannulation, the impella cp again stopped and would not restart. the ecmo was started and the impella was explanted. two percloses provided hemostasis. the patient was transferred to the other facility for more advanced care. the physicians reported that there were many factors complicating this patient case. | [5 empty] | The impella cp, white connector cable and purge cassette were returned to the manufacturer. the visual inspection of the returned products revealed they were in good condition. there was no damage to the repositioning unit that would cause access site bleeding or vascular damage. a conversation with the abiomed field representative conducted by engineering during the device evaluation revealed that disseminated intravascular coagulation (dic) was suspected by the physicians or another coagulopathic disorder based on the initial stent thrombosis and the heavy bleeding from the impella site. at this point, however, the physicians have not made a definitive diagnosis regarding the patient's coagulopathic state during the case. the root cause of the heavy oozing from the impella site was unable to be determined. there were no defects found on the device that would cause access site bleeding or vascular damage. no corrective action is recommended at this time because the failure mode was determined to have a low risk index based on very low occurrence and moderate severity. failures of this type will continue to be monitored and trended. the evaluation for the low pump flow and pump stopped was performed. data logs were reviewed and confirmed the clinical account of low pump flow and pump stop. the pump was run for 5.5 hours and after 2.5 hours, the purge pressure abruptly rises to >1300 mmhg and there are alarms for high purge pressure. the purge fluid is changed to include heparin and the purge pressure decreases to about 900 mmhg, but immediately following the high purge pressure, the motor current rises while the pump is running at a constant p-level. the pump first stopped 4 hours after insertion due to high motor current. the staff was able to restart the pump for just over an hour, but it stopped again after 5.5 hours of support. the pump was subsequently explanted and replaced with an ecmo. the pump was tested and would not start due to a "high motor current" alarm. the pump was torn down and blood was found in the rear bearing. there was also blood on the stator but no biomaterial buildup on the inner diameter of the sleeve bearing was seen. the root cause of the pump stop and high purge pressure was blood ingress into the rear motor bearing. heparin was not used in the purge solution for the first three hours of the case. the root cause of the blood ingress is unknown. no corrective action is recommended at this time due to the device being used outside of product claims. the ifu of this device instructs the user that "as soon as practical after placement, change the purge fluid to include heparin. the recommended heparin concentration is 50 iu/ml in 20% dextrose solution. failures of this type will continue to be monitored and trended. (b)(4)
38211 | 1220648-2017-00042 | 6621487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6621487&pc=OZD | 05/09/2017 | 05/09/2017 |  | 06/08/2017 |  |  | 05/09/2017 | 06/08/2017 |  |  | 04/30/2017 |  | 06/08/2017 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | AIOMED, INC. |  |  | US |  | 1193762 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Malfunction | Yes | Yes | Break | Entrapment of Device | Malposition of Device | Positioning Problem |  |  |  | Hematoma | Hemolysis | Device Embedded In Tissue or Plaque | [10 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | The impella 5.0 ((b)(4)) was implanted through left axillary site on (b)(6) 2017 to replace an impella 5.0 ((b)(4)) that was supporting a (b)(6), male patient. on (b)(6) echo revealed the impella was positioned too deep within the ventricle, and the ct surgeon stated the device would be repositioned and pulled back. a hematoma was noted to have developed at the left axillary insertion site. the patient status otherwise remained unchanged and there were no other issues or concerns at this time. on (b)(6) pump flows were low and the placement signal had drifted. when the sensor was re-zeroed the placement signal monitoring was discovered to be disabled, and when it was turned on an "outflow blocked" alarm was triggered. on (b)(6) the patient went to the or for possible decannulation secondary to the left axillary hematoma, and due to increased bilirubin (sign of hemolysis). an evaluation of the hematoma was performed in the or, and the impella remained in place. on (b)(6) the patient was taken to the or for removal of the device due to hemolysis. upon attempting to remove the device the distal portion of the catheter detached from the motor junction and became lodged in the carotid artery. during the attempt to snare the device the cannula slipped down into the aorta. the device was retrieved through the axillary sheath. the patient was stable at this time and did not require any further mechanical circulatory support, and no additional inotropic support was administered at the time. there was no indication of any further adverse effect to the patient following the surgical removal of the device. | [5 empty] | The investigation of this event included a review of the clinical case that was conducted previously, and the pump and console data have since been returned for evaluation. a review of the console data revealed "suction," "outflow blocked," and "impella position under valve" alarms that suggest the pump may have been sitting too deep in the left ventricle. this may have caused additional loading at the joint between the cannula and pump housing. this position of the outflow cage being under the valve as suggested in the console data logs may have contributed to the reported hemolysis, due to the recirculation of blood through the pump housing. the pump was returned with the cannula separated from the pump housing. the pump housing and impeller were checked for roughness and surface imperfections, however no abnormalities were found. no burrs or sharp edges were found at the inflow cage or pump housing, and the repositioning unit was intact. a closer inspection of the pump revealed silicone on the textured surface where the cannula bonds to the pump housing. this silicone prevented the bonding material from properly adhering. the root cause of the hemolysis is most likely the pump being positioned deep enough in the left ventricle that the outflow cage was under the aortic valve, as indicated in the console data. the root cause of the cannula detachment is silicone at the textured surface of the pump housing. there is a corrective action being pursued in regard to this issue. the root cause of the hematoma is unable to be determined. (b)(4). | Neither the pump nor the console logs were returned from the field for evaluation. according to the clinical account, an echo revealed the pump was sitting too deep within the left ventricle and was repositioned. the patient experienced symptoms of hemolysis on (b)(6) some suction events occurred. an echo was performed to assist in the repositioning of the device. the patient was brought to the or for removal of the device due to hemolysis. upon explant the cannula detached from the pump housing and became lodged in the carotid artery. while attempting to snare the cannula it slipped into the aorta. the cannula was recovered through the axillary sheath. although the pump has not been returned, a review of the photographs of the pump taken in the or post-explant revealed the detachment occurred at the joint between the outer pump and pump housing. the root cause of the cannula detachment could not be determined because the device was not returned for evaluation. several attempts were made to retrieve the patient and event information, the device, and the console data logs. multiple attempts have been unsuccessful, however if any additional information is obtained a supplemental report will be submitted. a review of the device and lot histories did not reveal any nonconformities. (b)(4). device not returned.
38212 | 1220648-2017-00038 | 6621645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6621645&pc=OZD | 05/08/2017 | 05/09/2017 |  | 06/08/2017 |  |  |  | 06/08/2017 |  |  | 12/31/2018 |  | 06/08/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1270545 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 74 YR | [5 empty] | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | Appropriate Term/Code Not Available | [5 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | An impella cp was emergently implanted into a (b)(6) female patient with a history of acute fibrillation. the impella was placed when the patient acutely decompensated during a repeat prosthetic aortic valve replacement procedure. the left femoral access site began to ooze and a notable hematoma developed during surgery. unfortunately, formal imaging of the groin access site was not performed. a cut down was performed and the femoral artery was repaired and closed off. the patient required 22 units of blood replacement products during the 28-hour case. the patient showed signs of leg ischemia. the leg was cool and grey while the patient was on levophed and epinephrine, however the leg was warmer the next morning after the patient had been weaned from the drug. however, there were still no pulses at that time, which is why they proceeded with surgical removal of impella later that day. the pump was weaned to performance level 2 in preparation for removal. the patient's active bleeding slowed and the pump was explanted. an iabp was placed. the abiomed field team reported an update on the patient post-explant. the patient was reported to be in stable condition, and the patient's leg was reported to be perfusing, with adequate csm (circulation, sensation, motor) responses with good pedal pulses. there was no indication of any further adverse effect to the patient following the surgical removal of the device. | [5 empty] | Neither the pump nor the console data logs have been returned from the field for evaluation. the pump was discarded at the facility. multiple attempts to retrieve the console data logs have been unsuccessful, however if additional information is obtained a supplemental report will be submitted. (b)(4). device was not returned for evaluation.
38213 | 1220648-2017-00045 | 6618440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6618440&pc=OJE | 05/12/2017 | 05/12/2017 |  | 06/07/2017 |  |  | 05/11/2017 | 06/07/2017 |  |  | 12/31/2018 |  | 06/07/2017 |  | RIGHT VENTRICULAR BYPASS ASSIST DEVICE |  |  | IMPELLA RP PUMP SET, EU |  |  | OJE |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1276464 |  | 0046-0011 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 40 YR | [5 empty] | Injury | Yes | Yes | Break | Device Handling Problem | [5 empty] | Blood Loss | [12 empty] | Required Intervention | [5 empty] | IMPELLA CP, SERIAL NUMBER (B)(4) | VA-ECMO | [21 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | On (b)(6) 2017 the physicians placed an impella cp in a (b)(6) female patient. the impella cp was placed in the common iliac artery. a cct scan was done which revealed severe pulmonary edema with massive pleura effusion on both sides. the physician made the decision to remove the pump and to then re-insert it in the left ventricle and then to add an impella rp for central perfusion. a tricky wireless re-insertion of the cp occurred with the aortic valve not opening, and with nearly no left ventricle pulsatility. impella rp placement was begun, but after removal of the 23 fr introducer occurred there was severe patient bleeding from a cracked valve assembly on the introducer. the physician was then able to perform a quick and easy insertion and placement of rp, and the physician was able to change to the repositioning sheath. after placing/starting the rp it was immediately necessary and crucial to place an additional av-ecmo cannula to provide oxygenated blood the central circuit. loss of patient blood occurred and was estimated not to be over 15 ml/min, with an approximate total of 300 to 500 ml. the patient was reported to have required replacement blood products (rbc, tc, ffp) before, during and after implantation of the impella cp and impella rp, as a result of the patient very critical condition. with the patient's heart doing well the impella rp was successfully weaned from the patient and the device was removed. the patient was successfully supported with both the impella cp and impella rp for 24 days. upon removal of both impella pumps the patient was reported to be in stable condition. | [5 empty] | The impella rp was not returned for evaluation, as it was discarded subsequent to the event. therefore, no device analysis could be performed. photos of the suspect 23fr introducer were forwarded to the manufacturer for review. a device history review was performed, as well as a lot analysis. impella rp pump set (0046-0011; s/n (b)(4)) contained 23fr introducer kits originating from batch c1-12705. the 23fr introducer kits originating from batch c1-12705 were inspected at the time of receipt. no relevant issues were found during the incoming inspection of the introducers. in addition, the lot analysis revealed that there were no relevant introducer complaints found that were linked to any other pump set containing introducer kits originating from batch c1-12705. the clinical narrative describes severe bleeding originating from a crack at the 23fr introducer sheath hub. this account was verified with the submission to the manufacturer of two photos showing complaint condition, and confirming the failure to have happened. because the device was discarded no detailed engineering analysis could be conducted for this failure. without the product returned for evaluation the root cause of this event is unable to be determined. because the root cause was unable to be determined, no corrective action is recommended at this time. the failure will continue to be monitored and trended. (b)(4).
38214 | 1220648-2017-00040 | 6616076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6616076&pc=OZD | 05/10/2017 | 05/10/2017 |  | 06/06/2017 |  |  | 05/10/2017 | 06/06/2017 |  |  | 01/31/2019 |  | 06/06/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP SET, US |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1276855 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | Yes | Entrapment of Device | Difficult to Remove | [5 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that a (b)(6) male patient had been admitted to the hospital with chest pains on (b)(6) 2017, and was found to be too unstable for catheterization procedures to be performed. the patient's troponins continued to climb, as well as his creatinine levels. the patient was showing signs of cardiogenic shock, and although his ejection faction had been normal upon admission it was 10% on (b)(6) 2017. on (b)(6) 2017 the patient's bi-lateral iliac disease was treated with iliac stents. the iliac on the right showed better access, so insertion of the impella cp was in the right common femoral artery. the impella sheath and device were inserted without issue. impella flows were maintained throughout the procedure. during the procedure the right carotid artery was found to be totally occluded. the left main artery, left ascending artery and the cq artery all had at least 90% stenosis present. all three vessels were treated with balloon angioplasty and stenting. the patient lost pulsatility multiple times, but maintained map of greater than 60 with the impella in place. at the conclusion of the case the staff determined that the peel away sheath needed to be removed and the repositioning sheath needed to be inserted. while inserting the repositioning sheath, the impella cp was advanced too far into the ventricle. when the physician pulled the device back, it inadvertently pulled out of the ventricle. the physician attempted to re-advance the impella across the valve, but the device folded over on itself (bending at the cannula) and would not unfold during this attempt to wirelessly re-cross the valve. the patient was very hemodynamically unstable and needed the impella support; however, they were not able to re-advance the device into the ventricle. it was determined that the impella should be removed. while pulling the device out, and when the device reached the iliac bifurcation, it became stuck on the iliac stent struts and was not able to be removed. the physician tried a snare from the brachial and also in the contralateral iliac, and pushed a terumo destination long sheath dilator to try to push the impella out. during this time the patient was coded and expired. at that point, the device was left in the patient. | [5 empty] | The device was not returned for evaluation as the impella cp remained with the patient following the patient expiring. the automatic impella console data logs were returned for analysis. the analysis of the data logs confirmed loss of impella cp positioning. the impella cp was unable to be returned to the manufacturer the analysis; consequently, the cause of this event could not be definitively determined. the following are the likely root causes of the event: · the root cause of the impella going too deep into the ventricle is most likely due to the pump being pushed forwards while advancing the repositioning sheath. · the root cause of the impella completely exiting the ventricle is most likely due to the pump being pulled back too far during a repositioning attempt. · the root cause of the cannula kink and being unable to re-cross the aortic valve is most likely due to attempting to wirelessly re-cross the valve. · the root cause of the device getting caught on the iliac stent is unknown. no corrective action was recommended because the failure mode has been determined to be low risk index due to very low occurrence and major severity. failures of this type will continue to be monitored and trended. (b)(4). discarded after use.
38215 | 1220648-2017-00046 | 6609392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6609392&pc=OZD | 05/03/2017 | 05/03/2017 |  | 06/02/2017 |  |  | 05/03/2017 | 06/02/2017 |  |  | 11/30/2018 |  | 06/02/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET, US |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1269152 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Human-Device Interface Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2017 following successful patient support for 48.80 hours with an impella cp the device was removed from the patient. the patient was weaned prior to pump removal , and the patient sat up in bed prior to access site bleeding. following the patient sitting up excessive bleeding at the impella access site began, resulting in the need for replacement blood products being administered. in addition, a fem stop was applied, and also manual pressure to the access site was applied for 60 minutes. the patient was reported to be in stable condition following this event. additional information regarding amount of blood loss, or amount of replacement blood administered to the patient was unable to be obtained. | [5 empty] | The impella cp was not returned to the manufacturer as it was discarded by the user subsequent to removal from the patient; consequently, a device evaluation was unable to be performed. in addition, the automatic impella console data logs were not returned for analysis. a device history record review was performed. the review revealed that one pump from this lot had the pump plug replaced as part of routine rework. no other issues or rework relevant to this failure mode have been documented in the device history record. a review of the complaint data base was also performed; no complaints relevant to this failure mode have been filed for other pumps in this lot. the root cause of the access site bleeding is most likely because the patient sat up in bed, although this could not be definitively determined as the device was not available for evaluation. no corrective action is recommended because the failure mode cannot be definitively determined. (b)(4). discarded by user.
38216 | 1220648-2017-00039 | 6595545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6595545&pc=OZD | 04/26/2017 | 04/26/2017 |  | 05/26/2017 |  |  | 04/26/2017 | 05/26/2017 |  |  | 03/28/2017 |  | 05/26/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP, |  |  | IMPELLA CP PUMP SET, US |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1277051 |  | 0048-0016 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | No | Malposition of Device | Device Handling Problem | [5 empty] | Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2017 the staff was treating an (b)(6) female patient. during a mitroclip procedure the patient decompensated due to an air embolism. an impella cp was then placed in the patient via the left femoral artery (lfa.) the pump was just turned on, but the fluoroscopy images revealed that the pump was too deep in the ventricle, resulting in the physician pulling it back. an echocardiogram revealed a decompressed left ventricle (lv) and the right ventricle (rv) was down. the patient also was found to have severe tricuspid regurgitation (tr), right to left shunting, right ventricle infarction and severe pericardial effusion. blood products and the inotropes were administrated. the patient was then taken to the operating room to explore possible patient bleeding. upon arrival in the or a swan was placed, and while waiting the patient began to dramatically improve. an echo that was performed which showed that the lv and rv were now back to baseline and the confidence interval cardiac index was now at 2.4. approximately 15 minutes later the patient decompensated, and had low native heart pulsatility. at this point the surgeons opened the patient chest and found a very large clot that was then removed. while exploring the patient's chest the surgeons felt a hole with a protrusion on the backside of the patient's heart. the heart was tilted and the impella cp's pigtail was found to have perforated through the ventricle. the surgeons tried to pull the pump back through, but then decided to cut the pigtail off and bring the pump back into the ventricle. the ventricle was successfully repaired. the patient was then weaned down on the bypass pump and put back up on the impella cp flow, although the a biomed representative had advised the surgeons to remove the pump due to the fact that the pigtail had been cut. an echo confirmed pump placement and color flow was up in the ascending aorta. although the surgeon wanted to continue impella cp support to keep the patient stable, the a biomed medical director was contacted, and it was recommended that the pump be removed as it could damage the patient's heart with the blunt tip exposed. upon arrival of the patient to the sicu the patient remained stable and all drips were off. the impella cp pump was weaned from the patient and successfully removed. the patient remained stable. the clinicians reported that they were aware that the vascular perforation was the result of their excessive force on the pump during placement. | [5 empty] | The impella cp was returned for this investigation. during the visual inspection of the pump it was found that the portion of the pigtail, which had been cut off by the physicians was not returned. there were no signs of kinking on the cannula and no other signs of defect on the pump during the evaluation. the device was started and ran properly. in conclusion, the physicians attributed the perforation to forces they applied during insertion of the pump. the device was returned; however, the portion of the pigtail which perforated the heart was not returned. there was no evidence of damage to the rest of the pump. no images taken during the procedure were available for review. without a complete device review, a definitive root cause for the perforation is unable to be determined. due to the fact that a definitive root cause could not be determined, there is no corrective action recommended at this time. the failure mode will continue to be monitored and tracked. the manufacturer will continue to investigate all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received (b)(4).
38217 | MW5069951 | 6585550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6585550&pc=PBL | 05/17/2017 | 05/22/2017 |  | 05/22/2017 |  |  |  | 05/22/2017 |  |  |  |  | 05/22/2017 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | PBL |  |  | 0048-000 |  |  | ABIOMED INC. |  |  | US |  | 123319 |  |  |  | Device was returned to manufacturer | [5 empty] | No | No Information |  |  | Voluntary report | [8 empty] | 82 YR | [5 empty] | Injury | Yes | No | Device Operates Differently Than Expected | [6 empty] | Complaint, Ill-Defined | [12 empty] | Life Threatening | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Impella catheter failed during an acute stemi procedure. the catheter functioned properly for 1 hour before an apparent malfunction. when the malfunction occurred, the patient's condition deteriorated quickly requiring intubation and cpr until the impella cp catheter and console could be exchanged for a new impella cp catheter and impella console. the second catheter and console functioned properly and the case was completed without further complications. impella pump intermittently shut off and turned on, on its own, before complete malfunction. representative was notified of equipment symptoms prior to failure.
38218 | 1220648-2017-00035 | 6570423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6570423&pc=OZD | 04/26/2017 | 04/26/2017 |  | 05/16/2017 |  |  | 04/26/2017 | 05/16/2017 |  |  | 08/31/2018 |  | 05/16/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1257926 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 85 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | Human-Device Interface Problem | [5 empty] | Ischemia | Patient Problem/Medical Problem | [11 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | On (b)(6) 2017 a (b)(6) male with a complex medical history was admitted for a high risk pci in the cardiac catheterization lab of the left anterior descending artery (lad) and right coronary artery (rca) and for a future there would be mitral valve repair as well. the patient's history of peripheral vascular disease (pvd), diabetes, coronary artery bypass, mitral regurgitation, and chronic renal failure prompted the use of impella cp for support. the cp was placed and the team left the oscor peel-away sheath in place, in the left femoral, and admitted the patient to the icu. the right femoral artery was not an option for placement due to the small vessel size, pvd, and previous access of the right femoral artery. the patient was admitted to the icu after a surgical repair of the rca and a planned staged revascularization of the lad the following day. this decision to leave the peel away sheath in place in the icu goes against the ifu for the cp which states the following: "to prevent failure of the peel-away introducer, remove the peel-away introducer prior to transport when activated clotting time (act) is less than 150 seconds. " on the next day, (b)(6), the left leg began to exhibit signs of ischemia. the leg was pulseless and vascular surgery was consulted. the patient's anticoagulation was assessed and the ptt was 38. later in the day of the (b)(6), the oscor sheath was peeled away and the repositioning sheath was advanced into the artery. the doppler assessment of the leg showed reduced blood flow at the popliteal artery and no measurable flow in the foot. there was no rigor or mottling present but, the decision was made to wean the patient from the impella cp pump and explant in the operating room. to treat the ischemic leg an endarterectomy was done. during the surgical procedure no thrombus was found at the access site or in the femoral or iliac arteries. | [5 empty] | The impella cp and introducer were not returned for evaluation, as they were discarded subsequent to the event; consequently no device analysis could be performed. in addition, the data logs were not reviewed as they did not pertain to the failure mode of this event. the pump lot was analyzed and no other complaints were leveled against the lot, nor were there any internal reworks for this (b)(4) manufactured pump. the root cause of the limb ischemia was a combination of the patient's known peripheral vascular disease, that was present in bilateral lower limbs, and the introducer being left in the patient's leg for an extended duration. this introducer being left in goes against the ifu recommendations and field training. the failure type will continue to be monitored and trended. the manufacturer will continue to investigate all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report, if additional information is received. internal reference (b)(4). device not returned.
38219 | 1220648-2017-00036 | 6570430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6570430&pc=OZD | 05/01/2017 | 05/01/2017 |  | 05/16/2017 |  |  | 05/01/2017 | 05/16/2017 |  |  | 12/31/2018 |  | 05/16/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1273198 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Inadequacy of Device Shape and/or Size | Patient-Device Incompatibility | [5 empty] | Blood Loss | [12 empty] | Required Intervention | [5 empty] | COOK 14FR X 30 SHEATH | COOK 16FR SHEATH | OCCLUSION BALLOON | [20 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | On (b)(6) 2017 a (b)(6) male was admitted for a high risk pci. while in the cardiac cath lab for the intervention, the patient developed a thrombotic burden in the left anterior descending artery and circumflex and this prompted placement of an impella for support. the patient's access of the femoral artery was not noted to be difficult one, though peripheral vascular disease was present. as the patient decompensated, while on support with the impella cp, he was intubated and had angiojet thrombectomy of both coronary vessels, followed by stents. for transport to the icu, the oscor sheath was peeled away, and bleeding was significant around the impella repositioning sheath. the team performed contralateral balloon inflation and exchanged to a cook 14 x30cm sheath. on the following day, (b)(6), there was still bleeding from the access site in the right groin, and so the sheath was again exchanged to a 16fr cook sheath. this upsizing resolved the bleeding and hemostasis was achieved. due to the blood loss, the patient received a total of 5 units of blood product. the impella cp was able to be weaned and explanted on the (b)(6). the patient is stable, though still intubated and following commands as requested. | [5 empty] | The impella cp pump, cable, and purge cassette were all returned for investigation. upon review of the products, there was an observed tear and scratch in the wound closure portion of the impella cp. the data logs were returned for analysis. they confirmed the clinical account of the pump being placed into "surgical mode" when the occlusion balloon is advanced to the access site and the sheath is replaced and inserted to the arteriotomy. the data logs revealed that the patient was successfully supported with optimal flows and the purge pressures were within specification throughout the support. there were no persistent alarms. the root cause of the access site bleeding was a tear in the tip of the wound closure. the lot of cp pumps has had no other complaints. there will be no corrective action for this adverse event. the failure was determined to be low risk. these types of failures will continue to be monitored and trended. internal reference (b)(4).
38220 | 1220648-2017-00041 | 6570449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6570449&pc=OZD | 05/09/2017 | 05/10/2017 |  | 05/16/2017 |  |  | 05/10/2017 | 05/16/2017 |  |  | 03/31/2019 |  | 05/16/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1279596 |  | 0048-0016 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 58 YR |  |  | 136 |  |  | Malfunction | Yes | No | Patient-Device Incompatibility | Device Handling Problem | [5 empty] | Perforation | Blood Loss | [11 empty] | Hospitalization | Life Threatening | Required Intervention |  |  |  | STANDARD .035 X 145CM WIRE | [22 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | On the (b)(6) a (b)(6) obese male was admitted to the hospital and found to have an acute in-stent thrombosis in the coronary artery due to non compliance with medications. the patient had an intervention (pci) in the cardiac catheterization lab to open up the left anterior descending artery. the pci was successfully completed and the cardiac team made the decision to place an impella cp for support while in recovery in the intensive care unit. of note listed in the ifu for the cp is a statement regarding "evaluation prior to inserting the impella catheter: observations: obesity". the physician began to prep the patient for the impella insertion by removing the 6fr sheath and introducing the 14fr sheath to the right femoral artery, using a non-abiomed provided .035 x145 cm wire. per the abiomed ifu the wire provided in the access kit is listed as a: ".035 inch stiff access guide wire". the patient was obese and so the access site was assessed. the femoral angiogram revealed the artery to be 9mm in diameter. when the team attempted to place the 14fr sheath into the artery there was an observation made of the "odd angulation" of the insertion and associated bleeding. due to the bleeding the impella insertion was aborted. the 14fr sheath was immediately replaced by a larger 16fr dryseal sheath to occlude the artery and decrease the blood loss. the patient was then transferred to the operating department where there was a vascular repair. the anterior stick and posterior laceration were surgically repaired. the physician did inform the abiomed representative that this stick was improper technique, rather than product malfunction. the patient is recovering from the massive acute mi and had a permanent pacemaker placed. | [5 empty] | The impella cp and 14fr sheath were not returned for evaluation. the cp was never inserted into the patient and the 14fr sheath was discarded subsequent to the event. consequently, no device analysis was performed. in addition, no data logs were returned as the console and cp were never utilized for support. a review of the clinical information revealed the patient to be in excess of (b)(6). due to the patient's large body habitus the angulation of access was atypical. the insertion angle and use of a non-abiomed provided wire contributed to the adverse event. the physician stated the combination of the patient's large habitus and his technique caused the bleeding and vascular injury. though the device was not returned for analysis, extreme angles of introducer insertion are known to facilitate this type of failure. the review of the lot of pump and introducer revealed no other complaints. the root cause of the event is user related and therefore no corrective action has been initiated. the failure will be monitored and trended. the manufacturer will continue to investigate all reasonable and obtainable sources of information, and will provide results and conclusions in a supplemental medwatch report, if additional information is received. internal reference (b)(4). device not returned.
38221 | 1220648-2017-00034 | 6571002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6571002&pc=OZD | 04/19/2017 | 04/19/2017 |  | 05/16/2017 |  |  | 04/19/2017 | 05/16/2017 |  |  | 12/31/2018 |  | 05/16/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET, US |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1272918 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Compatibility Problem | Missing Value Reason | [5 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that an (b)(6) male patient was admitted to the hospital with a non-st-elevation myocardial infarction. the patient presented with chest pain and shortness of breath after having cancelled his visit to his cardiologist office visit 2 weeks prior. a diagnostic exam revealed a 99% blockage of the left main and a wide open flailing mitral valve. the patient was also experiencing hemodynamic compromise with multiple rhythm changes and chest pain. an impella cp was then placed in the patient's left femoral artery without incident. the patient was accepted for a coronary artery bypass graft (cabg) and a mitral valve replacement surgery. the impella cp catheter position was confirmed under fluoroscopy and again was confirmed prior to the surgical procedure under transesophageal echocardiogram (tee). the patient was then placed on bypass and the aorta was soft cross-clamped over the impella. during heart manipulation the physician noted the impella cp pigtail had perforated the myocardium. when the perforation was discovered, the curved portion of the pigtail was visualized outside of the ventricle. at this point, the pigtail was not kinked and appeared as it normally does. after the evaluation, the physician opted to purse string suture around it before removing the impella to maintain proximity of the perforation. the impella was then retracted and removed. the sutured closure of the myocardium was completed successfully. when the left ventricle (lv) was perforated by the pigtail, the heart was not moving, and it wasn't actively full of blood; the impella was in surgical mode, and the physician was able to repair it almost immediately. three coronary artery bypass grafts were successfully placed and the mitral valve replacement was completed successfully. the patient then came off the pump without incident. the three physicians involved in the case reported that they believed that the mv prolapse over the many years resulted in a dilated myocardium and weakened walls. the combination of the patient's decreased myocardial integrity and advanced age made the event more likely to occur than normal. this issue was thus concluded by the physicians to be patient specific and incidental to the therapies being provided. the six hours of patient was reported as successful, and the patient was in stable condition with no additional adverse events. | [5 empty] | The impella cp was returned to the manufacturer for evaluation. in addition, the automatic impella console data logs were returned for analysis. the analysis of the data logs revealed that the pump provided good support of 2.1 liters per minute for several hours before it was placed in surgical mode for the cabg. purge pressure and flow were within specification for the duration of support and the placement signal showed an aortic waveform before the patient was put on bypass for the cabg. the returned pump, cable and purge cassette were returned in good condition. the pigtail was intact but showed a slight kink. there were no defects on the pigtail, or any other portion of the pump that could have caused the lv perforation. it is possible that the pigtail kinked when pressure from the myocardial wall was applied on the curled portion during heart manipulation. because the pump and aorta were cross-clamped the pump could not retreat back into the aorta when the pressure was applied, and the myocardium could have been punctured by the kinked pigtail. the evaluation of the device determined that the root cause of the lv perforation was a combination of decreased myocardial integrity due to mitral valve prolapse and the patient's advanced age, as well as the cross-clamping of the aorta, as that prevented the impella from giving way to resistance that occurred during heart manipulation. a review of the device history record revealed that there had been no other issues relevant to this failure mode for any other pumps in this lot. no corrective action is recommended at this time because the failure mode was determined to have a low risk index based on very low occurrence and major severity. failures of this type will continue to be monitored and trended. (b)(4).
38222 | 1220648-2017-00033 | 6571050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6571050&pc=OZD | 04/25/2017 | 04/25/2017 |  | 05/16/2017 |  |  | 04/25/2017 | 05/16/2017 |  |  | 12/31/2017 |  | 05/16/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET, EU |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1269031 |  | 0048-0002 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 65 YR | [5 empty] | Death | Yes | No | Improper or Incorrect Procedure or Method | Device Handling Problem | [5 empty] | Injury | Blood Loss | [11 empty] | Death | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2017 an impella cp was successfully placed in a (b)(6) male patient in preparation for a high risk percutaneous coronary intervention procedure (hrpci). insertion was uneventful and the physician had no issues gaining vascular access. the patient was successfully supported for approximately 75 minutes. the staff reported that following patient support the physician peeled away the 14 french oscor sheath, while the sheath was still inside the patient's femoral artery, rather than removing the introducer from the artery before separating the two halves. this resulted in a femoral vascular injury, and 500ml loss of blood. the injury necessitated vascular surgery to repair the artery. a balloon was placed in the patient to aid in the femoral artery repair. no replacement blood products were administered to the patient. following the surgical repair the patient was transported to the cath lab. approximately 7 hours later the patient expired. the complainant reported that the physician and his team were aware that it was a mistake for the physician to have peeled the sheath while it was still inside the vessel and understand that this was the reason of this patient event. the ifu specifically instructs users to "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding" when removing the peel away introducer and inserting the repositioning unit. the physician in this case separated the two halves of the sheath while still inside the vessel, causing this vascular damage in the process. | [5 empty] | The impella cp, white cable, and ½ of the 14fr oscor introducer were returned: no obvious defects were found during the visual examination of the introducer: a slight deformation was found at the introducer tip and there was a very small crack that extended <1mm up the sheath. there were no defects found on the introducer that would have caused a femoral artery injury. the automatic impella console data logs were analysed. the analysis confirmed the clinical account of this issue. the pump was successfully run for about 45 minutes. approximately 30 minutes after the pump was turned on, the patient's blood pressure spiked and then rapidly declined. during the time of the blood pressure abnormalities, the pump flow also decreased. before the femoral artery was damaged the pump provided good support of 3.1 l/min. there were no issues found during the manufacture and inspection of the pump, and no other complaints were reported for this impella cp lot number. in addition, the introducer packaged into this kit had been inspected and had passed all inspection criteria. in conclusion, : the root cause of the vascular damage was use error due to the physician separating the sheath while it was still inside the vessel rather than removing the introducer from the artery before separating the two halves. the corrective action of retraining is currently being scheduled in regard to the proper method of removing the peel-away introducer. (b)(4). device not yet returned to manufacturer.
38223 | 1220648-2017-00032 | 6565177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6565177&pc=OJE | 04/11/2017 | 04/11/2017 |  | 05/12/2017 |  |  | 04/11/2017 | 05/12/2017 |  |  | 08/31/2018 |  | 05/12/2017 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA RP SET |  |  | OJE |  |  | IMPELLA RP PUMP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1261980 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 72 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Appropriate Term/Code Not Available | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | ANGIOVAC | HEARTMATE II LVAD | [21 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2017 a (b)(6) patient was taken to the operating room for the placement of a heartmate ii left ventricle assist device (lvad.) following placement of the lvad the patient's right ventricle struggled to come off the heart lung pump. as a result the impella rp was placed. the pump was placed with ease, and was in good position with the pig tail deflected into the patient's left pulmonary artery (lpa), as confirmed with x-ray. the following day the intra-aortic balloon pump (iabp) was removed the patient continued to be successfully supported with the impella rp and the lvad. on (b)(6) 2017 the patient's chest was opened. heparin levels were reduced slightly due to bleeding concerns after the patient's chest was opened, and the activated clotting time (act) fell from recommended levels to 157. the patient's chest was closed on (b)(6) 2017; the act was at 151. the impella rp placement was verified via chest x-ray with no reported issues. the lvad remained in place. heparin was confirmed to be in the purge fluid, and the patient remained on epinephrine, milrinone and vasopressin. on (b)(6) 2017 a clot in the patient's right atrium was found during an echocardiogram. the patient was taken to surgery where the initial intervention to remove the clot with an angiovac was unsuccessful. clots external to the heart were washed out. it was noted by the physician that the rp was doing almost all of the work for the patient's right heart during the procedure. heparin in the purge fluid was changed to argatroban as there were heparin-induced thrombocytopenia (hit) concerns. following this procedure the patient was taken to the icu without event. on (b)(6) 2017 an increase in hematuria had been noted since the previous day, both the impella rp and the lvad were still in place. the patient hemoglobin lab results showed the pfhgb to be at 460. the patient was successfully weaned and the impella rp was removed from the patient. the impella rp pump was found to have clot at the inflow cage upon explant. no blood products were administered to the patient and no dialysis was performed as a result of the reported hematuria, and patient's impella 5.6 days of impella rp support was reported as successful. | [5 empty] | The impella rp was returned to this manufacturer for evaluation. a visual examination of the device was performed. the catheter was found to have been cut, and biomaterial was found to be wrapped around the proximal end of the impeller. the pump housing was removed and was soaked in tergazyme to remove dried blood and biomaterial. no defects were found on the pump housing surface. since it was also indicated that a clot was found in the right atrium, the biomaterial on the pump was most likely ingested and did not build up over time. the biomaterial could have been ingested by the pump when it was running or during explant. a review of the device history record showed no other devices from this lot with this or any other failure mode. in conclusion, the root cause of the hemolysis was unable to be determined, as the catheter had been cut and the console data logs were not returned for analysis. (b)(4).
38224 | 1220648-2017-00037 | 6556493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6556493&pc=OZD | 04/10/2017 | 04/10/2017 |  | 05/10/2017 |  |  | 04/10/2017 | 05/10/2017 |  |  | 12/31/2018 |  | 05/10/2017 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1268465 |  | 005060 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 58 YR | [5 empty] | Malfunction | Yes | No | Device Inoperable | [6 empty] | No Code Available | [12 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | On (b)(6) 2017 a (b)(6) year old male was admitted to a hospital in (b)(6) for acute myocardial infarction and cardiogenic shock with an ejection fraction of 30%. on (b)(6), the patient had an impella 5.0 placed due to shock via a cutdown access of the right axillary artery. on the (b)(6), the patient was neurologically disorientated and attempted to get out of the bed. during the fall to the floor the patient damaged the pump's cable and arterial line. the impella 5.0 was no longer functional for support, on this the 18th day of support. of note the ifu for the eu states the indication for cardiogenic shock to be 10 days. the patient decompensated and due to hemodynamics being life threatening, the patient was returned to the operating department for explant of the damaged 5.0 and installation of a new 5.0. with a new 5.0 inserted and supporting the patient, the hemodynamics stabilized and the patient was supported again successfully. in addition, the patient had a scan done to show that the fall did not cause any neurological harm or injury. a few days after placement of the second 5.0 the patient returned to the operating suite for explant and placement of a lvad. the patient is reported to be doing well. | [5 empty] | The impella 5.0 was not returned for evaluation, as it was discarded subsequent to the event, consequently no device analysis could be performed. the device breakage location could not be confirmed without the device return. in the ifu for the impella 5.0, there are recommendations for general patient care. the use of restraints and immobilization may not be adhered to due to cultural reasons and beliefs in parts of europe. these ifu recommendations pertain to the need for patient management to maintain successful support. the patient should not have gotten himself out of bed in a disorientated state and caused damage to the impella. the statements are inclusive of the following: "use knee immobilizer as needed to maintain access site straight. use care when moving or turning a patient; the impella® 5.0 catheter may move out of position and cause a positioning alarm. be careful not to pull on the impella® 5.0 catheter when transferring a patient from one bed to another." the console data logs were returned for analysis. the logs revealed that the device came "disconnected" but otherwise we are unable to determine the extent of what occurred. the root cause of the pump stop, that necessitated replacement, was the patient falling out of bed causing the pump to disconnect. no corrective action is recommended because the failure is determined to be low risk due to very low occurrence and moderate severity. the failure will continue to be monitored and tracked. upon review of the lot of pumps, there were no other complaints leveled against the lot nor any reworks that apply to this failure. the manufacturer will continue to investigate all reasonable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. (b)(4). device discarded at facility.
38225 | MW5069718 | 6562812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6562812&pc=PBL | 05/05/2017 | 05/10/2017 |  | 05/10/2017 |  |  |  | 05/10/2017 |  |  |  |  | 05/10/2017 |  | CP SET |  |  | IMPELLA CP SET |  |  | PBL | [5 empty] | ABIOMED |  |  | US | [5 empty] | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 62 YR | [5 empty] | Injury | No | Yes | Failure to Pump | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | During procedure, impella device was inserted and once inserted, the device would not pump. a new device was inserted and functioned fine. no harm to pt.
38226 | 1220648-2017-00029 | 6533766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6533766&pc=OZD | 04/07/2017 | 04/07/2017 |  | 05/01/2017 |  |  | 04/07/2017 | 05/01/2017 |  |  | 01/31/2019 |  | 05/01/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1276359 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Appropriate Term/Code Not Available | [6 empty] | Low Blood Pressure/ Hypotension | Perforation | Blood Loss | [10 empty] | Life Threatening | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | On (b)(6) 2017 a (b)(6) female was diagnosed with left main coronary artery disease. the patient was taken for surgery and two bypass grafts were placed. while in the icu the patient decompensated and active cpr was done. the patient did not recover with cpr and so was taken back to the operating department where the care was escalated from an iabp to the impella cp. during the surgical placement there was mention of "hard cardiac massage" to the ventricle with the impella in place. the patient's right heart status was also being monitored for right heart failure when admitted back to the icu. on the (b)(6) the patient was being monitored and given fluids and blood products to treat the impella cp's suction alarms . the patient's chest was still open from the surgical procedures and so anticoagulation was minimal. with the patient's chest open the team had difficulty visualizing the position of the impella. without good visualization the suction alarms and flows were being monitored as they were not ideal for successful support. later in the day of (b)(6), the physician attempted to blindly reposition the impella cp without echo guidance. per the ifu repositioning is to be: "echocardiography is a commonly used tool for evaluating the position of the impella® catheter relative to the aortic valve and other intraventricular structures post-placement. the best echocardiographic views for positioning the impella® catheter in the left ventricle are a long axis transesophageal echocardiogram (tee) or a parasternal long axis transthoracic echocardiogram (tte). these long axis views allow you to see both the aortic valve and impella® catheter inlet area." after repositioning of the impella cp a tte was performed. the tte showed position of the impella in relation to the aortic valve, however the apex of the heart was not seen. the patient suddenly decompensated and was taken back to the operating department where the surgeon observed a perforation of the heart wall. the impella was pulled back into the aorta and the surgeon patched the perforation. the patient did expire despite all efforts and cardiac massage. | [5 empty] | The impella cp was not returned for evaluation as it was retained within the patient and sent along to the funeral home. due to this, no device analysis could be performed and the root cause of the left ventricle perforation could not be determined. the data log review is not relevant to the failure investigation. based off of the clinical account, it appears that the perforation occurred while the physician was repositioning the device without echo guidance, whether tee or tte. abiomed recommends the use of echo guidance when repositioning the device, and states this in the ifu. the root cause of the left ventricular perforation is unable to be determined because the device was not returned for review. no corrective action is recommended because the failure mode was unable to be determined. this failure mode will continue to be monitored and trended. no other complaints have been made against this lot of cp impella pump. the manufacturer will continue to investigate all reasonable and obtainable sources of information and will provide results and conclusions in a supplemental medwatch if additional sources are obtained. (b)(4).
38227 | 1220648-2017-00031 | 6533786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6533786&pc=OJE | 04/02/2017 | 04/02/2017 |  | 05/01/2017 |  |  | 04/02/2017 | 05/01/2017 |  |  | 12/31/2017 |  | 05/01/2017 |  | RIGHT VENTRICULAR BYPASS ASSIST DEVICE |  |  | IMPELLA |  |  | OJE |  |  | IMPELLA RP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1230922 |  | 004334 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH | [10 empty] | Injury | Yes | No | Difficult To Position | Human-Device Interface Problem | [5 empty] | Hypoxia | Blood Loss | [11 empty] | Hospitalization | Life Threatening |  |  |  |  | ECMO | HEART MATE | INTERNAL JUGULAR SWAN | PLATINUM PLUS WIRE, .025 | V18 WIRE | [18 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A male patient of unknown age was admitted to the medical facility in right ventricular failure post placement of a left ventricular assist device (lvad) in (b)(6) of 2017. the patient was taken to the cardiac cath lab for placement of the impella rp for right sided heart support. the patient was accessed and a swan-ganz catheter was placed with some difficulty. the swan was exchanged out over an .018 v18 wire and the impella rp was advanced. of note in the ifu for the impella rp the wire recommendation is as such: "a 0.025 inch, 260 cm platinum plus or similar stiff placement guidewire is used for the placement of the impella rp® catheter" there is no recommendation to use a v18 wire for placement. there was placement difficulty as the wire was pulled back, and so the rp was removed and the swan re-advanced. the physician chose then to use a "buddy wire technique" and place the .018 v18 wire and platinum plus wire into the left pulmonary artery via the swan. the rp was successfully then positioned over the .025 wire and delivered appropriate flows. the lvad device also had improved flows with the rp support running concurrent. the patient was transferred to the icu for monitoring and placement of the rp was confirmed by x-ray and echo. during the next 2-4 hours the patient's condition deteriorated. there was question of bleeding into the patient's left lung. although there was no proof, the physician believed that the bleeding was due to a pulmonary artery perforation from guidewire insertion. on the following day, the patient had the impella rp explanted, and extra corporeal membrane oxygenation (ecmo) placed. ecmo and the lvad continue to support the patient. | [5 empty] | The impella rp was not returned for evaluation, as it was discarded subsequent to the event. therefore, no device analysis could be performed. the console logs were returned for analysis. the logs however do not contribute any information to the root cause of the event. there is no anomaly in the data logs, with the logs showing normal pump performance. the pulmonary bleed was most likely caused by repeated advancement and manipulation of guidewires into the pulmonary artery during the difficult positioning of the rp. no corrective action is recommended as the root cause was unable to be determined. failures of this type will be monitored and trended. the manufacturer will continue to investigate all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. (b)(4).
38228 | 1220648-2017-00028 | 6535472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6535472&pc=OZD | 04/04/2017 | 04/06/2017 |  | 05/01/2017 |  |  | 04/06/2017 | 05/01/2017 |  |  | 10/31/2018 |  | 05/01/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1266569 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Crack | Fracture | Patient-Device Incompatibility | Material Deformation |  |  |  | Perforation of Vessels | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | It was reported that on (b)(6) 2017, a (b)(6) year old male patient was being treated at the hospital. the patient was reported to have a history of aortic stenosis, severe 3 vessel disease, and peripheral vascular disease with right superficial femoral artery complete occlusion and iliac calcification. the patient had been brought to the facility for a right coronary artery (rca) and left main atherectomy and percutaneous coronary intervention (pci.) the physician chose to perform the rca intervention prior to impella cp. during the mid rca intervention the vessel was dissected requiring serial stent deployment. post pci rca, the patient presented with thrombolysis in myocardial infarction profusion grade 2. an angiogram of the right femoral artery (rfa) was performed using a 6fr sheath showing appropriate vessel size and puncture location. the physician reported that the oscor sheath was inserted without resistance using abiomed 0.035 wire. after the oscor sheath was inserted the patient's blood pressure dropped to the 50's and the patient became bradycardic. cpr was started. an angiogram revealed that the oscor sheath had split distally, showing right iliac dissection and staining suggesting the vessel had been perforated. the oscor sheath was removed over the wire, and the physician attempted the insertion of a 14 x 30 cook sheath but was not able to do so due to tortuosity. at that point vessel was closed using the 2 existing percloses. the patient was administered 6 units of replacement blood products and was put on support using contralateral groin percutaneous transluminal angioplasty (pta) ballooning to remedy the issue; besides rt iliac pta and manual compression no other intervention was necessary or performed to achieve hemostasis and the patient became stable. the pci of the left main (lm) was then performed. an echo confirmed that correct impella cp cannula positioning had been achieved and right ventricle (rv) contractility was also confirmed. no hematoma/bleeding occurred and distal pulses were confirmed. the patient support was reported as successful, there were no further adverse effect to the patient as a result of the oscor sheath splitting and ventricle perforation. the patient subsequently expired, the facility's heart team staff does not believe the impella cp contributed to the patient expiring, since the bleeding was controlled and blood was given to replace loss. the physician related that the patient outcome was most likely due to the right coronary low flow state post pci, which was performed prior to the initial implant attempt and incident. | [5 empty] | The 14frencnh introducer and the dilators used with the impella cp pump during this event were returned for analysis. upon receipt of these devices an initial inspection was performed. this inspection revealed multiple splits along the length of the introducer sheath; in addition, there was noticeable damage to each of the dilator tips. the sheath was examined under magnification and confirmed that the splits occurred along score-lines. the dilators were also examined under magnification, and the magnification revealed that each dilator tip was nicked outward. this type of damage is indicative of calcification/peripheral vascular disease (pvd) occluding the insertion path. a review of the patient clinical notes written by the physician confirmed pvd. the pvd likely caused the sheath splitting. the introducer "pacekenged" into this kit had been inspected and had passed all inspection criteria. in conclusion, the root cause of the patient bleeding was a result of the splitting that occurred along multiple score-lines of the introducer, and that the patient's pvd likely contributed to the sheath splitting. a corrective action has been initiated to address 14fr introducer sheath failures of this kind. (b)(4).
38229 | 1220648-2010-00030 | 6535496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6535496&pc=OZD | 04/04/2017 | 04/04/2017 |  | 05/01/2017 |  |  | 04/04/2017 | 05/01/2017 |  |  | 09/30/2018 |  | 05/01/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET, US |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1263038 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 86 YR |  |  | 65 |  |  | Injury | Yes | No | Appropriate Term/Code Not Available | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on the morning of (b)(6) 2017 an (b)(6) male patient was admitted to the facility to have an elective high-risk percutaneous coronary interventions (hrpci.) the patient exhibited a 20% ejection fraction. prior to stent placement an impella cp was placed. the impella cp was successfully placed per protocol. the physician then performed stent placement in the patient's right coronary artery, left anterior descending artery, and left main artery. during the left main manipulation there was loss of contractibility. the doctor also performed a balloon angioplasty on the patient's circumflex and obtuse marginal arteries. the repositioning sheath was inserted following the oscor sheath removal. the device remained in the patient to support the patient's diminished kidney function and to help unload the left ventricle during recovery. the patient was brought to the icu after the intervention. while in the icu, the patient coded, lost pulsatility and went into ventricular fibrillation, and his blood pressure dropped. pressors were administered and cpr was started. the patient's heart returned to a normal sinus rhythm and flows were restored. the physician reported that the echo that was performed on the patient revealed very minimal right heart and left heart function, and that the patient had severe mitral regurgitation. that evening the patient coded again and received chest compressions. later that night the patient's urinary output had turned pink. during the 25 hours of support the patient experienced many cardiac events, including atrial fibrillation. in response to these cardiac events cpr was performed, the patient was defibrillated, and cardioverted. minimal signs of hemolysis / very light pink urine was observed. an echo was performed and showed good pump position. there was one event of a suction alarm due to a central venous pressure drop from 12-8; and was adequately corrected and resolved. volume and flow improvements were attained without suction alarms or diastolic motor current waveform indicators of suction. there were no diagnostic tests performed to confirm hemolysis. there was only clinical diagnosis of pink urine and a drop in hemoglobin values. patient developed distended abdomen overnight and a ct was performed prior to removal of the impella cp to rule out the possibility of a bleed. the ct showed no vascular component related to abdominal distention. on (b)(6) 2017 the patient was weaned and the impella cp was removed. in response to the suspected hemolysis the patient was administered 7 to 8 units of packed red blood cells. on (b)(6) 2017 the patient placed on palliative care, and did expire on that same day. | [5 empty] | The device was not returned for evaluation as it had been discarded subsequent to the event. the automatic impella console data logs were returned for analysis. in addition, echo images taken during patient support were also returned for analysis. the analysis of the data logs revealed a number of suction alarms had occurred during patient support. the echo images were reviewed by the abiomed staff physician. the images revealed that the device was not positioned correctly. the inlet cage seen to be pressed up against the interior leaflet of the mitral valve. in addition, and as reported in the event description portion of this event, the patient had class iv mitral valve regurgitation. poor positioning and mitral valve regurgitation can increase the amount of shear stress experienced by red blood cells. in conclusion, the root cause of this event indicates that the patient's severe mitral regurgitation and poor pump placement most likely contributed to the hemolysis observed in the case; however, the root cause is unable to be definitively determined because the device was not returned for evaluation. as a corrective action customer retraining on recommended cannula placement during impella cp support has been scheduled to be performed with the staff at this facility during the week of (b)(6) 2017. (b)(4). device discarded by user.
38230 | 1220648-2017-00026 | 6508725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6508725&pc=OZD | 03/21/2017 | 03/21/2017 |  | 04/20/2017 |  |  | 03/21/2017 | 04/20/2017 |  |  | 10/31/2018 |  | 04/20/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1265807 |  | 0048-0003 |  | No | [5 empty] | No | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 54 YR | [5 empty] | Malfunction | Yes | No | Patient-Device Incompatibility | Device-Device Incompatibility | Missing Value Reason |  |  |  |  | Ischemia | Thrombus | [11 empty] | Required Intervention | [5 empty] | ECMO | IABP | [21 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A (b)(6) male suffered cardiac arrest while playing hockey, had an aed placed and cpr was performed. the patient underwent a pci of his left anterior descending artery, and an iabp was placed. the patient continued to decompensate, requiring an escalation of inotropes and vasopressors; and was experiencing intermittent episodes of vtach. the patient was brought to the cardiac cath lab with the plan to remove the iabp and place an impella cp. the iabp was removed and the vessel was dilated with 10 fr, 12 fr and 14fr dilators. the 14fr oscor sheath was unable to advance smoothly into the vessel, therefore the 14 fr oscor introducer was exchanged for a 14 fr long cook sheath. the impella cp was inserted on (b)(6) and flow was initiated. the physician performed a peripheral angiogram to assess distal limb flow. the patient was reported to be hypotensive, with unstable pulsatility, and required blood products during support. upon admittance to icu on the evening of (b)(6) the pump was repositioned using tte. the patient was transferred to a new facility later in the evening. according to the abiomed field representative the patient's antigoagulation was subtherapeutic at times. the patient was administered integrilin, in addition to heparin in the impella purge solution. the patient arrested on the (b)(6) and was placed on ecmo; and the impella performance level was decreased to p-level 2. the patient developed a hematoma at the ecmo, left groin insertion site on the (b)(6) and the impella insertion site was reported to have oozing that resolved. the patient stabilized over the next several days, and was taken off ecmo on the (b)(6). the patient remained on impella support for several more hours after which the impella was removed through a 14 fr cook sheath. upon removal there was visualization of thrombus on the pigtail. a thrombus was reported to obstruct flow to the extremity, until an embolectomy was performed on the 23rd; at which point flow was returned. the clinical team stated the impella saved the patient's life. there was no indication of any further adverse effect to the patient following the embolectomy. | [5 empty] | The console data was returned to the manufacturer, however the pump was not returned for evaluation. consequently, a device analysis was unable to be performed. the data logs revealed erratic flows while the pump was operated at p-level 2, however there was no evidence of biomaterial being ingested by the pump. according to the clinical team's account during the case, anticoagulation was sub-therapeutic at times, and the patient was receiving integrilin. the thrombus noted on the pigtail during explant was likely picked up when the pump was explanted through the 14fr long cook sheath. the clinical presentation of leg ischemia occurred immediately thereafter, including limb swelling and loss of pulse. the limb ischemia was resolved by surgical embolectomy. due to the sub-therapeutic anticoagulation it is likely the sheath which had been left inserted for six days, clotted off during support. the root cause of the thrombus observed on explant and limb ischemia is most likely a result of improper anticoagulation, which was sub-therapeutic during support. the manufacturer will continue to investigate all reasonable and obtainable sources of information and will provide results and conclusions if additional information is received.
38231 | 1220648-2017-00024 | 6494382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6494382&pc=OZD | 03/16/2017 | 03/17/2017 |  | 04/14/2017 |  |  | 03/17/2017 | 04/14/2017 |  |  | 12/31/2018 |  | 04/14/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1272918 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 57 YR | [5 empty] | Injury | Yes | Yes | Physical Property Issue | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2017, a (b)(6) male suffering from an acute myocardia infarction and in cardiogenic shock was determined to be in need of impella cp support. an angiography was performed on the patient's right femoral artery using a 6fr sheath prior to 14fr oscor sheath insertion. the physician confirmed adequate vessel size and the absence of vessel tortuosity or calcification. the right femoral artery puncture location was considered appropriate. the vessel was sequentially dilated. according to the physician the 14fr oscor was inserted over the abiomed 0.035" guidewire without any resistance or difficulty. no unusual bleeding was observed and the impella cp was inserted through the oscor sheath without issue, the pump was started, and the oscor sheath was peeled away. the repositioning sheath was flushed and inserted also without difficulty, then sutured in place, and the catheter was secured. tegaderm was applied covering the entry site. no unusual bleeding or hematoma was noted by the cath lab staff. the physician reported that there was only a minor ooze which was expected. the patient was transferred to the icu. the patient's act at this time was 185. after being in the icu for an hour, a nurse noted moderate oozing around the repositioning sheath. manual pressure was applied for 30 minutes. no hematoma was noted. the ooze progressed to a mild bleeding, and the bleeding was controlled using a femostop. the nurse stated the pressure on the femostop controller was set to 20 mmhg below the patient's systolic pressure. approximately 30 minutes later the nurse reported high purge alarms, trouble shooting was unable to determine the causes for the alarms. it was then thought that the excessive femostop pressure could be causing an obstruction in the catheter, thus leading to high purge pressure. the nurse repositioned the femostop, and the high purge pressure alarm was resolved. after repositioning the femostop, mild bleeding returned requiring more adjustments. the nurse was able to stop the bleeding, and stated that the patient received one unit of replacement blood, no hematoma was seen. the following morning the patient was transferred to another hospital to be evaluated for a possible heart transplant. upon arrival active bleeding was noted from the impella site. the physician and nurse related that bleeding was noted by the transport team prior to the patient leaving the first hospital. the estimated blood loss was reported to be at least 2 liters. the staff at the transfer facility attempted to remove the femstop, but the bleeding became pulsatile and excessive. the physician ordered that the anticoagulant be reversed, and after two hours the femstop was removed. bleeding was still an active pulsatile bleed and was then controlled with manual pressure. after 15 minutes of manual pressure, the doctor lightened pressure and bleeding resumed. the impella site was assessed and the repositioning sheath was found to be in the proper positon, sutured with no acute angles or issues. the femstop was placed back on the patient. vascular surgery and cardiology were both consulted and the decision was made to remove the impella cp. the pump was then successfully weaned and removed from the patient. the physician held pressure and hemostasis was able to be achieved. patient support was deemed successful; there were no additional adverse effects to the patient reported. | [5 empty] | The impella cp at issue in this event was returned to the manufacturer for evaluation; although the automatic impella controller data logs were also returned for analysis, data log review was not performed as a review was not relevant to this failure investigation. during the evaluation of the returned impella cp a tear was found on the end of the 9fr vygon sheath. no other issues or deformities were found with the pump. a device history record (dhr) review was completed. the review revealed that there were no issues or rework relevant to the failure mode and none were documented on the dhr. the 9 french repositioning vygon sheath was part of batch 7052802, and it passed incoming inspection. those units had also been visually and geometrically inspected. all inspected units passed inspection. in addition, a review of the complaint data base revealed no other complaints have been filed for other pumps from lot number 1272918 or for other 9 fr vygon sheaths from batch 7052802. (b)(4).
38232 | 1220648-2017-00022 | 6487725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6487725&pc=OZD | 03/13/2017 | 03/13/2017 |  | 04/12/2017 |  |  | 03/13/2017 | 04/12/2017 |  |  | 07/31/2018 |  | 04/12/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPLELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1253197 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 62 YR | [5 empty] | Death | Yes | Yes | Pumping Stopped | [6 empty] | Death | [12 empty] | Death | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | On (b)(6) 2017 a (b)(6) year old female patient was admitted to the hospital with chest pain and an atrial fibrillation heart rhythm. upon admission, the patient suddenly decompensated and became hypolic. the patient's ejection fraction was 20%, and she was in pulmonary edema. a left heart catheterization (lhc) was performed which revealed 90% left main disease. the patient left ventricular end diastolic pressure (lvedp) was at 42. an impella cp was placed for patient support successfully and without issue. the following day it was reported that the patient was suffering from a gi bleed and a "neck site" bleed. it was confirmed by the site that the bleeds were not related to impella. the patient's condition began to improve over the next few days as the doctor decide to rest the patient's heart over the weekend prior to performing a coronary artery bypass graft (cabg.) on the second day of patient support, bleeding was seen from the central venous pressure line and from a gi bleed. this bleeding was attributed to stress, and required blood transfusing. the bleeding was reported not to have been impella related. on (b)(6) 2017 a "placement signal" alarm occurred after the patient was slightly turned. the alarm resolved after the patient was turned back to the previous resting position. after 110 hours of successful patient support, the impella cp spontaneously stopped. multiple unsuccessful restart attempts were made, but the patient's status began to decline. the hospital staff was advised by abiomed support that the pump needed to be removed and replaced as soon as possible; however, the patient coded and expired approximately 45-50 minutes after the pump stopped, and before the patient was able to be transferred back to the cardiac catheterization lab. | [5 empty] | The impella cp and the automatic impella controller data logs were returned to abiomed for evaluation. the analysis of the data logs revealed that the pump stopped after 110 hours of successful patient support, but was not able to be restarted. the analysis further revealed that the initial pump stop was due to high motor current, and the following pump stops were due to high motor current and rpm deviation. the impella cp was returned intact. during the evaluation of the device the motor was disassembled, and it was found that there was visible wear to the motor bearing. this bearing wear was the result of 110 hours of impella cp use. the impella cp is indicated for 4 days or 96 hours of patient support. in conclusion, the root cause of the pump stop was bearing wear. there is no corrective action recommended as the failure occurred beyond the indicated use for the impella cp. the instructions for use and clinical reference manual for the impella cp states the following: the ... impella cp ..., in conjunction with the automated impella® controller, are temporary ventricular support devices intended for short term use (< 4 days for the ... impella cp, ... and indicated for the treatment of ongoing cardiogenic shock that occurs immediately (< 48 hours) following acute myocardial infarction or open heart surgery as a result of isolated left ventricular failure that is not responsive to optimal medical management and conventional treatment measures. the intent of the impella support systems therapy is to reduce ventricular work and to provide the circulatory support necessary to allow heart recovery and early assessment of residual myocardial function. | The supplemental medwatch report is being submitted to user mdr (b)(4) that was received by this manufacturer on 04/12/2017. (b)(4).
38233 | 1220648-2017-00021 | 6475003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6475003&pc=OZD | 03/11/2017 | 03/11/2017 |  | 04/10/2017 |  |  | 03/11/2017 | 04/10/2017 |  |  | 04/30/2018 |  | 04/10/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1243292 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 38 YR | [5 empty] | Injury | Yes | No | Malposition of Device | [6 empty] | Bradycardia | Low Blood Pressure/ Hypotension | Perforation | Blood Loss | [9 empty] | Required Intervention | [5 empty] | INTRA AORTIC BALLOON PUMP | THE .018 V18 WIRE BY BOSTON SCIENTIFIC | [21 empty] | Initial submission | [5 empty] | * | HOSPITAL |  |  | On (b)(6) 2017 a (b)(6) male was hiking and arrested. his family performed cpr as he was down in the field awaiting the ems team. upon the arrival of the ems team, the patient had 8 rounds of epinephrine and continued cpr. while being transported to the hospital the patient continued to code and the neurological status was not known. the patient was admitted to the cardiac cath lab and a temporary pacer was placed and angioplasty to the left anterior descending artery was performed. this stabilized the patient enough for the team to place an impella cp. the cp was placed over the v18 wire for additional support, rather than the abiomed provided .018 wire. the 018 wires, recommended as alternatives by abiomed, in the ifu are: boston scientific platinum plus¿ st 0.018 in, boston scientific v-18 control wire¿ st 0.018 in. even with this level of cardiac support with the impella, he continued to deteriorate and had runs of ventricular fibrillation and needed many and varied acls iv drips. it was discovered that the patient had a large effusion. the cardiac surgery team performed a pericardiocentesis and pericardial window in the cardiac cath lab. the patient's condition did not improve and so he was placed on cardiopulmonary bypass. the surgeon did observe blood coming from the left ventricle and so blood products were given and an intra-aortic balloon pump (iabp) was placed, in addition to the impella, for support. the thought was that the .018 wire had caused the ventricular perforation. the patient was transferred to the icu for monitoring and impella explantation on the (b)(6). the patient did expire on the (b)(6) while on cardiac support with the iabp. | [5 empty] | Neither the impella cp pump nor the .018 wire were returned for evaluation. the product was discarded subsequent to the event. the data log files were not deemed relevant to the failure investigation. the cp pump serial number and lot were reviewed and no other complaints have been filed against this lot of (b)(4) built pump. the patient's left ventricle was most likely perforated by the v18 guidewire. what led to the wire puncturing the ventricle is unable to be determined because the device was not returned for review. no corrective action is warranted as the root cause could not be determined. the manufacturer will continue to investigate all reasonable and obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. (b)(4).
38234 | 1220648-2017-00027 | 6460112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6460112&pc=OZD | 01/12/2017 | 01/12/2017 |  | 04/05/2017 |  |  | 01/12/2017 | 04/05/2017 |  |  | 09/30/2018 |  | 04/05/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1260599 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 81 YR | [5 empty] | Injury | Yes | No | Human-Device Interface Problem | [6 empty] | Hematoma | Low Blood Pressure/ Hypotension | Perforation | Blood Loss | [9 empty] | Hospitalization | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | An impella cp was placed on (b)(6) 2017 to support an (b)(6) female undergoing a complex valvuloplasty. the patient had many comorbidities inclusive of multi-vessel coronary artery disease. upon insertion of the cp there was a slight bloody ooze when the oscor sheath was removed and the repositioning sheath was placed. with propping up of the repo-sheath the ooze was remedied. at some point on the (b)(6) the patient coughed and a clot was dislodged from the access site. this clot was indicative of a hematoma at the groin access site in the left femoral artery. the plan of care was to leave the cp in for support for several days due her complexity and medical comorbidities, despite the groin ooze and hematoma development. on the second day of support, the (b)(6) , there was a point where the patient was rolled in the bed and the groin ooze began again and was only remedied with propping up of the repo-sheath again. on the (b)(6) there was also a noted drop in the patient's hematocrit and hemoglobin (h&h) levels. there was a present nose bleed that was not able to account for such a drop. the h and h was 6 and 17. the team chose to transfuse 2 units of blood during the evening and 3 units in the overnight to remedy the h&h drop. the patient's condition continued to deteriorate. the physician diagnosed a retro-peritoneal bleed at the left groin, and chose to treat the patient for it on (b)(6) by a placing a covered stent in the left femoral via the opposing leg access and deploying the stent in an "up and over" method. the cp was explanted and the patient was monitored in the icu. the patient did expire the following week due to multi organ failure. | [5 empty] | The data logs were reviewed for the investigation of this event. they confirm that the patient's blood pressure dropped on the second day of support. prior to that drop the pump was supporting the patient successfully with good flows and pressures. there were 15 minutes in which many suction alarms were triggered and then resolved. the pump itself was not returned for analysis despite many attempts at retrieving it. the root cause of the left femoral artery bleed and retroperitoneal hemorrhage was unable to be determined as no product was returned. the manufacturer is filing this report late due to an oversight. the manufacturer will continue to investigate all reasonable and obtainable sources of information and will provide a supplemental medwatch report if additional information is received. (b)(4).
38235 | 1220648-2017-00023 | 6460984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6460984&pc=OZD | 03/08/2017 | 03/10/2017 |  | 04/05/2017 |  |  | 03/10/2017 | 04/05/2017 |  |  | 12/31/2018 |  | 04/05/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  | 1271503 |  | IMPELLA CP |  | No | [5 empty] | No | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 73 YR | [5 empty] | Malfunction | Yes | No | Entrapment of Device | Difficult To Position | Use of Device Problem | Malposition of Device | Human-Device Interface Problem | Positioning Problem | Device Handling Problem | Perforation | Blood Loss | [11 empty] | Required Intervention | [5 empty] | ABBOTT PERCLOSE PROGLIDE X2 | BOSTON SCIENTIFIC ROTABLATOR | [21 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A (b)(6) male with severely calcified left anterior descending coronary and circumflex arteries presented to the hospital with 30% ejection fraction, hypertension and multi-vessel coronary disease. the patient was to undergo a roto blade and pci procedure while under the support of impella. the patient had pre- and post- procedural transesophageal echocardiograms, while on epinephrine and levophed. the patient's blood pressure was not responding to these medications. the impella cp was successfully inserted through the left groin using a 14fr cook sheath. after insertion of the impella the pump was repositioned, and came out of position from the left ventricle. the clinical team attempted to recross the aortic valve approximately ten times. during the attempts at repositioning, the pump became entangled with a papillary muscle. when the pump was successfully repositioned the patient began to decompensate. abiomed recommends using fluoroscopy to ensure the pump is properly positioned. the impella ifu states "fluoroscopy is required to guide placement of the impella catheter. the small placement guidewire must be reliably observed at all times." a transesophageal echocardiogram (tee) revealed the patient's myocardium was extremely thin with visible scarring. the tee revealed a pre-existing infarct, as well as pericardial effusion originating from the apex. there was no perforation or dissection of the aorta seen on angiogram. the patient was immediately tapped and 200 ml was aspirated from the effusion. the patient was weaned from impella support and the pump was explanted. the pci procedure was aborted and a pericardial window was performed instead. a total of approximately 800 ml of fluid was drained from the patient's pericardium. the 800 ml was auto-transfused back to the patient. the patient was extubated later the same day and reported to be stable without lasting harm or injury. | [5 empty] | The aic console data was returned for investigation, however the pump was discarded at the end of the clinical case. the console data revealed the pump was out of position during the beginning of the case. the position alarms resolve, however the patient's blood pressure subsequently drops with a simultaneous decrease in pump flow. the blood pressure then rises which is presumably when the clinical team aspirated the first 200 ml of blood from the effusion. the data logs show the patient's blood pressure stabilizes after the tap was performed, and reveals the pump was explanted approximately one hour later. a conversation between engineering and the abiomed field representative revealed the pump was being repositioning without the use of a guidewire, and the impella cp repeatedly became entangled with the papillary rather than settling in the apex. the field representative also reported the three physicians involved in the case reviewed the clinical images and case, and concluded they believed the reported tamponade was due to the patient's exceptionally thin myocardium combined with repeated prodding of the papillary tissue with the pigtail during the attempts to reposition the pump. the root cause of the tamponade and pericardial effusion is most likely a combination of the multiple attempts at wireless repositioning of the pump. this failure mode is considered to be low risk, and events of this type will continue to be monitored and trended. internal reference: (b)(4). device not returned.
38236 | 6449285 | 6449285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6449285&pc=KFM | 03/13/2017 | 03/16/2017 | 03/16/2017 | 03/31/2017 | 03/16/2017 | 03/16/2017 |  | 03/31/2017 |  |  |  |  | 03/31/2017 |  | CARDIOPULMONARY BYPASS SYSTEM |  |  | IMPELLA CP |  |  | KFM | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Device was returned to manufacturer | [5 empty] | No | NURSE |  |  | User Facility report | [8 empty] | 62 YR | [5 empty] | Death | Yes | No | Failure to Reset | Device Displays Incorrect Message | [5 empty] | Death | Ventricular Tachycardia | Loss Of Pulse | [10 empty] | Death | [28 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | A female cardiac patient was implanted with a percutaneously inserted ventricular assist device. a few days later, the lvad system console presented a "purge alarm" message instructing user to replace the purge cassette. purge pressure at this time was within normal limits (wnl). prior to cassette exchange patient was tilted with a pillow when console was presented a "position warning alarm" which self-resolved and displayed a "confirming this resolution" message. staff confirmed the resolution by noting the motor waveform, purge pressure, and placement waveforms were all wnl. shortly thereafter the console display alarmed "impella stopped" and would not respond to attempts to restart. super-user contacted manufacturer for assistance and was advised to re-boot as last part of troubleshooting. patient's pressure began to drop. stat physician consult and echo request ordered. echo imaging noted adequate impella position and physician slightly repositioned with no effects on impella performance or re-start. while scheduling stat cath lab procedure to replace impella, patient went into pulseless v-tach and expired after extensive resuscitation efforts. impeller device released to on-site manufacturer clinical consultant and impeller controller/console sequestered for biomedical logging and evaluation. manufacturer response for ventricular assist device, impella cp (per site reporter): manufacturer will report and disclose all findings and testing results to our facility. manufacturer response for impella controller, impella controller (per site reporter): manufacturer in process of scheduling inspection and analysis of controller function and retrieval of logs.
38237 | 1220648-2017-00025 | 6443433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6443433&pc=OZD | 03/03/2017 | 03/20/2017 |  | 03/29/2017 |  |  |  | 03/29/2017 |  |  | 11/30/2018 |  | 03/29/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP, PRODUCT CODE: OZD |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1268221 |  | 0048-0003 |  | Yes | [5 empty] | No | No Information |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | Device Operates Differently Than Expected | [5 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | WIRE AND BALLOON WERE USED TO TAMPONADE FLOW | [22 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | A (b)(6) obese male presented to the hospital with shortness in breath and changes in a recent ekg reading. the patient was brought to the cardiac cath lab for a coronary exam and was evaluated for a possible pci. the patient was hypotensive, tachycardic, in a state of cardiogenic shock and not responding to inotropes when the impella cp was inserted. the patient was stabilized following the impella insertion and a pci was completed on the proximal left anterior descending artery and mid-circumflex artery. the patient developed a hematoma following the exchange for the repositioning sheath, and manual pressure was applied. an angiogram through the contralateral groin revealed the patient had a very long insertion path, and the repositioning sheath was entering the vessel at a 70-90 degree angle. the ifu recommends the hub of the repositioning sheath should be at skin level. a wire and balloon was placed to tamponade flow and the patient was brought to the or for surgical removal. the patient was slowly weaned off of support over the course of one hour. the device was removed and the vessel was repaired. the clinical team stated the impella saved the patient's life. he was discharged and is currently at a rehab facility. there was no indication of any further adverse effect to the patient following the surgical removal of the device. | [5 empty] | As part of our normal process of reviewing the complaint file associated with this mdr it was discovered that all of the event problem and evaluation codes were not entered. the initial mdr included the patient code, and the supplemental mdr incudes the device, method, result and conclusion codes. | The device and aic console data were returned for evaluation. the console data does not provide insight into the reported hematoma event. the pump was returned with the repositioning unit cracked and kinked from the tip along the length of the repositioning sheath. the damage observed to the device is most likely the result of the patient's size, which necessitated a steep insertion angle. the impella ifu gives the following warning regarding the use of the device with an obese patient: when securing the repositioning sheath, vascular closure may be difficult in obese patients with extensive adipose tissue. the root cause of the vascular injury is repositioning sheath damage that was most likely caused by the steep insertion angle. a corrective action has been implemented to address the design of the repositioning unit, to reduce the occurrence of this failure mode. this failure mode is considered to be of very low occurrence, and events of this type will continue to be monitored and trended. (b)(4).
38238 | 6431656 | 6431656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6431656&pc=OJE | 03/02/2017 | 03/08/2017 | 03/08/2017 | 03/24/2017 | 03/08/2017 | 03/08/2017 |  | 03/24/2017 |  |  | 03/31/2017 |  | 03/24/2017 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA RP |  |  | OJE |  |  | 004334 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 004334 |  | Yes | [5 empty] | No | Physician |  |  | User Facility report | [14 empty] | Malfunction | No | Yes | Device Operates Differently Than Expected | [6 empty] | No Known Impact Or Consequence To Patient | [18 empty] | NO | NO OTHER THERAPIES | [21 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The impella rp device malfunctioned per the dr. and the product device representative.
38239 | 1220648-2017-00018 | 6424216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6424216&pc=OZD | 02/17/2017 | 02/21/2017 |  | 03/22/2017 |  |  | 02/21/2017 | 03/22/2017 |  |  | 11/30/2018 |  | 03/22/2017 |  | IMPELLA CP PUMP SET, US |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1269856 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Kinked | Device Issue | Patient-Device Incompatibility | Human Factors Issue |  |  |  | Blood Loss | [12 empty] | Required Intervention | [5 empty] | FEMOSTOP | INTRA AORTIC BALLOON PUMP | [21 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | An obese (b)(6) male was taken to the cardiac cath lab , and found to have a dissection of the right coronary artery, as well as multivessel disease. the patient was unstable and in need of 3 stents to his rca and an intra aortic balloon pump (iabp) was placed on (b)(6) 2017. the iabp was not sufficient support, and so care was escalated and an impella cp was placed. there were unconfirmed reports that the oscor sheath was cracked/removed while in the vessel, prior to placement of the repositioning sheath and transfer to the icu. the ifu states: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." in addition the ifu states: "when securing the repositioning sheath, vascular closure may be difficult in obese patients with extensive adipose tissue." the pump was seen to jump forward and out of desired position. during this movement, the pump kinked. the cp pump was explanted and a second cp was placed in the opposing leg. the patient was transferred to the icu while on cp support. the patient was noted to be unstable and receiving blood replacement products. in total the patient received 5 units of blood due to ooze and profuse bleeding at the groin access sites. the patient was then transferred to a second, tertiary, medical center for monitoring and care. at the second medical center the team applied a femostop to the groin site and consulted vascular surgery. the femostop pressure was released during cp repositioning in the left ventricle. with the pressure released from the femostop there was blood lost and 3 more units of blood were given. on the following day the bleeding had reduced and the patient was noted to be stable. the support continued for six days allowing cardiac iv meds to be discontinued and the patient to recover from the cardiogenic shock. the patient's ejection fraction rose from 10% to 30%. the team was also able to complete the coronary artery revascularization. the circumflex and ramus coronary arteries were stented with 4 stents in total. on the (b)(6) the impella cp was weaned and explanted. this was the 10th day of support. | [5 empty] | The first impella cp, which kinked with movement, was returned for evaluation. the first device was free of any damage that would have contributed to the access site bleeding. the cannula was visibly kinked near the outflow cage. the second impella cp was not returned for evaluation. it is unclear at this time if the bleeding was as a result of any issue with the introducer or impella cp pump. the necessary product was not available for review so the root cause of the access site bleeding was unable to be determined. the data logs were analyzed from the first cp pump. the data logs for the second device were not returned. upon review of the lot of cp pumps, there were no other complaints for this pump, nor any reworks that apply to the failure mode. there is no corrective action recommended because the root cause was unable to be determined. the failure will continue to be tracked and trended. should the manufacturer obtain any other sources of information, a supplemental medwatch will be filed with investigation results. (b)(4).
38240 | 1220648-2017-00016 | 6409047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6409047&pc=KFM | 02/21/2017 | 02/21/2017 |  | 03/16/2017 |  |  | 02/21/2017 | 03/16/2017 |  |  | 04/30/2017 |  | 03/16/2017 |  | IMPELLA 5.0 AIC PUMP SET |  |  | IMPELLA 5.0 |  |  | KFM |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1193898 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 77 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | Blood Loss | [12 empty] | Required Intervention | [5 empty] | 10 X 30 HEMASHIELD GRAFT | INTRA AORTIC BALLOON PUMP | [21 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient was a (b)(6) female post cabg who was not thriving and in need of escalated care. her care needed escalation due to an ef of 20%, multiple cariac iv drips, an iabp, and a cardiac index of 1.3. the patient was prepped for insertion of a 5.0 impella via the right axillary artery via the placement of a 10 x 30 graft. the 5.0 would not advance much beyond the anastomosis. attempted placement of the 5.0 continued for 45 minutes, when insertion was aborted and an impella cp was inserted instead. the cp also had difficulty inserting though the axillary graft and was not able to fully advance through the axillary. after 2 hours of attempted placement, a surgical cut down was made to the left axillary artery, and the 5.0 placement attempted again. once again, it would not pass through the anatomy, and so the cp was placed instead through the left axillary cut down. within a brief amount of time on impella cp support the cardiac index rose to 1.9 and on day two of support it rose again to 2.8. due to the frequent, and lengthy insertion attempts, there was over 1000ml of blood lost. some of the volume was re-transfused with the cell saver but, due to the drop in hematocrit and hemoglobin, the team transfused 5 units of blood and 1 unit of platelets. the cp supported the patient successfully. the iabp had been removed during cp support, and the patient was able to sit up and verbalize. | [5 empty] | The investigation of this complaint was inclusive of the product returned, but not the data logs as they were not relevant to the investigation. the 5.0 pump was returned and upon observation many kinks were found throughout the length of the cannula. in addition, the cannula was found to have delaminated. the pigtail was found to be damaged and bent in the wrong direction.the engineer found that there was no issue tracking the pump over an .018 guidewire. the pump damage observed was indicative of a very difficult insertion. however, no feature or deformity was seen on the pump that would have caused the difficult and ultimately unsuccessful insertion of the pump. the root cause of the inability to advance the 5.0 past the anastomosis could not be determined as the graft and guidewire were not returned. the failure mode could not be reproduced. there were no issues or reworks relevant to this failure mode. no other complaints have been filed for other pumps from this lot manufactured in (b)(4). due to the fact that the root cause could not be determined, there is no corrective action recommended at this time. the failure mode will continue to be monitored and tracked. internal reference (b)(4).
38241 | 1220648-2017-00017 | 6409048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6409048&pc=OZD | 02/22/2017 | 02/22/2017 |  | 03/16/2017 |  |  | 02/22/2017 | 03/16/2017 |  |  | 06/30/2017 |  | 03/16/2017 |  | IMPELLA 2.5 AIC PUMP SET |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1250869 |  | 005042 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | Fracture | Human-Device Interface Problem | [5 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [5 empty] | PERCLOSE | [22 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient was a (b)(6) male admitted with multivessel coronary artery disease. the surgical option presented to the patient was a lima graft. the patient declined this single vessel graft and proceeded to have a multivessel angioplasty. the left main, circumflex, and right coronary arteries were stented with support on board in the form of an impella 2.5 pump. there was noted difficulty inserting the introducer initially and to accommodate the 13fr introducer the skin knick was expanded. the patient habitus necessitated a 45 degree insertion angle of the 13 fr introducer. when consulting the ifu for the impella 2.5 there is a mentioned: "figure 5.14 inserting the peel-away introducer" which is not a 45 degree angulation. when the pump was weaned and explanted in the cath lab, a large hematoma was noted at the groin access site. upon the sheath being pulled back the physician noted a crack in the 13 french sheath. the sheath was removed and hemostasis was attempted by the placement of a perclose. the perclose failed and due to the bleeding, the team performed a contralateral balloon tamponade and administration of protamine. no further surgical intervention was needed. no blood products were given. with manual pressure the groin resolved. | [5 empty] | The data logs were not returned for analysis, however the 13fr product was; both the introducer and dilator. the data logs would not have been relevant to this failure mode. upon visual inspection, there was a crack found in the sheath approximately 39mm from the hub. several kinks were also found along the length of the sheath at various locations. the damage to the introducer was most likely caused during difficult insertion caused by the patient's size. of note in the ifu for the 2.5 impella there is a pre-support evaluation statement to observe for "obesity". all introducers inspected from this lot of manufacturing passed inspection. the root cause of the crack was most likely the difficult insertion caused by the patient size. no corrective action is recommended as this failure mode is low risk due to moderate severity and very low occurrence. tha manufacturer will continue to monitor and investigate all reasonable and sources of information regarding this failure mode, and the mode will continue to be tracked and monitored. (b)(4).
38242 | 1220648-2017-00015 | 6411132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6411132&pc=OZD | 02/15/2017 | 02/15/2017 |  | 03/16/2017 |  |  | 02/15/2017 | 03/16/2017 |  |  | 05/31/2018 |  | 03/16/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP, PRODUCT CODE: OZD |  |  | IMPELLA CP |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  | 1246851 |  | IMPELLA CP |  | Device was returned to manufacturer | [5 empty] | Y | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Malfunction | Yes | No | Material Deformation | Torn Material | [5 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A (b)(6) female with an ejection fraction of less than 20% presented to the cardiac cath lab in post arrest and in advanced cardiogenic shock. the patient had been administered cpr for 45 minutes. the impella cp was placed using wired insertion through the right femoral artery on (b)(6) 2017. the patient's act was reported to be 220. (the impella cp ifu recommends that after insertion of the catheter (and until explant), act should be maintained at 160 to 180 seconds.) the repositioning sheath was sutured in place and correct positioning was confirmed. there were no signs of bleeding or hematoma upon the patient's arrival to the icu, approximately two hours after implantation. after approximately five hours of support, blood was noted to be leaking from around the repositioning sheath at the insertion site. the patient was reported to have lost approximately 600 ml of blood and received 3 units of whole blood replacement product. the bleeding was unable to be controlled with manual pressure and the decision was made to explant the device. the patient was successfully weaned off of support. there was no indication of any further adverse effect to the patient following the surgical removal of the device. | [5 empty] | The device and aic console data were returned for evaluation. the data logs revealed a suction and "position unknown" alarm, however the logs show the pump provided good support otherwise. the console data does not provide additional insight into the reported bleeding event. the pump was returned with a tear in the wound closure, and a portion of the wound closure was slightly peeled away. additionally, there was a scratch mark on the surface of the wound closure that lined up with the tear. the root cause of the reported bleeding from the insertion site is likely a tear in the wound closure component of the repositioning unit. historically we have seen this issue when the repositioning unit can cause damage to the femoral artery which can lead to bleeding as seen in this case, however there is no evidence to support that this is a definitive root cause. a review of the device history record revealed there were no other issues reported for any other device with this lot number. the design and the manufacturing process has been modified to eliminate this issue with the wound closure. the manufacturer will continue to investigate all reasonable and obtainable sources of information and will provide results and conclusions if additional information is received. (b)(4).
38243 | 6405546 | 6405546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6405546&pc=KFM | 02/28/2017 | 03/03/2017 | 03/03/2017 | 03/15/2017 | 03/03/2017 | 03/03/2017 |  | 03/15/2017 |  |  |  |  | 03/15/2017 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS |  |  | IMPELLA |  |  | KFM | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Yes | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 74 YR | [5 empty] | Malfunction | No | Yes | Device Displays Incorrect Message | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The low purge pressure alarm on the impella device could not be resolved. the impella was removed from the patient and a new impella device was inserted with no additional issues reported. the manufacturer's representative was present in the operating room (or) and aware of the issue. there is no known patient injury.
38244 | 6398618 | 6398618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6398618&pc=KFM | 11/23/2016 | 02/16/2017 | 02/16/2017 | 03/13/2017 | 02/16/2017 | 02/16/2017 |  | 03/13/2017 |  |  |  |  | 03/13/2017 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA CP |  |  | KFM |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No | [5 empty] | No | Physician |  |  | User Facility report | [8 empty] | 59 YR |  |  | 102 |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | An impella pump was implanted during a catheterization procedure, with the patient experiencing a cardiac arrest at the time of implantation. an augmentation of co was not obtained. the device was removed and an intra-aortic balloon pump (iabp) was implanted. the patient was not harmed. i only learned of the issue with the pump recently. i understand the pump was returned to the field rep for assessment via the company's qi process. unfortunately i do not have details of where and when that occurred. i have not been contacted with follow up.
38245 | 1220648-2017-00019 | 6401187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6401187&pc=OZD | 02/15/2017 | 02/15/2017 |  | 03/13/2017 |  |  | 02/15/2017 | 03/13/2017 |  |  | 11/30/2017 |  | 03/13/2017 |  | IMPELLA 5.0 AIC PUMP SET US |  |  | IMPELLA |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1222121 |  | IMPELLA 5.0 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Human-Device Interface Problem | Compatibility Problem | [5 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [5 empty] | GELSOFT BY VASCUTEK/TERUMO | [22 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A (b)(6) female with cardiomyopathy was admitted and had an impella 5.0 placed via a right axillary cut down and gel soft graft x10mm. this is not the abiomed recommended graft. as mentioned in the ifu for the impella 5.0 "10 mm diameter x 20 cm length hemashield platinum graft" is recommended. on day 2 of support the patient went to the operating room for an exploration of the graft site. there was no active bleeding but, a significant hematoma at the cut-down site. the area was evacuated, and a jp drain placed. on the following day 4 units of replacement blood products were given. during these days of support the patient was hypercoagulable and there was excessive anticoagulation with heparin being given. when the jp was no longer draining anything a planned hematoma excavation was done. on the 5th day of support the team broke the luer proximal to the pressure reservoir and attempted to trouble shoot the issue with placement of 19g needle proximal to the pressure reservoir. the support was able to continue with this fix and an intact purge system. unfortunately, the following day the patient kicked the purge system and the impella moved back within the lv. during repositioning, the sterile sleeve was torn and the physician cleaned the area with antimicrobial solution. the field rep consulted on pump replacement, but the decision was made not to replace due to the risk/benefit and history of bleeding. during support there were blood cultures drawn, which were negative for staph, and she was found to be positive for c-diff. she was also evaluated for sepsis and hit. the hit panel was negative. on day 11 of support, the patient had her 5.0 explanted in the surgical suite as she was having a heart transplant. | [5 empty] | The impella 5.0 was returned for analysis. the data logs were reviewed as well. the 5.0 pump was made in the (b)(4) facility and there were no other complaints made against the pump lot. the data log review shows a record of ac power disconnection and heparin level changes that are also mentioned in the clinical narrative. the patient had over anti-coagulation due to the patient's low weight. there were no defects found on the pump that would suggest that the pump contributed to the bleeding. the bleeding and hematoma that developed were at the access/graft site. the cause of the bleeding was unclear and unable to be determined. no corrective action is recommended due to the low risk failure due to moderate severity and very low occurrence. the failure mode will be tracked. the manufacturer will continue to monitor and investigate all reasonable and obtainable sources of information and will provide results and conclusions in a supplement med watch report, if any additional information is received. the root cause of the bleeding and hematoma was a bleed from the graft site and an over-anticoagulation. (b)(4).
38246 | 1220648-2017-00014 | 6397426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6397426&pc=OZD | 02/09/2017 | 02/13/2017 |  | 03/10/2017 |  |  | 02/09/2017 | 03/10/2017 |  |  | 09/30/2018 |  | 03/10/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1262371 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 59 YR | [5 empty] | Injury | Yes | No | Improper or Incorrect Procedure or Method | Device Handling Problem | [5 empty] | Hematoma | Ischemia | Blood Loss | [10 empty] | Required Intervention | [5 empty] | CARDIOPULMONARY BYPASS | ECMO | IABP | [20 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2017 a (b)(6) male patient with a history of (b)(6), hypertension, coronary artery disease (cad), chronic obstructive pulmonary disease (copd) and gastroesophageal reflux disease (gerd) presented to the emergency room with chest pain and shortness of breath. on the morning of (b)(6) 2017 while patient was in the ccl, the patient coded and an impella cp was placed. the patient was then removed to the operating room (or) for an emergent coronary artery bypass graft procedure in which three grafts were placed. the peel-away sheath was left in place by the intervention cardiologist, despite recommendations from the abiomed representative to remove it. several tries to remove the patient from the bypass, with the impella at a low setting were attempted, but the patient continued to decompensate. the impella cp was left in place in the patient while attempting to wean the patient from the cpb. the patient was eventually removed from cpb. multiple blood products (prbc's, ffp, cryo) were administered to the patient in the or. the patient was also coagulopathic from liver cirrhosis. heparin was not added to the purge as the patient was still bleeding and blood products were continued to be administered. the nurse noticed that the patient's right leg with the peel-away sheath still in place was absent of pulses and was cold to the touch. on (b)(6) 2017 the patient's right impella leg was mottled, and had no pulses. the physician removed the peel-away sheath, and the repositioning sheath was placed without issue. following removal of the peel-away sheath, pulses returned to the patient's foot and his leg turned pink, in addition an observation of a hematoma was made. the impella site was reported to be soft with no active bleeding. on (b)(6) 2017 worsening bleeding into subcutaneous tissue to the impella site that extended to mid-thigh. the right leg was found to be tight and mottled with fluid filled blisters and labs drawn were indicative of rhabdomyolysis. the patient was transfused with 3 units of prbc's and given platelets. the patient continued to be supported by the impella cp until it was removed on (b)(6) 2017. at that time it was reported that the patient was in stable condition. | [5 empty] | The impella cp and the automatic impella controller logs were returned to the manufacturer for evaluation. data logs were reviewed and confirm the clinical account. the pump, white cable, and purge cassette were returned in good condition. the 14fr oscor introducer was not returned for review. the pump ran well with appropriate rpms and flow rates. the repositioning sheath was examined and there were no tears or defects found. per company medical personnel, although it cannot be definitively determined the bleeding may have been due to the 14fr peel away oscor sheath being left in place for 24 hours and the vessel had stretched out at the arteriotomy and therefore did not have enough rebound left when the smaller diameter repositioning unit was inserted. this can lead to leakage into the subcutaneous tissue in the patient's leg. (b)(4).
38247 | 1220648-2017-00012 | 6339371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6339371&pc=OZD | 01/19/2017 | 01/19/2017 |  | 02/17/2017 |  |  | 01/19/2017 | 02/17/2017 |  |  | 06/30/2018 |  | 02/17/2017 |  | IMPELLA 5.0 AIC PUMP SET US |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1251765 |  | 005062 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 46 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | [6 empty] | Thrombosis | [12 empty] | Required Intervention | [5 empty] | CARDIOHELP SYSTEM | HEMASHIELD DACRON GRAFT, 10MM | [21 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The patient was a (b)(6) year old female admitted with cardiomyopathy to a community hospital. upon admission, the patient was sent to the cardiac cath lab for evaluation of the coronary arteries, which were found to be without disease. subsequently, the patient's left main coronary artery dissected which prompted the placement of an intra-aortic balloon pump (iabp), cardiohelp system, and transfer to a tertiary center for escalation of care. when admitted to the tertiary center va ecmo (extracorporeal membrane oxygenation) was placed. on day 6 of the patient's cardiac care the ecmo and iabp were removed and the impella 5.0 pump was placed via the right axillary artery through a 10mm hemashield dacron graft. there was minimal difficulty placing the pump within the left ventricle. the sternotomy was closed and the patient was transferred to the icu on impella support without any heparin in the purge due to bleeding concerns. this lack of heparin goes against the ifu which states the following: "dextrose solution (typically 20% dextrose in water with 50 iu/ml of heparin) is used as the purgefluid through the impella® 5.0 catheter." during the evening hours of the first day of 5.0 pump support, the patient was assessed by vascular surgery due to the loss of pulses in the right arm. the vascular surgery team performed a brachial thrombectomy which did restore radial and ulnar pulses. the team started to infuse heparin through the purge solution at this time as well. the plan was to reach a goal of the partial thromboplastin time (ptt) being 80 seconds. per the 5.0 ifu the anticoagulation is to be monitored with the act value, not a ptt, as stated in the following: "after insertion of the catheter (and until explant), act should be maintained at 160 to 180 seconds." the patient was still on multiple cardiac iv drips, and impella support, when two days later there was an observation made of an enlarged pupil which was indicative of an ischemic stroke. the patient's diagnosis was confirmed with ct scan imaging. the team at this time decided to infuse heparin systemically via an iv drip. the following day, (b)(6) 2016 the impella 5.0 was weaned and explanted. the patient expired the next day from neurological and systemic comorbidities. | [5 empty] | The impella 5.0 was not returned for evaluation, as it was discarded subsequent to the event. consequently, no analysis could be performed. the console logs were returned for analysis. the logs showed a suction event approximately 40 minutes before the stroke is diagnosed. suction events are able to be caused by the ingestion of thrombus. the root cause of the brachial thrombus and ischemic stroke suffered by the patient are unable to be determined because the device was not returned for review. no corrective action is recommended because the root cause could not be determined. this failure mode will continue to be monitored and trended. in review of other complaints, none have been filed for other pumps from this lot for relevant failure modes. the manufacturer will continue to monitor and investigate all reasonable and obtainable sources of information and will provide results and conclusions in a supplemental medwatch report, if any additional information is received. (b)(4). device discarded at user facility.
38248 | 1220648-2017-00011 | 6335507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6335507&pc=OZD | 01/19/2017 |  |  | 02/16/2017 |  |  | 01/19/2017 | 02/16/2017 |  |  | 05/31/2018 |  | 02/16/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1241928 |  | 0048-0002 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 69 YR | [5 empty] | Injury | Yes | No | Fluid/Blood Leak | [6 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2017 a (b)(6) male patient was admitted to the facility's emergency room. an ecmo was placed and the patient was transferred to the cath lab, where the patient underwent a cabg to the left anterior descending artery. following the cabg being performed the patient was admitted to the icu department. in the icu the patient became hemodynamically unstable. cpr was performed and a re-angioplasty was also performed. it was then determined that the bypass was insufficient and the patient was stented. the physician then decided that for patient recovery an impella cp should be placed. the device was placed and the patient was successfully supported for 12.50 hours. during patient support low purge pressure occurred and visible leakage was seen emanating from the red plug at the impella cp connector. during trouble shooting for the leakage the patient's insertion site was examined and a massive hematoma at the groin was seen. packed red blood cells were then administered to the patient to help remedy the blood loss. the physician then removed the device and another impella cp was placed in the same right femoral artery. the 14fr introducer sheath was left in place in order to stop bleeding. the bleeding was halted and the patient was successfully supported for 12.5 hours with the second impella cp. the patient was reported to have expired while on support with the second impella cp. the clinicians reported that the patient outcome was not caused or contributed to by the reported event that occurred with the first impella cp used. | [5 empty] | A definitive root cause of the hematoma in the groin and the bleeding from the goin could not be determined as the introducer used with the impella cp in this case was discarded by the staff subsequent to use. the manufacturer will continue to investigate all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. internal reference: (b)(4).
38249 | 1220648-2017-00010 | 6316065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6316065&pc=OZD | 01/09/2017 | 01/10/2017 |  | 02/09/2017 |  |  | 01/10/2017 | 02/09/2017 |  |  | 11/30/2018 |  | 02/09/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1259825 |  | 0048-0015 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Appropriate Term/Code Not Available | [6 empty] | Hematoma | Blood Loss | Vascular Dissection | [10 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2016 a (b)(6) male patient came into the emergency department complaining of crushing chest pains. an ekg was performed which revealed lateral and anterior lead st elevation. the patient's blood pressure and oxygen levels began to deteriorate requiring intubation. the patient was immediately taken to the cath lab where the cardiologist placed an impella lp2.5, in preparation of the repair to the left main artery. as the impella catheter was being prepped the scrub nurse forced the impella cable onto the impella catheter, bending the wires. the hospital staff decided to fix the wires and to "got the impella catheter to work." procedures were performed including the placement of an intra-aortic balloon pump (iabp) and the placement of a stent to the left main. following those procedures the peel away sheath was taken off and the repositioning unit was sutured to the patient. at this point the cardiologist recommended that the patient be transitioned to the impella cp for increased hemodynamic support. prior to the sheath exchange a large hematoma was noted to the right groin. the staff mashed out the hematoma, but continuous oozing was noted. the patient began to experience decrease flows and suction alarms, requiring increased vasopressors and inotropes. once in the icu the patients groin site began to bleed requiring continuous pressure. the patient was also bleeding from all iv sites. surgery was consulted and the patient was removed to the or, where the repair to the right groin was performed. the surgeon stated the artery had dissected around the repositioning sheath. there was a 600 cc blood loss during the or procedure resulting in the administration of 2 units of packed red blood cells. during the surgical procedure the impella 2.5 was pulled back across the valve until resistance was met. the impella lp 2.5 was removed. the impella cp was then inserted. the impella cp was placed in the patient via the right lateral circumflex femoral artery. during the impella cp placement slack was removed, the peel away sheath removed, and the repositioning sheath was sutured in place. the cardiologist then decided to transport the patient to another facility. the patient had been successfully supported for approximately 47 hours when the impella cp was removed. following pump removal the patient was reported to be in stable condition, with no additional adverse effects. | [5 empty] | A definitive root cause of the this event could not be determined, as the impella cp was discarded by the staff following the explant; consequently, no device analysis could be performed. the console logs were also not returned for analysis. the manufacturer will continue to investigate all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. (b)(4). discarded by user.
38250 | 1220648-2017-00008 | 6305914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6305914&pc=OZD | 01/07/2017 | 01/09/2017 |  | 02/06/2017 |  |  | 01/09/2017 | 02/06/2017 |  |  | 07/31/2018 |  | 02/06/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1255489 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2017 a (b)(6) morbidly obese female patient with triple vessel disease was treated in the facility's cath lab. during patient treatment an impella cp was placed in the patient. following placement the impella insertion site exhibited bleeding. pressure was held at the groin's bleeding site. a femstop was applied to the site to maintain control and the patient did receive 2 units of replacement blood products. the physician reported that it is believed that the bleeding occurred as a result of rolling the patient during transfer to the icu causing the groin to get bent. the patient was reported to have been successfully supported for approximately 62 hours and was in stable condition following the removal of the impella cp. | [5 empty] | The lmpella cp was not returned for evaluation, as it is the facility's policy not return devices. the automatic impella console logs were returned for analysis. the analysis revealed that the case proceeded for 62.5 hours, and that there were no notable alarms relevant to the reported bleeding. a review of the device history record was performed and revealed that there were no issues reported for any other device with this lot number. without an evaluation of the device used in this case, the root cause of the bleeding was unable to be determined. it is most likely that the bending of the patient at the groin during transfer to the icu caused the bleeding to start, although this was unable to be definitively determined. in addition, the large patient size may have also played a role in the cause for the patient bleeding. the impella cp instructions for use advises the user of the following: ** when securing the repositioning sheath, vascular closure may be difficult in obese patients with extensive adipose tissue. internal reference: spr 17-0024. not returned per hospital policy.
38251 | 1220648-2017-00013 | 6296985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6296985&pc=OZD | 09/27/2016 | 09/27/2016 |  | 02/02/2017 |  |  | 09/27/2016 | 02/02/2017 |  |  | 04/30/2018 |  | 02/02/2017 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA LP 2.5 |  |  | OZD |  |  | IMPELLA LP2.5 |  |  | ABIOMED, INC |  |  | US |  | 1240965 |  | 5042 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 68 YR |  |  | 111 |  |  | Injury | Yes | No | Patient-Device Incompatibility | Deformation Due to Compressive Stress | Difficult to Advance |  |  |  |  | Perforation | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The customer reported that on (b)(6) 2016 they were treating a (b)(6) male for an elective angioplasty to known diseased left main and circumflex arteries in the cardiac cath lab. the decision was made to utilize an impella 2.5 for support during this high risk pci. the first impella pump would not function after being inserted into the left ventricle. the pump was explanted to swap out for a new 2.5 pump. the second 2.5 pump was setup and inserted through the originally placed abiomed 13fr introducer sheath. the device was unable to be advanced through the sheath and upon fluoroscopic examination there was an observed kink in the sheath. this kink is suspected to have been due to peripheral vascular disease, specifically calcified iliofemoral vasculature. the team was unable to remove the 2.5 device through the sheath without sheath movement so, the team peeled away the 13fr sheath and placed an 8fr sheath. the .018 wire was used to maintain vascular access. there were multiple wire exchanges done by the team to facilitate advancement to no avail. the team then did a fluoroscopy run with contrast and discovered a perforation. to treat the perforation, and tack up the arterial wall, the physician performed a contralateral ballooning five times. after the fifth balloon inflation there was a fine angiographic result. in addition to the ballooning the patient did receive 2 units of blood product. the patient had a fine recovery without the need for any additional support or intervention. | [5 empty] | The impella 2.5 was returned for analysis. the data logs were not, as the pump was never turned on and there was therefore no data collected. the investigation of the device, introducer, and dilator were conducted and completed. the device was compared to another 2.5 sample product to compare and evaluate the 2.5 in question's rigid length. the length was confirmed to be within specification. the cannula of the 2.5 impella did have some damage at the outflow cage. the sheath had been peeled away, but there were some signs of damage along the length of the sheath. the damage could have been caused during the multiple manipulations and product exchanges. the failure to advance the impella 2.5 may have been due to the presence of peripheral vascular disease. the facility did share a fluoroscopic image with abiomed as part of the investigation and there is tortuosity present. this degree of tortuosity had prompted the physician to ask for another sheath, a cook 14fr x 30cm, but the facility had none in stock. the physician had shared with the abiomed representative that the interaction of the tortuosity, motor, and sheath may have cause the perforation. in conclusion, there were no issues observed with the returned impella 2.5 which could have contributed to the reported issue. the perforation was most likely due to the tortuous native anatomy and inclination of the sheath. the failure mode is low risk due to very low occurrence, but will be monitored and tracked. this medwatch report is being submitted as a replacement report, as the first report was inadvertently submitted without being sequentially numbered, but was submitted on october 26, 2016 as 1220648-2016-0000, instead of 1220648-2016-00034. the october 26, 2016 mdr was resubmitted with the correct sequentlail mdr number of 1220648-2016-00034, but was submitted to the emdr test environment. (b)(4).
38252 | 1220648-2017-00009 | 6298463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6298463&pc=OZD | 01/09/2017 | 01/09/2017 |  | 02/02/2017 |  |  |  | 02/02/2017 |  |  | 08/31/2018 |  | 02/02/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP, PRODUCT CODE: OZD |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1258301 |  | 0048-003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 50 YR | [5 empty] | Malfunction | Yes | No | Difficult or Delayed Positioning | Difficult to Fold, Unfold or Collapse | Retraction Problem | Activation, Positioning or Separation Problem | Positioning Problem |  |  | Patient Problem/Medical Problem | Device Embedded In Tissue or Plaque | [11 empty] | Required Intervention | [5 empty] | ECMO | [22 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | A (b)(6) male with ongoing chest pain suffered since a prior cabg procedure performed on (b)(6) 2016, presented to the hospital for diagnostic catheterization. the patient coded during the procedure, resulting in the physician attempting to place an impella cp for support. during device placement and while exchanging for the repositioning sheath, the impella cp moved back out of the left ventricle (lv.) while trying to re-advance the pump into the lv, the catheter buckled over on itself. the physician tried to pull the device back into the abdominal aorta (ao) to unfold the cannula, in order to then advance it down to the iliacs. however after multiple attempts to reposition the device, the decision was made by the physician to leave the device in the abdominal ao, and place a second impella cp. the second impella cp was placed successfully and supported the patient for approximately eight hours before support was exchanged for an ecmo. the first impella cp also remained in the patient until ecmo was placed. the first impella cp was reported to have been surgically removed, without the use of a snare, as it was reported to have straightened out while being pulled down. the patient was later transferred to another facility for long term ecmo support due to biventricular and respiratory failure. there was no indication of any further adverse effect to the patient following the surgical removal of the device. | [5 empty] | A definitive root cause of the reported positioning issue, and surgical removal of the device could not be determined as the impella cp was discarded by the staff following the explant. at the present time it is unclear if the issue occurred as a result of an interaction with the introducer, lmpella cp or guidewire; or if the issue occurred as a result of the patient's anatomy, device placement or the procedure that was performed. the manufacturer will continue to investigate all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received.
38253 | MW5067613 | 6297551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6297551&pc=PBL | 01/24/2017 | 01/31/2017 |  | 01/31/2017 |  |  |  | 01/31/2017 |  |  | 12/31/2017 |  | 01/31/2017 |  | ABIOMED IMPELLA CP PERCUTANIOUS LVAD |  |  | ABIOMED IMPELLA CP PERCUTANIOUS LVAD |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  |  |  | 0048-0003 |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 43 YR |  |  | 127 |  |  | Malfunction | No | Yes | Device Displays Incorrect Message | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Other | [5 empty] | IMPELLA CP PERCUTAEOUS LVAD | IMPLANTED: | [21 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Impella cp did not go about 0.9 l/min or flow after insertion into the patient. suction alarms also started to show. impella was repositioned several times without resolution of the problem. impella pump pulled out and tested in a basin filled with heparinized saline. suction alarms continued to show during testing. gross examination of the pump did not reveal any clots caught in the pump cage. new impella catheter was prepped and inserted with no problems encountered. diagnosis or reason for use: cardiogenic shock.
38254 | MW5067618 | 6297833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6297833&pc=KFM | 01/24/2017 | 01/31/2017 |  | 01/31/2017 |  |  |  | 01/31/2017 |  |  | 11/30/2018 |  | 01/31/2017 |  | IMPELLA CP, PERCUTANEOUS HEART PUMP |  |  | IMPELLA CP, PERCUTANEOUS HEART PUMP |  |  | KFM |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  |  |  | 0048-0003 |  | No | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 56 YR |  |  | 102 |  |  | Injury | No | Yes | Device Operates Differently Than Expected | Improper Flow or Infusion | [5 empty] | Chest Pain | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Impella cp was inserted in cath lab room 04 and started at 15:47. impella cp flows only went up to 0.9 to 1.0lpm on auto. another impella cp was opened and started at 15:57 second impella cp was successful. the malfunctioned impella cp pump was saved & will be sent back to the company, abiomed. diagnosis or reason for use: chest pain with moderate risk of acute coronary syndrome. is the product compounded: no. is the product over-the-counter: no.
38255 | 6282609 | 6282609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6282609&pc=KFM | 12/14/2016 | 01/05/2017 | 01/05/2017 | 01/27/2017 | 01/05/2017 | 01/05/2017 |  | 01/27/2017 |  |  |  |  | 01/27/2017 |  | PUMP, BLOOD |  |  | IMPELLA CP |  |  | KFM | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | 004B-0003 |  | No answer provided | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 51 YR | [5 empty] | Malfunction |  | Yes | Device Operates Differently Than Expected | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The impella device malfunctioned.
38256 | 1220648-2017-00007 | 6267972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6267972&pc=OZD | 12/14/2016 | 12/22/2016 |  | 01/20/2017 |  |  | 12/22/2016 | 01/20/2017 |  |  | 08/31/2018 |  | 01/20/2017 |  | TEMPORARY CARDIA SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMP;ELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1256704 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 68 YR | [5 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | Ventricular Tachycardia | [11 empty] | Death | [28 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The complainant reported that a (b)(6) male patient with a non-st-elevation myocardial infarction (nstemi) was catheterized in another facility's cath lab on (b)(6) 2016. during treatment at that facility it was discovered that the patient had severe distal left main/proximal left anterior descending (left main/plad) and mid right coronary ascending (mrca) chronic total coronary occlusion (cto). in response to that discovery the patient was transferred to this facility on the evening of (b)(6) 2016 to have a high risk percutaneous coronary intervention for the following day. the patient was reported to have a history of congestive heart failure, coronary artery disease, chronic obstructive pulmonary disease and an ejection fraction of 10%. the patient was intubated; the impella cp was then inserted via right femoral artery through a 14fr cook introducer sheath (due to the patient's vessel tortuosity.) a rotoablation of the distal left main/plad was done and a drug eluding stent (des) was placed, then a des was applied to the mrca cto for complete revascularization. throughout procedure, the patient mean arterial pressures (map's) did dip into the 50's, especially during rotoablation. a neosynephrine iv push was given periodically to keep the patient's blood pressures up. at the conclusion of the intervention, the doctors were unable to successfully wean from the impella cp, therefore the decision was made to leave impella cp in overnight and to continue to rest the patient's heart, with a plan to wean and remove the pump the following day. levophed gtt was started to provide stability prior to the patient leaving the cath lab. in addition and prior to leaving the cath lag the purge fluid was switched to d5 with heparin 50 units/ml, systemic heparin gtt was started, and the impella cp was transferred to standard configuration. upon arrival to the unit a bedside echo was performed to confirm cannula placement. no impella related alarms occurred. a central line was placed to monitor the central venous pressure (cvp.) the impella cp successfully supported the patient without complication for a total of approximately 12 hours. during the late evening of (b)(6) 2017 the patient went into a ventricular tachycardia heart rhythm and then asystole and expired. the physician reported that the patient outcome was as a result of severe aortic insufficiency. | [5 empty] | The impella cp was not returned for evaluation, as it was discarded subsequent to the event, therefore no device analysis could be performed to assist in confirming that the patient outcome was the result of aortic insufficiency and that the impella cp used during this event caused or contributed to patient aortic insufficiency. in addition, the requested images from the event were unable to be retrieved from the complainant. an analysis of the returned console data logs was performed. the data logs reveal a declining placement signal that suggests deteriorating patient condition. spikes seen in the placement signal following the clinical timeline of the onset of tachycardia and asystole were most likely caused by chest compressions. the pump was not returned from the field. as the product was not returned from the field, and no clinical imaging was provided, the root cause of the reported aortic insufficiency was not able to be determined. a review of the device history record revealed that there were no issues reported for any other device with this lot number. no corrective action is recommended as the root cause was unable to be determined. the manufacturer will continue to investigate all reasonable and obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. (b)(4). | On april 25, 2017 additional clinical imaging from this event was received from the complainant. an analysis of the imaging was performed. the following is the detailed evaluation of the imaging. the clinical imaging that was returned from the field was reviewed with abiomed physicians. the patient was more hypotensive and hemodynamically unstable than would normally be expected after impella support was initiated. the initial imaging study that occurred about 11:40 am showed that the pump was sitting much deeper in the ventricle than recommended, with the inflow cage sitting about 7.6 cm from the aortic valve and the outlet cage sitting under the valve. doppler flow imaging showed plumes of flow back into the left ventricle that appeared to be evidence of aortic insufficiency (ai.) a later imaging study showed that the pump had been pulled back slightly, and aortic regurgitation was more apparent. this was likely due to the reduced interference after the pump was repositioned; however, the pump was still too deep and patient pressures had not improved. while the outflow cage was sitting under the aortic valve in the images returned from the field, the outflow cage would not normally damage the valve in the few hours it was in that position. the automatic impella controller data logs were again reviewed in conjunction with the received clinical images. the data logs suggest that the patient's condition was deteriorating. the diastolic suction alarms that largely occurred at the end of the case were most likely caused by low volume due to patient condition. the signal spikes at the end of the case appear to be caused by chest compressions. the aortic insufficiency reportedly began about 10 hours into the impella cp support. in conclusion, the data logs suggest that patient condition was deteriorating throughout the case. the pump was not returned. clinical imaging showed aortic insufficiency, but no imaging was provided from before the pump was placed or post-explant so it could not be determined whether the aortic insufficiency was pre-existing or related to the pump. the root cause of the aortic insufficiency was unable to be determined. no corrective action is recommended as the root cause was unable to be determined. failure modes of this type will continue to be monitored for trends. (b)(4).
38257 | 6259971 | 6259971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6259971&pc=KFM | 11/03/2016 | 11/25/2016 | 11/25/2016 | 01/18/2017 | 11/25/2016 | 11/25/2016 |  | 01/18/2017 |  |  |  |  | 01/18/2017 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE, CATHETER |  |  | IMPELLA |  |  | KFM | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003 |  | No answer provided | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 76 YR | [5 empty] | Malfunction |  | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL | PERCUTANEOUS 14 FRENCH |  | Leak noted in catheter after insertion, crack noted on inspection.
38258 | 1220648-2017-00005 | 6260115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6260115&pc=OZD | 12/20/2016 | 12/20/2016 |  | 01/18/2017 |  |  | 12/20/2016 | 01/18/2017 |  |  | 08/31/2018 |  | 01/18/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP, PRODUCT CODE: OZD |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1258949 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 82 YR | [5 empty] | Injury | Yes | Yes | Break | Difficult or Delayed Positioning | Physical Resistance | Patient-Device Incompatibility |  |  |  | Blood Loss | [12 empty] | Required Intervention | [5 empty] | THE 14 X 28 FR SENTRANT SHEATH | [22 empty] | Initial submission | Followup |  |  |  |  | * | HOSPITAL |  |  | An (b)(6) female with a history of peripheral vascular disease, severe microvascular disease, coronary ostial stenosis and chronic total occlusion of the right coronary artery was scheduled for a high-risk pci procedure on (b)(6) 2016. the plan was to implant impella cp for support during the rotational atherectomy portion of the hrpci. rescue angioplasty with stenting of the left main and left anterior descending was performed. the patient began to decompensate at 5 am and the decision was made to implant impella cp through the right axillary insertion site. a 14 x 28 fr sentrant sheath was inserted with difficulty due to heavy calcified and tortuous vasculature. the vascular surgeon inserted the sheath into the right axillary using ultrasound, and the impella cp was advanced. the pump was started without issue, however when the pci procedure began the sentrant sheath slipped out of the artery. to prepare for repositioning of the device, the pump was pulled back across the valve into the axillary artery to the access site; and the decision was made to reposition the device wirelessly. when the clinical team attempted to pull the pump back into the sheath the pump became lodged in calcium. as the sheath and pump were attempted to be maneuvered around the calcium, the inflow cage separated from the catheter. the separated inflow cage was noted to be securely in the subclavian artery and not in the carotid artery, therefore the decision was made to implant a second impella cp through the left femoral artery in order to complete the pci procedure. the second impella cp was implanted and the pci was successfully completed, however the patient was pulseless throughout the procedure. the clinical team attempted to remove the inflow cage through the right subclavian sheath using a guide wire snare and 5 mm balloon, however they were unsuccessful. a 14 fr introducer was inserted into the right femoral artery where the inflow cage and pigtail were successfully removed. the right femoral insertion site was unable to be successfully closed using a perclose closure device, therefore a covered stent was placed for hemostasis. throughout the course of the procedure the patient received 4 units of blood products, however it was unclear if the blood replacement was needed due to the pump separation or due to access of the patient's anatomy. the patient was transported to ccu in stable condition and on impella cp support. the patient subsequently expired as a result of multi-system organ failure following 25 hours of successful impella cp support | [5 empty] | The investigation of this event included a review of the clinical case and console data that was conducted previously, and the pump has since been returned for evaluation. the pump was returned with the inflow cage separated from the cannula, and the pigtail was deformed and kinked. the cannula is made of a nitinol coil surrounded on both sides by polyurethane. the polyurethane coating had ripped and the nitinol coil had been pulled out of the remaining coating. the bond between the cannula and inflow cage did not fail. there was evidence of a kink right where the polyurethane coating had ripped. (b)(4). | The investigation of this event included a review of the clinical case and console data, however the pump has not yet been returned for analysis. the console data confirmed the clinical account. the pump primed successfully and provided support for two minutes before the device was stopped by the user. the pump was never turned back on due to the reported issue with the introducer and separated inflow cage. according to the clinical description insertion of the pump was difficult due to the patient's heavily calcified and tortuous vasculature. as listed in the ifu under contraindications: "severe peripheral arterial disease that precludes the placement of an impella cp catheter." as the pci procedure began the sentrant introducer slid out of the artery; requiring repositioning. in order to re-insert the introducer the impella was pulled back through the introducer, however the inflow cage became separated from the catheter at this time. the root cause of the inflow cage separation from the catheter is most likely due to the patient's peripheral vasculature disease, and heavily calcified and tortuous vasculature. the impella cp has not yet been returned for evaluation as it is being held in the facility's risk management department. the manufacturer will continue to investigate all reasonable and obtainable sources of information and will supply results and conclusions in a supplemental medwatch report if the device and/or additional information is received. (b)(4). product not yet received.
38259 | 1035166-2016-00201 | 6250598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6250598&pc=DYB | 08/03/2016 | 01/13/2017 |  | 01/13/2017 |  |  | 08/31/2016 | 01/13/2017 |  |  | 06/01/2019 |  | 01/13/2017 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-12562 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 61 YR | [5 empty] | Injury | Yes | No | Leak/Splash | [6 empty] | Stenosis | Blood Loss | [11 empty] | Hospitalization | Required Intervention |  |  |  |  | IMPELLA CP | PERCLOSE SYSTEM | [21 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The customer reported on (b)(6) 2016 at 1:00 pm, an impella cp was implanted to support a (b)(6) male patient who became unstable during pci (percutaneous coronary intervention). the patient was of "average" size, and there were no problems noted with the patient's vasculature. bleeding through the 14 fr introducer valve was resolved when the repositioning sheath was advanced while the patient was still in the cardiac catheterization lab. some bleeding around the sheath was also noted. the oscor introducer was left in place so the physician could have access to use the perclose system during the intended explant on (b)(6) 2016. the procedure was completed successfully and the patient was moved to the icu. there, significant bleeding developed around the sheath. the patient was brought back to the cardiac catheterization lab, and the pump was explanted and bleeding ceased. | [5 empty] | The device was in use for treatment. there were no manufacturing rejects or anomalies of this type recorded in the device history record. a review of complaints against this lot number identified two (2) additional reports for different reported events. the investigation of the returned product found the device was manufactured per specification. upon evaluation of the returned product, it was found that the hemostasis valve was torn. this was likely caused by off center dilator/device insertion during use out in the field. the leaking experienced by the user may have occurred from the torn hemostatic valve. the potential causes of the seal leaking may be seal damaged during device insertion, improper seal molding/slitting, and improper slitting of blades or improper introducer sheath assembly. the potential effects to the user may be bleeding or prolonged intervention. a review of the risk for this failure mode identified the risk to be as low as practicable. based on the investigation, a capa is not required. oscor will continue to monitor this device for complaint trends and risk. a leak test is performed on 100% of the introducer sheaths by qa per procedure. the introducer sheath passed all in-process and qa final inspection steps before shipping to the customer including visual, dimensional and leak testing. the instructions for use (ifu) provides these precautions: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. infusion through the sideport can be done only after all air is removed from the unit. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only.
38260 | 1220648-2017-00006 | 6248173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6248173&pc=OZD | 12/13/2016 | 12/13/2016 |  | 01/12/2017 |  |  | 12/13/2016 | 01/12/2017 |  |  | 08/31/2018 |  | 01/12/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1259779 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 64 YR | [5 empty] | Injury | Yes | No | Malposition of Device | Unintended Movement | [5 empty] | Perforation | Blood Loss | [11 empty] | Required Intervention | [5 empty] | ECMO | FOLEY | SWAN-GANZ CATHETER | [20 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | Complainant reported that on (b)(6) 2016 a (b)(6) old female patient was admitted to cardiac care lab for a left heart catheterization (lhc). at that time a proximal lesion was noted to a large ramus. the physician then placed a stent to the ramus, but blood flow was then lost and the patient coded. patient cpr was begun and an impella cp was placed without issue via the left femoral artery. the patient was then taken to the unit, but upon moving the patient to a bed, the impella cp was pulled out resulting in the patient coding. cpr was initiated again, and the patient was taken back to the cath lab. the impella was then advanced back across the valve and cpr was continued. the patient was then escalated to an ecmo and the impella pump was left in the patient as a vent for the left ventricle. once stabilized on ecmo the patient was moved to a bed for transport. upon arrival in cardiovascular icu a foley and swan-ganz catheter were placed. intermittent suction alarms occurred. a tee probe was then placed by anesthesia. pericardial effusion was noted at this time as something to "watch"; images were reportedly difficult to obtain, but the images that could be seen did indicate that the catheter was deep and the left ventricle was empty. volume was being given due to the blood loss that occurred in the cath lab. there was a definite decrease in ecmo flows due to a decrease in the venous return. it was then decided to perform a pericardiocentesis resulting in 500 cc's of blood immediately being removed. the patients hemodynamics then improved. blood continued to accumulate, so the surgeon decided to do a pericardial window. at this point the physicians were sure that there was a definite hole somewhere, so a sternotomy was then performed at bedside. upon opening the chest trauma was first observed, and was attributed to the cpr that had been performed. the physician then observed that the impella pigtail had punctured the ventricle, and was situated outside the ventricle. the impella cp was pulled back and the hole to the ventricle was repaired. it was then decided to put impella back across the valve and continue to use it as the vent. it was reported that the general consensus of the physicians at the time of this event was that the puncture "probably" occurred during cpr. the impella cp continued to successfully support the patient, and on (b)(6) 2016 a cabg was performed and the ecmo was removed. the impella cp was left in place and continued to successfully support the patient for a total of 8.8 days. the device was removed from the patient on (b)(6) 2016, and the patient was reported to be in stable condition, with no continued adverse effects. | [5 empty] | The device was returned for evaluation. an analysis of the automatic impella controller was performed. the data from the first several hours of support revealed that "impella position in ventricle" alarms were recorded, and the automatic suction response was triggered for over an hour after cpr was stopped. this was due to the low volume caused by the perforation. the "real time" data from the first day of this case was not available as it had been overwritten. the pump was returned intact. the distal section of the impella cp pigtail was found to be deformed, in addition, the cannula was found to be deformed with evidence of kinking, the evaluation results concluded that after cpr and numerous "impella position in ventricle" alarms, motor current decreased and the automatic suction response was triggered. this corroborated the position and volume issues as reported by the complainant in the clinical account. the root cause of the left ventricle perforation is believed to have been the result of the administration of patient cpr with the pump in place, although this cannot be definitely determined. (b)(4).
38261 | 1220648-2017-00004 | 6232548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6232548&pc=OZD | 12/16/2016 | 12/16/2016 | 12/16/2016 | 01/09/2017 |  | 12/16/2016 | 12/16/2016 | 01/09/2017 |  |  | 08/31/2018 |  | 01/09/2017 |  | IMPELLA 5.0 |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED INC. |  |  | US |  | 1257201 |  | 005062 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Malposition of Device | Device Operates Differently Than Expected | Positioning Problem |  |  |  |  | Mitral Regurgitation | Rupture | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | HOSPITAL |  |  | The patient was a (b)(6) male with a history of heart failure and cardiomyopathy who needed an impella 5.0 placed, via the right axillary. upon the insertion the flows were good and the team discharged the patient to the icu on (b)(6) 2016. the following morning an echo was done that demonstrated mitral regurgitation (mr) due to the position of the impella 5.0 being below the mitral valve. the team attempted troubleshooting and repositioning in the operating room unsuccessfully so, at that time the 5.0 was explanted an impella cp pump was placed through a new access site at the left femoral artery. after implantation of the impella cp the mr did not immediately resolve due to an apparent chordae rupture. the team did not need to perform any intervention to remedy the mr or cord rupture. the patient was stable with the cp support and the acls iv medications were weaned over the following two days. the cp was successfully explanted on the second day of support without any lasting harm or injury on (b)(6) 2016. | [5 empty] | The impella 5.0 pump was not returned for evaluation, as it was discarded subsequent to the event. consequently, no device analysis could be performed. the console logs were returned for analysis. the logs showed that the pump ran normally during support. several placement alarms were triggered when the pump was pulled into the aorta at the end of the case. the physician did share echo images of the 5.0 pump and ventricle. the echo on the morning of (b)(6) 2016 did reveal severe mitral regurgitation, which was thought to be due to the pump's position near the valve. the positioning was due to the anatomy of the axillary anastomosis and the bicuspid aortic valve. within the instructions for use of the 5.0 pump positioning of the pump is stated to be: "using fluoroscopic imaging to properly position the inlet area in the left ventricle no more than 3.5 cm below the aortic valve." upon the explant of the 5.0 pump and insertion of the replacement, impella cp pump, there was no further intervention done to remedy the mitral regurgitation and cordae rupture. the patient was weaned successfully from the cp pump and discharged from in-hospital care with consideration of hospice. the manufacturer will continue to monitor and investigate all reasonable and obtainable sources of information and will provide results in a supplemental medwatch report, if any additional information is received. no corrective action is required as the root cause of the valve damage was contact with the pump due to poor positioning caused by the patient's anatomy. the failure will continue to be monitored and tracked. (b)(4) device discarded by user facility.
38262 | 1220648-2017-00003 | 6232581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6232581&pc=OZD | 12/13/2016 | 12/13/2016 |  | 01/09/2017 |  |  | 12/14/2016 | 01/09/2017 |  |  | 10/31/2018 |  | 01/09/2017 |  | IMPELLA 2.5 AIC PUMP SET |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED INC. |  |  | US |  | 1259686 |  | 005040 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | No | Device Operates Differently Than Expected | [6 empty] | Hemolysis | Low Blood Pressure/ Hypotension | Blood Loss | [10 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | HOSPITAL |  |  | On (b)(6) 2016 a (b)(6) male was admitted to a (b)(6) hospital for a high risk percutaneous coronary intervention (pci). the pci was to treat and stent three coronary arteries with drug eluting stents, the left main, left anterior descending, and circumflex. an impella 2.5 pump was placed for support. during the stenting of the arteries there was hypotension noted with a decrease in the blood pressure. the physicians were mentioned to be satisfied with the support provided by the imepella and the patient was discharged to the icu on support. after 7 hours of support there were signs of hemolysis and the impella was explanted. the lab values of the hemoglobin were falling and the lactate levels were rising. in addition to the hemolysis there was a retroperitoneal bleed. to remedy both the hemolysis and bleed the patient received 4 units of blood replacement products. there was no need for further intervention in the form of dialysis or surgical repair. the patient stabilized. | [5 empty] | The impella 2.5 pump was not returned for evaluation despite all attempts to retrieve it from the german facility. consequently, no device analysis could be performed. the console logs were returned for analysis. the logs were reviewed and showed the mentioned hypotension during the pci procedure and ballooning. they further showed that there was good pressure and flow with the pump in correct position throughout the support. the manufacturer will continue to monitor and investigate all reasonable and obtainable sources of information and will provide results and conclusions in a supplemental report, should any additional information be received. the root cause of the observed hemolysis and retroperitoneal bleed were unable to be determined because the product was not returned. no corrective action is required at this time. failures of this type will continue to be monitored and trended. (b)(4).
38263 | 6232975 | 6232975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6232975&pc=PBL | 12/14/2016 | 12/19/2016 | 12/19/2016 | 01/09/2017 | 12/19/2016 | 12/19/2016 |  | 01/09/2017 |  |  | 06/30/2018 |  | 01/09/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP SET |  |  | PBL |  |  | 14F |  |  | ABIOMED, INC |  |  | US |  | 21 109432 17 180630 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | User Facility report | [8 empty] | 79 YR |  |  | 88 |  |  | Malfunction |  | Yes | Fluid/Blood Leak | [6 empty] | Hematoma | [18 empty] | CARDIAC DRUGS | [22 empty] | Initial submission | [5 empty] | Y | HOSPITAL | 0048-2604/B |  | Patient in cardiogenic shock. dr placed abiomed impella cp catheter. after pump was started, low purge pressure was noted. the impella rep noticed a leak in the catheter housing which is located outside the body. impella catheter was replaced with a new impella sheath and catheter. at the end of procedure, patient developed a hematoma at the right femoral sheath site. manual pressure was held and then femstop applied. patient is still in hospital. the current medical regimen is effective. he is hemodynamically stable and free of bedside evidence of chf except for his trace leg edema. additionally, he has nothing to suggest any angina and his blood pressure is well controlled. his platelets are coming down, likely a heparin effect; dr will change this to argatroban. otherwise, continue present medications. perhaps he can be weaned off impella in am. device was given to the rep.
38264 | 1220648-2017-00002 | 6233489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6233489&pc=OZD | 12/12/2016 | 12/13/2016 |  | 01/09/2017 |  |  | 12/13/2016 | 01/09/2017 |  |  | 08/31/2018 |  | 01/09/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC |  |  | US |  | 1257333 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 90 YR | [5 empty] | Injury | Yes | No | Fracture | [6 empty] | Perforation | Rupture | Blood Loss | [10 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2016 a (b)(6) male patient was being treated with an elective left main pci. the physician planned on attempting a left circumflex (lcf) approach along with impella cp support. initially there were no problems with access and dilation when performing the impella cp placement. the patient was found to have tortuous and stiffer than usual vessels, and upon inserting the 14fr oscor sheath the femoral artery was perforated. the oscor sheath was replaced with a 14fr cook sheath. a large amount of bleeding ensued from the perforation and was estimated to be a loss in excess of 200cc's. a vascular surgeon was called to the cath lab, and successfully repaired the artery. the patient was administered 2 units of packed red blood cells (prbc). the patient had become bradycardic and required the administration of atropine and epinephrine injections and was eventually placed on a levophed drip. the impella cp was then placed. the patient was administered 2 units of packed red blood cells (prbc), and the physician decided to let the patient rest and to perform the pci later. the patient was transferred to the icu. the following day the pci was successfully performed on the left main and left anterior descending arteries and successful stenting to both arteries. the patient was weaned from the pump support and was hemodynamically stable. successful perclose was completed of the surgically repaired site. the physician was reported to be very happy with the impella cp support and stated that he believed that the oscor sheath had fractured at the tip (but did not become detached) during insertion causing the perforation. | [5 empty] | A definitive root cause of the perforation could not be determined, as the impella cp was discarded by the staff following the explant. as stated in the event description the patient had tortuous and stiffer than usual vessels, which could have contributed to the reported perforation, although at the present time it is unclear if the perforation occurred as a result of any issue with the introducer, lmpella cp pump guidewire or if the issue occurred as a result of the patient's anatomy, device placement or the procedure that was performed. the manufacturer will continue to investigate all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. a review of the device history record was unable to identify any other devices from this lot with this failure mode. internal reference (b)(4). discarded by user.
38265 | 6228926 | 6228926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6228926&pc=PBL | 11/08/2016 | 11/22/2016 | 11/22/2016 | 01/06/2017 | 11/22/2016 | 11/22/2016 |  | 01/06/2017 |  |  | 09/30/2018 |  | 01/06/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP LMR |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No answer provided | [5 empty] | No | Physician |  |  | User Facility report | [8 empty] | 65 YR |  |  | 60 |  |  | Injury | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Hospitalization | Required Intervention |  |  |  |  | CARDIAC DRUGS | OTHER, ANTICOAGULANTS | [21 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Patient presented for a left ventricular descending coronary artery (lad) intervention, complicated by a left main dissection. an impella device was implanted, the patient sustained a perforation to the left ventricle requiring surgery to repair the perforation. please note: i will obtain more details about the specific device, i do not have access to this information at the time i am entering this report, but will obtain. therefore, the device specific information provided will be limited at this time.
38266 | 1220648-2017-00001 | 6230439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6230439&pc=OZD | 09/11/2016 | 12/08/2016 |  | 01/06/2017 |  |  | 12/08/2016 | 01/06/2017 |  |  | 06/30/2018 |  | 01/06/2017 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1245945 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 29 YR | [5 empty] | Death | Yes | No | Missing Value Reason | [6 empty] | Blood Loss | [12 empty] | Death | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that the physicians were treating a 29 year old female patient who was in cardiogenic shock. an impella cp was placed, and the patient was supported for over 11 hours. the device was then removed because of reportedly severe bleeding occurring from an aortic leak. during patient support many alarms occurred warning the user of positioning issues and aspirations. the patient subsequently expired as a result of multi-organ failure and hemorrhagic shock. | [5 empty] | The event was not reported to abiomed by the customer, until december 8, 2016. the lmpella cp, 14french introducer and guidewire were not returned for evaluation, as they were discarded following explant. it is unclear at the present time if the aortic leak occurred as a result of any issue with the introducer, lmpella cp pump guidewire or if the issue occurred as a result of the patient's anatomy, device placement or the procedure that was performed. an incomplete set of data logs were returned for this investigation. the logs show periods of time where the device was positioned incorrectly, as evidenced by numerous impella "position wrong" alarms. these events show that the device had a drop in motor current pulsatility and a placement signal that was representative of an aortic waveform. the analysis of data logs further revealed periods of mis-positioning, most likely as a result of the device was sitting almost entirely in the patient's aorta. the positioning of the device with the inlet cage near the aortic valve could have contributed to the observed aortic insufficiency; however, without a complete evaluation of the device this cannot be confirmed. the root cause of the aortic valve damage was unable to be determined because the product was not returned. a review of the device history record was unable to identify any other devices from this lot with this failure mode. the manufacturer will continue to investigate all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. internal reference spr (b)(4).
38267 | MW5067045 | 6224972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6224972&pc=KFM | 10/14/2016 | 01/03/2017 |  | 01/03/2017 |  |  |  | 01/03/2017 |  |  |  |  | 01/03/2017 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | KFM | [5 empty] | ABIOMED |  |  | US |  | 1247056 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 62 YR |  |  | 89 |  |  | Malfunction | No | Yes | Device Displays Incorrect Message | [6 empty] | No Information | [41 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | "low purge pressure" alarm requiring removal and re-insertion of new product.
38268 | 1220648-2016-00045 | 6210607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6210607&pc=OZD | 11/29/2016 | 11/29/2016 |  | 12/28/2016 |  |  | 11/29/2016 | 12/28/2016 |  |  | 08/31/2018 |  | 12/28/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1257206 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 79 YR | [5 empty] | Injury | No | No | Device Operates Differently Than Expected | Human-Device Interface Problem | Material Integrity Problem |  |  |  |  | Low Blood Pressure/ Hypotension | Perforation | Blood Loss | [10 empty] | Life Threatening | [5 empty] | FEMOSTOP | [22 empty] | Initial submission | [5 empty] | * | HOSPITAL |  |  | On the (b)(6), a (b)(6) male was admitted with a non-stemi and was sent to the cardiac cath lab on the (b)(6) for a procedure that revealed severe multivessel disease inclusive of the left main. cardiothoracic surgery was consulted and the impella cp pump was placed for support, although of note prior to the placement of the impella the patient was hypotensive. while the physician was attempting to remove the oscor sheath and advance the repositioning sheath there was a considerable amount of blood lost. the physician was removing/cracking the peel away oscor while still in the patient's native vasculature. according to the ifu the removal of the sheath a procedure as follows: "flush the sidearm of the repositioning sheath located on the catheter shaft. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. grasp the two "wings" and bend back until the valve assembly comes apart. continue to peel the two wings until the introducer is completely separated from the catheter shaft." the patient's response to the removal of the oscor and the blood loss was hypotension requiring medical iv support, a placement of the femostop, and a mattress suture to control the bleeding. the femostop application did occlude flow and so the pressure was adjusted to prevent limb ischemia, as the limb did become mottled. a total of four units of blood replacement products and albumin were infused to the patient. in the following days whenever the femostop pressure was adjusted the ooze of blood loss continued at the access site. the surgical team did turn down the patient for escalated care in the form of a cabg and so, the patient's family declined any further intervention. on the (b)(6) the patient was taken to the or for removal of the impella cp pump with hemostasis achieved by contralateral balloon inflation. the angiography of the access site revealed a dissection which was intervened upon with balloon and stenting. the dissection was successfully treated. upon the visual review of the impella cp the physician determined the repositioning sheath portion to be damaged. after the explant of the impella and access site intervention the patient did experience pulseless electrical activity (pea) and did expire in the icu. | [5 empty] | The investigation of this event included the review of the product returned as well as the data logs. the data logs did not contribute any true information regarding this event. they are not applicable to the failure mode. the pump was returned and upon visual review, the repositioning sheath was frayed and split. upon the break down of the pump there were areas of uneven surface. this pump damage resulted in the bleeding that required the blood products, due to the constant ooze. the bleeding was initiated when the physician peeled away the oscor sheath while still in the femoral artery. therefore it can be concluded that the repositioning sheath damage alone did not cause the bleeding. the tip fraying cause could not be determined without evidence of what the wound closure portion looked like prior. there was no photographs provided. there were no other complaints for any other pumps within this manufacturing lot. the failure mode was determined to be low risk based on very low occurrence and moderate severity. retraining of the user is recommended. this training should cover proper use of the oscor 14fr sheath. the root cause of the bleeding was peeling away the sheath while it was still in the patient's femoral artery, contrary to the ifu. the failure mode will continue to be monitored and tracked. (b)(4).
38269 | 1220648-2016-00043 | 6192898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6192898&pc=OZD | 11/21/2016 | 11/21/2016 |  | 12/21/2016 |  |  | 11/21/2016 | 12/21/2016 |  |  | 07/31/2018 |  | 12/21/2016 |  | IMPELLA 5.0 AIC PUMP SET |  |  | IMPELLA 5.0 |  |  | OZD |  |  | 005062 |  |  | ABIOMED INC. |  |  | US |  | 1254063 |  | IMPELLA 5.0 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 54 YR | [5 empty] | Injury | Yes | No | Malposition of Device | Human-Device Interface Problem | [5 empty] | Aortic Insufficiency | Mitral Regurgitation | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | HOSPITAL |  |  | A (b)(6) male patient was transferred to (b)(6) for heart failure management from an outlying hospital on (b)(6) 2016. the patient had a surgical axillary implant of the impella 5.0 and was monitored for concurrent liver and kidney failure. the patient was receiving dialysis and the liver and kidneys were seen to be shutting down. the patient was unable to be weaned from the 5.0 on day 9 of support even though a wean attempt was made. the 5.0 pump was left in and a plan for possible afib ablation was discussed due to the patient having atrial tachycardia. heart rates were mentioned to be in the 140's beats/minute. patient continued to decompensate and need respiratory support. during these days of decompensation, the team did repeated echocardiography (echo) studies. on day 26 of support the mitral valve cordae was seen to be damaged on echo. the decision was made to wean and explant the impella 5.0 and repair the mitral valve. while in the surgical suite, the 5.0 appeared to be positioned with the inlet cage of the pump right under the aortic valve and it was probable that the pigtail was stuck in the mirtral valve leaflets with ensuing mitral regurgitation. while in the operating suite the team also chose to replace the aortic valve as well. aortic insufficiency was observed upon explant of the 5.0 pump. upon review of the echo films, dating back to (b)(6) 2016, the inlet cage of the 5.0 pump had been roughly 1 cm below the aortic annulus the entire time. training and the ifu state proper positioning as follows: "for optimal positioning of the impella® 5.0 catheter, the inlet area of the catheter should be about 4 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve". upon the double valve replacement, the patient was stable and without any further harm or injury. | [5 empty] | The investigation of the event took place without the return of any product. the only analysis done was that of the data logs, as well as a review in-house of the echo images shared by the team at the hospital. the investigation data points were also supplemented by conversations with the physician and the field representative. the data logs were analyzed and showed normal flows. the 5.0 pump was operating well. it is believed that bad positioning and atrial tachycardia led to the ingestion of a mitral chordae and eventual damage to it. this trauma most likely led to the need for a replacement of the mitral valve from mitral regurgitation. the sustained bad positioning was known to the physician and most likely led to the damage which caused the observed aortic insufficiency. this positioning was observed via echo and the poor position (at a noted 1cm below the av annulus) was not rectified for several days. the patient was in poor condition throughout the majority of the case. a full review of the device used would be needed to confirm conclusively if the impella 5.0 was the sole cause of the reported damage. this investigation is not possible as the product was discarded. the root cause of the damaged mitral and aortic valves was deemed to be bad positioning of the inlet at the annulus exacerbated by pre-existing conditions of atrial tachycardia and frail mitral valve leaflets. the failure mode will continue to be monitored and tracked but, no corrective action is recommended as this is the first occurrence of this failure. (b)(4).
38270 | 1035166-2016-00190 | 6196143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6196143&pc=DYB | 09/12/2016 | 05/04/2017 |  | 12/21/2016 |  |  | 11/17/2016 | 12/21/2016 |  |  | 02/01/2019 |  | 12/21/2016 |  | INTRODUCER, CATHETER |  |  | ABIOMED 23 F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3001 |  |  | OSCOR INC. |  |  | US |  | C1-12322 |  | 0052-3001 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 76 YR | [5 empty] | Malfunction | No | Yes | Break | [6 empty] | Myocardial Infarction | [18 empty] | IMPELLA RP | [22 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | On (b)(6) 2016, an impella rp was placed to support a (b)(6) female patient suffering from a rv infarct and cgs. the pump was implanted "smoothly." the doctor then noticed that the 23 fr oscor sheath had cracked inside the patient, and pressure was held to control possible bleeding. mattress sutures were placed without bleeding, and the patient's pressures improved with the rp at p-6. later, the patient's condition declined and care was withdrawn. | [5 empty] | Visual: the returned introducer passes visual inspection except for 1 split seal that did not tear along slit line. a close inspection of the formed tip area using the vertex micro vu equipment shows an inward bend that is at the torn score-line. if a force great enough to bend the tip took place during the procedure, it may have caused the formed tip to split along the length of the sheath. this may have happened post-surgical procedure. a close inspection of the torn valve (using the vertex micro vu equipment) shows a jagged tear that runs parallel to the slit in the center of the bottom seal. the area of the seal that did not tear has material (by design) that is not slit/cut in order to give the seal strength during component insertion. occasionally, depending on if the component used is inserted off center, it may weaken or tear the seal. this may not be apparent until the hub is broken in two. one of the seals or both of the seals may not tear along the predetermined slit/tear-line. both seals appear to be slit properly and oriented correctly in the introducer hub. evidence of stress on the seal using the micro vu equipment shows a component most likely being inserted off center. dimensional: sheath measured: usable length 29.7cm /o.d. measured i.d and measured. tip i.d. (could not be measured). length is in specification (per drawing). summary of final analysis: returned device analysis reveals one 23f abiomed introducer sheath that is within manufacturing specifications. the potential causes of the reported failure may be related to: device not used by adequately trained personnel and/or does not follow instructions for use. the potential effect to the user for the reported failure may be related to: device damage, tearing of seal, leaks-device replacement and procedure delay. a review of the risk for this failure mode identified that the risk remains in the (medium risk). based on the investigation information, a capa is not required. oscor will continue to monitor this device for complaint trends and risk. root cause: upon evaluation of the returned product, it was found that the introducer may have sustained blunt force at the formed tip. this force applied during the procedure likely exceeded the design of the introducer which could cause the score-line to tear causing failure or crack as described in the customer's return report. evidence of stress on the seal (using the vertex micro vu equipment) shows the component was most likely inserted off center. this scenario often either tears the seal during insertion or weakens the seal and causes it to tear erratically. the reason for return was confirmed, however, no manufacturing defects were found. according to procedure (abiomed introducer sheath in-process and final inspection), qa performs this inspection 100%: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. instructions for use (ifu), abiomed adelante s2 - hemoslaticp peel away & sideport. under precaution: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. never advance or withdraw guide wire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. dilators, catheters, and pacing leads should be removed slowly from the sheath. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. there are no manufacturing rejects or anomalies of this type recorded in the device history record. the introducer and dilators passed all in-process and qa final inspection steps before shipping to the customer including visual, dimensional and leak testing. | Conclusion not yet available, evaluation in progress.
38271 | 1220648-2016-00044 | 6190639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6190639&pc=OJE | 11/23/2016 | 11/23/2016 |  | 12/20/2016 |  |  | 11/23/2016 | 12/20/2016 |  |  | 02/28/2018 |  | 12/20/2016 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA RP |  |  | OJE |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 1239146 |  | 004334 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 69 YR | [5 empty] | Death | Yes | No | Difficult To Position | [6 empty] | Death | Perforation | [11 empty] | Death | [5 empty] | IMPELLA CP | [22 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that on (b)(6) 2016 a (b)(6) female patient with a history of indeterminate stress echo, presented to the emergency room with cardiac pain. an ekg was performed which revealed non-specific t-wave elevation changes, and possible apical hypokinesis. the patient further presented with atrial fibrillation (afib) and a vaso-vagal event with chest pain. ventricular septal defect closure surgery was performed. the patient then went into cardiogenic shock while being supported with an intra-aortic balloon pump and while being administered with maximum pressors and inotropes. the patient was then brought into the cardiac care laboratory for an emergent impella rp placement, as there were escalating pulmonary blood pressures. the patient's hemodynamics were evaluated and resulted in an impella cp also being placed. the patient was successfully supported for 22 hours, as the physician planned to continue to allow the heart to rest., and then to wean the pressors, as tolerated. the following evening during the physician's attempt to reposition the impella rp the right ventricle was perforated. the patient coded and the patient's chest was immediately opened in an attempt to repair the rv perforation, but the patient expired. | [5 empty] | The products involved in this event were discarded by the user subsequent to the event. the manufacturer will continue to investigate all reasonable obtainable source information, and will provide them on a supplemental medwatch report if they become available. device discarded by user.
38272 | 1220648-2016-00042 | 6186015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6186015&pc=OZD | 11/08/2016 | 11/17/2016 |  | 12/19/2016 |  |  | 11/17/2016 | 12/19/2016 |  |  | 09/30/2018 |  | 12/19/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1255865 |  | 0048-0015 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR |  |  | 60 |  |  | Injury | Yes | No | Coiled | Difficult or Delayed Positioning | [5 empty] | Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup | Followup |  |  |  | * | Invalid/Incomplete |  |  | Complainant reported that a (b)(6) female patient was admitted to the hospital. the patient was reported to have had a previous history of coronary artery disease and a percutaneous transluminal coronary angioplasty had been previously performed on the patient's left anterior descending artery (lad) and the left main artery (lm). during treatment the patient was found to have a difficult anatomy and a tight left main and left anterior descending artery lesion. when the physician placed an intra-aortic balloon pump in the patient, a lm/lad dissection occurred. the patient then became very unstable. the physicians then quickly stented the proximal lad, but accidentally closed off the circumflex artery. the patient then coded and the impella was placed via the right femoral artery and the patient stabilized. the physicians were then able to intervene to repair the circumflex artery and to stent the lad and lm. the patient remained stable throughout the procedure. during the procedure the patient was found to have a small left ventricle (lv) resulting in the catheter slipping out of the aorta, and requiring several repositionings of the catheter. at the conclusion of the case the clinicians attempted to peel away the oscor sheath and insert the repositioning sheath. at that time they saw under fluoroscope that the catheter had coiled in the lv. upon pulling the device back, the patient became unstable and an echo was performed which confirmed the lv perforation and revealed the pump's inlet cage protruding through the lv. the area of the perforation was at the site of the infarct when the circumflex was down. the clinicians then performed an emergent pericardial drain and patient was transferred to another facility were the patient's chest was opened, and the impella cp was removed and left ventricle was sutured and patched. the patient was reported to have been discharged from the hospital with a "normal ventricle" and is doing well. the physicians further reported that the patient would not have survived had the impella cp not been placed at the time it was placed. | [5 empty] | The impella cp was discarded subsequent to the event, and therefore was unable to be returned for evaluation. the console data logs were returned and were evaluated. the analysis of the data logs revealed that the impella cp pump's position was erratic throughout the case. drops in p-levels were found to have occurred and were due to automatic suction response. the analysis also revealed that a dip in motor current, pump flow, and pressures occurred in the middle of the case and were due to the pump entering the ventricle. decreases in the placement signal pulsatility happened two hours into the case and coincided with when the perforation occurred. a pump stop was found to have occurred after the perforation, but the pump was restarted successfully. the evaluation revealed fluctuations in placement signal were due to positioning issues, numerous instances of automatic suction responses, and a drop in placement signal pulsatility marked when the perforation occurred. the perforation occurred when the pump was advanced into the portion of the lv wall.. this area became non-viable after its blood supply was compromised when the circumflex artery was blocked. the lv perforation occurred in the patient's non-viable lv wall tissue which coincided with the site of the infarct that occurred during the procedure. no corrective action is recommended as this failure is due to patient condition. failures of this type will continue to be monitored and tracked. (b)(4). device discarded by user. | The patient's weight as reported in the user medwatch number is (b)(4) was reported as (b)(6). the date of event as reported in user medwatch number is (b)(4) was reported as 11/08/2016. | This supplemental mdr is being filed to provide an explanation of why the initial report that was due to be filed on 12-16-2016, was instead filed on 2016-12-19.
38273 | 1035166-2016-00187 | 6182458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6182458&pc=DYB | 09/28/2016 | 11/17/2017 |  | 12/16/2016 |  |  | 09/28/2016 | 12/16/2016 |  |  | 12/01/2018 |  | 12/16/2016 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | CR-01981 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION | [10 empty] | Malfunction | No | Yes | Kinked | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | Followup |  |  |  |  | N | Invalid/Incomplete |  |  | The customer reported a documented account of a 14f introducer sheath kink initiated this investigation. | [5 empty] | Conclusion not yet available, evaluation in process. | Upon evaluation of the returned product, it was found that the introducer most probably had extreme diagonal force (determined by the stress marks) to the introducer. the diagonal force translated stresses to the mid-section of introducer which caused it to kink. the stress likely also caused the structural integrity of the length of the sheath along the score-line to give way and collapse. in addition, there are scrape marks on the introducer tip along with deformation of i.d. calcification of the patient may have also stressed or weakened the introducer tip causing it to fail. force applied during the procedure likely exceeded the capacity of the design of the introducer. the reason for return was confirmed, however, no manufacturing defects were found. a review of risk for this failure mode identified the risk to be medium. based on this investigation a capa is not required. oscor will continue to monitor this device for complaint trends and risk. according to procedure (abiomed introducer sheath in-process and final inspection), qa performs this inspection 100%: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. there were no manufacturing rejects or anomalies of this type recorded in the device history record. the introducer and dilators passed all in-process and qa final inspection steps before shipping to the customer including visual, dimensional and mechanical testing. instructions for use (ifu), abiomed adelante s2 - hemoslaticp peel away & sideport. under precaution: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. never advance or withdraw guide wire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. dilators, catheters, and pacing leads should be removed slowly from the sheath. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action.
38274 | 1035166-2016-00185 | 6177261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6177261&pc=DYB | 08/03/2016 | 01/09/2017 |  | 12/15/2016 |  |  | 11/17/2016 | 12/15/2016 |  |  | 06/01/2019 |  | 12/15/2016 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-12532 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Deformation Due to Compressive Stress | [6 empty] | Myocardial Infarction | [12 empty] | Required Intervention | [5 empty] | HEPARIN | IMPELLA CP | INTRA-AORTIC BALLOON PUMP | [20 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | On (b)(6) 2016 at 8:15 pm, an impella cp was prepared to support a (b)(6) male patient suffering from ami induced cgs. the patient coded, an iabp was placed bedside, and the patient was intubated in the er. cpr was performed on the way to the cardiac catheterization lab. there was no indication of arterial issues shown in imaging performed prior to insertion. the low flow state at this time may have increased the risk of clot formation. impella support was prepared, and the 14 fr oscor sheath was inserted after difficulty advancing the dilators and the sheath over the 0.035" guidewire. the final dilator was removed so the pump could be inserted "as quickly as possible." the cp was implanted before act was checked. the pump was explanted to troubleshoot low flow, and then reimplanted. the 14 fr sheath was examined and the sidearm was unable to be flushed, which led the team to believe the sheath had clotted off. va ecmo was initiated and the iabp was removed. the 9000 units of heparin were given during the surgery to place ecmo, and act was 261 s. the physician decided to reinsert the same cp for lv venting with ecmo in place. when the original 14 fr sheath was removed, it was "noticed to be separated." the impella was reimplanted and support continued until care was withdrawn on the night of (b)(6) 2016. | [5 empty] | The device was in use for treatment. there were no manufacturing rejects or anomalies of this type recorded in the device history record. a review of complaints against this lot number identified one (1) additional report for a different reported event. returned device analysis reveals one 14f abiomed introducer sheath that is within manufacturing specifications. upon evaluation of the returned product, it was found that the introducer had most probably sustained extreme diagonal force (determined by multiple kink/stress marks). the diagonal force translated stress to the mid-section of the introducer which caused it to kink. the stress likely caused the structural integrity at the score-line to fracture and tear at the kink point toward the tip. force applied during the procedure likely exceeded the design of the introducer. based upon the analysis investigation the root cause was determined to be force applied during the procedure which likely exceeded the design of the introducer. the potential effect to the patient for the reported failure may be related to: device is difficult to use and the procedure is delayed, and possible tissue damage. potential cause of this failure may be: improper or damaged extruded parts, improper introducer manufacturing, and improper or damaged packaging. a review of risk for this failure mode identified the risk to be medium. based on this investigation a capa is not required. oscor will continue to monitor this device for complaint trends and risk. according to introducer sheath in-process and final inspection procedure, qa performs a 100% inspection with the naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. they ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded foreign material, or any other damages. the introducer and dilators passed all in-process and qa final inspection steps before shipping to the customer including visual, dimensional and leak testing. the instructions for use (ifu), under "precaution" informs the user: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. never advance or withdraw guide wire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. the ifu also informs the user dilators, catheters, and pacing leads should be removed slowly from the sheath. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. | Conclusion not yet available, evaluation in process.
38275 | 1035166-2016-00186 | 6179270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6179270&pc=DYB | 09/21/2016 | 03/24/2017 |  | 12/15/2016 |  |  | 09/28/2016 | 12/15/2016 |  |  | 06/01/2019 |  | 12/15/2016 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-12600 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION | [10 empty] | Malfunction | No | Yes | Kinked | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The customer reported a documented account of a 14f introducer sheath kink initiated this investigation. | [5 empty] | Conclusion not yet available, evaluation in process. | The device was returned for analysis. the device history record (dhr) for this lot was reviewed. the dhr review identified three (3) damaged boxes during final packaging of this lot. a review of complaints against this lot number identified one (1) additional complaint. returned device analysis reveals one (1) 14f abiomed introducer sheath that is within manufacturing specifications. upon evaluation of the returned product, it was found that the introducer most probably had extreme diagonal force (determined by the stress marks) to the introducer. the diagonal force translated stresses to the mid-section of the introducer which caused it to kink. the stress likely also caused the structural integrity of the length of the sheath along the score-line to give way and collapse. in addition the scrape marks on the introducer tip denote calcification of the patient which may have also stressed or weakened the introducer causing it to fail. force applied during the procedure likely exceeded the capacity of the design of the introducer. there are no manufacturing rejects or anomalies of this type recorded in the device history record. the introducer and dilator passed all in-process and qa final inspection steps before shipping to the customer including visual, dimensional and mechanical testing. the potential effect to patient for the reported failure may be related to: device is difficult to use and the procedure is delayed, and possible tissue damage. potential cause(s) for the reported failure may be related to: improper or damaged extruded parts, improper introducer manufacturing, improper or damaged packaging. a review of the risk for this failure mode identified the risk to be medium. based on this investigation a capa is not required. oscor will continue to monitor this device for complaint trends and risk. according to the abiomed introducer sheath in-process and final inspection procedure, qa performs this inspection 100%: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages.
38276 | 1220648-2016-00030 | 6159291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6159291&pc=OZD | 05/22/2016 | 08/25/2016 |  | 12/09/2016 |  |  | 08/25/2016 | 12/09/2016 |  |  | 02/28/2016 |  | 12/09/2016 |  | PUMP, BLOOD CARDIOPULMONARY BYPASS NON-ROLLER TYPE |  |  | IMPELLA 5.0 IMC PUMP SET EU |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC |  |  | US |  | 1229773 |  | 005060 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG | [10 empty] | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Thrombus | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | HOSPITAL |  |  | The complainant reported that on (b)(6) 2016 axillary placement of an impella 5.0 was performed on a pregnant female patient with myocarditis and in cardiogenic shock. the patient was reported to be in her last trimester of pregnancy. the patient was successfully supported with the impella 5.0 for over 3 days (89 hours.) upon patient recovery impella 5.0 removal was begun, but the physician felt resistance when attempting to perform an axillary removal of the pump. the device was unable to be removed despite multiple attempts to remove it. a fluoroscopy confirmed that the pump was caught at the height of the beginning of the right subclavian artery. it was decided to gain access to the brachiocephalic trunk. a sternotomy was performed with the opening of the sternum via a vertical inline incision. preparation of the brachiocephalic trunk was begun, during which it was found that the patient had an unusually long (almost 4 to 5cm) vessel. the clinicians were able to manage the removal by gaining control of the right carotid, the right subclavian and the brachiocephalic trunk. the device was removed pulling it through the axillary access and by manually helping the impella to pass the ostium of the subclavian. during the removal process the patient developed a thrombosis at the outflow of the right subclavian artery, impacting the right upper limb resulting in blood pressure asymmetry and numbness. the clinician opened the thrombosed vessel and no further complex intervention was needed. the surgeon reported that the impella removal problem resulted from a very long vessel, and an acute angle between the axillary and brachiocephalic trunk. on (b)(6) 2016 it was reported that the "the whole family (mother and baby) is doing well." | [5 empty] | The user facility was unable to provide an explanation to the manufacturer of why this event was not reported to the manufacturer at the time of the event. the device was discarded on may 22, 2016; therefore no evaluation of the impella 5.0 was able to be performed. in addition, no photographs of the impella 5.0 pump were provided for analysis and no fluoroscopy imaging was available for review. a review of the complaint data base revealed that there have not been any other complaints for impella 5.0's with the lot number of 1229773. the root cause of the failed explant is believed to be most likely the result of the patient's anatomy, as the patient had a long vessel with an acute angle between the axillary and brachiocephalic trunk, although without the device for evaluation or images taken during the event this cannot be definitively determined, because this failure mode is related to patient condition, no corrective action is suggested. this failure mode will continue to be monitored and trended. the manufacturer will continue to investigate all reasonable obtainable source information and will provide evaluation results and updated conclusions in a supplemental report if they become available. since july of this year abiomed has had emdr submission problems due to technical firewall issues. as a result abiomed was unknowingly submitting to the test environment since july resulting in 13 mdr's not being submitted to the production environment and are subsequently past the 30 day window. with the help of the emdr team this problem has been resolved. on 11/16/2016 a conference call was held between abiomed and fda and due to the uniqueness of the situation it was agreed abiomed will re-submit the 13 mdr's in the production environment and note in section h10 the reason for being beyond the 30 day window is due to the issues described above. in our first attempt. on 09/23/2016 to send this report unforseen emdr submission problems occurred, the report was finally successfully sent on 09-27-2016, but it was subsequently learned that this mdr was submitted to the production environment. this mdr was one of the 13 and was originally submitted in the test environment on 09/27/2016 and the reasons stated above are the reasons it is being submitted beyond the 30 day window. internal reference: (b)(4). device discarded by user facility.
38277 | 1220648-2016-00046 | 6159525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6159525&pc=OZD | 11/11/2016 | 11/11/2016 | 11/11/2016 | 12/09/2016 |  | 11/11/2016 | 11/11/2016 | 12/09/2016 |  |  | 09/30/2017 |  | 12/09/2016 |  | IMPELLA 5.0 AIC PUMP SET |  |  | IMPELLA |  |  | OZD |  |  | 005062 |  |  | ABIOMED INC. |  |  | US |  | 1216804 |  | IMPELLA 5.0 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 74 YR | [5 empty] | Injury | Yes | Yes | Loss of or Failure to Bond | [6 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [5 empty] | GELOSOFT PLUS 10X30 GRAFT BY (B)(4) | [22 empty] | Initial submission | [5 empty] | * | HOSPITAL |  |  | The patient was transferred to the hospital for escalation of care on (b)(6) 2016. the patient had a history of cardiomyopathy and cabg. the impella 5.0 was placed successfully via the axillary access cutdown. the graft that the physician used was not the normally placed one, but rather a knitted graft, the gelsoft plus 10x30. this was not the abiomed recommended hemoshield plus. the support with the 5.0 pump continued for four days as the patient was stabilized for a planned angioplasty. an intervention was done on the saphenous vein graft to the diagonal branch. three days after the angioplasty the patient returned to the operating department for the explantation of the 5.0 pump. the patient's pump had been pulled back inadvertently during the patient support and the suture wing tab was separated from the sleeve of the pump. there had been noted a report of "excessive force" being applied and causing inadvertent separation of the blue butterfly tab. a hematoma had developed during the surgical day underneath the dressing and so blood products were given. the vascular surgery team removed the hematoma and applied vessel loops to control the bleeding. the patient had been supported by the 5.0 for 6 days successfully and there was no lasting harm or injury from the hematoma nor the broken blue butterfly tab. | [5 empty] | The investigation of this event was done with what data was possible to obtain from the case report and a photograph shared of the access site. the product and data logs were not returned for analysis. in the photo the 5.0 repositioning unit suture tab appears to be separated from the decontamination sleeve during use. this separation could have taken place from an inadvertent pull with excessive force. in addition the type of graft used could have contributed to the hematoma and blood loss. the graft was not the one abiomed recommends for axillary 5.0 support. abiomed recommends a woven graft while the one used in this procedure was knitted. the ifu for the graft states: "clamping may damage any vascular prosthesis.... excessive force should be avoided, as it will damage the polyester fibers and the gelatin impregnation." no corrective actions are to be taken as the root cause could not be determined without product being returned for analysis. the failure mode is determined to be low risk, no corrective action is recommended and the failure mode will be monitored and trended. internal reference:(b)(4). device not returned.
38278 | 1220648-2016-00039 | 6156470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6156470&pc=OZD | 11/07/2016 | 11/08/2016 |  | 12/08/2016 |  |  | 11/08/2016 | 12/08/2016 |  |  | 06/30/2018 |  | 12/08/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1248372 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Device Handling Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that a (b)(6) male patient was taken by ems to the hospital after suffering chest pains for over 6 hours. the patient was reported to have a long cardiac history, and was recently diagnosed with renal insufficiency. the patient was admitted to the hospital. during catheterization, an occlusion of the left ascending artery (lad) was found. once the lad was opened the patient went into a ventricular fibrillation (vf) arrest. the patient coded for several minutes, but was then stabilized. it was reported that "massive doses of blood thinners were administered to the patient during treatment." an impella cp was then placed in the patient without issue. during the exchange of the sheath to the repositioning sheath the patient again coded, resulting in the administration of 1 amp of epinephrine. as a result the patient's blood pressure rose to 270/110. at this point profuse bleeding from the impella insertion site began, requiring the physician to hold manual pressure at the site. shortly thereafter a large hematoma developed from the insertion site down to the patient's knee along the medial side causing the leg to become mottled. the patient was taken to the operating room for femoral artery repair, where several layers of suture were wrapped around the artery to successfully stop the bleeding. no replacement blood products were administered to the patient. impella cp support was continued, and the patient was successfully supported for over 59 hours. the impella was removed from the patient on (b)(6) 2016. the patient subsequently expired, but the physician reported that the patient outcome was not caused or contributed to by the reported issue at the insertion site. | [5 empty] | The impella cp and the automatic impella controller data logs were returned for evaluation. the review of the data logs identified an event at the start of the case where the patient's aortic pressures increased significantly, as indicated by the placement signal, and was followed by a gradual decrease. this dramatic increase in the aortic pressure aligned with the administration of 1 amp of epinephrine. the placement signal sustained at the high pressure for some time. the gradual decrease in placement signal after it peaks indicates that patient bleeding was occurring. the initial inspection of the returned impella cp revealed tht the tip of the outersheath on the repositioning sheath was damaged; when viewed under magnification a number of splits were observed on the tip of the outer sheath. this damage contributed to the patient bleeding. the increase in the patient's blood pressure initiated the bleeding, and the damage on the outer sheath of the repositioning unit facilitated it. in conclusion, the root cause of the bleeding was a combination of the outer sheath damage and the increase in the patient's blood pressure. a corrective action has been initiated to address repositioning sheath damage of this type. the increase in the patient's blood pressure was related to patient condition and care, thus no corrective action is recommended in regard to the blood pressure portion of the issue. this failure mode will continue to be monitored and trended. (b)(4).
38279 | 1220648-2016-00040 | 6156716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6156716&pc=OZD | 11/03/2016 | 11/10/2016 |  | 12/08/2016 |  |  | 11/10/2016 | 12/08/2016 |  |  | 04/30/2016 |  | 12/08/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1237345 |  | 0048-0002 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 19 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The complainant reported that a (b)(6) year old female patient was exhibiting cardiomyopathy upon entry in the hospital. the patient coded and cpr was performed. the patient was intubated and was diagnosed with pneumothorax. the impella cp was placed in the patient and she was taken to the intensive care unit where she again coded, again requiring cpr to be performed. the patient was supported with the impella cp for over 69 hours, during which time the patient was reported to have exhibited more and more pulmonary problems. while on impella cp support the patient was diagnosed with suspected hemolysis; the facility was unable to confirm the hemolysis as 2 plasma free hg levels were unable to be drawn because the hospital lab was unable or not equipped to perform the testing. it was reported that despite hemodynamic stabilization the patient developed renal failure, requiring dialysis to negate renal impairment. the impella cp was removed from the patient and ecmo support was then given, and the patient was transferred to another facility. | [5 empty] | The lmpella cp was not returned for evaluation, as it was discarded subsequent to the event, consequently no device analysis could be performed to assist in finding a definitive root cause of the hemolysis. the automatic impella controller data logs from the event were analyzed. the analysis revealed that in the beginning of impella cp patient support there were a number of issues with catheter placement and suction. it is possible that incorrect placement and/or suction could have contributed to the reported hemolysis; however, without the availability the impella cp for evaluation this is unable to be definitively determined. the manufacturer will continue to investigate all reasonable and obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. (b)(4). discarded by user following event.
38280 | 1220648-2016-00041 | 6142323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6142323&pc=OZD | 09/12/2016 | 11/08/2016 | 11/08/2016 | 12/02/2016 |  | 11/08/2016 | 11/08/2016 | 12/02/2016 |  |  | 03/31/2018 |  | 12/02/2016 |  | IMPELLA 2.5 AIC PUMP SET |  |  | IMPELLA |  |  | OZD |  |  | 0046-3061 |  |  | ABIOMED INC. |  |  | US |  | 1224898 |  | IMPELLA 2.5 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Positioning Problem | [6 empty] | Aortic Insufficiency | Perforation | [11 empty] | Required Intervention | [5 empty] | ECMO | [22 empty] | Initial submission |  |  |  |  | 11/08/2016 | * | HOSPITAL |  |  | In (b)(6) 2016 a patient was admitted to a hospital in (b)(6) with myocardial infarction and a known ventricular-septal defect. the patient was taken to the surgical suite to repair the vsd after being stabilized on ecmo for 2-3 days. after the surgical repair, an impella 2.5 pump was placed into the left ventricle for more support beyond the ecmo. the positioning of the impella 2.5 did need two different wires to place the pump within the ventricle, and the position was checked by fluoroscopy and support initiated. the impella and ecmo supported the patient for approximately 5 days. during the 5 days of support the 2.5 pump had been repositioned multiple times without the assistance of fluoroscopic or echocardiographic imaging, i.e. "blind manipulation". after the fifth day of support, aortic insufficiency was observed under echo. the pump was pulled back slightly and the ai persisted. subsequently, the 2.5 was removed from the left ventricle. at that time valvular damage was observed. the valve was surgically replaced by a tissue valve. the patient remained under care and support with ecmo and the addition of a pulmonary vein drainage cannula. after approximately 2 -3 days the patient did expire due to multi organ failure. | [5 empty] | The impella 2.5 pump was not returned for evaluation, as it was discarded subsequent to the event. consequently, no device analysis could be performed. the console logs were returned for analysis. the logs showed numerous position unknown alarms in the beginning of the case/support. in addition during the last two days of support there were numerous wrong position alarms. the physician did share photographs of the extracted native valve. the photographs were observed and did show a damaged valve. the damage was most likely caused by a combination of patient condition, being the known vsd, as well as the repositioning of the device. the instructions for use of the impella 2.5 do site the following regarding repositioning of the pump: "use fluoroscopy for placement" impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve." the manufacturer will continue to monitor and investigate all reasonable and obtainable sources of information and will provide results and conclusions in a supplement medwatch report, if any additional information is received. no corrective action is recommended at this time because the failure is determined to be low risk. the failure will continue to be monitored and tracked. (b)(4).
38281 | 1035166-2016-00177 | 6140102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6140102&pc=DYB | 10/27/2016 | 01/18/2017 |  | 12/01/2016 |  |  | 10/31/2016 | 12/01/2016 |  |  | 12/01/2018 |  | 12/01/2016 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | OR-04736 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION | [10 empty] | Malfunction | No | Yes | Material Integrity Problem | [6 empty] | No Known Impact Or Consequence To Patient | [18 empty] | IMPELLA CP | [22 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The customer reported tip split was observed, confirming the clinical account. in addition to this, the dilator was noticeably bent at the tip. | [5 empty] | The device was in use for treatment. there are no manufacturing rejects or anomalies of this type recorded in the device history record. a review of complaints against this lot number identified no additional reports. the returned device analysis reveals one (1) abiomed 14f introducer sheath with dilator that is within manufacturing specifications. upon evaluation of the returned product, it was found that the dilator most probably exerted extreme diagonal force, (determined by the bent tip) to the introducer. the diagonal force translated stresses to the tip of the introducer which caused it to split at 2 adjacent score lines. in addition the scrape marks on the introducer that run parallel to the tear may have also weaken the introducer in the area of the tear. force applied during the procedure likely exceeded the design of the introducer. the potential effect to patient for the reported failure may be related to: bleeding, potential product damage, prolonged intervention, and inability to provide support. potential causes may be: insufficient dilator stiffness/radial strength, improper extrusion (wall thickness), improper molding, material too fragile, and excessive force applied during dilator/device insertion/withdrawal. a review of risk for this failures modes identified the risk to be medium. based on this investigation a capa is not required. oscor will continue to monitor this device for complaint trends and risk. according to the introducer sheath in-process and final inspection procedure, qa performs this inspection 100%: with naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded foreign material, or any other damages. the introducer and dilator passed all in-process and qa final inspection steps before shipping to the customer including visual, dimensional and mechanical testing. the instructions for use (ifu) provides these precautions: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. never advance or withdraw guide wire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. dilators, catheters, and pacing leads should be removed slowly from the sheath. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. | A follow-up report will be sent upon completion of evaluation.
38282 | 1220648-2016-00032 | 6125639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6125639&pc=OJE | 09/06/2016 | 09/06/2016 | 09/06/2016 | 11/23/2016 |  | 09/06/2016 | 09/06/2016 | 11/23/2016 |  |  | 12/18/2016 |  | 11/23/2016 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA RP |  |  | OJE |  |  | IMPELLA RP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1181611 |  | 0046-0011 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 56 YR | [5 empty] | Death | Yes | Yes | Device Operational Issue | [6 empty] | Cardiopulmonary Arrest | Low Blood Pressure/ Hypotension | Hypoxia | Respiratory Distress | [9 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 09/06/2016 | * | HOSPITAL |  |  | A (b)(6) male was admitted to facility in (b)(6) with a st elevation myocardial infarction. the right coronary artery was occluded. revascularization was attempted, however, this attempt was not successful. the patient was monitored for two days. the patient continued to deteriorate and became hemodynamically unstable, with a low blood pressure, increasing lactate levels, and an echo showing right ventricular failure. the decision was made to place an impella rp. the placement of the rp was difficult, however successful, and support was initiated. within 30 minutes of support, the impella console did alarm for low purge flow and increased purge pressure. troubleshooting did not reveal any kink in the pump. the team further troubleshot by exchanging the purge cassette. the alarm did not resolve, but support remained on and over the following hours the flow issue did resolve. approximately six hours later, the rp stopped running due to high motor current. once the rp stopped, the patient did experience severe shock with reduced oxygenation and low blood pressure. the rp was restarted and an echo was performed to observe the patient and pump function. the echo did show some movement of the rp, though the position was still across the pulmonic valve. there was evidence of massive right to left shunting (at the known pfo), poor arterial oxygenation, and rising lactic acidosis. during this time, the rp pump was flowing at 3-3.5l/minute with almost absent pulsatility. the team chose at this time to change the therapy/support to veno-arterial (va) ecmo. unfortunately, the patient did expire prior to the va ecmo loop being placed for support. according the physician the patient expired after the subsequent failure to restore pulmonary flow. | [5 empty] | The definitive root cause of the impella rp pump stop was ingestion of biomaterial. the ingested biomaterial wrapped around the impeller. the data log investigation and product testing found that spikes in motor current and high purge pressure were evident. the product was run in the lab and no pump stops occurred. this could have been due to different conditions than found in the procedure. the biomaterial was most likely interfering with the purge flow, creating the behavior/alarms seen in the case, and the eventual pumps stops. the team did not make the choice to replace the pump after experiencing the pump stops and restarting the pump. the patient expired while the va-ecmo was being setup for support. within the ifu there are instructions on what to do upon a pump stop alarm. they are: replace white connector cable. switch to backup controller. replace impella catheter. the ingested biomaterial could have been due to the patient's underlying condition of a pfo and the blood shunting. the pfo and the associated shunting could have contributed to clot formation, but this is unknown. no corrective action is recommended as the failure is due to patient condition. failures of this type will continue to be monitored and tracked. (b)(4). this mdr was one of the 13 and was originally submitted in the test environment on 10/03/2016 and the reasons stated above are the reasons it is being submitted beyond the 30 day window. (b)(4).
38283 | 1220648-2016-00033 | 6120616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6120616&pc=OZD | 09/15/2016 | 09/16/2016 |  | 11/22/2016 |  |  | 09/15/2016 | 11/22/2016 |  |  | 03/31/2018 |  | 11/22/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 103737 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 45 YR | [5 empty] | Injury | Yes | No | Difficult or Delayed Positioning | Device Issue | Patient-Device Incompatibility |  |  |  |  | Blood Loss | Patient Problem/Medical Problem | [11 empty] | Required Intervention | [5 empty] | ANGIOJET PERIPHERAL THROMBECTOMY SYSTEM | [22 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | Complainant reported that during transfer from one sister hospital to another a (b)(6) morbidly obese female patient required advanced cardiac life support (acls), which entailed 5 defibrillations. upon entry into the second hospital on (b)(6) 2016 the patient was diagnosed with an acute myocardial infarction (ami) and was in cardiogenic shock (cgs.) an angiogram was performed which revealed a clot in the left anterior descending artery. a peripheral thrombectomy was performed with an angiojet and one 4.0 bare metal stent was placed without issue. the physician attempted to place an lmpella cp via access through the left femoral artery (lfa) while another physician worked on the patient's coronary arteries. it was discovered during this emergent treatment that the patient's profunda was actually accessed, not the lfa. the physician stated that the patient's obesity played a role in accessing the arteries and coronaries. as the 14fr oscor sheath was peeled away and as the repositioning sheath was advanced the physician continued to encounter issues advancing the repositioning sheath. and found that the sheath would not advance. the physician stated that the issues with advancing the sheath were as a result of the patient's extreme obesity. upon the removal of the repositioning it was found that the sheath was peeled away and was shredded at the tip. the physician went contralateral and performed an angiogram to the lfa and the superficial femoral artery (sfa) and found that the profunda had been perforated. the team performed a contralateral tamponade to control bleeding, and then they tacked up the perforated profunda with 2 icast stents. no replacement blood products were needed to be administered to the patient. the patient continued to be administered norepinephrine. the team elected not to replace the lmpella cp as the patient's pressures had improved. | [5 empty] | The lmpella cp was returned for evaluation. the automatic lmpella controller data logs were also returned. the analysis of the data logs showed normal patient support for over 3 hours. the evaluation of the returned repositioning sheath revealed a large amount of tearing on the sheath, as well as a bent tip on the sheath. the damage observed to the sheath was most likely the result of patient's morbidly obese condition, also confirmed by the physician's report that sheath access was difficult because of the paitent's size. the perforation of the profunda most likely occurred while the physician was attempting to advance the repositioning sheath. the root cause of the split repositioning sheath was the patient's obese condition. the root cause of the patient's perforated profunda was the physician's attempt to advance the repostioning sheath despite the split in the sheath. a corrective action has been initiated to address this failure mode. the abiomed directions for use for the impella cp gives the following alert regarding the use of the device with an obese patient: when securing the repositioning sheath, vascular closure may be difficult in obese patients with extensive adipose tissue. internal reference: (b)(4). | The device has been returned for evaluation. the automatic lmpella controller data logs have not yet been received for the analysis to be performed in conjunction with the lmpella cp evaluation. upon the return of the automatic lmpella controller data logs the analysis and evaluation will be performed. upon completion a supplemental medwatch report containing the results and conclusions will be submitted. (b)(4).
38284 | 1220648-2016-00035 | 6121003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6121003&pc=OZD | 08/31/2016 | 08/31/2016 |  | 11/22/2016 |  |  | 08/31/2016 | 11/22/2016 |  |  | 04/30/2018 |  | 11/22/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1240020 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Malposition of Device | [6 empty] | Cardiac Tamponade | Cardiac Perforation | Blood Loss | [10 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | HOSPITAL |  |  | The complainant reported that a (b)(6) year old female patient was being treated for cardiogenic shock after having three days of chest pain prior to coming for medical attention. upon admission to the emergency room the patient's ekg did display an st segment elevation mi. upon fluoroscopy imaging in the cardiac cath lab, the patient was seen to have a chronic 100% occlusion of the left anterior descending coronary artery, a 99% occlusion in the mid circumflex artery, and a diffusely diseased right coronary artery. the left ventricle also showed diminished function. due to the 3 vessel disease and lv dysfunction, an impella cp was placed for support. the impella cp support commenced successfully, although within approximately five minutes suction was noted. in addition, due to patient hypotension acls medications were administered intravenously. the patient continued to decompensate and a full code was called with cpr being initiated. an echo and intubation were called for immediately. the echo revealed an effusion in the pericardium that necessitated escalated care. the patient was taken to the operating room. once she was placed on bypass support the physician noted the pigtail of the cp pump had perforated through the ventricle. some patient bleeding was noted, but no replacement blood products were necessary. the perforation was repaired and 3 bypass grafts were placed to the coronary arteries. the surgical and cardiac teams could not determine if the perforation could have occurred during cpr or from friable necrotic ventricular tissue. the tissue was thought to be necrotic due the delay in the patient seeking care. the patient was successfully weaned from the impella and the pump was explanted on day 3 post op. unfortunately, the patent's neurological status was determined to be impaired due to subdural infarctions and right sided swelling. the patient's care was withdrawn and she did expire on day 3 post op. | [5 empty] | The impella cp was not returned for evaluation, as it was discarded subsequent to the event, consequently, no device analysis could be performed to assist in finding a definitive root cause of the perforation the physicians were unable to determine if the perforation was caused by the narcotic portion of the left ventricle where the impella just happened to be placed or as result of the cpr that was performed, or if the perforation resulted in a combination of both. the physicians did note that all parties involved in the patient's care felt the impella played a major role in giving her a chance for survival. the impella was able to successfully support the patient post bypass surgery. the patient outcome was determined to be due to her poor neurological prognosis, which was a result of her delay in care and stroke history. the manufacturer will continue to investigate all reasonable and obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. (b)(4). device discarded by user.
38285 | 1220648-2016-00036 | 6121063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6121063&pc=OJE | 09/29/2016 | 09/29/2016 | 09/29/2016 | 11/22/2016 |  | 09/29/2016 | 09/29/2016 | 11/22/2016 |  |  | 03/31/2018 |  | 11/22/2016 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA RP US HDE PUMP SET |  |  | OJE |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 1248732 |  | 00434 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Physician |  |  | Manufacturer report | No | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Improper or Incorrect Procedure or Method | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | Unknown | Initial submission | Followup | [5 empty] | HOSPITAL |  |  | Additional information obtained from the physicians treating the patient reported that the patient subsequently expired. the physicians further reported that the diminished patient brain function was not caused or contributed to by this reported event. | Additional information obtained from the physicians treating the patient reported that the patient subsequently expired. the physicians further reported that the diminished patient brain function was not caused or contributed to by this reported event. | Complainant reported that on (b)(6) 2016 the clinicians placed an impella rp in a (b)(6) male patient following the patient's transfer to another facility for a cabg, but had then received a surgical turn down. on (b)(6) 2016 bi-ventricular assist devices (vad) were placed. between (b)(6) 2016 and (b)(6) 2016 the patient was reported to have had multiple medical issues, including multiple codes and clostridium difficile colitis (c-dif infection.) on (b)(6) 2016 the patient was taken to the cath lab were the clinicians placed stents in the lm, lad, rca and the circumflex artery. then on (b)(6) 2016 a heartwear left ventricular assist device (lvad) was placed in the patient, and on (b)(6) 2016 the right ventricular assist device (rvad) was removed from the patient. on (b)(6) 2016 the impella rp was placed in the left femoral artery, as right ventricular failure was present. impella flows were achieved. the impella rp placement was reported to have been normal and standard, although when closing the groin site the physician needed coaching on how to perform a mattress suture, and the mattress sutures were performed. following the impella rp placement the patient was also administered vasodilators and epinephrine, and sent the icu department. on the night of (b)(6) 2016 flows dropped in both the impella rp and the lvad after the patient was turned in bed. the patient's hemoglobin/hematocrit had dropped, so the patient was taken for a cat scan where a large retroperitoneal bleed of unknown origin was found. the patient was given 5 units of replacement blood, and after fluid resuscitation both impella rp and lvad flows returned to normal, and the patient's condition was stable. the plan was to stabilize the patient and leave on impella and lvad support. on (b)(6) 2016 the patient was reported to have been decompensating; the patient's hemoglobin/hematocrit had dropped and the patient was not oxygenating. the decision was made to remove the impella rp and place patient back on ecmo for oxygenation and right ventricle decompression. the impella rp was discontinued and removed. the patient required intervention of 18 units of blood product, coil embolization, and 2 surgical evacuations of the hematoma at the left femoral artery. the 18 units of replacement blood were given over a 24 hour period. additional information received from the complainant revealed that the patient was unstable due to multiple procedures and surgical interventions. |  |  |  | This supplemental report was originally reported on mislabeled medwatch report 1220648-2016-00026-01. this report is being submitted to report the information on a correctly labeled supplemental medwatch report. additional patient information was subsequently obtained and was added to section "b5" of this report. that information reported that the physicians treating the patient reported that the patient subsequently expired. the physicians further reported that the diminished patient brain function was not caused or contributed to by this reported event. section d3 was changed to reflect the correct manufacturer of the suspect medical device. the device was returned for evaluation. the automatic lmpella controller (aic) data logs were not able to be analyzed as they were unable to be returned, as they had been written over during subsequent and different patient use at the user facility. upon the return of the lmpella rp an inspection of the device was performed, which revealed that the pump to be in good condition. there were no sharp protrusions found on the repositioning sheath. the guidewire was unable to be examined as it was not returned with the lmpella rp and repositioning sheath. during the examination of the pump biomaterial was found on the impeller. the returned lmpella rp was performance tested with an aic. the data from the test revealed that the pump flow was in specification for most of the test, but was below specification for 10 seconds when the pump was initially set to p-8; during this time the pump flow was at 2.8 l/min when the expected range was 3.4-4.2 umin. flows improved to their expected range after a few seconds when the biomaterial was removed from the impeller. a review of the device history record was performed that did not reveal any anomalies, and that there had not been any other reported issues for other pumps in this lot relevant to the failure mode. the root cause of the low flow during the case was most likely low volume due to the reported bleeding. the evaluation of the returned lmpella rp was unable to definitively determine the root cause of the bleed that only became apparent after the patient was turned with the rp and lvad in place. the evaluation was unable to definitively determine the root cause of the low flow was most likely low volume due to the retroperitoneal bleed. the root cause of the bleed was most likely a femoral vein puncture during the initial stick. internal reference (b)(4). | Additional patient information was subsequently obtained and was added to section "b5" of this report. that information reported that the physicians treating the patient reported that the patient subsequently expired. the physicians further reported that the diminished patient brain function was not caused or contributed to by this reported event. section d3 was changed to reflect the correct manufacturer of the suspect medical device. the device was returned for evaluation. the automatic impella controller (aic) data logs were not able to be analyzed as they were unable to be returned, as they had been written over during subsequent and different patient use at the user facility. upon the return of the impella rp an inspection of the device was performed, which revealed that the pump to be in good condition. there were no sharp protrusions found on the repositioning sheath. the guidewire was unable to be examined as it was not returned with the impella rp and repositioning sheath. during the examination of the pump biomaterial was found on the impeller. the returned impella rp was performance tested with an aic. the data from the test revealed that the pump flow was in specification for most of the test, but was below specification for 10 seconds when the pump was initially set to p-8; during this time the pump flow was at 2.8 l/min when the expected range was 3.4-4.2 l/min. flows improved to their expected range after a few seconds when the biomaterial was removed from the impeller. a review of the device history record was performed that did not reveal any anomalies, and that there had not been any other reported issues for other pumps in this lot relevant to the failure mode. the root cause of the low flow during the case was most likely low volume due to the reported bleeding. the evaluation of the returned impella rp was unable to definitively determine the root cause of the bleed that only became apparent after the patient was turned with the rp and lvad in place. the evaluation was unable to definitively determine the root cause of the low flow was most likely low volume due to the retroperitoneal bleed. the root cause of the bleed was most likely a femoral vein puncture during the initial stick. internal reference (b)(4). | The impella rp has not yet been returned for evaluation; consequently no device analysis could be performed to assist in finding a definitive root cause of the event. the manufacturer will continue to investigate all reasonable and obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if the device and/or additional information is received. since july of this year abiomed has had emdr submission problems due to technical firewall issues. as a result abiomed was unknowingly submitting to the test environment since july resulting in 13 mdr's not being submitted to the production environment and are subsequently past the 30 day window. with the help of the emdr team this problem has been resolved. on 11/16/2016 a conference call was held between abiomed and fda and due to the uniqueness of the situation it was agreed abiomed will re-submit the 13 mdr's in the production environment and note in section h10 the reason for being beyond the 30 day window is due to the issues described above. this mdr was one of the 13 and was originally submitted in the test environment on 10/28/2016 and this is the reason it is being submitted beyond the 30 day window. (b)(4).
38286 | 1220648-2016-00027 | 6116663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6116663&pc=OZD | 08/17/2016 | 08/17/2016 |  | 11/21/2016 |  |  | 08/17/2016 | 11/21/2016 |  |  | 04/30/2018 |  | 11/21/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1239232 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | No | Fluid/Blood Leak | Device Displays Incorrect Message | [5 empty] | Perforation of Vessels | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | HOSPITAL |  |  | Complainant reported that on (b)(6) 2016 the hospital staff was treating a (b)(6) male who was admitted in an acute myocardial infarction. upon admission to the cardiac cath lab, the fluoroscopy showed left main (lm) disease and a totally occluded proximal left anterior descending (lad) artery. as the patient was hypotensive, and needed a bridge to surgery (cabg), the impella cp was placed. the cp pump was prepped in the usual fashion, the peel away oscor sheath was removed, and the repositioning sheath was advanced. shortly afterwards, the impella console alarmed low purge pressure. trouble shooting began and the discovery of a slow leak at a fissure within the 9f portion of the catheter was discovered. the team made the choice to remove the cp and replace it with another impella, this time the model selected was the 2.5. to exchange to the 2.5 impella, the team pulled the cp to the patient's ascending aorta, the anticontamination sleeve was withdrawn completely, and the cp was cut. this would allow the same access site to be utilized, thereby not accessing the contralateral leg. the repositioning sheath was then removed from the cut proximal end and pressure was held at the groin. in an effort to gain sheath access the team attempted access with a 14fr by 30cm cook, and then they attempted to place an 11fr sheath over the cp, both failed. after failing at attempting to place a .014 guidewire between the 11fr sheath and the impella cp, the cp was then removed. the team placed a third sheath, the 14fr by 13cm cook, over the .014 wire, before exchanging the wire for the abiomed .018 wire. the 2.5 impella was then successfully placed in the left ventricle for support. the repositioning sheath of the 2.5 was advanced into the 14fr sheath, but there was observed oozing of blood at the sheath and fluoroscopy showed a significant bleed/vessel damage. vascular surgery was consulted, and the decision was made to access the contralateral leg, and place the impella 2.5 there so that a vascular repair could be done. the surgeon was noted to successfully repair the artery and no lasting harm or injury to the leg. | [5 empty] | The products used in this complaint have been retained by the facility. the risk management team has not released the products for engineering analysis to abiomed. without the product being returned no evaluation could be made of the fissure noted in the 9fr portion of the impella cp catheter. no photographs of the cp pump were provided either to analyze, nor any imaging of the vascular bed and injury. the manufacturer will continue to investigate any obtainable source information from the complainant, should that become available. (b)(4). device not returned to manufacturer.
38287 | 1220648-2016-00028 | 6116867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6116867&pc=OZD | 08/19/2016 | 08/19/2016 |  | 11/21/2016 |  |  | 08/19/2016 | 11/21/2016 |  |  | 04/30/2018 |  | 11/21/2016 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1240149 |  | 005042 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 79 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Blood Loss | [11 empty] | Life Threatening | Required Intervention | [27 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | On (b)(6) 2016 a (b)(6) female was placed on support with an impella 2.5, after the cardiothoracic surgeon declined to treat the patient, and the decision was made to perform a high risk percutaneous angioplasty intervention (hrpci). prior to placement of the impella 2.5 the physician needed to balloon the iliacs at the bifurcation to allow placement and advancement of the 2.5. the patient had bilateral peripheral vascular disease. furthermore, there was difficulty in ascending through the patient's aorta. the patient was known to have a porcelain aorta and be heavily calcified. with the 2.5 in place for support, the physician was able to stent the lad and left main. the pump position was adjusted prior to discharge to the icu. upon admission to the icu, a bedside echo was done which revealed the 2.5 to be located within the papillary muscle of the left ventricle. the pump therefore was readjusted. also, noted on the echo was the presence of an effusion. the team chose to perform a pericardiocentesis (tap) procedure bedside. before the tap procedure could begin the patient decompensated and needed intubation for support. cardiothoracic surgery was called to perform a bedside pericardial window. due to the critical nature of the decompensating patient, the team instead took the patient to the or for the window procedure. due to the patient's continued decline, the team chose to open the chest and to directly assess the location of the bleeding. the surgical team discovered a small hole at the apex of the heart. they attributed the hole to the v18 wire, manufactured by boston scientific. the v18 wire is an abiomed approved wire that can be used in conjunction with the impella 2.5. the hole was repaired prior to discharge from the or. the impella 2.5 pump continued to support the patient successfully throughout their icu stay. during the overnight stay in the icu blood replacement products were infused for the patient. the patient was then successfully weaned from the 2.5. the patient was hemodynamically stable with no further need for the pump nor any inotropic support. | [5 empty] | The root cause of the left ventricle perforation can not be definitively attributed to the impella pump. the size of the hole/perforation was smaller than the components on the impella. it is unlikely that the impella 2.5 pump caused the perforation. the team at the facility believes the v18 wire caused the perforation. the patient did not suffer any symptoms until after the patient was discharged to the icu on support. the symptoms did not present during the intervention nor during the placement of the impella 2.5. the root cause of the perforation is unable to be determined. it is unlikely that the impella caused the perforation, but this failure mode will be monitored and trended. no corrective action is necessary. (b)(4).
38288 | 1220648-2016-00029 | 6118513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6118513&pc=OZD | 08/23/2016 | 08/24/2016 |  | 11/21/2016 |  |  | 09/09/2016 | 11/21/2016 |  |  | 06/30/2018 |  | 11/21/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 124-923 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 50 YR |  |  | 99 |  |  | Injury | Yes | Yes | Fracture | Material Separation | Deformation Due to Compressive Stress |  |  |  |  | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | HOSPITAL |  |  | The complainant reported that on (b)(6) 2016, a (b)(6) year old male patient presented in the emergency department with pulmonary edema, and shortly thereafter suffered a 2 minute code blue and developed cardiogenic shock. the patient was then transferred to the facility's cath lab for impella cp placement. the pump was placed in the patient's right groin, and although the patient was sedated he became agitated and began thrashing his leg. the interventional cardiologist reported that he then advanced the impella forward and it went across the aortic valve and back into the lv. the patient again moved and displaced the device into the aorta. the impella alarm then sounded, and the clinicians found that the cannula had kinked. the physician then attempted to remove the device from the right groin without first trying to straighten out the prolapsed catheter. the catheter was then pulled all the way down to the common femoral artery (cfa). during this attempt, everything exited the patient's groin except the inflow cage and pigtail which had become detached from the rest of device at the location of the kink. the portion of the device remaining in the patient appeared to be in the profunda. wire access was maintained to right groin while the physician advanced a 14fr sheath from a second cp and placed it in the common femoral artery (cfa). pressers and inotropes were restarted in an effort to manage pressure while the patient was without impella support. the doctor then tried to snare the pigtail from the right groin, but without success. he then accessed the left groin with a 6fr sheath in an attempt to snare the detached portion of the cannula contralaterally. at this point snares were manipulated bilaterally to orient the inflow and pigtail portion of the cannula, which resulted in successful removal of the detached cannula via the oscor 14fr sheath. the patient experienced bleeding as a result of the staff moving the oscor sheath in and out of the groin to manipulate the pigtail and inflow cage back into the iliac where it could better be manipulated for removal. the blood loss was reported to be from non-occlusive proximal pressure being held and released to manipulate the groin. it was not a continuous blood loss but was brisk incrementally. the patient received 2 units of replacement blood. a second cp was then successfully placed and secured, and the patient's right leg was immobilized, and the patient was reported to be in stable condition. the patient was later transferred to another hospital for vad support to bridge to heart transplant. | [5 empty] | The impella cp and the automatic impella console data logs were returned to the manufacturer for evaluation. the analysis of the data logs showed that the motor current suddenly lost pulsatility, which was never regained, which revealed when the initial kink occurred, and that the kink worsened shortly after when the motor current and calculated flow fell. the pump was returned without the distal end of the cannula, or the inflow cage and pigtail. the cannula was found to be torn and the nitinol coil had unraveled at the torn end. two kinks were found along the cannula closer to the pump housing. the cannula became kinked when the cardiologist attempted to wirelessly re-cross the atrioventricular after the pump came out of position due to patient movement. the cannula then became torn and separated when the pump was being explanted with the cannula severely kinked. the explant was reported to have been attempted faster and more forcefully than would be expected with a severely kinked cannula. the root cause of the cannula separation was determined to be the result of excessive force being applied during explant of a pump with a severely kinked cannula, which had become kinked during a failed attempt to wirelessly re-cross the aortic valve. a corrective action has been initiated to address this failure mode. in addition, re-training of the facility staff was performed on august 24, 2016. (b)(4).
38289 | 1220648-2016-00031 | 6119035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6119035&pc=OJE | 08/31/2016 | 08/31/2016 |  | 11/21/2016 |  |  | 08/31/2016 | 11/21/2016 |  |  | 10/31/2018 |  | 11/21/2016 | 11/21/2016 | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA RP US HDE PUMP SET US |  |  | OJE |  |  | IMPELLA RP US HDE PUMP SET |  |  | ABIOMED, INC |  |  | US |  | 1212833 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 63 YR | 63 YR |  |  |  |  | Malfunction | No | Yes | Detachment Of Device Component | Material Separation | Device Issue | Physical Resistance |  |  |  | Death | Thrombus | Device Embedded In Tissue or Plaque | [34 empty] | IMPELLA 5.0, S/N (B)(4) |  |  |  |  | Initial submission | Followup |  |  |  |  | * | HOSPITAL |  |  | The complainant reported that they were treating a very large patient with significant amounts of adipose tissue between the vein and skin. this patient had known cardiac disease and was in the hospital for mitral valve repair. on (b)(6) 2016 an impella 5.0 was placed in the patient for left side support. on (b)(6) 2016 it was determined that the patient was in need of escalated support due to right ventricle failure; consequently, an impella rp was also placed. access and placement of the rp was done in the usual fashion. pump flows were not optimal due to an extremely low pulsatility and low mean arterial pressures. to augment support and improve flow the patient received multiple transfusions of red blood cells (prbc), plasma (ffp), and albumin. flows improved, as did mean arterial pressure. the physician chose not to remove the oscor peel away sheath at this time, so that they would have arterial access should the patient require escalation of care with an ecmo at a later time. the patient was then discharged to the icu for care. in the overnight, the patient was monitored to see if more blood and volume would be needed. the chest tubes were also monitored for output, and the patient did demonstrate progress on support. the iv medications were able to be weaned down. due to this fact, the team decided to peel away the oscor sheath and to position the impella rp's repositioning sheath, as an ecmo was no longer felt to be necessary. the patient was provided dialysis during the day for volume control. during the second day the patient demonstrated improvement & was provided with dialysis. during the second evening of impella rp support, the pump flows were reducing. more volume was given in the form of prbcs and albumin. unfortunately, flows did not improve and the patient was deteriorating with rising lactate numbers (ldh), and was becoming coagulopathic and acidotic. the team also performed a chest x-ray to confirm proper position within the right ventricle. because the pump was in the proper position, the decision was made to explant the impella rp. upon explant, there was resistance felt by the physician. upon inspection it was discovered that the pump had come apart into two pieces, with the pump and motor portion remaining lodged in the femoral vein. as the patient was at risk with the coagulopathy to undergo a surgical removal of the rp, the team decided to wait until the patient's bleeding risk improved. while the team was awaiting an improvement in the bleeding risk, the patient's family chose to withdraw the patient's care. the patient expired and the detached portion of the rp cannula remains within the patient post-mortem. | [5 empty] | The impella rp was returned without the cannula, as the cannula was apparently discarded subsequent to use. a visual inspection of the returned portion of the device was performed. the inspection revealed that the delo adhesive had filled the etchings of the laser-etched bonding surface, but there was minimal adhesive coverage over the rest of the surface. the evaluation of the returned portion of the impella rp revealed that the pump most likely separated during explant due to the pressure applied at the venous access site and the steep insertion angle required by the patient's size. the patient had been reported by the complainant as very large, and with significant amounts of adipose tissue between the vein and the patient's skin. manual pressure is applied during explant from arterial sites, and not as regularly at venous sites. the level of pressure applied was increased due to the depth of the femoral vein due to the patient's size. the depth of the patient's vein would also require a steeper insertion angle. the increased load required to explant the catheter against the manual pressure through the sharp angle of venous access may have caused the failure of the delo adhesive joint, which caused the cannula to detach from the pump. the root cause of the cannula detachment was most likely the load created by the steep insertion angle and high manual pressure applied during explant due to the patient's size. corrective action: a capa was opened to address rp cannula detachments. failures of this type will continue to be monitored and tracked. (b)(4) supplement. | The device, with the exception of the detached portion of the cannula, has been returned for evaluation; the evaluation is currently being performed. upon the completion of the evaluation of the impella rp a supplemental medwatch report containing the results and conclusions will be submitted. resolved. (b)(4).
38290 | 1220648-2016-00024 | 6114602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6114602&pc=PBL | 08/04/2016 | 08/04/2016 | 08/04/2016 | 11/18/2016 |  | 08/04/2016 | 08/04/2016 | 11/18/2016 |  |  | 02/28/2018 |  | 11/18/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1233151 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Device Operates Differently Than Expected | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/04/2016 | * | HOSPITAL |  |  | Complainant reported that on (b)(6) 2016 the hospital staff was treating a (b)(6) admitted with cardiomyopathy, hypertension, and acute renal failure who now presented with elevated labs, edema, and shortness of breath. patient was treated with lasix and an intraaortic balloon pump. as there was a need for escalated care due to symptoms worsening, the iabp was removed and an impella cp was placed at the same groin access site. support was seen be successful with good flows and the patient having an easier time breathing. on the next day of support, after a restless night, there was an observed change in the urine output color. a urinalysis was performed and labs pointed to hemolysis. to troubleshoot this issue the catheter position was checked under echo and repositioned. later in the evening of (b)(6), the staff did remark that the patient was being ruled out for an ileus due to abdominal pain and distension. the patient's condition was monitored and care was again escalated with an increase in inotropes on the following day, (b)(6). also, the physician decided to give the patient three units of blood replacement products for the hemolysis. the physician did note that the patient's condition could contribute to the hemolysis, along with the use of impella support. the decision was made to wean the patient off of impella support and explant the device. due to the patient being made a dnr and not being a candidate for a vad device, there is only care and comfort being provided. the patient is being monitored and supported with multiple cardiac medications, including dopamine, levophed, dobutamine, and milrinone. | [5 empty] | The investigation of product, returned by the hospital, included running the pump through internal hemolysis testing in porcine blood twice over four days. serial plasma free hemoglobin labs were drawn during these three hour tests. the lab results all were below the value which is used to determine hemolysis on the first day of testing. when the hemolysis testing was repeated days later, the results were positive for hemolysis. in addition to the results being of a value that confirmed hemolysis, a defect was seen at the outflow cage that could have contributed to or caused the hemolysis. the defect could protrude and cause obstructed flow. the delaminated area from the sleeve bearing was raised into the outflow cage window. of note, no plasma free hemoglobin labs were drawn at the hospital. hemolysis was only documented via a standard cbc and urinalysis, as well as an observation of the urine changing in color and output. in addition to the porcine blood testing, the data logs were analyzed for signs of hemolysis. the logs show the device provided good flow throughout the support. the logs did show suction alarms, however each of these alarms rectified quickly, and can be accounted to the patient's pressure lowering in response to inotropic support. the use of the impella was the most likely cause of the hemolysis. however, the physician stated to the representative that the patient condition could contribute to the hemolysis. as stated in the event description, the patient had multiple contributing cardiac and renal conditions. no corrective action is recommended. the outflow cage defect that prompted the patient to experience hemolysis, with the delamination failure, is of low risk and rare. (b)(4).
38291 | 1220648-2016-00025 | 6115018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6115018&pc=OZD | 08/06/2016 | 08/06/2016 | 08/06/2016 | 11/18/2016 |  | 08/06/2016 | 08/06/2016 | 11/18/2016 |  |  | 09/30/2017 |  | 11/18/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 1216120 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 61 YR | [5 empty] | Malfunction | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 08/06/2016 | * | HOSPITAL |  |  | Complainant reported that following neck surgery performed on (b)(6) 2016 a (b)(6) male patient, with a history of heart disease, coded in the unit. the patient was then brought to the cardiac catheterization laboratory (ccl) where stents were placed while the patient was being supported by an impella cp after insertion into the patient's right groin. the following day the patient was transferred to another facility. the aic console was transferred to the 2nd facility's console, and the impella cp continued to support the patient. on (b)(6) 2016 the patient was brought to the operating room where a veno-arterial ecmo (va-ecmo) was placed in the patient's left groin. during the placement of the va ecmo the patient's leg on the impella side was observed to be ischemic and mottled. the ecmo flows were recorded as 4.31 liters/minute and the impella flows were recorded as 4.0 liters/minute. the impella was left in the patient to vent. during the weekend ((b)(6) 2016) the patient's partial thromboplastin time (ptt) was 90. on (b)(6) 2016 the ptt reading was 25u/ml. on (b)(6) 2016 the patient's status was unchanged. on this date it was also confirmed that the patient was neurologically intact. at this time the ecmo flows were at 3.0 liters/minute and the impella flows were at 1.3 liters/minute. the patient's impella leg was cold and still mottled. the staff planned to wean the ecmo and remove the impella cp "within the next couple of days." the vascular surgeons were hesitant to remove the impella cp at that time, as they felt that ecmo support alone would lead to many fasciotomies. on the night of (b)(6) 2016 the ecmo and impella were removed from the patient. the patient was reported to be in stable condition following the removal of the ecmo and pump. it was also reported that the impella cp had successfully supported the patient. on (b)(6) 2016 the physicians performed a "high above the knee amputation" of the patient leg. the physician reported the patient was "lucky to be alive", as the patient was "very sick on pressors and the impella cp." | [5 empty] | The impella cp was not returned for evaluation, so no device analysis could be performed. it is unclear at this time if the ischemia was as a result of any issue with the impella cp, or if the ischemia occurred as a result of the patient's condition. as reporeted in this report's event description, the physician reported that this patient was "lucky to be alive" and was "very sick" on pressors and the impella. further determination of the device use and the malfunction could not be determined as the pump was not returned. a review of the complaint history data revealed that there have been no other complaints reported for this impella cp, lot number 1205033. the root cause of this event was unable to be determined because the impella cp was not returned for evaluation. because the root cause could not be determined, no corrective action was implemented. the manufacturer will continue to investigation all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. since july of this year abiomed has had emdr submission problems due to technical firewall issues. as a result abiomed was unknowingly submitting to the test environment since july resulting in 13 mdr's not being submitted to the production environment and are subsequently past the 30 day window. with the help of the emdr team this problem has been resolved. on 11/16/2016 a conference call was held between abiomed and fda and due to the uniqueness of the situation it was agreed abiomed will re-submit the 13 mdr's in the production environment and note in section h10 the reason for being beyond the 30 day window is due to the issues described above. this mdr was one of the 13 and was originally submitted in the test environment on and this is the reason it is being submitted beyond the 30 day window. internal reference: (b)(4). device discarded by customer.
38292 | 1220648-2016-00026 | 6115053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6115053&pc=OZD | 08/11/2016 | 08/11/2016 | 08/11/2016 | 11/18/2016 |  | 08/11/2016 | 08/11/2016 | 11/18/2016 |  |  | 03/31/2018 |  | 11/18/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP PUMP SET, CANADA |  |  | ABIOMED, INC |  |  | US |  | 1235314 |  | 0048-0004 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 75 YR | [5 empty] | Death | Yes | No | Kinked | Improper or Incorrect Procedure or Method | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Death | Perforation of Vessels | [11 empty] | Death | [28 empty] | Initial submission | Followup | Followup |  |  |  | * | HOSPITAL |  |  | Additional information obtained from the physicians treating the patient reported that the patient subsequently expired. the physicians further reported that the diminished patient brain function was not caused or contributed to by this reported event. | The complainant reported that on august 11, 2016 a (b)(6) woman with known severe coronary artery disease (cad) was undergoing an elective high risk percutaneous coronary intervention (hrpci.) the impella cp was placed prior to the angioplasty procedure beginning. during the wireless insertion of the impella cp there was an kink deformity at the pigtail, that was never remedied. the angioplasty procedure then commenced. the left main, left anterior descending, and circumflex arteries were ballooned and stented. the patient decompensated and became pump dependent during the angioplasty. an echo was immediately performed which revealed the patient in tamponade. the left ventricle was found to have a perforation. a pericardiocentisis was reported to have been successfully performed. the patient continued to bleed, then thrombosed and finally arrested prior to an ecmo being placed. |  |  |  |  | Additional patient information was subsequently obtained and was added. that information reported that the physicians treating the patient reported that the patient subsequently expired. the physicians further reported that the diminished patient brain function was not caused or contributed to by this reported event. the device was returned for evaluation. the automatic impella controller (aic) data logs were not able to be analyzed as they were unable to be returned, as they had been written over during subsequent and different patient use at the user facility. upon the return of the impella rp an inspection of the device was performed, which revealed that the pump to be in good condition. there were no sharp protrusions found on the repositioning sheath. the guidewire was unable to be examined as it was not returned with the impella rp and repositioning sheath. during the examination of the pump biomaterial was found on the impeller. the returned impella rp was performance tested with an aic. the data from the test revealed that the pump flow was in specification for most of the test, but was below specification for 10 seconds when the pump was initially set to p-8; during this time the pump flow was at 2.8 l/min when the expected range was 3.4-4.2 l/min. flows improved to their expected range after a few seconds when the biomaterial was removed from the impeller. a review of the device history record was performed that did not reveal any anomalies, and that there had not been any other reported issues for other pumps in this lot relevant to the failure mode. the root cause of the low flow during the case was most likely low volume due to the reported bleeding. the evaluation of the returned impella rp was unable to definitively determine the root cause of the bleed that only became apparent after the patient was turned with the rp and lvad in place. the evaluation was unable to definitively determine the root cause of the low flow was most likely low volume due to the retroperitoneal bleed. the root cause of the bleed was most likely a femoral vein puncture during the initial stick. (b)(4). | The impella cp was returned for evaluation. the automatic impella controller data logs were also returned. the analysis of the data logs showed normal patient support until the end of the logs. at that point the patient's pressure decreased, showing worsening of the patient's condition. a fluoroscopy dvd of the case was provided by the hospital for analysis. the dvd provided much information pertaining to the angioplasty procedure and impella usage, as it revealed that upon placement of the impella cp there was on an obvious kink and deformity of the pigtail portion of the impella. although some repositioning was done, the kink was never remedied. the pump was not pulled back nor was it inserted over a guidewire. the dvd further revealed the actual perforation and the associated escape of contrast. the team did attempt to rectify the perforation by performing the tap procedure to pull off the volume of blood from around the heart and relieve the associated pressure. of note, on the fluoroscopy, the coronary arteries no longer had good perfusion and flow, and the team was actively performing cpr. the inspection of the returned pump identified a kink in the pump's cannula pigtail. upon the review of the fluoroscopy video it was obvious that this was a very complicated case. multiple wires and stents were used during this procedure; however, it cannot be ruled out that the impella cp may have contributed to the perforation. (b)(4). | This supplemental report is being submitted to correct the information that was inadvertently submitted on medwatch supplemental report 1220648-2016-00026-01. that supplemental medwatch report's information did not apply to medwatch report 1220648-2016-00026-01, but should have been reported and submitted on supplemental medwatch report 1220648-2016-00036-01. this supplemental medwatch report is requesting that supplemental medwatch report 1220648-2016-00026-01 be disregarded.
38293 | 1220648-2016-00038 | 6110868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6110868&pc=OZD | 10/18/2016 | 10/18/2016 | 10/18/2016 | 11/17/2016 |  | 10/18/2016 | 10/18/2016 | 11/17/2016 |  |  | 06/30/2018 |  | 11/17/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1249784 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 51 YR | [5 empty] | Injury | Yes | No | Material Frayed | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | HOSPITAL |  |  | The complainant was treating a (b)(6) female admitted for coronary artery bypass (cabg) surgery and mitral valve repair. the patient had a previous non-stemi myocardial infarction. the impella cp was inserted successfully via the left femoral artery and the patient was sent to the icu. upon admission to the icu there was observed oozing of blood from the two groin sites, the bypass access site as well as the impella site. of note, the patient had been on aggrastat, a blood thinner, for multiple days. the bleeding/oozing was controlled with manual compression as well as repositioning of the impella repositioning sheath. in the overnight the bleeding did cease, though started again on day 2 of support. the physician chose at this time to place a stitch to stop the bleed and the blood loss slowed to a slight ooze. the ooze continued til day 4 of support when the cabg was performed. to intervene and remedy the blood loss, the patient did receive two units of blood products. in the subsequent days of support with the impella cp pump the patient was taken back to the operating room for washout due to bleeding from the chest post-cabg. on day six of support the patient was successfully weaned from the impella cp and the device was explanted. the groin access site was noted to be dry with no bleeding or ooze present. | [5 empty] | The impella cp and the data logs were returned for evaluation. the inspection of the impella cp revealed a tear ti be present within the outer/second layer of the repositioning sheath. this tear was the root cause of the bleeding at the impella insertion site. the data logs were reviewed and matched with the clinical account of the case and patient condition. the investigation findings did not warrant a corrective action as the failure mode is of a low risk. these type of failures will continue to be monitored and tracked. (b)(4).
38294 | 6105541 | 6105541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6105541&pc=KFM | 08/20/2016 | 09/29/2016 | 09/29/2016 | 11/16/2016 | 09/29/2016 | 09/29/2016 |  | 11/16/2016 |  |  |  |  | 11/16/2016 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA |  |  | KFM |  |  | 5.0 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No answer provided | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 64 YR | [5 empty] | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | Device Displays Incorrect Message | Mechanical Jam |  |  |  | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Perfusion team responded due to impella alarm for high purge pressure (1098). team followed corrective procedures as indicated by company. purge cassette changed and was working properly. purge pressure came down but still found to be high. perfusion team responded a second time when the main pump stopped. team performed several hard resets on the device, but could not get pump to return to normal function. cardiologist notified and determined best course of action was to replace device. patient take to cath lab for impella exchange which was successful. team believes that the pump motor seized due to particulate matter.
38295 | 1220648-2016-00037 | 6090474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=6090474&pc=OZD | 10/10/2016 | 10/10/2016 | 10/10/2016 | 11/09/2016 |  | 10/10/2016 | 10/10/2016 | 11/09/2016 |  |  | 06/30/2018 |  | 11/09/2016 |  | IMPELLA |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 1248599 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Malfunction | Yes | Yes | Detachment Of Device Component | Kinked | Device-Device Incompatibility |  |  |  |  | Ventricular Fibrillation | Ventricular Tachycardia | Irregular Pulse | [10 empty] | Death | Life Threatening | Required Intervention |  |  |  | COOK 14FRX 30CM CHECK-FLOW PERFORMER | [22 empty] | Initial submission |  |  |  |  | 10/10/2016 | * | HOSPITAL |  |  | The complainant reported that on (b)(6) 2016, a (b)(6) year old male was being treated for an acute myocardial infarction and cardiogenic shock. the patient was brought to the cardiac cath lab for care and an impella cp was prepped in the usual fashion. upon insertion into the cook 14fr x 30 cm introducer sheath, there was an inability to pass the cp through the sheath and into position for support. due to the need for support, a second cook sheath was then inserted. once again there was difficulty inserting the cp. two attempts were made to position, and during these attempts the cook sheath was observed to be kinked. the physician pulled back upon the impella cp to remove it, and with the exerted force, did fracture the impella cp into two distinct pieces. the portion of the cp with the inlet cage and pigtail remained lodged within the cook sheath. at this time due to the intraprocedural instability, the physician attempted to access the contralateral leg with a new sheath and place a new impella cp. the patient was unstable during this time with runs of ventricular tachycardia (vt) and ventricular fibrillation (vf). the patient expired during this time due to cardiac instability. the patient had lost all rhythm and was exhibiting pulseless electrical activity (pea). the team did relay to the representative that had the patient survived, the team would have utilized a snare to retrieve the portion of the impella cp lodged within the kinked cook sheath. due to the patient's demise, the portion was not retrieved for analysis. | [5 empty] | The portion of the impella cp that was retrieved was returned for evaluation. the portion that was reported to be lodged within the cook sheath was not returned for evaluation. the entire second impella cp that was prepped for insertion into the contralateral access/leg was also returned. the data logs were not returned for evaluation as the impella cp had never been turned on and support was never initiated. the first cook sheath that was placed in the patient was returned for inspection. the visual inspection of the fractured impella cp confirmed the report that separation had occurred at the inlet cage. the first cook sheath was returned for inspection, and was inspected revealing a kink in two separate locations and extensive damage along the length. the second impella cp was found to be in good condition. after inspection and testing the root cause of the impella cp's inability to pass through the cook sheath was determined to be the kink within the sheath. the root cause of the kink could not be determined. the kink location was confirmed by analysis of the fluoroscopy image provided by the hospital. the patient was not known to be morbidly obese or have any abnormality in the vascular path/anatomy that may have caused or contributed to the kinking. as the root cause of the kinked cook sheath could not be determined there is no corrective action being initiated. the failure mode will continue to be monitored. (b)(4).
38296 | 1035166-2016-00160 | 5970983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5970983&pc=DYB | 08/08/2016 | 10/20/2016 |  | 09/22/2016 |  |  | 08/31/2016 | 09/22/2016 |  |  | 04/01/2019 |  | 09/22/2016 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-12451 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 93 YR | [5 empty] | Injury | Yes | Yes | Backflow | Break | [5 empty] | Blood Loss | No Known Impact Or Consequence To Patient | [11 empty] | Required Intervention | [5 empty] | IMPELLA CP | [22 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The customer reported valve damage during use of this device which caused the bleeding observed in the case. | On september 22, 2016 the following new information was received from the customer: how much bleeding was observed? noted to be significant but hard to quantify as under the drape. how did the physician proceed when the bleed was recognized? hematocrit had fallen from 43 to 27 so blood products given (2 units). was the device exchanged for another unit? explanted in cath lab. did the doctor experience any difficulty with insertion of the dilator? no issues noted. what type of procedure was being performed? hrpci as turned down for cabg. rotablator of the lad and stent x2. age and sex of patient (b)(6) year old female. |  |  |  |  | The device was in use for treatment. a review of the device history record identified these rejects during final packaging one (1) foreign material and one (1) ripped outer bag. a review of complaints against this lot number identified no additional reports. the investigation for this event is in process. a follow-up report will be sent upon completion of the investigation. additional information has been requested. | On september 22, 2016, the following new information was received from the customer: how much bleeding was observed? noted to be significant but hard to quantify as under the drape. how did the physician proceed when the bleed was recognized? hematocrit had fallen from 43 to 27, so blood products given (2 units). was the device exchanged for another unit? explanted in cath lab. did the doctor experience any difficulty with insertion of the dilator? no issues noted. what type of procedure was being performed? hrpci as turned down for cabg. rotablator of the lad and stent x2. age and sex of patient (b)(6) year old female. date received by manufacturer: 09/22/2016. type of report: follow-up # 1. type of reportable event: serious injury. if follow-up, what type: correction. device evaluated by manufacturer: yes. device manufacture date: 04/22/2016. labeled for single use: yes. corrected data: upon evaluation of the returned product, it was found that the hemostasis valve was torn. it is believed this was caused by off center dilator/device insertion during use in the field. the potential effect to patient for the reported failure may be related to: bleeding, prolonged intervention. the potential cause(s) for the reported failure may be related to: improper overmolding, improper extrusion, damaged during packaging/transit. a review of risk for this failure mode identified the risk to be medium. based on the investigation information, a capa is not required. per qa inspection procedure, during in process inspection, the introducer sheaths are 100% leak tested. in addition, the cap and seal are 100% inspected and an occlusion test is performed. the instructions for use (ifu) informs the user: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. infusion through the sideport can be done only after all air is removed from the unit. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheter, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. in addition, the directions for use state inform the user to aspirate all air from the sheath valve assembly by using a syringe connected to the sideport. flush the sheath with 5 cc of saline immediately before peeling the sheath away in order to minimize back bleeding.
38297 | 1035166-2016-00161 | 5971586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5971586&pc=DYB | 08/12/2016 | 10/20/2016 |  | 09/22/2016 |  |  | 09/06/2016 | 09/22/2016 |  |  | 04/01/2019 |  | 09/22/2016 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-12451 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 85 YR | [5 empty] | Injury | Yes | No | Backflow | [6 empty] | Aortic Valve Stenosis | Blood Loss | [11 empty] | Required Intervention | [5 empty] | IMPELLA CP | [22 empty] | Initial submission | Followup |  |  |  |  | * | Invalid/Incomplete |  |  | The customer reported an impella cp was prepared to support an (b)(6) year old female patient with moderate aortic stenosis for elective hrpci. during preparation, the dilators were inserted slowly through the hemostatic valve of the 14 fr oscor introducer to ensure proper valve lubrication. the hemostatic valve began leaking once the impella cp was inserted. bleeding was minimized by advancing/repositioning the sheath through the valve. by the end of the case, the patient's hemoglobin level was down to 80, from 105 at the beginning. one unit of packed red blood cells was administered. | [5 empty] | The device was used in treatment. a review of the device history record identified these rejects during final packaging one (1) foreign material and one (1) ripped outer bag. a review of complaints against this lot number identified one additional complaint for the same failure. | (b)(4). upon evaluation of the returned product, it was found that the hemostasis valve was torn. it is believed this was caused by off center dilator/device insertion during use in the field. the potential effect to patient for the reported failure may be related to: bleeding, prolonged intervention. the potential cause(s) for the reported failure may be related to: improper overmolding, improper extrusion, damaged during packaging/transit. a review of risk for this failure mode identified the risk to be medium. based on the investigation information, a capa is not required. per qa inspection procedure, during in process inspection the introducer sheaths are 100% leak tested. in addition, the cap and seal are 100% inspected and an occlusion test is performed. the instructions for use (ifu) informs the user: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. infusion through the sideport can be done only after all air is removed from the unit. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheter, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. in addition, the directions for use state inform the user to aspirate all air from the sheath valve assembly by using a syringe connected to the sideport. flush the sheath with 5 cc of saline immediately before peeling the sheath away in order to minimize back bleeding.
38298 | 1035166-2016-00163 | 5972176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5972176&pc=DYB | 08/25/2016 | 08/09/2018 |  | 09/22/2016 |  |  | 09/07/2016 | 09/22/2016 |  |  | 03/01/2019 |  | 09/22/2016 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-12403 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION | [10 empty] | Malfunction | Yes | No | Backflow | Break | [5 empty] | Blood Loss | [12 empty] | Required Intervention | [5 empty] | IMPELLA CP | IMPELLA CP | [21 empty] | Initial submission | Followup | Followup | Followup |  |  | * | Invalid/Incomplete |  |  | A cp was placed for an er patient in cgs. the cp was on auto and running when the rep arrived. the icu check in was completed and then the patient was flown to the large center. the md and staff had stated that the oscor sheath separated during impella insertion and there was significant blood loss. a medtronic 14f x 28 cm sheath was used and the rep reminded the clinical team that a cp can only go through an oscor or cook 14f x 30 cm. | [5 empty] | The customer informed oscor the orange split cap that came off during use by the physician was not returned. the device was initially received by the customer with the other orange split cap on the device, and was removed for their own internal investigation. this is possibly how the damage/crack occurred. adhesive was observed on the returned pieces (both the hub and orange split cap). based upon the investigation, a corrective and preventive action has been initiated to address this issue. oscor will continue to monitor this device for complaint trends and risk. the event will be re-evaluated if additional information becomes available. | The customer informed oscor the orange split cap that came off during use by the physician was not returned. the device was initially received by the customer with the other orange split cap on the device, and was removed for their own internal investigation. this is possibly how the damage/crack occurred. adhesive was observed on the returned pieces (both the hub and orange split cap). based upon the investigation, a corrective and preventive action has been initiated to address this issue. oscor will continue to monitor this device for complaint trends and risk. the event will be re-evaluated if additional information becomes available. | The device was used in treatment. a review of the device history record identified two (2) rejects during manufacturing, for sheath out of specification against this lot number. a review of complaints against this lot number identified no additional complaints. the investigation for this event is in process. a follow-up report will be sent upon completion of the investigation. additional information has been requested. | (b)(4). the device has not been returned for analysis; as a result, the allegations against the device (bleed back) cannot be confirmed. the potential effect to patient for the reported failure may be related to: bleeding, prolonged intervention. the potential cause(s) for the reported failure, may be related to: improper overmolding, improper extrusion, damaged during packaging/transit. a review of risk for this failure mode identified the risk to be medium. based on the investigation information, a capa is not required. per qa inspection procedure, during in process inspection the introducer sheaths are 100% leak tested. in addition, the cap and seal are 100% inspected and an occlusion test is performed. the instructions for use (ifu) informs the user: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. infusion through the sideport can be done only after all air is removed from the unit. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheter, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. in addition, the directions for use state inform the user to aspirate all air from the sheath valve assembly by using a syringe connected to the sideport. flush the sheath with 5 cc of saline immediately before peeling the sheath away in order to minimize back bleeding. | Returned device analysis reveals one 14f abiomed introducer sheath with the caps and hemostasis valve detached from the hub. there was evidence of adhesive present on the caps and hub post cleaning with enzymatic detergent and disinfectant. it was, however, difficult to conclude whether or not there was enough adhesive applied to the hub/cap of the introducer sheath post cleaning/disinfecting. there was also a crack in the side of one of the orange caps indicating that excessive force was applied to the sheath. according to abiomed introducer sheath in-process and final inspection: verify split caps are properly placed and secured onto sheath hub and free of cracks/damages or excessive adhesive. qa performs this inspection on 100% of the product. the reason for return was confirmed, however, it is uncertain whether this issue was a manufacturing defect or end user damage. there were no manufacturing rejects or anomalies of this type recorded in the device history record. the introducer sheath passed all in-process and qa final inspection steps before shipping to the customer including visual, dimensional and mechanical inspection.
38299 | 1035166-2016-00151 | 5935692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5935692&pc=DYB | 07/26/2016 | 08/11/2016 |  | 09/08/2016 |  |  | 08/11/2016 | 09/08/2016 |  |  | 12/01/2018 |  | 09/08/2016 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | OR-04736 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 83 YR | [5 empty] | Injury | Yes | No | Backflow | [6 empty] | Hematoma | Blood Loss | [11 empty] | Required Intervention | [5 empty] | IMPELLA CP | [22 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The customer reported that during a percutaneous coronary intervention, bleeding was noticed at the end of the case when the drape was being removed. it was further reported that the patient received two (2) units of blood after a hematoma developed from changing out the introducer to a different sheath. | [5 empty] | The device was in use for treatment. a review of complaints for this lot number identified no additional complaints. returned device analysis reveals one 14f abiomed introducer sheath that is within manufacturing specifications. upon further evaluation of the returned product, a slow leak was found when tested with a syringe where the sideport attaches to the hub. this area of the sheath was inspected under a microscope and the adhesive was found to be cracked. a leak test is performed on 100% of the introducer sheaths by qa. the underlying cause of the failure could not be determined. the introducer sheath passed all in-process and qa final inspection steps before shipping to the customer including visual, dimensional and leak testing. the potential causes of the sideport leaking may be related to excessive manipulation during use, not enough adhesive or improper adhesive material. the potential effects to the user are bleeding, prolonged intervention, potential product damage or inability to flush the sheath. a review of risk for this failure mode identified the risk to be as low as practical. based on the investigation, a capa is not required. oscor will continue to monitor this device for complaint trends and risk. the instructions for use (ifu) provide these precautions: aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. infusion through the sideport can be done only after all air is removed from the unit. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only.
38300 | 1035166-2016-00150 | 5930791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5930791&pc=DYB | 05/25/2016 | 08/11/2016 |  | 09/06/2016 |  |  | 08/11/2016 | 09/06/2016 |  |  | 04/01/2019 |  | 09/06/2016 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-12436 |  | 0052-3000 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 83 YR | [5 empty] | Injury | Yes | No | Backflow | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [5 empty] | IMPELLA CP | [22 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | This case was a percutaneous coronary intervention, in a patient suffering from cardiogenic shock post-ami. during the procedure, the bleeding was noticed from the sideport. in response to the bleeding, the doctor placed a kelly clamp on the 14f sheath. it was reported that the patient lost 100cc of blood; however, no blood product was administered. the patient was able to recover and was weaned off the impella, which was then explanted. | [5 empty] | The device was in use for treatment. one 14f a biomed introducer sheath was returned from the customer. there were no other accessories. blood was found on and inside the sheath, and the hemostasis valve was found torn. a review of the device history record for this lot number identified one (1) reject for inner label backwards. a review of complaints for this lot number identified no additional complaints. upon evaluation of the returned product, no leakage was found when tested with a syringe where the sideport attaches to the hub. this area of the sheath was inspected under a microscope and the adhesive appeared to be thorough and completely intact. the hemostasis valve was torn likely due to off center insertion of the device used in conjunction with the sheath. a leak test is performed on 100% of the introducer sheaths by quality assurance (qa). no manufacturing defects were found. the reason for return cannot be confirmed. the introducer sheath passed all in-process and qa final inspection steps including visual, dimensional and leak testing. the potential effects to the user are bleeding or prolonged intervention. a review of risk for this failure mode identified the risk to be as low as practicable. based on the investigation, a capa is not required. oscor will continue to monitor this device for complaint trends and risk. the instructions for use (ifu) provide these precautions: when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. infusion through the sideport can be done only after all air is removed from the unit. indwelling introducer sheaths should be internally supported by a catheter, lead, or dilator. dilators, catheters, and pacing leads should be removed slowly from the sheath. rapid removal may damage the valve membrane resulting in blood flow through the valve. never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only.
38301 | 5909835 | 5909835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5909835&pc=PBL | 08/23/2016 | 08/25/2016 | 08/25/2016 | 08/29/2016 | 08/25/2016 | 08/25/2016 |  | 08/29/2016 |  |  |  |  | 08/29/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Yes | [5 empty] | No | Physician |  |  | User Facility report | [8 empty] | 50 YR |  |  | 99 |  |  | Malfunction |  | Yes | Break | Kinked | Device Displays Incorrect Message |  |  |  |  | No Code Available | [18 empty] | NO OTHER THERAPIES | [22 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | In the cath lab, an impella device was placed in the right groin. although sedated, the patient became agitated and thrashed his leg. the device at that point began to alarm and it was found to be kinked in the patient. upon trying to remove the device, the device broke at the point of the kink and remained in his body. the broken piece was able to be removed completely with snare. post fluoroscopy showed no retained pieces and no lacerations or punctures from the device. the patient tolerated the incident well and a new device was successfully placed.
38302 | MW5064435 | 5913430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5913430&pc=DSQ |  | 06/23/2016 |  | 08/26/2016 |  |  |  | 08/26/2016 |  |  |  |  | 08/26/2016 |  | IMPELLA TEMPORARY LVAD |  |  | IMPELLA TEMPORARY LVAD |  |  | DSQ | [14 empty] | No answer provided | [5 empty] | Invalid/Incomplete | No Information |  |  | Voluntary report | [14 empty] | Other | No | No | Fail-Safe Design Failure | [6 empty] | No Known Impact Or Consequence To Patient | [41 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | No error has occurred, but there is a risk. the impella temporary lvads require d5w + heparin as purge fluid. the machine software does not allow our hospital's standard heparin concentration (40 units/ml), rather it gives options for 25 units/ml and 50 units/ml. patient also requires systemic heparin during infusion. the company rep/liaison says that it will not update software to accommodate other concentrations (even though we are not the only hospital with this issue). this leaves us with 3 solutions: change the standard concentration house-wide, which has been used for decades and is ingrained in staff; run 2 different heparin concentrations on the same patient, and run risk of purge fluid running out because pharmacy must specially compound it whereas our standard concentration is available in omnicell for the rns; run a 40/unit ml bag that the machine believes is 50 units/ml and calculates heparin infusion based on what it "thinks" is infusing. the software should be adjusted. follow-up: our director is considering changing our housewide heparin standard to 50 units/ml for other reasons, which may eliminate the problem (at least for our hospital). the impella lvad is basically a temporary lvad (a step up from an aortic balloon pump) put in via a femoral line when a patient is in cardiogenic shock as a temporizing measure after a severe, acute mi (these only stay in approx 24-48 hours). imagine a small boat propeller in someone's heart to help propel blood through when the heart cannot squeeze. this happens in the cardiac cath lab, then those patients come to the icu as a 1:1 nursing ratio because they have such high acuity (basically, they are train wrecks). they also can turn around amazingly well, the one that prompted my report was extubated, off the impella and all centra lines, and walking around the unit a few days later. my institution only uses these things about once a quarter (they cost approx (b)(6) per use, and at least the past 2 times i've dealt with one the rep comes to our hospital when it gets put in and is available in the room), so it is not something each nurse is very familiar with. once set up, the impella adjusts its own flow rate based on pressure needs to help the person's heart pump, and it has a "purge fluid" to prevent the machine from seizing up or clotting. in the past the purge fluid had to be d20 with a certain heparin concentration (i can't remember what the concentration was), but several months back they updated and said we can start with d5w, but are supposed to change to a heparinized d5w fluid of basically any concentration as soon as possible. the patient also has to be fully systemically anticoagulated while on the pump. pharmacy can only adjust the systemic infusion, but since the purge flow rates can change an account for approx 1/3 of the heparin the patient is getting, we keep tabs on what it is running at. when you look at the impella monitor, it gives infusions in ml and as units heparin (which are based on the concentration of heparin that the impella "thinks" is in the bag, which in our case is not the same concentration as what is actually in the bag). pharmacy is doing a work around and only looks at the mls infusion off the machine's display, but for a (b)(6)/machine cost, it really shouldn't take too much effort to have someone write code and allow more than those two concentrations. if we were a big hospital system, the rep said they would probably do it, but because we are a single independent hospital it is unlikely they would do so. medication administered to or used by the patient: no. when did the event occur: hospital. patient counseling provided: unk. severity: circumstances or events have capacity to cause error. ismp medication errors reporting program. (b)(4).
38303 | 1035166-2016-00144 | 5892494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5892494&pc=DYB | 07/21/2016 | 07/22/2016 |  | 08/20/2016 |  |  | 07/22/2016 | 08/20/2016 |  |  | 06/01/2019 |  | 08/20/2016 |  | INTRODUCER, CATHETER |  |  | ABIOMED 14F INTRODUCER KIT FOR IMPELLA |  |  | DYB |  |  | 0052-3000 |  |  | OSCOR INC. |  |  | US |  | C1-12562 |  | 0052-3000 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | DISTRIBUTOR | HEALTH PROFESSION |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | No | Obstruction of Flow | [6 empty] | Hematoma | Low Blood Pressure/ Hypotension | Blood Loss | [10 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | * | Invalid/Incomplete |  |  | The customer reported a complaint about an impella cp device being "unable to pass through the oscor sheath". it was noted in the case that after the 14 fr oscor sheath was placed, the cp was unable to be inserted through the sheath. significant bleeding and hematoma at the insertion site were noted post insertion. sheath was then removed and required a femstop to stop the bleeding. four hours later, the patient began to have hypotension and acute blood loss from the femoral artery and had to be sent to the or for emergency femoral artery repair. the affected product has not been returned for investigation yet. | [5 empty] | The device was in use for treatment. the device has not been returned for analysis; as a result, the allegations against this device cannot be confirmed. a review of the device history records identified no rejects for this lot number. a review of the complaints against this lot number identified one additional complaint. since the device was not returned, a root cause of the failure could not be determined. the potential effect to patient for the reported failure may be related to: device is rendered useless or the procedure is delayed. potential cause: improper or damaged extruded parts, improper introducer/dilator manufacturing, introducer/dilator is out of tolerance, or wrong size introducer used. a review of risk for this failure mode identified the risk to be low. based on this investigation a capa is not required. oscor will continue to monitor this device for complaint trends and risk. per the quality assurance in-process and final inspection procedure, it instructs the operators to use 10x magnification for visual inspection unless specified otherwise. visual inspection - verify the hub is smooth, has no cracks, no short shots, sinking or unfilled areas, check for foreign material, check hub shape, french size embossed on hub, and for excessive flash. using 10x magnification look down into hub for irregularities; verify no splitting at break lines and a smooth transition between sheath and hub on inside of hub. verify score lines are spaced evenly around the circumference of sheath tube; look down into the hub and verify score lines on sheath align with break line of the hub. measure dimensions - verify sideport opening diameter using pin gauges; visually inspect the sideport hole is without any flash or blockage. measure usable sheath length after overmolding from end of hub to end of tube, using metric ruler. verify i.d. and inner clearance of sheath by inserting a calibrated pin gauge into hub end, all the way through to the tip. use go-no-go gauge for o.d. verification; verify the full length passes through the gauge. using 10x magnification, verify the tip is round, no flash, no splitting, cracks or other damages. insert a tipped dilator through the tipped sheath and check if dilator inserts without excessive force or obstruction. with the naked eye, inspect where the sheath and dilator meet and check gap and that a smooth transition is present. measure usable sheath length after tipping; from end of hub to tip of tube, using metric ruler. measure tip i.d. of sheath using pin gauge. visual inspection with naked eye, ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded foreign material or any other damages. verify split caps are properly placed and secured onto sheath hub and free of cracks/damages or excessive adhesive. perform pressure and vacuum leak test and perform occlusion test. assembly fit check, insert dilator into sheath, lock and unlock dilator hub lock nut for verification of proper fit; verify no gaps between sheath tip and dilator. the abiomed adelante s2 hemostatic peel-away introducer system with infusion sideport instructions for use (ifu) lists the following adverse effects and possible complications: air embolism, bleeding, hematoma formation and vessel damage. precaution - aspiration and saline flushing of the sheath, dilator, and valve should be performed to help minimize the potential for air embolism and clot formation. also, the ifu instructs the user to aspirate all air from the sheath valve assembly by using a syringe connected to the sideport. flush the sheath with 5 cc of saline immediately before peeling the sheath away in order to minimize back bleeding.
38304 | MW5064289 | 5900575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5900575&pc=KFM | 08/17/2016 | 08/19/2016 |  | 08/19/2016 |  |  |  | 08/19/2016 |  |  | 04/30/2018 |  | 08/19/2016 |  | IMPELLA |  |  | IMPELLA CP |  |  | KFM |  |  | 0048-003 |  |  | ABIOMED |  |  |  |  | 1005505 |  |  |  | Yes | [5 empty] | Y | Health Professional |  |  | Voluntary report | [8 empty] | 70 YR |  |  | 77 |  |  | Injury | No | Yes | Leak/Splash | [6 empty] | No Information | [12 empty] | Other | Required Intervention | [27 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | After the impella cp was placed, the introducer sheath removed, a leak alarm was noted on the console. investigation revealed leak in the main impella tubing. impella support was contacted. recommendation was to replace the impella device.
38305 | 1220648-2016-00022 | 5883951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5883951&pc=PBL | 07/19/2016 | 07/19/2016 | 07/19/2016 | 08/17/2016 |  | 07/19/2016 | 07/19/2016 | 08/17/2016 |  |  | 04/30/2018 |  | 08/17/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1240716 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | Yes | Physical Resistance | Material Deformation | [5 empty] | Hematoma | Low Blood Pressure/ Hypotension | Blood Loss | [10 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/19/2016 | * | HOSPITAL |  |  | Complainant reported that on (b)(6) 2016 a (b)(6) obese female was being treated for cardiogenic shock and an acute anterior mi in the cath lab. she was noted to need the vasopressor, levophed, to maintain a systolic blood pressure above 90mm hg, and her anticoagulation was regulated by iv infusion of angiomax. the physician completed his angioplasty of the left anterior descending artery and had slow flow, concurrent with the cardiogenic shock, which prompted the decision to place an impella cp for support. the physician did the recommended serial dilations utilizing an 8 then 10 then 12 french dilator to prep the artery for installation of the 14fr oscor sheath. the sheath was placed. when the physician attempted to insert the impella cp through the 14fr sheath there was resistance felt and blood was flowing back from the sheath and a hematoma was forming at the insertion site. the team examined the groin and sheath under fluoroscopy and noted that the sheath had deformed. the decision was made to remove the sheath. two staff members applied manual pressure to achieve hemostasis. the manual pressure was stopped and the femostop (manufactured by st. jude medical/abbott) was placed prior to discharge from the cath lab suite and transfer to the icu for continued monitoring. the groin was resolving at the time and noted to be "soft". the complainant monitored the patient in the icu for 4 hours and then chose to begin the weaning of the femostop . five minutes after the femostop was removed the patient again had a spontaneous and significant blood loss from the arterial access site and hypotension. the patient was taken to the operating room for a vascular repair. the repair was an interposition graft performed by vascular surgery. the post-operative recovery was a good one and the patient was discharged home with no permanent damage. | [5 empty] | The impella cp and 14fr introducer sheath were not returned for evaluation, as they are currently being held in the complainant's risk management department. therefore, to date, no analysis can be performed. it is unclear at this point if the artery and any abnormalities of patient anatomy in the vascular bed could have contributed to the blood loss and noted "deformed" sheath. furthermore, it is unclear if patient habitus could have contributed to the issue. of note in the ifu there is a warning to observe for obesity and severe peripheral arterial obstructive disease before initiating the procedure, as obesity and pvd are factors that can prevent successful placement of the impella catheter. the patient involved in this event was documented to be obese. the physician had noted resistance which could have been a sign of pvd. the manufacturer will continue to investigate this complaint and all reasonable sources of information. upon release and return of the device an evaluation and investigation will be conducted and the ensuing conclusions will be reported in a supplemental medwatch report.
38306 | 1220648-2016-00021 | 5883961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5883961&pc=PBL | 07/19/2016 | 07/19/2016 | 07/19/2016 | 08/17/2016 |  | 07/19/2016 | 07/19/2016 | 08/17/2016 |  |  | 01/31/2018 |  | 08/17/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | PBL |  |  | 0048-0003 |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1241332 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 74 YR | [5 empty] | Malfunction | No | Yes | Material Separation | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/19/2016 | * | HOSPITAL |  |  | The complainant reported that on (b)(6) 2016 a (b)(6) male patient presented at the hospital with a totally occluded right coronary artery (rca) and left anterior descending coronary artery (lad). the patient's left main (lm) was also found to be diseased. an impella cp was placed to allow hemodynamic support to the patient during stenting. following successful stenting of the lm and lad, the rca was opened, but only one stent was able to be applied to the rc. the physician was unable to fully complete the repair of the rca, as radiation usage was reported to have become too high. the patient was then transferred to the intensive care unit (icu) and was reported to be doing well with the impella cp support successfully continued. the patient was administered heparin and integrin, blood seepage was then observed around the repositioning sheath and the right angioseal site. a femstop was then applied and the seepage was resolved. after over 22 hours of successful support the patient was reported to have been both looking and feeling better. it was then decided to remove the impella cp. during explant of the device the pump cannula separated into 2 pieces. the separated portion was able to be retrieved by hand at the femoral insertion site after the application of slight pressure to the surrounding tissue. following removal of the impella pump hemostasis was obtained. there were no adverse effects to the patient and the patient was reported to be in stable condition following the removal of the impella cp. | [5 empty] | The impella cp has not yet been returned for evaluation as it is currently being held in the complainant's risk management department; consequently, no device analysis could be performed. upon the release and return of the device an evaluation will be performed and the results and conclusions will be reported in a supplemental medwatch report. (b)(4). held in user's risk management department.
38307 | 1220648-2016-00020 | 5841523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5841523&pc=KFM | 07/09/2016 | 07/09/2016 | 07/09/2016 | 08/02/2016 |  | 07/09/2016 | 07/09/2016 | 08/02/2016 |  |  | 03/31/2018 |  | 08/02/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | KFM |  |  | IMPELLA CP PUMP SET, US |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1236971 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 70 YR | [5 empty] | Injury | Yes | Yes | Fracture | [6 empty] | No Known Impact Or Consequence To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 07/09/2016 | * | HOSPITAL |  |  | The complainant reported that on (b)(6) 2016 a (b)(6) male patient had cardiac arrested at him home. the patient was admitted to the hospital where an impella cp was placed using a 30cm cook sheath. during the placement of the pump the physician observed some prolapse of the cannula as it was pulled back into the sheath. following device placement low pump flows occurred. the pump flows would not surpass a 2liter maximum at all "p" levels. the cannula was repositioned under x-ray several times, but the flows continued to be low. the physician then attempted to remove through cook sheath and as he pulled the inflow and pigtail separated from device. the detached portion of the device was successfully removed via the aid of a snare. the pigtail was snared in one "smooth" attempt from the patient's vasculature. there was no damage to the native vasculature. impella support was then aborted. no further intervention or support devices, including impella support were necessary as at that point the patient was stable condition. | [5 empty] | The impella cp and automatic impella controller data logs were returned for evaluation and analysis. the analysis of the data logs showed evidence of cannula prolapse during repositioning. the pump flow then decreased sharply to approximately 0 liters/minute. when the cannula prolapsed, blood was prevented from flowing across the impeller. this resulted in decreased motor current and pump flow as confirmed in the data logs; therefore, the sharp decrease in flow likely occurred when the cannula prolapsed during repositioning. during the inspection of the returned impella cp it was revealed that the inflow cage and cannula exhibited separation that occurred because the cannula failed and tore. evidence of extreme high force application was found on the catheter. there was also excessive extension due to stretching of the catheter just proximal to the metal end cap. in addition, damage was found on the tip of the introducer sheath. an inward dent on the tip opposite where the sheath was cut was identified. this dent could have occurred when the device was pulled through the sheath. because the cannula was prolapsed, the clinician would have been unable to pass the cannula back through the sheath. the physician's attempt to pull the device through the sheath with the tip of the cannula prolapsed may have allowed the inflow cage to wrap over the tip of the sheath and cause the dent. the increased force required to pull the catheter back likely resulted in the separation of the inflow cage from the introducer and the extension of the catheter near the metal end cap. this could not be definitively determined. in conclusion, the most likely root cause of the separation of the inflow cage and cannula was entanglement of the prolapsed cannula with the tip of the sheath. because of the entanglement, force was applied to the cannula which exceeded the ultimate tensile strength and caused failure of the polyurethane. retraining of the physician was performed the day of this event to ensure extreme force were not applied to the catheter during removal of the device. (b)(4).
38308 | 1220648-2016-00019 | 5834825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5834825&pc=KFM | 06/30/2016 | 06/30/2016 | 06/30/2016 | 07/29/2016 |  | 06/30/2016 | 06/30/2016 | 07/29/2016 |  |  | 03/31/2018 |  | 07/29/2016 |  | PUMP, BLOOD, CARDIOPLMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA LP2.5 |  |  | KFM |  |  | IMPELLA LP2.5 |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1237580 |  | 005042 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 78 YR | [5 empty] | Injury | Yes | Yes | Positioning Failure | Fracture | [5 empty] | Skin Tears | Blood Loss | Foreign Body In Patient | [10 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/30/2016 | * | HOSPITAL |  |  | Complainant reported that a (b)(6) patient was scheduled for a primary percutaneous coronary intervention (ppci) (b)(6) 2016, but both femoral arteries were found to be completely occluded, resulting in the patient being rescheduled for a collaborative axillary impella cp (serial number (b)(4)) insertion on (b)(6) 2016. a recent catheterization revealed severe left main disease and the left anterior descending (lad) artery with 95% proximal stenosis and left ventricular ejection fraction (lvef) of 20%. the patient was also scheduled to have a transcatheter aortic valve replacement (tavr) to be performed at a later date. the patient was brought to the cath lab and an 8mm graft was sewn on to the right axillary artery without issue. the 14fr oscor sheath was inserted in the graft with an 0.018 wire into the left ventricle. using the 14fr oscor sheath during an impella cp axillary insertion technique is contrary to the instructions for use for both the impella cp and impella lp2.5 pumps. the 14fr introducer is not suitable for use during an axillary approach. the instructions for use instructs the user of the following when an axillary insertion is undertaken: impella® axillary insertion kit (impella® 2.5, 5.0 and impella cp®). the impella® axillary insertion kit contains the following: 23 fr diameter x 6 cm length peel-away introducer; 2 graft locks used to attach a graft onto the introducer (note: only one graft lock is required when used with the recommended hemashield platinum graft; a back-up is provided.); 8 fr silicone-coated lubrication dilator; 2 silicone plugs. it is recommended that the impella® axillary insertion kit be used in conjunction with a 10 mm diameter x 20 cm length hemashield platinum graft. the impella cp was placed in the graft and into the artery where great resistance was found after making the turn into the ventricle. after many attempts to work past the resistance wire access was lost and the impella cp was removed. removal also incurred a great deal of resistance. this resistance was most likely the result of the off label use of the impella 14fr oscor introducer. contrast dye used to assess the artery path revealed that the artery was of adequate size and without stenosis. access was then obtained after much difficulty getting the guidewire and 0.018 wire across the atrial ventricular valve (av). at this point it was felt that the patient's anatomy would not allow insertion of the impella cp. the impella cp insertion was aborted and an impella lp2.5 was obtained, as it was hoped that the smaller cannula would be able to make a turn in the axillary anatomy. the impella lp2.5 was placed in the artery, while again using a 14fr introducer (rather than the 23fr introducer), but again got "hung up" in the exact same area of the subclavian artery as the impella cp. continued attempts to advance the pump were unsuccessful. the procedure was then aborted due to the patient becoming unstable as the blood pressure began to drop into the 60's. during the physicians attempts to remove the impella 2.5 much resistance was encountered. when the pump was pulled from the artery, the catheter was found to be broken in half, separating just below the impeller motor. one section of the fractured cannula remained in the oscor sheath, while the other portion of the cannula was out of the sheath. both sections of the separated cannula were successfully removed from the vasculature and removed from the 8mm graft. the patient did require 3 total units of replacement blood, to replace the blood that is believed to have been lost as a result of much manipulation in trying to place the pump. as a result of the manipulation the physician was also required to perform an arterial repair, which consisted of sewing more stiches at the anastomosis site. the patient was successfully discharged home in stable condition and will return to the hospital for an open procedure, rather than the previously scheduled procedures of a tavr and a coronary artery bypass graft (cabg.) the patient's cardiologist believes that the inability to pass and place the device, the fracture of the device and any injury suffered by the patient to be the fault of the patient's extensive vascular calcified anatomy. | [5 empty] | The device was returned was returned for evaluation and analysis. the inspection of the returned impella lp2.5 revealed separation of the cannula, and it appeared that the maximum force to the cannula was exceeded causing the cannula to fracture into 2 distinct pieces. the returned device was compared to a retained impella lp2.5 sample, and no differences were identified in either the diameter or length of the rigid components. this reported failure occurred because of large removal forces required by a difficult patient anatomy. the instructions for use of this product lists among the contraindications for the use of this device the following: contraindications: patients with aortic stenosis or other abnormal aortic valve performance may be compromised by the use of the impella® catheter. patients with aortic valve disease should be observed for aortic insufficiency; aortic valve stenosis/calcification (equivalent to an orifice of 0.6 cm2 or less); moderate to severe aortic insufficiency (echocardiographic assessment of aortic insufficiency graded as [?] +2); severe peripheral arterial disease that precludes the placement of the impella® 2.5. in addition, and as reported in the description of the event, the instructions for use instructs the user of the following when an axillary insertion is undertaken: impella® axillary insertion kit (impella® 2.5, 5.0 and impella cp®). the impella® axillary insertion kit contains the following: 23 fr diameter x 6 cm length peel-away introducer; 2 graft locks used to attach a graft onto the introducer (note: only one graft lock is required when used with the recommended hemashield platinum graft; a back-up is provided.); 8 fr silicone-coated lubrication dilator; 2 silicone plugs; it is recommended that the impella® axillary insertion kit be used in conjunction with a 10 mm diameter x 20 cm length hemashield platinum graft. also as stated in the description of the event the patient's cardiologist believes that the inability to pass and place the device, the fracture of the device and any injury suffered by the patient to be the fault of the patient's extensive vascular calcified anatomy. in conclusion the root cause of the cannula separation was patient anatomy that resulted in higher than normal removal forces. (b)(4).
38309 | 5778311 | 5778311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5778311&pc=OJE | 06/27/2016 | 06/28/2016 | 06/28/2016 | 07/08/2016 | 06/28/2016 | 06/28/2016 |  | 07/08/2016 |  |  |  |  | 07/08/2016 |  | RIGHT VENTRICULAR BYPASS (ASSIST) DEVICE |  |  | IMPELLA RP |  |  | OJE | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | 0042-000-US |  | Yes | [5 empty] | No | No Information |  |  | User Facility report | [14 empty] | Malfunction |  | Yes | Device Operates Differently Than Expected | [6 empty] | Thrombus | [18 empty] | NO. | [22 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Abiomed impella device inserted about three days ago in cardiac cath lab at our hospital for decompensated chf and severe mitral valve regurgitation as patient not responsive to medical management after myocardial infarction (mi). ejection fraction (ef) 10-15% and cardiogenic shock. on removal in operating room, clot found in device. device failed prior to arrival to operating room and trouble-shooting with abiomed not successful, leading to need for device removal and placement of new rvad.
38310 | 1220648-2016-00016 | 5765728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5765728&pc=KFM | 06/01/2016 | 06/02/2016 | 06/02/2016 | 07/01/2016 |  | 06/02/2016 | 06/02/2016 | 07/01/2016 |  |  | 11/30/2017 |  | 07/01/2016 |  | PUMP, BLOOD, CARDIOPULMONARY, BYPASS, NON-ROLLER TYPE |  |  | IMPELLA LP5.0 |  |  | KFM |  |  | IMPELLA LP5.0 |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1223628 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 37 YR | [5 empty] | Injury | Yes | Yes | Fracture | Kinked | [5 empty] | Bone Fracture(s) | [12 empty] | Required Intervention | [5 empty] | ECMO | [22 empty] | Initial submission |  |  |  |  | 06/02/2016 | * | HOSPITAL |  |  | The complainant reported that on (b)(6) 2016 a (b)(6) male patient presented with and atypical chronic pain. the patient had a long standing history of cardiac artery disease (cad), and had a coronary artery bypass graft:(cabg) "a few years ago." on (b)(6) 2016 the patient had ventricular tachycardia (vt) runs, and then sustained vt earlier that day. the patient was urgently transferred to the cath lab for placement of an impella lpo5.0. during placement of the pump the patient was found to have a very short aortic root that required an intervention/ballooning to ascend through the iliacs. the physician employed an abiomed 23fr peel away sheath system after performing pre-dilation. the physician reported that the biggest problem/issue was the patient's very short ascending aortic root and made positioning a challenge. achieving "reasonable position" was difficult as the "elbow did not sit well" and the pump was "jumpy" in the ventricle. this resulted in the pump's motor housing/catheter junction becoming kinked as indicated in fluoro images taken during impella placement. the pump was successfully placed and was supporting the patient. mixed venous labs came back in the 40s, which resulted in the staff opting to place an ecmo in the patient. the patient was taken to the operating room for cannulation, while the impella 5.0 remained in place for left ventricle (lv) venting. on (b)(6) 2016 the patient continued on impella support with a venous access ecmo in place. the patient was being administered nitro and sedation medications, and was reported to have good impella flow down the impella leg and had a warm pink foot. the patient's ecmo side was reported to be cold with no pulse, resulting in cardiac surgery at bedside with a retrograde stick down the leg that was connected to the ecmo. the ecmo cannula was now providing in good perfusion. the patient continued to be successfully supported with the impella lp5.0. on (b)(6) 2016 impella required repositioning, which was done under echo. during the repositioning the impella cannula appeared to become separated from the motor. the patient was taken to the hybrid lab and after over 6 days of successful patient support the cannula was snared out of the lv without further issue. the patient remained on ecmo support. on (b)(6) 2016 patient care was withdrawn and the patient expired. the staff reported that no one at the facility attributied the patient outcome to the issue with this device. | [5 empty] | The pump, the detached cannula and a portion of the catheter were returned from the customer. the proximal end of the cannula was separated from the pump housing at the hydex ring/pump-housing interface. the pump plug had been severed from the catheter and was not returned. visual examination of the returned components revealed a significant kink in the steering catheter just proximal to the motor housing. the exterior of the motor and pump housing appeared undamaged. the distal end of the pump housing was deformed at the hydex ring interface, exhibiting a "flattened" non-circular geometry along with a slight indent on one side. the cannula was deformed and the delaminated at the proximal end near the hydex ring and delaminated at the distal end. there was a thrombus attached to the flow straightener on the inside of the pump housing. the mating surfaces of the hydex ring and pump housing were covered with varying amounts of adhesive. the amount of adhesive in the hydex ring interface with the pump housing appeared in the normal range based on a previous survey of other lp 5.0 pumps. a 23f abiomed sheath, normally supplied with the impella rp introducer was used for a femoral insertion of the impella 5.0. the impella 5.0 instructions for use recommends a cut-down insertion into the femoral artery or an axillary insertion using the supplied axillary insertion kit. because this was a non-standard use of the introducer/pump testing was done to evaluate potential damage to the pump resulting from insertion through the 23f sheath. an impella 5.0 was inserted over a guidewire through a 23 fr sheath through bends of decreasing radii of curvature to simulate any femoral/iliac tortuosity. no damage to the pump or the connection between the cannula and the hydex ring, and the hydex ring and the pump housing was observed. as part of a previous investigation of an impella 5.0 cannula separation, a test was done to evaluate the effect of pump housing deformation on the strength of the bond with the cannula. six previously used impella 5.0 pumps were tested. deformation of the bond area appeared to weaken the joint in two of six assemblies tested, allowing the housing and cannula to separate with a small bending force applied to the junction. the relatively small length to diameter ratio of the interface between the cannula and pump housing probably contributed to the separation. the cause of the damage to the catheter and cannula, and the deformation of the cannula-pump housing interface may have been frequent repositioning performed on (b)(6) 2016 because of the patient's very short ascending aortic root, which made achieving and maintaining pump position difficult. at the time it was reported the catheter was kinked at the junction with the motor housing indicating high forces were probably used for repositioning. deformation of the pump housing probably weakened the cannula-pump bond and when the pump was repositioned on (b)(6) 2016; the additional stress was probably enough to cause the separation of the cannula. in conclusion, the root cause of the deformation of the pump housing caused by manipulation of the catheter on (b)(6) probably weakened the cannula-pump bond and when the pump was repositioned on (b)(6) the additional stress was probably enough to cause the separation of the cannula. a review of the complaint database revealed no other pumps from this manufacturing lot were involved in field failures. internal reference: (b)(4).
38311 | 1220648-2016-00017 | 5765740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5765740&pc=PBL | 06/03/2016 | 06/03/2016 | 06/03/2016 | 07/01/2016 |  | 06/03/2016 | 06/03/2016 | 07/01/2016 |  |  | 04/30/2018 |  | 07/01/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP, |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1234487 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 75 YR | [5 empty] | Injury | Yes | No | Patient-Device Incompatibility | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/03/2016 | * | HOSPITAL |  |  | Complainant reported that on (b)(6) 2016 the hospital staff was treating a (b)(6) year old male patient with an ejection fraction of 30%, multi vessel disease and a coronary chronic total occlusion (cto).of the left circumflex artery (lcx). the patient was scheduled to undergo a high-risk percutaneous coronary intervention (hrpci.) the impella cp was placed for patient support on (b)(6) 2016, but approximately 5 minutes after insertion the patient's pressures dropped. following the drop in pressures, the physician determined that the patient's left ventricle was perforated. the patient was taken to the operating room where the impella cp was removed and the ventricle was repaired and 2 perclose sutures were applied. the physician indicated that the patient's left ventricle wall was calcified and thinner than the average person's ventricle, which resulted in the perforation. the patient was reported to be in stable condition following the surgical repair of the left ventricle. | [5 empty] | A definitive root cause of the perforation could not be determined. the pump was discarded subsequent to the event and was consequently unavailable for analysis. the exact cause could not be determined from the available data. the analysis of the data revealed that the patient's pressure decreased after the pump insertion and pulsatility decreased substantially. this behavior suggested that the left ventricle was perforated. as stated in the event description according to information from the clinician the patient's left ventricle wall was calcified and thinner than the average person's ventricle, which resulted in the perforation. (b)(4). device discarded by user.
38312 | 1220648-2016-00018 | 5765747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5765747&pc=PBL | 06/06/2016 | 06/06/2016 |  | 07/01/2016 |  |  | 06/06/2016 | 07/01/2016 |  |  | 03/31/2018 |  | 07/01/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1231124 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 74 YR |  |  | 87 |  |  | Injury | Yes | No | Patient-Device Incompatibility | [6 empty] | Injury | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 06/06/2016 | * | HOSPITAL |  |  | The complainant reported that the facility was treating an obese (b)(6) female patient in cardiogenic shock and who had an acute myocardial infarction. the patient had undergone an alcohol ablation while under support of an intra-aortic balloon pump (iabp). following the ablation the patient's pressures dropped. the iabp was removed from the patient and an impella cp was placed without issue. after two arterial sticks were made access was gained femorally, and the cp was inserted via the second stick site. impella support was initiated and the oscor sheath was removed and the repositioning sheath was placed. the physician noted patient's size and fat tissue dissuaded the peel away sheath from stopping any blood loss at the second arterial stick site. following 2 hours of successful impella support a hematoma developed on the patient's impella side and the patient's hemaglobin level had dropped from 11.8 to 5.3. the patient was then taken to the operating room for evacuation of the hematoma and exploration. during this procedure the physician found a nick in the patient's artery at the location of the first stick site. the site was oozing and the artery was closed with a suture. the surgeon then closed the surgical wound site and the patient was administered three units of blood product. the impella cannula was repositioned under echo guidance and impella cp support continued. over the following 3 days of impella cp support the patient's hemoglobin rose to 9.4. after being successfully supported for over 2 days the impella cp was removed from the patient. there was no harm to the patient other than the repaired nick. | [5 empty] | The lmpella cp was not returned for evaluation, as it was discarded subsequent to the event; consequently no device analysis could be performed. it is unable to be definitively established if the arterial nick and bleeding were as a result of any issue with the introducer, lmpella cp pump, guide-wire or if the issue occurred solely as a result of the patient's anatomy. the physician reported that he believed that the large patient size prevented the repositioning sheath from providing hemostasis and the patient size also resulted in the hematoma. the console logs were returned for analysis. the logs however did not contribute any information to the root cause of this event. an examination of the event determined that the root cause of the hematoma and associated blood loss was most likely due to the patient's obesity there is therefore no corrective action to be taken. the instructions for use of the impella cp warns the user of the following: when securing the repositioning sheath, vascular closure may be difficult in obese patients with extensive adipose tissue. before initiating the procedure, evaluate the patient for factors that may prevent successful placement of the impella catheter.... "obesity"...is listed as one of those factors. the manufacturer will continue to investigate all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. (b)(4). device discarded by user.
38313 | 1220648-2016-00015 | 5729317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5729317&pc=PBL | 05/17/2016 | 05/17/2016 |  | 06/16/2016 |  |  | 05/17/2016 | 06/16/2016 |  |  | 02/28/2018 |  | 06/16/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1234144 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Failure to Advance | [6 empty] | Perforation | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/17/2016 | N | HOSPITAL |  |  | The complainant reported that on (b)(6) 2016 a (b)(6) old morbidly obese male patient with systolic congestive heart failure was admitted to the coronary care lab (ccl) to undergo a primary percutaneous coronary intervention (ppci) for complex multi-vessel coronary artery disease (which included a long occlusion of the left anterior descending artery [lad], a circumflex coronary artery (cx-om) lesion and a severely diseased right coronary artery). the patient also had a compromised ejection fraction of 20%. as a result of the patient's large body mass index and the hemodynamic burden that would occur during this complex pci the physician chose to place the impella cp in the patient. the impella cp was successfully placed in the patient. shortly following pump placement it was noted that there was pulsatile blood jetting back from the hemostatic valve of the oscor sheath. the impella was moved back 2cm, as it was out of the apex, but significant blood loss continued. another clinician moved above the arteriotomy and held pressure while the physician removed the peel away sheath from artery, then peeled away and flushed the repositioning sheath. multiple attempts were made to advance the repositioning sheath into femoral artery, but it would not pass. another attempt was made to slide the repositioning sheath into femoral artery without success. the clinician was holding pressure above insertion stick but the patient continued to bleed profusely. at this point another physician assisted the doctor and performed contralateral balloon tamponade from the rfa to the lfa and ballooned alongside the impella catheter at the left femoral artery (lfa) access site. this helped with the bleeding. the first physician then attempted to insert repositioning sheath pulled on the impella catheter and pushed the repositioning sheath at the same time for increased pushability, but this did not work. he then cut the impella catheter and attempted to place a 10fr, 12 fr, 14fr cook sheath over the impella catheter, but none fit onto impella catheter, as each sheath just crumpled up at entry site without the dilator. the patient had now lost a significant amount of blood, estimated to be between 750cc's and 1 liter and was hemodynamically unstable. the patient was then started on levophed at 20mcg/min and dopamine at 10mcg/kg/min. the patient was also administered 3 units of blood during this time. his blood pressure got as low as 60/40. the patient became more responsive with better arterial control with the balloon tamponade and manual pressure along with levophed , dopamine and blood products and was alert and oriented. the patient was brought to surgery to remove the device and close the artery. a cut down to femoral artery was performed which revealed an extremely deep femoral artery. the physician reported that he felt the deep femoral artery explained why the repositioning sheath would not advance, as it was straight down and the strength of the impella catheter was not enough to carry the repositioning sheath all the way into the common femoral artery. the femoral artery repair was completed, and it was reported that there was no permanent patient harm or injury. | [5 empty] | The impella cp kit, which includes the 14 fr oscor sheath, dilator and repositioning sheath were all returned for evaluation. the inspection of the impella cp kit revealed visible damage to the impella cp, 14 fr oscor sheath, dilator, and repositioning sheath. the pump catheter had been cut in 2 places, with approximately 70cm of the catheter missing. the repositioning sheath was found to be severely damaged. the console logs were also returned for analysis. the log analysis revealed that the pump was moved out of position 10 minutes after it was started and remained in the wrong position until it was disconnected after 25 minutes. the visual examination of the pump, repositioning sheath and oscor sheath determined that the cause of this event was the impella cp being placed in a morbidly obese patient. after the pump was implanted, blood was leaking rapidly from the 14 fr oscor sheath. the oscor sheath was peeled back and the physician tried inserting the repositioning sheath. the femoral artery was very deep and the access angle was steep. the repositioning sheath would not enter the artery, and when greater force was applied, the wound closure on the repositioning sheath failed and was severely damaged resulting in the patient bleeding. the root cause of this failure was excessive force applied due to patient condition. as the failure was caused do to patient condition, no corrective action is recommended at this time. the instructions for use of the impella cp warns the user of the following: when securing the repositioning sheath, vascular closure may be difficult in obese patients with extensive adipose tissue. the instruction for use of this device also instructs the user of the following: before initiating the procedure, evaluate the patient for factors that may prevent successful placement of the impella® catheter {such as} obesity the investigation also included an investigation of the continued patient bleeding. the inspection of the 14fr oscar sheath revealed that the hemostatic valve was torn in an irregular pattern, and evidence was found of the silicone having reconnected along the valve slits, causing an insufficient seal. the root cause of the bleeding was a valve defect on the top and bottom valves in the 14 fr oscor sheath. no corrective action is recommended at this time due to the low risk of this failure mode. (b)(4).
38314 | 1220648-2016-00012 | 5727576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5727576&pc=PBL | 05/16/2016 | 05/16/2016 | 05/16/2016 | 06/15/2016 |  | 05/16/2016 | 05/16/2016 | 06/15/2016 |  |  | 07/31/2017 |  | 06/15/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP, PRODUCT CODE: PBL |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1209437 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 56 YR | [5 empty] | Injury | No | Yes | Device Operational Issue | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 05/16/2016 | * | HOSPITAL |  |  | The complainant reported that a (b)(6) year old male patient had been admitted to the hospital on (b)(6) 2016 as a result of chest pains and shortness of breath. this patient also had an increase in his troponin levels. this patient had a history of multiple myocardial infarctions and stent placements. an implantable cardioverter defibrillator/pacemaker (icd) had been previously placed. the patient was reported to have had runs of ventricular tachycardia (vtach) for a week with the icd shocking him into a normal heart rhythm. upon admittance into this facility the patient an ejection fraction of 15, and was intubated, but became worse necessitating pressors /inotropic support. the patient's lactic acid had also increased. following the placement of the impella cp an angiography was performed which revealed a mid-right coronary artery (rca). a stent was successfully placed at the rca. after stent placement the pump was found to need repositioning; the physician repositioned under echo guidance, but each time he pulled back further the pump would come back into the aorta. the patient was in stable condition, and it was planned to allow the patient to rest for a couple of days. the following day the impella again needed repositioning. the device was repositioned. the patient was reported to have been in stable condition during the previous night, but on this day the patient's urine was found to be "bloody." repositioning was again performed. that night the patient had a run of vtach and had coded. cpr was performed on the patient. after the code the impella required repositioning and the patient was found to have developed a large hematoma at the groin site, occurring as a result of patient movement during the code; pressure was held on the groin and the hematoma and oozing resolved quickly without incident and harm. it was then reported that after the repositioning of the pump earlier that same day the patient's labs had continued to hemolyze. the patient remained stable throughout the remainder of the night, without any more incidents of vtach, although the patient's troponins continued to rise and the patient remained "very diaphoretic." the patient was then transferred to another facility on (b)(6) 2016. upon arrival at this hospital a transesophageal echocardiography (tee) was performed which revealed that the impella cp was positioned in the mid-ventricular. the patient's urine had lightened a bit and the first sets of labs performed at the new facility were not hemolyzed. a second set of labs using "special tubes" was performed this set of labs did show hemolysis. this prompted staff's decision to explant the device after more than 42 hours of impella cp support. an ecmo was placed. the patient was reported to be stable on the ecmo and on dialysis treatment. the patient did recover and was discharged home and was no longer on dialysis. | [5 empty] | The impella cp was not returned for evaluation as the pump was discarded by the staff following explant. the console logs were returned for analysis, the analysis of the logs suggested that the impella cp was out of position repeatedly which could have contributed to the patient's hemolysis; however because the impella was not returned the root cause of the hemolysis was unable to be definitively confirmed. because the root cause of this event could not be determined, no corrective action is recommended. (b)(4). device discarded subsequent to use.
38315 | MW5062909 | 5735750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5735750&pc=KFM | 06/13/2016 | 06/15/2016 |  | 06/15/2016 |  |  |  | 06/15/2016 |  |  |  |  | 06/15/2016 |  | CATHETER |  |  | IMPELLA 5.0 |  |  | KFM | [5 empty] | ABIOMED |  |  | US | [5 empty] | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 53 YR | [5 empty] | Injury | No | Yes | Device Operates Differently Than Expected | Improper Device Output | [5 empty] | No Consequences Or Impact To Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Implanted device was not showing a correct wave form requiring the mfr's rep to provide manual calculation for the proper wave form. the device functioned appropriately following the manual correction and no direct harm to the pt. dates of use: (b)(6) 2016. diagnosis: coronary artery disease.
38316 | 5705333 | 5705333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5705333&pc=PBL | 05/31/2016 | 06/06/2016 | 06/06/2016 | 06/08/2016 | 06/06/2016 | 06/06/2016 |  | 06/08/2016 |  |  |  |  | 06/08/2016 |  | PUMP, CATHETER BASED |  |  | IMPELLA CP |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Yes | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 51 YR | [5 empty] | Malfunction |  | Yes | Device Operates Differently Than Expected | [6 empty] | Death | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL | 4124615077947 |  | This middle-aged male patient arrived to the cath lab for a left heart catheterization, coronary angiography and a left ventriculogram. the patient was transferred from the cath lab to the main or for an emergent cardiac salvage coronary artery bypass x 3. the patient arrived to the operating room with an impella pump. however, during the case, the impella pump did not register like it did in the beginning of the case. there was a question about whether or the impella pump was working correctly. the surgeon then placed a balloon pump in the patient. the patient subsequently expired during the surgery. it is unknown at this time if this event contributed to the patient's outcome. the rep for a biomed was present during the procedure.
38317 | 1220648-2016-00014 | 5686283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5686283&pc=PBL | 04/29/2016 | 04/29/2016 |  | 05/27/2016 |  |  | 04/29/2016 | 05/27/2016 |  |  | 05/31/2017 |  | 05/27/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP PUMP SET, US |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1197967 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | Fracture | Kinked | [5 empty] | Laceration(s) | Injury | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/29/2016 | * | HOSPITAL |  |  | The complainant reported that on (b)(6) 2016 an aortic and mitral valve replacement procedure was performed on a (b)(6) male patient; following this procedure the patient required multiple inotropes, pressors and required an intra-aortic balloon pump to be placed. the patient continued to struggle hemodynamically after surgery; consequently, on (b)(6) 2016 patient support was escalated to the placement of an impella cp. the 14fr introducer was placed in the patient's left groin without difficulty. during the impella placement procedure a .035 guidewire and a pigtail catheter were placed across the aortic bioprosthetic valve. the .035 guidewire was exchanged for a.018 guidewire and the impella cp was back-loaded onto the .018 guidewire. the physician reported that he was unable to advance the pump across the aortic valve. while pulling the impella catheter and the .018 guidewire back thru the 14fr introducer, part of the shaft of the sheath pulled back through the hemostatic valve of the sheath, and the wire become entangled in the sheath shaft section that had prolapsed out of the hemostatic valve. upon further inspection it appeared that the sheath had somehow split open from the distal tip and had pulled back up into the hemostatic valve as the doctor was pulling the impella catheter and wire out of the sheath. the physician then attempted to re-wire the sheath in order to exchange it out for a 14fr cook brand, but was unable to advance the wire into the sheath and back into the femoral artery. at this point, it was obvious that the distal tip of the 14fr introducer was no longer in the artery and the lumen was no longer patent, and the patient was found to have incurred a vascular injury necessitating surgical repair of the femoral artery. after the successful repair of the left femoral artery, the patient was still hemodynamically unstable and required continued mechanical circulatory support. a new impella cp was placed percutaneously in to the left femoral artery above previous site. the physician reported that the patient was in stable condition following the arterial repair and the placement of the second impella cp. there were no further issues reported following patient support with the replacement impella cp. | [5 empty] | The impella cp, sheath, guidewire and console data logs were returned to the manufacturer for evaluation and analysis. a device history review was performed. there were no issues found for any pump in this pump set lot, nor were any records found of any failed inspection criteria relevant to this failure mode. in addition, there have not been any other relevant issues reported to have occurred for any other 14french introducers from this lot. the review of the console data logs for this event did not provide any detailed information, as the logs showed that the pump was inserted for a short period of time. the pump was then removed and never turned on. the pump, 14french introducer and guidewire were returned for investigation. the pump was found to be in good condition with no obvious defects. it was examined to ensure it met outer diameter specification, and was found to be within specification. the introducer was also inspected. the tip of the sheath was found to be split and part of the shaft was pulled back on itself into the sheath. marks were identified on the tip where it split. additionally it appeared to have bent inwards when it split. the guidewire was inspected and found to have two severe kinks. during the evaluation an attempt was made to reproduce the reported issue using the pump and guidewire from the case with a new 14fr introducer sheath. the guidewire was placed through the pump and the pump was inserted through the sheath. the kink on the guidewire was placed between the outflow cage and the tip of the introducer. attempts were made to pull the pump through the sheath tip while the test setup was inside simulated vasculature. the failure mode was not able to be reproduced during this test. in conclusion, because the failure mode was unable to be reproduced, a definitive root cause for this event was unable to be determined. from the case notes it is apparent that the femoral artery repair was necessary because the sheath no longer provided access after the pump was removed. most likely the sheath tip split because of interaction between the sheath tip, outflow cage, and guidewire. no corrective action was initiated, as a definitive root cause was unable to be established. this failure mode will continue to be monitored. (b)(4).
38318 | 1220648-2016-00013 | 5686403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5686403&pc=PBL | 04/24/2016 | 04/27/2016 |  | 05/27/2016 |  |  |  | 05/27/2016 |  |  | 07/31/2017 |  | 05/27/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1209230 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 66 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Fasciitis | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/27/2016 | * | HOSPITAL |  |  | The complainant reported that on (b)(6) 2016 a (b)(6) female patient was admitted to the hospital with a total heart block. the physician performed the placement of bare metal stents to the circumflex and the left ascending artery (lad) and repaired a chronic total occlusion. during the 3.5 hour procedure the patient was successfully supported via the right femoral artery with an impella cp. following the procedure the patient continued to be supported by the impella pump and was administered pressors and inotropes. the patient was supported with the impella cp for a total of 38 hours. the patient was reported to have did have "deplorable" pulses in the right leg, but that the leg was cool when admitted to the icu. on (b)(6) 2016 the physician found that the patient's right impella leg was still cold and the pump was removed. a fasciotomy was reported to have been performed on the patient's right leg. following the fasciotomy the patient was reported to be in stable condition. | [5 empty] | The products involved in this event were discarded by the user subsequent to use. the console logs were returned to the manufacturer for analysis. an analysis of the console logs was performed, but was unable to reveal the cause of this issue. an examination of the reported event details revealed that this patient was on pressors, and that he was also exhibiting decreased arterial flow to the leg. with those issues the placement of an impella cp could have exacerbated the reduced flow to the patient's leg. without the product being returned no evaluation could be done on the repositioning sheath; therefore a definitive root cause for this event could not be determined. it is unclear if this reported issue was related to the impella cp's repositioning sheath or if the issue was the result of the physician's device placement/management, or the patient's condition. the manufacturer will continue to investigate all reasonable obtainable source information, and will provide them on a supplemental medwatch report if they become available. (b)(4). device discarded by user facility.
38319 | 5670954 | 5670954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5670954&pc=KFM | 03/27/2016 | 04/01/2016 | 03/30/2016 | 05/23/2016 | 03/30/2016 | 03/30/2016 |  | 05/23/2016 |  |  |  |  | 05/23/2016 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA INTRODUCER KIT |  |  | KFM |  |  | 0052-0001 |  |  | ABIOMED, INC. |  |  | US |  | NOT KNOWN |  | 0052-0001 |  | Yes | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 72 YR | [5 empty] | Death | Yes |  | Adverse Event Without Identified Device or Use Problem | [6 empty] | Death | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [5 empty] | CARDIAC DRUGS | [22 empty] | Initial submission | [5 empty] | Y | HOSPITAL | 14FR |  | A high risk percutaneous coronary intervention (pci) procedure was performed with impella support in the cardiac cath lab. the patient arrived in cardiac care unit and was noted to be bleeding from the femoral sheath site. the patient was administered protamine for heparin reversal. a stitch was placed in the patient's thigh. hemostasis was not achieved. manual pressure seemed to accelerate bleeding so the decision made to remove all devices from patient's groin.
38320 | 1220648-2016-00011 | 5652306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5652306&pc=PBL | 04/13/2016 | 04/14/2016 | 01/14/2016 | 05/13/2016 |  | 01/14/2016 | 04/14/2016 | 05/13/2016 |  |  | 12/31/2017 |  | 05/13/2016 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1226951 |  | 0048-0003 |  | Yes | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | Material Integrity Problem | Positioning Problem | Torn Material |  |  |  |  | Extravasation | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 04/14/2016 | * | HOSPITAL |  |  | The complainant reported that they were treating a (b)(6) male patient on postoperative day 1, after having received 4 bypass grafts. the patient was presenting with worsening cardiogenic shock, despite having multiple drips and the placement of an intra-aortic balloon pump (iabp.) the patient was brought to the cardiac catheterization laboratory (cath lab) to have the iabp replaced with an impella cp. during the impella cp placement in the left femoral artery the sheath was placed without issue. the pump was inserted and auto-flow was successfully initiated. the sheath was removed, and the repositioning sheath was advanced; however, the physician was unable to fully advance the repositioning sheath. while holding pressure, the repositioning sheath was removed and assessed. it was reported that a small longitudinal tear near the distal end of the sheath was seen, and was believed to have prevented the repositioning sheath's advancement. manual pressure was held to the impella's 9 french catheter. an angiogram of the left groin was performed, which revealed extravasation around the site. the physician placed a covered stent for hemostasis. as the covered stent deployment system was advanced from contralateral side, the impella to lfa was removed. the covered stent was deployed to the lfa with good results. the replacement impella cp was prepped and placed in the patient via the right femoral artery, and auto flow was initiated. the sheath was left in place and the repositioning sheath was successfully advanced through the sheath. the patient was reported to have required 2 units of replacement blood products. the patient was successfully supported with the replacement impella cp for approximately 7 hours, when the patient coded, and the family withdrew care. the patient outcome was reported not to have been the result of any issue that occurred with the first impella. | [5 empty] | The impella cp is currently under evaluation. the manufacturer will provide the evaluation results and conclusions in a supplemental medwatch report upon the completion of the analysis. | The impella cp was returned for evaluation. an inspection of the repositioning sheath revealed that the wound closure device had been badly damaged, and that the tip of the sheath had been cut off. the tear in the wound closure appeared to start in an area where a routine rework is performed during the manufacturing process. flash removal may weaken the sheath at this poing and if the sheath encounters resistance during insertion, then these small tears may act as initiation points for a larger tear. because the tip of the sheath had been cut off, it was impossible to determine the exact origin of the tear; however, it appeared that the tear may have originated in an area where the routine rework is done during the manufacturing process. if the sheath encounters resistance during insertion, then these small tears can act as initiation points for a larger tear. the root cause of the damaged repositioning sheath was most likely a combination of the patient condition and the manufacturing process. small tears in the wound closure as a result of flash removal during manufacturing contributed to a larger tear in the wound closure. in addition, the larger tear may have been due to contact with a rigid structure (e.g. calcification) during insertion of the repositioning sheath. a corrective action has been opened to address this repositioning sheath failure. internal reference: (b)(4).
38321 | 1220648-2016-00009 | 5635857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5635857&pc=PBL | 04/05/2016 | 04/05/2016 | 04/05/2016 | 05/05/2016 |  | 04/05/2016 | 04/05/2016 | 05/05/2016 |  |  | 03/03/2017 |  | 05/05/2016 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1215341 |  | 0048-0004 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | FOREIG |  |  |  |  | 67 YR | [5 empty] | Injury | Yes | No | Migration or Expulsion of Device | [6 empty] | Blood Loss | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/05/2016 | * | HOSPITAL |  |  | The complainant reported that a (b)(6) year old "heavy set" male patient in cardiogenic shock due to an occluded left main (lm) was transferred to this facility for the placement of an impella cp. the physician reported that the patient was in cardiogenic shock due to thrombus in the left main. the patient's lvedp was elevated at 40 and he was on levophed of 15ug/minute. the physician completed revascularization by removing a clot in the lm and no stents were needed. the 14 french peel away sheath was inserted effortlessly, but the physician was observed having difficulty with the repositioning sheath, and appeared to be "fussing over it." the physician performed the placement of the repositioning sheath without the aid of fluoro coverage. within 3 minutes of placement the impella catheter migrated out of the ventricle. adjustments were made and catheter fluoro images were then taken of the groin. the images indicated that the catheter was in the patient's abdomen. ten to fifteen minutes later, the patient became hemodynamically unstable and started bleeding. anesthesia, vascular surgery and another interventionalist were called to assist. they did see under fluoro what appeared to be two puncture sites that were sites of blood loss. one was seen to be above the access site and was bleeding profusely, while the one below was not bleeding out to that extent. the fluoro images indicated that the cannula tracked along the side of the artery and had never entered the artery. a balloon catheter was inserted via the left femoral artery and occluded the right artery proximally. unfortunately, the patient expired before the vascular surgeon could intervene. the staff reported that the cause of the patient outcome cause was reported as "stemi"; furthermore, the team stated that the patient would have expired without the impella device. | [5 empty] | The device was returned to the manufacturer for evaluation. the impella cp was evaluated and the aic logs were analyzed. a review of the data logs showed that the pump was inserted at approximately 14:16 on (b)(6) 2016. the pump was then started one minute later. at 14:45 there was an "impella position unknown" alarm. this may have been the result of the pump having been pulled out of the ventricle at this point. the evaluation of the pump revealed that there was a kink at the 56cm mark, and the wound closure sheath had been damaged. the tip of the repositioning sheath did not appear to be damaged. the catheter was cut by the evaluating engineers so that the repositioning sheath could be examined. the wound closure sheath was found to be cracked and torn in the area where it transitions to the sheath. a fluoroscopy image was also submitted for analysis. the image from the case showed the position of the repositioning sheath and was interpreted as showing the repositioning sheath positioned alongside of the artery. it appears that this was responsible for pulling the catheter back out of the heart and into the patient's abdomen. this could have contributed to the kink and the damage sheath that was observed during the failure analysis. failures of the wound closure sheath in the past have been attributed to the manufacturing process of flash removal which could result in small cracks and tears in the wound closure sheath and can contribute to failure of the device when under stress. in conclusion, the root cause of the damage to the repositioning sheath was that the sheath and catheter were pushed up alongside the artery during placement. this event may have been avoided with the careful use of fluoro guidance. the impella cp instructions for use warns the user of the following: never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action. fluoroscopy is required to guide placement of the impella® catheter. the small placement guidewire must be reliably observed at all times. the manufacturing process of flash removal was corrected with a product change. (b)(4).
38322 | 1220648-2016-00010 | 5620883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5620883&pc=PBL | 03/30/2016 | 03/30/2016 |  | 04/29/2016 |  |  | 03/30/2016 | 04/29/2016 |  |  | 12/31/2017 |  | 04/29/2016 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1226190 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 48 YR | [5 empty] | Injury | Yes | No | Compatibility Problem | [6 empty] | Foreign Body In Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 03/30/2016 | * | HOSPITAL |  |  | The complainant reported that a (b)(6) male patient was admitted to the facility as a stemi. the physician performed a repair of the patient's lad. the patient was reported to have decompensated during the night and the following day the patient was taken to the cath lab for the possible placement of an iabp and a further pci. due to low ejection fraction it was decided to place an impella cp. the facility did not have any 14fr 30cm (b)(4) sheaths in inventory, so placement was attempted with an unknown manufacturer's 14fr 13cm sheath. placement with this sheath was unsuccessful. despite recommendations from the abiomed representative that placement should be performed with the 14fr (b)(4) sheath that is included in the impella cp kit, the staff selected an (b)(4) 14fr 38cm sheath, which has no history with impella. when the wire was positioned within the left ventricle the patient had multiple runs of ventricular fibrillation and required repeated defibrillations. the physician then attempted to wirelessly to insert impella cp, but was unable to cross the valve despite multiple attempts. a decision made to remove the pump from the patient, but during removal the pump became caught on the (b)(4) sheath and the inlet cage of the pump fractured and detached from the device. the physician was able to successfully remove the device with the aid of 2 snares. another impella cp was successfully and without issue placed in the patient. the patient was successfully support and the complainant reported that there was no harm to the patient as a result of the reported issue. | [5 empty] | The automatic impella controller data logs and the impella cp were returned for evaluation. the analysis of the data logs revealed that the pump had run briefly with "impella position wrong" alarms occurring prior to the device being explanted by the physician. during the inspection of the returned impella cp pump it was found that the distal portion of the cannula, including the inflow cage, were not returned from the customer. the portion of the cannula that was returned exhibited a fracture in the nitinol coating, and the polyurethane coating was found to have deformed inwards. this condition was reproduced during bench testing while attempting to pull a prolapsed cannula through an introducer. the cannula separated from the inflow cage (with a portion of the cannula remaining on the inflow cage) and remained connected only by the nitinol coil. two pumps were tested, both failed in the same way. as shown with bench testing a potential root cause of the cannula separation was a prolapse of the cannula, with the cannula doubling over on itself, and the subsequent attempted removal through the introducer. this resulted in generating abnormally high forces on the cannula and its connection to the pump. this prolapse may have been caused by attempting to cross the aortic valve without a wire, which is not a recommended practice per the instructions for use of the impella cp. the instructions for use advises the user of the following: "advance the catheter through the hemostatic valve into the femoral artery and along the placement guidewire and across the aortic valve using a fixed-wire technique. follow the catheter under fluoroscopy as it is advanced across the aortic valve, positioning the inlet area of the catheter 3.5 cm below the aortic valve annulus and in the middle of the ventricular chamber, free from the mitral valve chordae. be careful not to coil the guidewire in the left ventricle." (b)(4). | Following this event the abiomed representative conducted retraining of the staff on the insertion of the device and the requirements in placing the device in a patient. the impella cp and introducers were not returned for evaluation. continued efforts are being made to obtain the devices at issue. upon the return and completed investigation the manufacturer will provide results and conclusions in a supplemental medwatrch report. it is important to note that the abiomed instructions for use of this device advices the user that the only alternative introducer sheath for use with this device is the (b)(4) check-flo performer introducer (non peel-away) 14fr 30cm. additionally it is not recommended to cross the aorta valve with an impella cp without a guidewire. device not yet returned.
38323 | 1220648-2016-00008 | 5621066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5621066&pc=PBL | 03/31/2016 | 03/31/2016 | 03/31/2016 | 04/29/2016 |  | 03/31/2016 | 03/31/2016 | 04/29/2016 |  |  | 10/31/2017 |  | 04/29/2016 |  | IMPELLA |  |  | IMPELLA RP |  |  | PBL |  |  | IMPELLA RP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1218232 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 68 YR | [5 empty] | Injury | Yes |  | Difficult to Remove | [6 empty] | Foreign Body In Patient | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/31/2016 | * | HOSPITAL |  |  | The complainant reported that a (b)(6) male post lvad patient had been successfully supported by an impella rp for over 3 days when the physician attempted to remove the pump at the patient's bedside. the sutures were removed and the device was pulled back, but only the portion of the pump, up to the inlet cage, was removed from the patient. the patient was then taken to the cath lab where fluoro images were taken. the fluoro images revealed that the rest of the device was intact and was located in the iliac area. the physician then successfully removed the detached portion of the device from the iliac area with the aid of 2 snares and a balloon through a 24 french sheath. no cutdown was necessary and just manual pressure was necessary following the device removal. the complainant reported that there was no harm to the patient as are result of the reported issue. | [5 empty] | The root cause of this spr was determined to be a high cumulative (3.8 day implant duration) load imparted to the cannula. the high cumulative loading of the cannula was due to a challenging device placement; along with the fact that once the pump was in place, a constant high loading force was imparted on the cannula during the days of support. this resulted in the deformation of the connection interface between the cannula and the pump housing. the deformation weakened the epoxy bond between the cannula and pump housing, resulting in the failure observed during pump removal. corrective action retraining: training currently includes education of users on the potential for a high residual load on the cannula during challenging device placements. the abiomed representative reported that following this event a review was performed with the staff on how to manage the use of the product during high residual load on the cannula during device placements. (b)(4).
38324 | 1220648-2016-00006 | 5591533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5591533&pc=PBL | 02/12/2016 | 03/25/2016 | 03/25/2016 | 04/20/2016 |  | 03/25/2016 | 03/25/2016 | 04/20/2016 |  |  | 10/17/2017 |  | 04/20/2016 |  | IMPELLA LP2.5 |  |  | IMPELLA LP2.5 |  |  | PBL |  |  | IMPELLA LP2.5 |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1217967 |  | 5042 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 65 YR | [5 empty] | Injury | Yes | No | Entrapment of Device | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/25/2016 | * | HOSPITAL |  |  | The complainant reported that on (b)(6) 2016 a (b)(6) patient entered the facility's emergency room complaining of shortness of breath and chest pains. the patient was examined and was found to have elevated troponin levels. the patient was taken to the catheterization lab where a diagnostic catheterization was performed that revealed multi-vessel disease and weak left ventricle (lv) function. an impella lp2.5 pump was placed in the patient without issue. the patient brought to the operating room on (b)(6) 2016 for 5 coronary artery bypass grafts (cabg.) the physician performed the procedure using the "cross clamp" technique. when the doctor clamped the aorta and manipulated the heart looking for targets the catheter pigtail was seen be sticking out of the ventricle. the catheter was immediately pulled back and the physician stitched the perforation. with the repair successfully completed the doctor completed the grafts. the clamps were then removed and the impella lp2.5 was repositioned back across the valve. the physician stated that the patient had "friable tissue." when stitching the aorta he found that the tissue was thin and could be torn easily. the clinical team believed that the left ventricle perforated when the heart was manipulated during the cabg procedure or possibly during cross clamping. following the procedure the patient was returned to the intensive care unit in stable condition. the patient had been successfully supported with the impella lp2.5 for over 67 hours when the impella lp2.5 was removed from the patient. the patient was reported to be "doing very well." | [5 empty] | The device was not returned for investigation; consequently a device failure evaluation could not be performed. the aic console logs were returned for analysis. the data logs provided information on the pump's performance and the patient's condition during the case. the long time logs showed that there was good patient pressure throughout the case. the normal patient pressures that were recorded for almost 2 days before the surgery make it unlikely that the ventricle was perforated during pump insertion. cardiac tamponade symptoms can manifest in as little as 30 minutes after perforation of the ventricle wall. no symptoms were noted leading up to the procedure making it unlikely that the ventricle was perforated before the cabg began. in addition, the logs show good flow and there were no suction alarms before the surgery. if a leak was present during support, low volume would cause suction alarms to be triggered. good flow suggests the ventricle was intact and completely sealed. it is likely that the ventricle was perforated during the surgery. the physician noted that the patient's cardiac tissue was prone to tears. the pump could have perforated the ventricle at two points during cabg. it is possible that when the pump was cross clamped it could have been moved and pushed through the wall, or it could have occurred when the physician manipulated the heart to look for targets. in conclusion, without the ability to evaluate the device in an effort to obtain additional information a definitive root cause for this event could not be determined. it is most likely that the root cause of the perforation was the friable cardiac tissue in the patient combined with manipulation of the heart during bypass with the impella in place. device was discarded subsequent to use.
38325 | 1220648-2016-00005 | 5583096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5583096&pc=PBL | 02/18/2016 | 03/18/2016 | 03/18/2016 | 04/15/2016 |  | 03/18/2016 | 03/18/2016 | 04/15/2016 |  |  | 01/31/2017 |  | 04/15/2016 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1182559 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 20 YR | [5 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Perforation | [12 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/18/2016 | * | HOSPITAL |  |  | The complainant reported that a hypotensive (b)(6) male patient who had been in gi distress for the past 10 days was admitted to the hospital. the patient had no past medical history. the patient was found to have an ejection fraction of 30-35% of unknown etiology. the patient was being managed on pressor support. the patient decompensated which resulted in intubation and mechanical support. the impella cp was placed in the patient without issue; the coronaries were not visualized during placement. the patient was then transferred to another facility for continued management and the potential need for escalation of support. the patient was reported to have been successfully supported and had "greatly improved." his ejection fraction had improved to 50%, and weaning was begun. after 4 days of pump support the impella cp was removed from the patient. the removal was reported as smooth with no resistance met. there was also no obvious "tissue" debris noted on the inlet cage or catheter apparatus following pump removal. two days post explant, the patient became critically unstable requiring reintubation. an echo was performed which revealed a flail posterior mitral valve leaflet and severe mitral regurgitation, which necessitated surgical mitral valve repair. the patient was referred for emergent mitral valve repair. the repair was reported as successful and the patient was released from the hospital a week later. | [5 empty] | A definitive root cause of the damage to the mitral valve could not be determined. the pump was discarded after the case and consequently unavailable for analysis. the exact cause could not be determined from the available data. according to reports from the site the positioning of the device was incorrect per the manufacturers' instructions, and the device was running while removing it from the heart, which is not per the manufacturers' instructions. device discarded by complainant facility.
38326 | 5569462 | 5569462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5569462&pc=PBL | 03/30/2016 | 04/01/2016 | 04/01/2016 | 04/12/2016 | 04/01/2016 | 04/01/2016 |  | 04/12/2016 |  |  |  |  | 04/12/2016 |  | BLOOD PUMP, TEMPORARY CARDIAC SUPPORT |  |  | IMPELLA CP |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003 |  | Yes | [5 empty] | No | Physician |  |  | User Facility report | [8 empty] | 47 YR |  |  | 83 |  |  | Malfunction |  | Yes | Fracture | Kinked | [5 empty] | No Code Available | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella device fractured during deployment. device was placed through a edwards 14f/35cm sheath. the typical sheath utilized 14f/30 cm cook was out of stock. device did kink after being placed into and out rv multiple times. this repositioning was necessary because of incessant vt episodes. fractured device was retrieved utilizing 2 snares. a new device was successfully placed through a 16f/30 cook sheath. this resulted in 2 hours of additional case time. patient was returned to the icu after successful coronary artery intervention in critical condition. lack of standard sheath (14f/30cm cook) and frequent repositioning of catheter were factors that contributed to this fracture.
38327 | 5482597 | 5482597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5482597&pc=DWB | 12/06/2015 | 12/18/2015 | 12/18/2015 | 03/06/2016 | 12/18/2015 | 12/18/2015 |  | 03/06/2016 |  |  |  |  | 03/06/2016 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA CP SET |  |  | DWB | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003 |  | Yes | [5 empty] | No | No Information |  |  | User Facility report | [8 empty] | 59 YR |  |  | 94 |  |  | Injury | Yes | No | Crack | Device Operates Differently Than Expected | [5 empty] | No Information | [41 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Abiomed impella lv support device was placed in patient in cath lab. device developed crack and failed to supply support. patient brought back to cath lab and device replaced.
38328 | 1220648-2016-00004 | 5466430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5466430&pc=KFM | 01/30/2016 | 01/30/2016 | 02/08/2016 | 02/29/2016 | 02/08/2016 | 01/30/2016 | 01/30/2016 | 02/29/2016 |  |  | 07/31/2017 |  | 02/29/2016 |  | IMPELLA LP5.0 |  |  | IMPELLA LP5.0 |  |  | KFM |  |  | 005062 |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1208235 |  | 005062 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 77 YR | [5 empty] | Death | Yes | No | Failure to Power Up | [6 empty] | Sudden Cardiac Death | [12 empty] | Death | [28 empty] | Initial submission | Followup | Followup |  |  | 01/30/2016 | Y | HOSPITAL |  |  | The complainant reported that a (b)(6) male patient was treated at a first hospital where he had undergone a cabg, and was reported to have spent 13 hours in the operating room. an iabp had been placed in the patient, but was unable to provide the necessary patient support. an impella lp5.0 was then placed. the patient was too unstable to initially move from or, and had to be defibrillated multiple times for a ventricular tachycardia. the patient was eventually stable enough to be transferred to intensive care unit. the patient was successfully supported for 11.4 hours and was then transferred to another hospital. upon entry at the second hospital the staff found evidence that the impella placement had not been done per recommended insertion technique, as the first clinical team had not used the components provided to implant the device. the clinicians at the facility where the patient was transferred also found that the graft had been placed deep under the patient's skin and was barely exposed. in addition, the sensor had been damaged, as a "placement signal unreliable" had been triggered. upon entry into the 2nd hospital the staff disconnected the white cable (that connects the impella lp5.0 to the automatic impella controller [aic]), from the first facility's aic and plugged it into the aic at the 2nd hospital. at this point the aic alarmed with an "impella defective" message, and the impella lp5.0 would not restart. the cable was then disconnected from the second aic and re-connected to the first aic, but the same alarm message was displayed. the patient had no underlying cardiac function, and internal cardiac massage was being performed during this time. the patient then decompensated and expired before a second imepella lp5.0 could be obtained and placed. | [5 empty] | Please provide the results for any investigation, evaluation, and/or failure analysis, including underlying cause identification, relevant to the reported event. please include: (1) an explanation for the reason for this occurrence based on the follow-up with the reporting facility or individual response: the complainant reported that a (b)(6) male patient was treated at a first hospital where he had undergone a cabg, and was reported to have spent 13 hours in the operating room. an iabp had been placed in the patient, but was unable to provide the necessary patient support. an impella lp5.0 was then placed. the patient was too unstable to initially move from or, and had to be defibrillated multiple times for a ventricular tachycardia. the patient was eventually stable enough to be transferred to intensive care unit. the patient was successfully supported for 11.4 hours and was then transferred to another hospital. upon entry at the second hospital the staff found evidence that the impella placement had not been done per recommended insertion technique, as the first clinical team had not used the components provided to implant the device. also, upon entry into the 2nd hospital the staff disconnected the white cable (that connects the impella lp5.0 to the automatic impella controller [aic]), from the first facility's aic and plugged it into the aic at the 2nd hospital. at this point the aic alarmed with an "impella defective" message, and the impella lp5.0 would not restart. the cable was then disconnected from the second aic and re-connected to the first aic, but the same alarm message was displayed. the patient had no underlying cardiac function, and internal cardiac massage was being performed during this time. the patient then decompensated and expired before a second impella lp5.0 could be obtained and placed. a complete description of investigation and analysis methodology(ies) used, response: the impella lp5.0 pump and the automatic impella console (aic) logs from the 2 consoles used during this event were returned for evaluation. the logs from the 2 consoles were analyzed. the analysis of the logs of the 2 consoles confirmed that "impella defective" alarms occurred, an "impella defective" alarm will prevent the pump from being recognized properly and will lead to a pump stop as was observed during this case. the pump plug was examined for any defects that could have led to the "impella defective" alarms. the examination revealed bent pins in the pump plug. these pins most likely became bent when the pump was disconnected during the transfer of the pump to the console at the second facility. once the pins were bent they triggered "impella defective" alarms to occur upon plugging into either console as seen in the aic data. the locking key on the pump plug was also examined and found to have rounded edges. this is indicative of the user twisting the plug upon disconnection. an identification of the specific failure mode(s) and or mechanism (s) and the associated device component(s) involved and response: when the patient was transferred to a 2nd hospital, the user disconnected the white cable and connected the red pump plug to the white cable of the 2nd hospital's aic. they received "impella defective" alarms from the aic indicating that there was a possible communication issue. the failure investigation determined that the 5.0 pump plug contained broken pins leading to the aic not recognizing the pump. the user twisting the impella 5.0 pump plug upon disconnection to transfer the impella 5.0 to the second site's console most likely led to the broken pins at the red pump plug. any conclusions reached based on the investigation and analysis results. response: the root cause of the pump stop was broken pins at the red pump plug, likely from when the pump was disconnected to transfer the impella 5.0 to the second site's console, resulting in "impella defective" alarms and ultimately a pump stop. the root cause of the broken pins at the red pump plug most likely occurred when the pump was disconnected to transfer the impella 5.0 to the second facility's console, which resulted in "impella defective" alarms and the pump stop. please provide any evaluation of other information used by your firm to determine whether the events described in the medical device report are or are not attributable to the device. response: there was no other information provided that has not been detailed in this report please identify if the patient or device problems (including any failure analysis) noted in this report are part of a trend analysis. please provide the basis for their inclusion in the trend, e.g. common clinical presentation/observation, common component, common manufacturing process, known failure mode, etc. response: the issue noted in this report is part of ongoing trend analysis. we conduct trend analysis of complaints by individual products based on failure mode. | The device is currently undergoing evaluation. upon the conclusion of the evaluation a supplemental medwatch report will be submitted. (b)(4) | This supplemental mdr is being filed to attach user mdr 2300530000-2016-08014 and to report the evaluation of the returned impella lp5.0 and to report the evaluation results of the returned impella lp5.0. the impella lp5.0 pump and the automatic impella console (aic) logs from the 2 consoles used during this event were returned for evaluation. the logs from the 2 consoles were analyzed. the analysis of the logs of the 2 consoles confirmed that "impella defective" alarms occurred, an "impella defective" alarm will prevent the pump from being recognized properly and will lead to a pump stop as was observed during this case. the pump plug was examined for any defects that could have led to the "impella defective" alarms. the examination revealed bent pins in the pump plug. these pins most likely became bent when the pump was disconnected during the transfer of the pump to the console at the second facility. once the pins were bent they triggered "impella defective" alarms to occur upon plugging into either console as seen in the aic data. the locking key on the pump plug was also examined and found to have rounded edges. this is indicative of the user twisting the plug upon disconnection. the root cause of the pump stop was broken pins at the red pump plug, likely from when the pump was disconnected to transfer the impella 5.0 to the second site's console, resulting in "impella defective" alarms and ultimately a pump stop. the root cause of the broken pins at the red pump plug most likely occurred when the pump was disconnected to transfer the impella 5.0 to the second facility's console, which resulted in "impella defective" alarms and the pump stop. internal reference: (b)(4).
38329 | 1220648-2016-00003 | 5464208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5464208&pc=DWA | 01/28/2016 | 01/28/2016 | 01/28/2016 | 02/26/2016 |  | 01/28/2016 | 01/28/2016 | 02/26/2016 |  |  | 08/31/2017 |  | 02/26/2016 |  | IMPELLA RP |  |  | IMPELLA RP |  |  | DWA |  |  | 004334 |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1206932 |  | 004343 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 44 YR | [5 empty] | Injury | No | Yes | Difficult to Insert | Suction Problem | Device Displays Incorrect Message | Positioning Problem |  |  |  | Hemolysis | Hematuria | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 01/28/2016 | * | HOSPITAL |  |  | The complainant reported that the physician was treating a (b)(6) male patient who had been transferred to the facility. on (b)(6) 2016 an impella cp ((b)(4)) was initially placed in the patient for bivad support. on (b)(6) 2016 the impella rp was also implanted. it was then decided to also replace the impella cp with an impella 5.0 ((b)(4)) to escalate the level of support to the patient. the impella cp was removed from the patient on (b)(6) 2016 after 27.35 hours of patient support, and the impella 5.0 was placed. the impella rp had been placed on (b)(6) without issue. the following day suction alarms occurred, and the performance level was reduced and the suction alarms resolved. on (b)(6) the impella rp pump was removed from the patient as a result of continued suction alarms and in an effort to obtain better pump positioning. a new sheath was obtained and the same impella rp was again placed in the patient. during this second placement of the device difficulties were incurred in getting the pump in.the proper location. the pump was successfully placed. following this second placement of the device the patient exhibited severe hemolysis requireing 3 to 5 units of replacement blood products to be administered to the patient. this impella rp was removed from the patient and a new impella rp was placed. this pump functioned without issue and the patient hemolysis was resolved. due to the patient's medical conditions the patient's family withdrew medical support on (b)(6) 2016 and the patient expired. the physician reported that the patient outcome was unrelated to the patient's hemolysis. | [5 empty] | The impella rp was returned for evaluation. in addition, the console logs were also returned for analysis. the evaluation of the pump revealed a significant amount of thrombus around the leading edge of the impeller, as well as the seal around the purge gap. after attaching a new catheter and plug to the pump and connecting it to a console the impella rp pump passed functional testing. the pump lost position multiple times due to patient hiccupping and the impella rp pump had to be repositioned under x-ray. hemolysis testing was also conducted to determine if the patient's hemolysis was due to defects of the impella rp pump. the pump was attached to another impella rp pump plug and the hemolysis bench test was performed. the results of the test revealed that the pump had an mih value of 6.2 at p9, which is a low normal value. this result indicates that this impella rp pump did not contribute to hemolysis in the patient. there were two issues for this case including: hemolysis, elevated ps. patient records show that his pfhg was at elevated levels prior to implantation of pump. the analysis of the console logs reveal that the increase pfhg levels during patient support align with suction alarms and improper pump positioning. after the pump was replaced the patient's hemolysis level was reduced. the hemolysis testing on the pump revealed that there was no abnormal pump related hemolysis. based on the symptoms, there is high probability that biomaterial was suctioned in ph, resulting in impeller rotational interference. post explant inspection discovered ingested clots. the following has been determined to be the root cause of this event: patient records show that pfhg was at elevated levels prior to implantation of pump. based on the symptoms, there is a high probability that biomaterial was ingested. post explant inspection discovered clots. performance in a standard hemolysis test in the lab showed pfhg level at 6.2 proving that the impella rp pump was not the cause of hemolysis. as a conclusion, the hemolysis was caused by possible pump position during the suction periods with elevated pfhg. it is not clear that there is a link between improper positioning and biomaterial ingestion. the hemolysis was not pump related. (b)(4). | The impella cp is still under evaluation. additional information of the details of this event are currently being sought. it is unclear if this reported issue was related to the impella rp, or if this issue was the result of the physician's device placement/management, or the patient's condition. the manufacturer will continue to investigate all reasonable obtainable source information and will provided the evaluation results and updated conclusions in a supplemental report. (b)(4). device currrently undergoing evaluation.
38330 | 5459889 | 5459889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5459889&pc=PBL | 01/11/2016 | 02/19/2016 | 02/19/2016 | 02/25/2016 | 02/19/2016 | 02/19/2016 |  | 02/25/2016 |  |  | 08/31/2017 |  | 02/25/2016 |  | TEMPORARY CARDIAC SUPPORT PUMP |  |  | IMPELLA 2.5 |  |  | PBL | [5 empty] | ABIOMED, INC |  |  | US | [5 empty] | Device was returned to manufacturer | [5 empty] | No | NURSE |  |  | User Facility report | [8 empty] | 81 YR |  |  | 63 |  |  | Malfunction | Yes |  | Improper or Incorrect Procedure or Method | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Myocardial Infarction | [10 empty] | Other | [28 empty] | Initial submission | [5 empty] | Y | HOSPITAL | 92903 |  | The patient who had been transferred from another facility following an acute chest pain episode. on arrival troponin levels were 5.83 and 6.89; diagnosed with non-st-segment elevation myocardial infarction (nstemi). patient became acutely hypotensive with worsening altered mental status and required an emergent cardiac catheter with an impella device placement. patient was critically ill and after two days of medical management, family decided to allow nature death for patient and not pursue any further treatment. impella device was to be removed two days later. during that process, nurse had pushed power off button and indicated to physician that they could remove device. as the impella catheter was being removed blood began to spray out, which no one expected. it was later discovered that the nurse had not pushed "enter" after putting the power level to power off. preliminary assessment was that the device did not malfunction; it was use error because nurse had not finished sequence when powering off device.
38331 | MW5060386 | 5454708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5454708&pc=OZD | 02/14/2016 | 02/19/2016 |  | 02/19/2016 |  |  |  | 02/19/2016 |  |  |  |  | 02/19/2016 |  | ABIOMED IMPELLA CP |  |  | ABIOMED IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 98318 |  |  |  | No answer provided | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 58 YR | [5 empty] | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Consequences Or Impact To Patient | [41 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Pt in operating room having lvad placed. implella device would not hold pump pressure high enough. should run between 350-600. pressure would not go higher than 200. no harm came to pt and there was no bad outcome.
38332 | 1220648-2016-00002 | 5427992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5427992&pc=PBL | 01/11/2016 | 01/11/2016 |  | 02/10/2016 |  |  | 01/11/2016 | 02/10/2016 |  |  | 06/01/2017 |  | 02/10/2016 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1203967 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 90 YR |  |  | 81 |  |  | Injury | Yes |  | Detachment Of Device Component | Improper or Incorrect Procedure or Method | [5 empty] | Foreign Body In Patient | [12 empty] | Death | Required Intervention | [27 empty] | Initial submission | Followup |  |  |  | 01/11/2016 | * | HOSPITAL |  |  | The complainant reported that a (b)(6) male cardiomyopathy patient had an impella cp placed using a perclose on (b)(6) 2016. the patient was reported to have clean arteries, and was diagnosed as takotsubo. the patient was supported with the pump for 66.13 hours. the patient was reported to have been doing well. all drips were stopped and the weaning process was begun on (b)(6) 2016. the impella cp patient support was reported as successful. on (b)(6) 2016 the pump was explanted from the patient. the physician explained that during explant the nurse helping in the procedure paused while pulling the device out, as she had observed blood coming from the outlet area and then stopped pulling. at that moment the physician was cinching down the sutures. after much force the device broke free and he was able to complete the pre-close process, and again there was bleeding observed from the site. the patient was reported to have remained in stable condition during the removal of the impella cp. following removal of the impella cp the physician noticed a "spring like" wire coming from the femoral artery. it was also noticed that the pigtail was not present. the physician explained that it appeared that the cannula had come apart leaving the pigtail, the inlet and part of the cannula still in the common femoral artery. the patient was then taken to surgery for removal of the remaining portions of the catheter. all remaining portions or the cannula were successfully removed from the patient. the patient was reported to have subsequently expired on (b)(6) 2016. the physician stated that the patient outcome was not as result of any issue with the device, but was due to the underlying heart issues that brought the patient to the hospital. | [5 empty] | Following the reported issue removing the device from the patient, and the subsequent required surgery to remove the detached portions of the impella cp cannual, the physician reported that he did not clearly understand that attempting to remove an impella through the percolse sutures is not recommended by abiomed or abbott, the manufacturer of the perclose device. the abiomed representative has provided retraining of the physician and the cath lab manager and staff. in addition, the abiomed representative has set up future in-service dates to help educate any new staff members and to provide ongoing education to the team. the impella cp was not returned for evaluation, so no device analysis could be performed. the manufacturer will continue to investigate all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. | Following the submittal of the initial mdr by abiomed for this event, the customer submitted user mdr report 199273664-2016-0001. this user mdr provided additional information not submitted on the initial mdr, and reported different information than was originally reported on abiomed's initial mdr. this supplemental mdr is correcting information previously submitted, adding additional information reported by the customer in their user mdr and additional information obtained during meetings with the complainant, and an abiomed's clinical review of the event. corrected information: death: (b)(6) 2016. relevant tests/laboratory data, including dates: (b)(6) 2016 @ 05:17 hgb 8.7 hct 26.2, (b)(6) 2016 @ 14:45 hgb 8.0 hct 25.9, (b)(6) 2016 @ 15:10 lactic acid 5.6, (b)(6) 2016 @ 11:14 lactic acid 1.8, (b)(6) 2016 @ 11:14 ck 631 trop 41.72 all high, but down from high values of (b)(6) 2016 @ 05:17 ck 998 ck-mb 134.6. other relevant history, including preexisting medical condition: nstemi, history of hypothyroidism, no known allergies,. this visit-septic shock , acute coronary syndrome, cad with stenosis of lad, atrial fibrillation, cardiogenic shock, takotsubo cardiomyopathy, acute coronary syndrome. concomitant products: left internal jugular 3 lumen central venous catheter (b)(6) 2016. clinical review of the event this was an implant of an impella cp on (b)(6) 2016 at 20:33 for takotsubo cardiomyopathy. in addition to takotsubo, the (b)(6) year old patient presented with septic shock, acute coronary syndrome, coronary artery disease (cad) with stenosis of the left anterior descending (lad) artery, atrial fibrillation, cardiogenic shock and a non-st elevated myocardial infarction (nstemi). the patient had a previous medical history of hyperthyroidism. the impella cp was implanted along with a perclose suture mediated closure (smc) system (abbott laboratories), utilizing the preclose technique for post-explant hemostasis. after implant on (b)(6) 2016, the patient was reported as doing well. all drips were stopped and the weaning process was begun on (b)(6) 2016. the following day, the pump was explanted. during removal of the impella, the physician reported the nurse pulling the impella out hesitated when she saw blood coming from the arteriotomy site. at the same time, the physician was cinching down the sutures on the perclose smc device. it was reported that removal of the impella pump required great force before the physician could complete the perclose procedure. following removal of the impella pump, bleeding continued from the arteriotomy site. in consultation with an abiomed clinical representative, the physician reported seeing a "spring-like" wire coming from the patient's femoral artery. he also stated there was no pigtail present on the explanted impella catheter. it was concluded that the cannula had come apart during the removal of the impella, leaving the pigtail, inflow cage and distal part of the cannula in the patient's common femoral artery. the patient was taken to vascular surgery to have the pump components removed. after surgery the patient was transferred to the cardiac icu where he suffered a cardiac (pea) arrest and expired at 15:20 on (b)(6) 2016. following the case, an abiomed clinical representative met with the physician and the catheter lab manager to discuss the case. it was reported that the physician did not understand that attempting to remove an impella through the perclose sutures is not recommended by abiomed or abbott. the abiomed representative has provided retraining of the physician and the cath lab manager and staff. the abiomed representative also set up future in-service dates to providing ongoing education to the team and any new staff members. (b)(4).
38333 | 5415836 | 5415836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5415836&pc=KFM | 01/30/2016 | 02/01/2016 | 02/01/2016 | 02/08/2016 | 02/01/2016 | 02/01/2016 |  | 02/08/2016 |  |  |  |  | 02/08/2016 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA |  |  | KFM |  |  | 0042-0000 |  |  | ABIOMED, INC. |  |  | US |  | NOT AVAILABLE |  |  |  | Device was returned to manufacturer | [5 empty] | No | No Information |  |  | User Facility report | [14 empty] | Death | Yes |  | Device Displays Incorrect Message | [6 empty] | Death | Low Blood Pressure/ Hypotension | [11 empty] | Death | [28 empty] | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Pt arrived via emergency medical services (ems) from outside hospital with impella in place. impella monitor alarming flow sensor disabled. attempted to change consoles, and upon switching consoles, new console alarmed flow detection disabled. we attempted to switch back to original console and the same flow detection disabled alarm message displayed. pt's blood pressure decompensated, code was called and patient expired. impella console taken to cath lab and impella catheter taken by a biomed rep.
38334 | 5403774 | 5403774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5403774&pc=KFM | 01/11/2016 | 01/13/2016 | 01/13/2016 | 01/14/2016 | 01/14/2016 | 01/14/2016 |  | 01/14/2016 |  |  | 06/30/2017 |  | 01/14/2016 |  | LEFT VENTRICULAR ASSIST DEVICE |  |  | IMPELLA |  |  | KFM |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  |  |  | REF# 0048-0003 |  | Yes | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 90 YR |  |  | 81 |  |  | Death | Yes |  | Nonstandard Device | [6 empty] | Cardiopulmonary Arrest | Death | Foreign Body In Patient | [10 empty] | Death | [5 empty] | LEFT INTERNAL JUGULAR 3 LUMEN CENTRA VENOUS | [22 empty] | Initial submission |  |  |  |  | 01/11/2016 | Y | HOSPITAL | 90106 |  | On (b)(6) 2016 removed impella lvad from patient, it was noted that distal end of catheter had broken and remained in patient's femoral artery. on (b)(6) 2016 12:51- 13:59 arteriotomy performed to remove distal portion. on (b)(6) 2016 at 1400 patient transferred to cardiac icu. on (b)(6) 2016 at 1520 patient suffered cardiopulmonary code pea arrest and expired.
38335 | 1220648-2016-00001 | 5358815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5358815&pc=PBL | 12/13/2015 | 12/13/2015 | 12/13/2015 | 01/11/2016 |  | 12/13/2015 |  | 01/11/2016 |  |  | 08/01/2017 |  | 01/11/2016 |  | IMPELLA CP, PRODUCT CODE: PBL |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1213544 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 58 YR | [5 empty] | Injury | Yes | No | Insufficient Information | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Blood Loss | [11 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 12/13/2015 | * | HOSPITAL |  |  | The complainant reported that a (b)(6) male patient presented with a segment elevation myocardial infarction (stemi:) resulting in an extracorporeal membrane oxygenation (ecmo) being placed in the patient. the patient also had three vessel disease (3vd) and his left ventricle was distended. the physician placed an intra-aortic balloon pump (iabp), as the patient's groin arteries were very small. the iabp did not decompress, resulting in an impella cp being placed in the patient. an antegrade sheath was also placed to provide flow down the patient's leg from the ecmo circuit. the impella cp was placed without issue in the patient. following patient treatment, bleeding was observed to be emanating from the impella insertion site. the pump angle was adjusted, but the bleeding continued, so a femstop was applied. during patient transport a significant amount of blood was lost. the patient blood loss was reported to be in excess of 15m/l of blood per minute. a vascular repair was performed that was reported to have resolved the bleeding at the impella site. the patient was also reported to have significant bleeding from the ecmo site, and was unrelated to the bleeding from the impella pump. the patient was reported to have been administered 50 units of replacement blood, but it was reported that the replacement blood was "mostly due to the groin bleeding from the ecmo site." the patient was reported to have been successfully supported for over 8 days with the impella cp, and was reported to be in stable condition at the conclusion of the impella cp support. | [5 empty] | The impella cp was not returned for evaluation, so no device analysis could be performed. it is unclear at this time if the artery bleeding was as a result of any issue with the introducer, impella cp pump, guide-wire or if the issue occurred as a result of the patient's anatomy, device placement or the procedure that was performed. the manufacturer will continue to investigate all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received. device discarded by user.
38336 | 1220648-2015-00009 | 5303488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5303488&pc=PBL | 11/19/2015 | 11/19/2015 |  | 12/16/2015 |  |  |  | 12/16/2015 |  |  | 08/31/2017 |  | 12/16/2015 |  | IMPELLA, CP PRODUCT CODE: PBL |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1212974 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | COMPANY REPRESENTATIVE | HEALTH |  |  |  |  | 76 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [33 empty] | 11/19/2015 | * | HOSPITAL |  |  | The complainant reported that a (b)(6) male patient was admitted to the (b)(6). it was determined that the patient has a 100% occlusion of the right coronary artery (rca) and a 99% occlusion of the left anterior descending artery (lad.) the rca received a full metal jacket (fmj: overlapping drug eluding stents >60mm in length) and the lad was stented. an intra-aortic balloon pump (iabp) was then placed. the patient continued to decompensate and the iabp it was decided to exchanged the iabp for an impella cp pump. the iabp sheath was exchanged for a 14x30cm cook brand sheath and the impella cp was inserted without issue into the correct position. the pump flows were set to auto and were 3.2-3.3 l/min and the waveforms were reported to be perfect. the sheaths were sutured into place. the impella displayed a "suction" alarm and the pump flows were 2.2 and decreasing. the p-level was then dropped to p5 and the alarms continued. the p-level then dropped to p2, and remained at that level. a fluid bolus was given and an stat echo was ordered to confirm catheter placement. the physician pulled the impella back, but it is thought to have been caught around the chorda, and when pulled, flung out of the left ventricle. the physician then attempted to re-cross, but the catheter prolapsed, or folded over on itself. the catheter was then pulled back further in the aortic arch, but the catheter would not uncurl. the plan was then to remove the device and reinsert it over a wire. as the device was removed from the sheath, it was noticed that the pigtail was not longer on the catheter. a fluoroscopy revealed that the pigtail tip/inflow cage had detached and remained the aortic arch. the pigtail was reported to have been successfully retrieved with the use of a snare. the patient was stable. the complainant reported that there were no adverse effects and no harm to the patient as a result for the reported issue. another impella cp was placed and the patient was reported to have been successfully supported. | [5 empty] | The device was returned in with the inflow cage and pigtail separated from the cannula. through bench testing it was found that attempting to pull a prolapsed cannula through the introducer would produce a similar failure. the cannula would separate from the inflow cage and be connected only by the nitinol coil. two pumps were tested, both failed in the same way, one at 41n and one at 32n, both higher than the specification of 30n. the evaluation of the returned device determined that the cannula likely prolapsed because the physician attempted to reinsert the pump wirelessly. when the physician attempted to remove the prolapsed cannula through the introducer the inflow cage likely got stuck on the introducer tip and separated when the physician applied significant force. bench testing at abiomed showed that attempting to remove a prolapsed cannula through the introducer with significant force (32-41 n) would cause the cannula to separate from the inflow cage. in conclusion and as stated above, the root cause of this event is most likely due to the cannula prolapsing. when the physician attempted to remove the prolapsed cannula through the introducer the inflow cage likely got stuck on the introducer tip and separated when the physician applied significant force. bench testing at abiomed showed that attempting to remove a prolapsed cannula through the introducer with significant force (33-41 n) would cause the cannula to separate from the inflow cage. this physician and staff are being provided with additional training of the use of the guidewire during insertion to prevent prolapsing the cannula. due to the low occurrence of this failure mode there is no other corrective action recommended. this failure mode will continue to be monitored. a review of complaint histories revealed that no other pumps from this same lot number have been returned for a similar issue. in addition, a device history record review was performed and did not reveal any issue that would have caused or contributed to this complaint issue. the ifu for this device use advises the user of the following: if the impella catheter is out of the ventricle, do not attempt toe reposition the catheter across the valve without a guidewire.
38337 | 1220648-2015-00008 | 5243210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5243210&pc=PBL | 09/10/2015 | 10/23/2015 | 09/22/2015 | 11/23/2015 | 09/22/2015 |  |  | 11/23/2015 |  |  |  |  | 11/23/2015 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | PBL |  |  | IMPELLA CP |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | UNK |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | OTHER | [5 empty] | 84 YR | [5 empty] | Injury | Yes | Yes | [20 empty] | Required Intervention | [33 empty] | 10/29/2015 | Y | HOSPITAL |  |  | During emergency cardiac catheterization procedure the md was attempting to insert an introducer for a left ventricular assist device. as he was doing so, it was noted that the tip of the sheath "accordioned" or bunched up, which prevented it from being threaded into the right femoral artery. subsequently, the femoral artery was perforated, "which may or may not have been the result of the sheath tip deformity." | [5 empty] | The reported issue was not reported to abiomed by the customer. neither the impella cp, 14french introducer nor guidewire were returned for evaluation. it is unclear at the present time if the femoral artery perforation was as a result of any issue with the introducer, impella cp pump, guidewire or if the issue occurred as a result of the patient's anatomy, device placement or the procedure that was performed. the manufacturer will continue to investigate all reasonable obtainable sources of information and will provide results and conclusions in a supplemental medwatch report if additional information is received.
38338 | MW5057223 | 5161871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5161871&pc=KFM | 10/02/2015 | 10/14/2015 |  | 10/14/2015 |  |  |  | 10/14/2015 |  |  |  |  | 10/14/2015 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | KFM | [5 empty] | ABIOMED |  |  | US | [5 empty] | No answer provided | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 55 YR |  |  | 95 |  |  | Injury | Yes | Yes | [20 empty] | Required Intervention | [28 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Abiomed impella circulatory assist device shut off without warning/precipitating event; dull beep could be heard but only if actually in room at bedside. no adverse patient event occurred. a biomed representative was contacted and removed device from service.
38339 | MW5056366 | 5096394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=5096394&pc=KFM | 09/06/2015 | 09/17/2015 |  | 09/17/2015 |  |  |  | 09/17/2015 |  |  |  |  | 09/17/2015 |  | IMPELLA 2.5 |  |  | IMPELLA 2.5 |  |  | KFM |  |  | IMPELLA 2.5 |  |  | ABIOMED |  |  | US |  |  |  | 005042 |  | No answer provided | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 72 YR |  |  | 68 |  |  | Injury | No | Yes | [20 empty] | Required Intervention | [5 empty] | CB RX 3.0 X 10 MM | [22 empty] | Initial submission | [5 empty] | I | NO INFORMATION |  |  | Stemi pt with emergent ptca and insertion of impella left ventricular assist device. after insertion, pump mechanism on impella catheter not working. vendor present and diagnostic testing and checks completed. unable to resolve malfunction so device removed and new impella catheter inserted.
38340 | 1220648-2015-00006 | 4943485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4943485&pc=DWA | 06/27/2015 | 06/27/2015 | 06/27/2015 | 07/27/2015 |  | 06/27/2015 | 06/27/2015 | 07/27/2015 |  |  | 05/31/2017 |  | 07/27/2015 |  | IMPELLA RP |  |  | IMPELLA RP |  |  | DWA |  |  | 004334 |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1193394 |  | 004334 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 54 YR | [5 empty] | Death | Yes | Yes | [20 empty] | Death | Required Intervention | [27 empty] | Initial submission |  |  |  |  | 06/27/2015 |  | HOSPITAL | 565404779 |  | The complainant reported that the clinicians were treating a (b)(6) male patient who was nitric oxide dependent, and was 2 weeks post left ventricular assist device placement. the patient was prepped for impella rp placement in order to provide primary support. a swann was placed in the left pulmonary artery with a 0.035 wire. an oscor rp sheath was placed. the patient's act reading was 202, and he was given and additional heparin bolus of 5000 units prior to the removal of the dilator. the swan-ganz catheter was placed deep in the lpa followed by the 0.025 wire. at this point the patient began bleeding quite a bit from the sheath's valve before the impella rp was placed. attempts to reset the valve were unsuccessful. the impella rp catheter was primed and back-loaded, and was placed into the oscor rp sheath, and then advanced to the inferior vena cava and zeroed. during the advancement, the catheter passed the tricuspid to the pulmonic valve, but then the wire prolapsed and was pulled from the left pulmonary artery. the physician was able to advance the wire while still on the impella rp into the right pulmonary artery. the patient bleeding had continued throughout the impella rp insertion. at this time the physician ordered a transfusion of blood, and 4 units were administered to the patient. due to patient bleeding the staff decided to accept a right pulmonary artery position for the device. the impella rp placement was completed, the introducer was removed and the physician performed a deep mattress suture for hemostasis, no additional bleeding issues occurred from the oscor sheath. the patient was supported by the impella rp for over 16 days, and the patient continued decompensating due to his underlying condition of pulmonary hypertension, sepsis and being vasodilated. the impella rp was removed from the patient on (b)(6) 2015 as the patient's family withdrew care. on (b)(6) 2015 the patient expired. the patient outcome was not as result of any issue with the impella rp. | [5 empty] | The device was returned for evaluation for the introducer failure. the clinical information and introducer were analyzed for the failure investigation. the clinical information described the patient bleeding from the introducer valve after insertion of the swan ganz catheter into the left pulmonary artery. this bleeding was not resolved before or during insertion of the impella rp into the patient. the examination of the valve on the sheath revealed an off-centered tear, likely occurring during insertion of the swan ganz. this off-centered insertion resulted in a torn and open valve condition that continued during the remainder of the case, including during the insertion fo the impella rp. the root cause of this reported event has been determined to be the off-centered insertion of the swan-ganz catheter through the oscor sheath which led to the bleeding from the introducer. additional staff inserviceing has been performed.
38341 | 1220648-2015-00005 | 4846681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4846681&pc=PBL | 06/05/2015 | 06/12/2015 |  | 06/12/2015 |  |  | 06/12/2015 | 06/12/2015 |  |  | 01/10/2017 |  | 06/12/2015 |  |  |  |  | IMPELLA CP |  |  | PBL |  |  | 0048-0003 |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  | 1180899 |  | 0048-0003 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | [20 empty] | Life Threatening | Required Intervention | [27 empty] | Initial submission | [7 empty] | 00813502011258 |  | Abiomed received a report of a (b)(6) female pt with a history of pt foramen ovale (pfo) and complaining of shortness of breath who returned to the facility's cath lab for a scheduled percutaneous coronary intervention (pci). the pt had an ejection fraction of 20%. two days previous the pt had had a coronary angiogram performed that revealed severe disease in the left anterior descending coronary artery (lad), diagonal artery, circumflex artery and the first obtuse marginal branch (om). the physician's plan was to support the pt with the impella cp, to perform a coronary stent implantation (csi) of both main vessel, and then to stent the main and branch vessels. shortly after arrival in the operating room and being sedated and draped the pt began to exhibit signs of cardiogenic shock, with mean blood pressures in the 40s. radial access was obtained and p-pressures were stabilized and a bilateral iliac angiogram was performed to assess the best placement for the impella sheath. perclose was performed and the sheath was placed in the left groin. the impella cp was placed using jr4 and .035 j wires to cross the valve then it was exchanged out with the provided .018 wire. the impella cp was inserted over wire. during the removal of the guidewire the physician grabbed a few centimeters from behind the peel-away sheath (rather than pinning the catheter further back and in one motion pulling the sheath out) after the dr broke the oscor sheath away, the repositioning sheath was flushed, and he had a struggle advancing it. he ended up taking both the repositioning sheath and sleeve together and inserted it in place. the impella cp, including the pigtail was visualized within the cardiac silhouette throughout the intervention. dopamine was titrated back up and the pt remained labile for the rest of the procedure, but required occasional epinephrine boluses to maintain pressure while the impella was in place. due to the pt's instability the choice was made to not perform the csi on the pt's arteries, but to predilate with an angiosculpt balloon. stents were placed to the lad and circumflex artery. the guide wires were removed from the coronary arteries and the impella catheter was pulled back some, as it didn't appear as far out to the apex as it did during the intervention. final shots were obtained of the left coronary system, as was a shot of the right coronary artery (rca). at this point in the case the pt is still very labile and the decision is made to keep the impella device in and transfer to icu. one of the next steps was inserting the repositioning sheath. the physician held pressure on the groin while a fellow proceeded with the exchange. he started by grabbing just a few centimeters from behind the peel away sheath and proceeded to do a push/pull method to pull the oscor sheath out, rather than just pinning the catheter further back and in one motion pulling the sheath out. then after he broke the oscor away, he flushed the repositioning sheath and had a little struggle advancing it. rather than disconnecting it from the repositioning sleeve taking the repositioning sheath in by itself, he took both the repositioning sheath and sleeve together and inserted them in place. they then began working towards prepping the right ij, because they also wanted to place a swan and evaluate right heart function. while they were working towards prepping and getting venous access, hemodynamics were changing, and pulse pressures were becoming narrowed. impella flow quickly decreased. the physician observed that the borders of the heart wall were not moving. an echo confirmed pericardial tamponade. a pericardiocentesis was then performed. impella flows increased, but pt's condition continued to diminish. there was approx 1 liter of blood obtained from the pericardiocentesis and it didn't appear to be stopping. the pt continued to bleed, and blood was transfused. the pt was intubated at some point and was also shocked. a decision was made to transfer the pt to another medical facility to provide the pt with emergent surgical support. upon the pt's arrival she was quickly prepped and draped and a sternotomy was performed. the cardiovascular surgeon quickly recognized the pigtail portion of the impella catheter exiting the heart wall. the impella catheter's pigtail was clipped off by the surgeon and the rest of the device was then pulled back across the aortic valve. the pt was transitioned to the cardiohelp, heart-lung support system, the ventricular repair was performed, and multiple transfusions were given. the impella cp was turned off, and it remained in the pt aorta until it was eventually explanted and perclosed by the surgeon. the pt was then transferred to a third facility on cardiohelp support. (b)(4). | [5 empty] | Neither the impella cp pump, repositioning sheath nor introducer were returned for eval. it is unclear at the present time if the reported issue resided with the impella cp pump, the repositioning sheath, the 14 french introducer, or if the issue occurred as a result of the device placement or the procedure that was performed. the mfr will continue to investigate all reasonably obtainable sources of info and will provide results and conclusions in a supplemental medwatch report if additional info is received. (b)(4).
38342 | 1220648-2015-00004 | 4727909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4727909&pc=KFM | 03/27/2015 | 03/28/2015 | 03/27/2015 | 04/24/2015 |  | 03/27/2015 | 03/27/2015 | 04/24/2015 |  |  | 12/01/2016 |  | 04/24/2015 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | KFM |  |  | IMPELLA LP5.0 |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 2015221440 |  | 004680 |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 67 YR | [5 empty] | Malfunction | Yes | Yes | [20 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 03/27/2015 |  | HOSPITAL |  |  | The complainant reported that on (b)(6) 2015, a (b)(6) year old male patient about to undergo a catheter ablation of ventricular tachycardia (vt ablation.) in preparation for the procedure an impella cp was placed in the patient. during the impella placement the perclose was deployed and the 14fr oscor sheath was inserted. with the insertion of the sheath the patient's pedal pulses were reduced, and it was noted that the patient's vessels were narrow, but placement was reported as uneventful. once the impella cp was placed the introducer sheath was cracked, peeled and removed. the perclose was cinched against the 9fr catheter body and hemostasis was achieved. the repositioning sheath was not advanced at this point. the ep procedure was performed, and during the procedure the patient coded. the patient responded quickly to the cpr that was performed, and the vt ablation procedure was aborted and could not be completed. the patient was prepared for transfer to the icu. as a result of the patient coding the physician opted to continue impella cp for the patient. the repositioning sheath was advanced approximately 1/3 of its length until the sheath was just in the femoral artery, in order to limit limb ischemia. the patient continued to be successfully supported by the impella for the entire, and remained in stable condition. there was no patient bleeding, and there were no issues with the impella. on (b)(6) 2015, the attending physician retracted impella cp up to and including the pump housing. at this point the catheter slipped from the patient's groin site without the cannula assembly. a spurt of blood from wound site was noted. the attending physician then pinched the cannula to stop the bleeding. the perclose sutures were held tight around catheter/repositioning sheath to gain hemostasis. the pump's pigtail and inlet cage were found to be stuck in the patient's femoral and he was taken to surgery for removal of the pump components and the application of a vascular patch. following successful removal of the pump's pigtail and inlet cage the surgeon reported that the pigtail had been caught in the calcification and that there was worsening calcification in the artery. there were no further patient issues related to the impella cp. | [5 empty] | The impella cp is still under evaluation. it is unclear if this reported issue was related to the impella cp, or if this issue was the result of the physician's device placement/management, or the patient's condition. the manufacturer will continue to investigate all reasonable obtainable source information and will provided the evaluation results and updated conclusions in a supplemental report.
38343 | 1220648-2015-00003 | 4727919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4727919&pc=KFM | 03/01/2015 | 03/02/2015 |  | 04/24/2015 |  |  | 03/02/2015 | 04/24/2015 |  |  | 12/01/2016 |  | 04/24/2015 |  | BYPASS BLOOD PUMP |  |  | IMPELLA LP5.0 PUMP |  |  | KFM |  |  | 004680 |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 2015221440 |  | 004680 |  | Device was returned to manufacturer | [5 empty] | No | Other |  |  | Manufacturer report |  |  | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR | [5 empty] | Malfunction | Yes | Yes | [20 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | The complainant reported that an impella lp5.0 had been placed via the subclavian route in post-cardiotomy in a (b)(6) male patient. the patient was successfully supported by the impella for over 5 days. the device was surgically removed form the patient without effort. during removal of the pump it became cut into two pieces, leaving one piece in the distal subclavian artery. this portion was recovered via femoral access, with the aid of a femoral endovascular loop. following this event the patient was reported to be in stable condition, with no adverse effects as a result of the reported issue. | [5 empty] | The impella lp5.0 is currently under evaluation. in addition, abiomed is currently attempting to connect with the physicians who were responsible for the case to discuss with them what they did in the handling of the pump, and how this was discovered since the pump ran normally during the course of treating the patient. the manufacturer will continue to investigate all reasonably obtainable sources of information, and will provide results and conclusions in a supplemental medwatch report when the evaluation is completed, and will include in the report any additional information obtained from the treating physicians.
38344 | 4745088 | 4745088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4745088&pc=KFM | 04/22/2015 | 04/22/2015 | 04/22/2015 | 04/22/2015 | 04/22/2015 | 05/04/2015 |  | 04/22/2015 |  |  |  |  | 04/22/2015 |  | CATHETER, INTRAVASCULAR, DIAGNOSTIC |  |  | IMPELLA CP |  |  | KFM |  |  | * |  |  | ABIOMED, INC. |  |  | US |  | * |  | * |  | Device was returned to manufacturer | [6 empty] | Invalid Data |  |  | User Facility report | [8 empty] | * | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | Impella device malfunctioned.
38345 | MW5042181 | 4714064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4714064&pc=KFM | 03/31/2015 | 04/15/2015 |  | 04/15/2015 |  |  |  | 04/15/2015 |  |  |  |  | 04/15/2015 |  | IMPELLA |  |  | IMPELLA 2.5 DEVICE |  |  | KFM |  |  | 2.5 LITERS |  |  | ABIOMED |  |  | US |  | 2015 - 5042 | [9 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 67 YR |  |  | 74 |  |  | Injury | No | Yes | [20 empty] | Life Threatening | [28 empty] | Initial submission | [9 empty] | Post cvl procedure - impella began to alarm that it had stopped. impella rep at bedside, console changed and same alarm continued to occur. unable to get impella motor to start. problem with the actual impella inside patient body. impella was removed by physician. rep took equipment with her to test.
38346 | 1220648-2015-00002 | 4639087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4639087&pc=KFM | 02/28/2015 | 02/28/2015 |  | 03/27/2015 |  |  | 03/06/2015 | 03/27/2015 |  |  | 08/01/2016 |  | 03/27/2015 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | KFM |  |  | 0048-0003 | [7 empty] | 1170654 |  | 0048-0003 |  | Device was returned to manufacturer | [6 empty] | Physician |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation | [10 empty] | Yes | Yes | [20 empty] | Death | Required Intervention | [27 empty] | Initial submission | [6 empty] | HOSPITAL |  |  | Complainant reported that a (b)(6) male patient was undergoing the placement of an impella cp. the physician attempted to perform a wireless insertion of the device. as the doctor was traversing the aorta the cannula prolapsed onto itself, and the pigtail became caught on the cannula. the physician then tried to grab it using a snare, but was unable to free the pigtail and to pull the pump back into the sheath. a vascular surgeon was called and the pump was successfully removed from the patient. the patient subsequently expired, but the clinician reported that the patient outcome was not as a result of any issue with the impella cp, as the impella placement was a "last ditch effort" to save this patient. | [5 empty] | The impella cp is still under evaluation. additional information of the details of this event are currently being sought. it is unclear if this reported issue was related to the impella cp, or if this issue was the result of the physician's device placement/management, or the patient's condition. the manufacturer will continue to investigate all reasonable obtainable source information and will provided the evaluation results and updated conclusions in a supplemental report. (b)(4).
38347 | 4584509 | 4584509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4584509&pc=DSQ | 02/28/2015 | 03/04/2015 | 03/04/2015 | 03/04/2015 | 03/04/2015 | 03/04/2015 |  | 03/04/2015 |  |  |  |  | 03/04/2015 |  | VENTRICULAR ASSIST DEVICE |  |  | IMPELLA CP |  |  | DSQ | [5 empty] | ABIOMED |  |  | US | [5 empty] | Yes | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 69 YR |  |  | 94 |  |  | Death | Yes | Yes | [20 empty] | Death | Life Threatening | Required Intervention | [26 empty] | Initial submission |  |  |  |  | 02/28/2015 | Y | HOSPITAL | 78536 |  | A (b)(6) year old male playing tennis on (b)(6) 2015 and experienced a cardiac arrest with subsequent fall and head trauma. he was resuscitated in the field. upon admission to the emergency room the ct was negative for intracranial bleed, ph was 6.93. he had st elevation acs on ecg and was taken to the cardiac cath lab for emergent treatment. an impella ventricular assist device was inserted but it folded upon itself. the cardiologist was unable to free the catheter from the right external iliac artery. the device and sheath were sewn in place. the patient was taken to the icu remaining in cardiogenic shock with high pressor and inotropic needs. the right lower extremity was ischemic with no doppler signals in the right foot. patient returned to the operating room for retrieval of the device and implant of another impella. a diagnostic iliac arteriogram was performed demonstrating tortuosity of the right external iliac artery. the previous retained device was found in pieces and removed. dissection of the artery required stent placement. patient continued to deteriorate despite maximum pressor and inotrope support and expired.
38348 | 4687030 | 4687030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4687030&pc=KFM | 02/11/2015 | 02/17/2015 | 02/17/2015 | 02/17/2015 | 02/17/2015 | 04/14/2015 |  | 02/17/2015 |  |  |  |  | 02/17/2015 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA CONTROLLER |  |  | KFM |  |  | * |  |  | ABIOMED, INC. |  |  | US |  | * |  | * |  | Device was returned to manufacturer | [6 empty] | Invalid Data |  |  | User Facility report | [8 empty] | 76 YR | [5 empty] | Malfunction |  | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | The patient had left-sided impella placed. the next day, the impella began with catheter positioning alarms the motor current waveform went flat. physicians confirmed appropriate catheter location - waveform did not return and the console continually alarmed. since the physicians determined the catheter placement was correct (through visual inspection), no additional actions were needed. the rep was notified and will take the device for analysis. there was no detectable harm to patient noted due to alarm failure.======================manufacturer response for vad, impella contorller (per site reporter).======================requested that abiomed share findings of their analysis.
38349 | 1220648-2015-00001 | 4485103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4485103&pc=KFM | 01/06/2015 | 01/06/2015 |  | 02/04/2015 |  |  | 01/06/2015 | 02/04/2015 |  |  | 09/01/2016 |  | 02/04/2015 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | KFM |  |  | 0048-0003 |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 1171855 |  | 0048-0003 |  | No | [5 empty] | No | Physician |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 40 YR | [5 empty] | Injury | Yes | No | [20 empty] | Life Threatening | Required Intervention | Disability | [26 empty] | Initial submission |  |  |  |  | 01/06/2015 | N | HOSPITAL | 00813502011289 |  | Complainant reported that a (b)(6) female patient presented at (b)(6) hospital with a stemi (st segment elevation myocardial infarction). while the staff was in the process of preparing the patient for air-transport to one hospital, the patient coded. the patient was administered 3 pressors and an impella cp was successfully placed in the patient's lfa (left femoral artery.) the impella was turned on, and was reported to have been working well. the patient transport was then changed to (b)(6) medical center. during check-in at (b)(6) medical center intensive care unit immediate suction alarms occurred when the suction control was disabled. during advancement of the repositioning sheath the catheter came out of the left ventricle. the physicians attempted re-advancement of the catheter multiple times, but they were unsuccessful. a second impella was inserted, into the right femoral artery (rfa) and was successfully placed into lv and turned on without issue. the physicians pulled the impella cp in the lfa down into the descending aorta/lfa and left it in place, but with the power off. the patient was sent to the icu with the impella cp in the patient's rfa and in auto mode. it was determined that the patient had no pulse in the right foot, but in the icu a dopplered pulse was found in the patient's left foot. a vascular surgeon was consulted and saw the patient on (b)(6) 2015 and continued to follow the patient. the patient was unstable, and was on multiple pressors and with blood sugars above 300. the following day the surgeon agreed to accept the patient at (b)(6() hospital for possible impella 5.0 or ecmo placement due to the patient's hemodynamic instability and the need for a higher level of care. the patient arrived at (b)(6) hospital where no pulses were palpated in either foot, and the legs were found to be stiff with no sensation. the impellas explanted and bilateral fasciotomies were performed. during that night, the patient required bilateral amputation below the knee (bka). additional information received from the physician at (b)(6) medical center revealed that this physician was present when the patient was transferred from (b)(6) medical center this physician evaluated the patient's bilateral lower extremity ischemia. he felt that there was no flow in either extremity for a significant amount of time, and subsequently his team amputated both the patient's lower extremities. he reported that unequivocally that this event was unrelated to impella cp, and wast related more to the fact that the previous facility's team had left the device off and not running for the transfer, and had indeed pulled the device back into the pelvis and didn't leave it in the descending aorta. both of these actions are in distinct opposition to the standard recommended procedure of leaving the device in the descending aorta and leaving the device running at a low p level. despite the fact that this occurred with the device in place, he felt very strongly that this was a combination of a vasculopath as well as a patient who was not handled well medically from the decision-making standpoint. the physician did report that this patient's ejection fraction (ef) improved to 35%. the doctor further stated that this patient's recovery was likely due to the impella cp support. | [5 empty] | The impella cp was not returned for evaluation, as it was discarded subsequent to use; consequently an evaluation of the device and event was unable to be performed. the physician evaluated the patient's bilateral lower extremity ischemia. he felt that there was no flow in either extremity for a significant amount of time, and subsequently his team amputated both the patient's lower extremities. he reported that unequivocally that this event was unrelated to impella cp, and wast related more to the fact that the previous facility's team had left the device off and not running for the transfer, and had indeed pulled the device back into the pelvis and didn't leave it in the descending aorta. both of these actions are in distinct opposition to the standard recommended procedure, (as outlined in the impella cp instructions for use) of leaving the device in the descending aorta and leaving the device running at a low p level. despite the fact that this occurred with the device in place, the physician felt very strongly that this was a combination of a vasculopath, as well as a patient who was not handled well medically from the decision-making standpoint. device not returned for evaluation.
38350 | 1220648-2014-00009 | 4337474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4337474&pc=KFM | 11/21/2014 | 11/21/2014 |  | 12/17/2014 |  |  | 11/21/2014 | 12/17/2014 |  |  | 03/22/2016 |  | 12/17/2014 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | KFM |  |  | 0048-0003 |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 201520675 |  | 0048-0003 |  | No | [5 empty] | Y | Physician |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 78 YR | [5 empty] | Death | Yes | Yes | [20 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 11/21/2014 |  |  | 00813502011289 |  | Complainant reported that after the clinicians performed 3 hours of patient cpr in the cath lab on a (b)(6) year old male patient who had a history of coronary artery disease, coronary artery by-pass graphs, left internal mammary artery (graft) to the left anterior descending artery patent. the impella cp as placed in the patient and a stent was placed. the patient remained in the cath lab for 5 hours. the patient was moved to the icu, where he again coded, but the impella kept the mean arterial pressure (map) at 64. an echo was performed, which resulted in 4 cm's of catheter slack removed, and the impella was in a more optimum position. the impella cp was successfully supporting the patient. after > 21 hours of successful impella support, the pump suddenly stopped without any previous indication that a pump stop was imminent. attempts to restart the impella cp were unsuccessful, and the patient arrested. cpr was begun on the patient, but the family decided not to continue therapy and the patient expired. | [5 empty] | The impella cp pump was not returned for evaluation. the automatic impella controller (aic) log data were returned for analysis. the analysis determined that impella cp support was initiated at 16:15 on (b)(6) 2014. early problems with a blocked placement signal lumen and pump placement were resolved after an echo was performed and 4 cm's of catheter slack was removed resulting in more optimum impella positioning. the analysis of the long term logs revealed a sudden pump stop had occurred approximately twenty-three hours after the pump was implanted. the pump stop was caused by a abrupt high motor-current spike. both "impella stopped - motor current high" (error-13) and "impella stopped" (error-119) alarms were logged by the console when the pump first stopped, and during subsequent attempts to restart. the error-13 alarm was triggered by 1850 ma motor-current spikes while the error-119 alarm indicated that a motor current imbalance of greater than 12% had occurred. while the pump was not returned for analysis, the combination of high motor current and a phase imbalance is strongly associated with a motor coil short and was the likely cause of the pump stop in this case. a corrective action has been initiated to address this problem.
38351 | 4336984 | 4336984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4336984&pc=KFM | 11/21/2014 | 12/03/2014 | 12/04/2014 | 12/04/2014 | 12/04/2014 | 12/04/2014 |  | 12/04/2014 |  |  |  |  | 12/04/2014 |  | AUTOMATED INTERNAL CONTROLLER (AIC) |  |  | IMPELLA |  |  | KFM | [5 empty] | ABIOMED, INC. US HEADQUARTERS |  |  | US | [5 empty] | No | [5 empty] | No | Health Professional |  |  | User Facility report | [8 empty] | 78 YR |  |  | 73 |  |  | Death | Yes | No | [20 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 11/21/2014 | Y | HOSPITAL |  |  | Pt on hemodynamic support with abiomed impella device. device alarmed "low flow" and ceased functioning. pt decompensated followed by cardiac arrest. pt expired. pt on hemodynamic support with abiomed impella device.
38352 | 1220648-2014-00008 | 4281357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4281357&pc=KFM | 10/21/2014 | 10/23/2014 |  | 11/20/2014 |  |  | 10/23/2014 | 11/20/2014 |  |  | 02/13/2016 |  | 11/20/2014 |  | LP2.5 |  |  | IMPELLA RECOVERY LP2.5 PERC CARDIAC SUPPORT |  |  | KFM | [5 empty] | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 88183 |  | 0046-3061 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 42 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [7 empty] | 00813502011081 |  | The complainant reported that during a long ep procedure, the patient's left ventricle had been perforated. the physician reported that after the ep procedure was performed, heparin was reversed with protamine and this began an anaphylactic shock sequence so emergently that the doctor placed the impella lp2.5 in the patient while the patient was undergoing cpr. with the pump on auto flow the ccp reported that he saw 2 liters of flow was obtained it was reported. the patient was eventually sent to the operating room. upon the patient's arrival in the operating room, the impella lp2.5 had already been removed. the ventricle was repaired and the patient was put on bypass. an ecmo v-a was placed. at this time, the surgeon mentioned that he thought the impella pigtail had kinked and that was what had perforated the left ventricle, and which he believed occurred during cpr in the cath lab. the surgeon was not present during the impella procedure. during a follow-up meeting with the abiomed sales representative, the surgeon first reported that the pump pigtail had come apart, but then conjectured instead that during the cpr the catheter must had kinked and that was how it perforated the lv. during this meeting, the surgeon retrieved the used catheter and bent the pigtail back in on itself to demonstrate the kink, and how this was enough of an edge to be sticking out of the lv perforating the wall. the patient subsequently expired. there was no indication that any issue with the impella lp2.5 contributed to the outcome. | [5 empty] | The impella lp2.5 pump was not returned for evaluation, as it is currently being held in the facility's risk management department. it is unclear if this reported issue was related to the impella lp2.5 pump, or if the issue occurred as a result of the procedure that was performed prior to the impella placement. the manufacturer will continue to investigate all reasonably obtainable sources of information, and will provide results and updated conclusions in a supplemental report if additional information is received. (b)(4).
38353 | MW5039210 | 4270851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4270851&pc=PBL | 11/01/2014 | 11/17/2014 |  | 11/17/2014 |  |  |  | 11/17/2014 |  |  |  |  | 11/17/2014 |  | HEART PUMP |  |  | IMPELLA |  |  | PBL | [5 empty] | ABIOMED, INC. | [14 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 83 YR |  |  | 50 |  |  | Injury | Yes | Yes | [20 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | Icu manager reports: impella device failed on the pt. she came from the cath lab with the hernia. per physician, her prognosis was very poor and family felt she would not have wanted further aggressive treatment. the device was not replaced and was removed and returned to the rep. he indicated results should be available within 6 weeks. he suspected blood clot vs motor failure.
38354 | 4168781 | 4168781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4168781&pc=KFM | 06/15/2014 | 10/08/2014 | 10/08/2014 | 10/08/2014 | 10/08/2014 | 10/14/2014 |  | 10/08/2014 |  |  |  |  | 10/08/2014 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP SET |  |  | KFM |  |  | 0048-0003 |  |  | ABIOMED INC |  |  | US |  | 71393A |  | 0048-0003 |  | Yes | [5 empty] | No | UNKNOWN |  |  | User Facility report | [8 empty] | 21 YR |  |  | 53 |  |  | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | Abiomed impella device was being inserted. staff obtained an abiomed impella cp set. at the beginning of the procedure, it was noted that the peel away sheath was cracked. it was removed and a new sheath put in place with no difficulties or complications. believed the peel away sheath was cracked prior to use. packaging was intact: only the sheath had a visual defect.
38355 | 1220648-2014-00007 | 4115810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4115810&pc=KFM | 08/19/2014 | 08/19/2014 |  | 09/18/2014 |  |  | 08/19/2014 | 09/18/2014 |  |  | 03/01/2016 |  | 09/18/2014 |  |  |  |  | IMPELLA RECOVER LP2.5 PERC CARDIAC SUPPORT |  |  | KFM |  |  | LP2.5 |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 88772 |  | 0046-3061 |  | Yes | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 58 YR | [5 empty] | Death | Yes | No | [20 empty] | Death | [28 empty] | Initial submission | [7 empty] | NA |  | The complainant reported that a (b)(6) yr old male pt was admitted to the hosp through the emergency room. the pt presented with chest pains and shortness of breath. the cardiac catheterization laboratory (ccl) placed stents in the left main artery and the left anterior descending artery. following the revascularization, the pt started to crash, resulting in the placement of an internal aortic balloon pump (iabp). the pt was then air lifted to this facility where the staff decided to remove the iabp and place an impella lp2.5 pump. the impella was prepped and inserted into the left ventricle. the impella was started and but then stopped several times. the team then began to remove the device in preparation for inspection and to then reinsert into the pt, but the pt expired, prior to completing this process. | [5 empty] | The motor power control (mpc) console data logs were reviewed. the review of the logs revealed that the start of the case was completed normally and the pump was successfully placed in the pt. as was reported, multiple attempts to start the impella were unsuccessful. based upon the data log review, the impella was never able to be started. a physical examination of the returned impella lp2.5 pump revealed a severe kink in the steering catheter just proximal to the catheter-motor junction. in addition, the outflow cage was significantly deformed, resulting in contact between the impeller motor and pump housing this contact prevented the impeller from rotating. during the eval, the outflow cage was straightened which allowed the impeller to rotate freely. the pump then operated normally in a functional bench test. in conclusion, the cause of the reported issue was the result of damage to the cage which most likely occurred during implantation of the pump. there was a shark kink in the steering catheter, just proximal to the motor housing, indicating that the pump saw high forces during insertion. the root cause of this event has been determined to be excessive force exerted by the user during device placement. internal (b)(4).
38356 | 1220648-2014-00006 | 4097977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4097977&pc=KFM | 08/07/2014 | 08/07/2014 |  | 09/05/2014 |  |  | 08/07/2014 | 09/05/2014 |  |  | 06/30/2015 |  | 09/05/2014 |  |  |  |  | IMPELLA CP |  |  | KFM |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  |  |  | 80453 |  | 0048-0003 |  | Yes | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 71 YR | [5 empty] | Death | Yes | No | [20 empty] | Death | [5 empty] | COOK 14FR X 30 CM INTRODUCER, LOT NUMBER 5038939 | [22 empty] | Initial submission | [7 empty] | NA |  | The complainant reported that the physician was treating an emergent (b)(6) year old male pt who was in cardiogenic shock (cs). the pt presented with only a stub for a left main/left anterior descending/circumflex artery (lm/lad/cx). because of peripheral vascular concerns the impella cp was inserted through a 14/x30cm cook sheath, lot number 5038939 (which is not supplied by abiomed. the impella was placed without issue was successfully set to auto flow. three stents were placed in the patient. the first two stents were placed at the lm/lad, with the third stent placed at the cx. during the procedure the clinicians notices leaking/bleeding coming from the cook introducer area during the case, but though it might just be some oozing around the area. following the procedure the pt was weaned off of dobutamine in the cath lab and the clinicians planned to continued impella cp support overnight, but as the ic fellow was suturing up the impella cp repositioning sheath and the cook 14frx30 introducer, the cook device broke cleanly in two, in the location of where it is fused to the blue sheath portion. the pt then arrested, and chest compressions ensued, and were successful. the impella cp and cook introducer were removed from the pt and the vascular surgeon successfully closed the groin. following the event the pt was administered 2 units of blood, and was reported to be in stable condition. the pt was reported to have expired later that evening. the physician was reported that the pt outcome was not as a result of the impella cp and that he was convinced that the impella cp worked properly and that there's not issue whatsoever with the device. | [5 empty] | The impella cp pump was returned for evaluation. the evaluation and inspection did not reveal any abnormalities related to the reported failure. the cook 14fr x 30cm introducer was not available for physical inspection, as it is being retained by the complainant facility. the complainant was able to provide several photos of the separated cook 14fr x 30cm introducer sheath. the photos revealed that the sheath separated from the hub at the joint. the sheath appears slightly flattened or elongated, and one edge of the wall was flipped back over the shaft. it cannot be determined which damage occurred during the procedure and which occurred after the sheath had separated. (b)(4). batch testing was performed on unused cook introducers. the hub-sheath joint consists of three components, the hub (male section), the sheath (flared portion) and the hub (female portion). this joint relies on the mechanical interface between the sheath and the hub. the flared portion of the sheath fits tightly between two halves of the hub. there is no adhesive or other type of fixation. thus, it is possible for the sheath to slide out of the hub without tearing the material. this is the type of failure seen in this case. a simple tensile force was evaluated as a potential model of failure. there was modeled as a force pulling along the sheath. the mean force required for separation was 62 plus/minus 4 n (n=9). by preconditioning the hub with several angled pulls, the tensile strength was reduced by 25%, to 45 plus/minus 5 n (n=5). to test whether these forces could reasonably by applied i the clinical setting, abiomed tested the maximum force that could be applied through the catheter. the maximum force that could be applied without bending the catheter was 25 n. with a pure tensile failure, the joint failed in a relatively symmetrical, clover-like shape. the geometry, and particularly the symmetry of the failure, is not consistent with that seen in the case. abiomed further considered twisting as a mechanism of failure. when the sheath is in the patient, the sheath motion is relatively restricted. qualitative tests found that twisting forces dislodged the hub far more easily than did pulling forces. this creates an asymmetrical failure. this damage is more similar to the failure seen in this case. specifically, it is an asymmetrical failure where one side of the sheath flare bends back over itself. in conclusion, and as stated previously, the sheath split from the hub at its attachment site. this joint relies solely on a friction interface between a flared portion of the sheath and the hub. the tensile strength of this joint was tested and found to be significantly higher than could be applied through the catheter sheath. instead, it is likely that some type of twisting force dislodged the catheter as the staff manipulated the hub into position for suturing.
38357 | 1220648-2014-00005 | 4096347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4096347&pc=KFM | 08/01/2014 | 08/01/2014 |  | 08/29/2014 |  |  | 08/01/2014 | 08/29/2014 |  |  | 02/01/2016 |  | 08/29/2014 |  |  |  |  | IMPELLA CP |  |  | KFM |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  | 88405 |  | 0048-0001 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 56 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [7 empty] | NA |  | The complainant reported that a (b)(6) year old female pt was admitted to the hospital's cath lab. during an intravascular ultrasound, the pt's right coronary artery was dissected, and the wire access was lost, and unable to be re-crossed. the clinicians implanted an impella cp, and the pt was transferred to another facility for an emergent coronary artery bypass graft (cabg). the cabg procedure was performed with the impella cp in place. during the off-pump coronary artery bypass graft procedure (opcabg), significant oozing was seen emanating from the repositioning sheath, which resulted in a large hematoma at the impella cp insertion site. the pt was reported to be "extremely stable" throughout the procedure. the pt was weaned from the impella cp and the pump was explanted in the operating room. the pt remained stable. an open vessel repair at the impella insertion site was performed. during the repair it was noted that the vessel appeared to be slightly torn at insertion point. the clinicians reported that the repositioning sheath was visually examined, but they were unable to verify any rough edges. neither the oscor sheath nor the catheter were able to be examined by the clinicians as they were discarded by the operating room personnel subsequent to the removal of the pump. no replacement blood products were required by the pt. the pt was transferred to the unit in stable condition and the impella support of the pt was reported as successful. | [5 empty] | The impella cp pump, repositioning sheath and oscor 14fr introducer were not returned for eval, as they were discarded subsequent to use. it is unclear if this reported issue resided with the impella cp, or if the issue occurred as a result of the cabg procedure that was performed. the mfr will continue to investigate all reasonably obtainable sources of info and will provide results and updated conclusions in a supplemental report if additional info is received. (b)(4).
38358 | 4173621 | 4173621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=4173621&pc=KFM | 08/03/2014 | 08/26/2014 | 08/26/2014 | 08/26/2014 | 08/26/2014 | 10/15/2014 |  | 08/26/2014 |  |  |  |  | 08/26/2014 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA |  |  | KFM |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | * |  | * |  | Device was returned to manufacturer | [5 empty] | No | Physician |  |  | User Facility report | [8 empty] | 72 YR | [5 empty] | Malfunction | No | Yes | [26 empty] | NO OTHER THERAPIES | [22 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | Surgeon checked function of impella 5.0 device and found it would not function.======================manufacturer response for impella 5.0 device, impella 5.0 (per site reporter).======================rm spoke to manufacturer's rep. he advises it may take as long as 8-10 to have the device analyzed and reported upon.
38359 | 1220648-2014-00004 | 3987000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3987000&pc=KFM | 06/27/2014 | 06/27/2014 |  | 07/24/2014 |  |  | 06/27/2014 | 07/24/2014 |  |  | 02/27/2016 |  | 07/24/2014 |  |  |  |  | IMPELLA CP |  |  | KFM |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  | 88566 |  | 0048-0001 |  | Yes | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 75 YR | [5 empty] | Death | Yes | No | [20 empty] | Death | [28 empty] | Initial submission | [7 empty] | NA |  | The complainant reported that a (b)(6) patient was transferred for further evaluation on (b)(6) 2014 to this facility from another health facility. the patient evaluation conditions included shortness of breath, past medical history of hypertension, severe aortic stenosis, myocardial infarction, and effusion rate (ef) of 10%, a 70% stenosis in the right coronary artery, as well as other conditions. an intra-aortic balloon pump was placed in the patient, but he continued to decompensate, his color was ashen gray upon arrival in the cath lab. in addition, the patient became anuric. it was decided to place the impella cp and it was placed without issue. the iabf was continued with impella, which is not recommended. on the night of (b)(6) 2014, the patient had a respiratory event, with a respiratory rate of 40, during which the patient flailed. an impella position wrong alarm occurred, which was immediately resolved with the repositioning under echo. it is not clear what forces were applied to the device during this incident. a bedside echo revealed an ef on 10%. the left ventricle had extremely poor function. an ekg revealed afib/aflutter and cardioversion failed. the patient was also having intermittent runs of ventricular tachycardia (vt) and pauses. on the morning of (b)(6) 2014, the patient remained impella dependent with the iabp still in when the iabp paused as a result of low native heart activity. the patient had declined to a greater degree overnight. on the night of (b)(6) 2014, after over 53 hours of successful support the pump stopped. the clinicians tried to restart the device numerous times without success. the patient was coding and an iabp was placed emergently. the patient support was reported as successfully, although the patient subsequently expired. | [5 empty] | The device and device data logs were returned for evaluation. the evaluation revealed that the failure of the pump was an imbalance in the device motor coils resistance that developed in use. measurement of the phase resistances confirmed this failure mode. (b)(4).
38360 | 3979083 | 3979083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3979083&pc=PBL | 06/27/2014 | 07/03/2014 | 07/03/2014 | 07/03/2014 | 07/03/2014 | 08/05/2014 |  | 07/03/2014 |  |  |  |  | 07/03/2014 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | PBL |  |  | * |  |  | ABIOMED, INC. |  |  | US |  | 73861 |  | 0048-0003 |  | Device was returned to manufacturer | [6 empty] | Invalid Data |  |  | User Facility report | [8 empty] | 75 YR |  |  | 109 |  |  | Death | Yes | Yes | [20 empty] | Death | [5 empty] | CONNECTED TO THE DEVICE. MODEL ZIMC 0042-0000-US. | CONTROLLER UNITS (IC 2056 OR 2059) EXTERNALLY | DIALYSIS | THERE WAS ONE OF 2 HOSPITAL OWNED ABIOMED IMPELLA | CARDIAC DRUGS | [18 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | The patient had an impella cp implanted via right femoral access for cardiogenic shock (placed 2 days prior). the impella "red alarm" sounded. the nurse adjacent to the patient bedside noted a blood pressure and pulse (sinus tachycardia/atrial fibrillation on monitor). the device console indicated that flow had stopped. the 1- 800 number was phoned from the patient bedside. the levophed drip was increased for a rapidly dropping blood pressure while trouble shooting the device over the phone. norepinephrine was increased and additional nursing and physician support came to the bedside. multiple actions were taken under the phone direction of the impella representative but the device was not restored to proper functioning. the patient went into ventricular tachycardia and cardiopulmonary resuscitation (cpr) initiated while still speaking with the representative. following unsuccessful advanced cardiovascular life support (acls) attempts, the patient was pronounced dead.
38361 | 3951915 | 3951915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3951915&pc=KFM | 06/18/2014 | 06/27/2014 | 06/27/2014 | 06/27/2014 | 06/27/2014 | 07/23/2014 |  | 06/27/2014 |  |  |  |  | 06/27/2014 |  | TEMPORARY CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | KFM |  |  | * |  |  | ABIOMED, INC. |  |  | US |  | * |  | * |  | No | [5 empty] | No | Invalid Data |  |  | User Facility report | [8 empty] | 86 YR | [6 empty] | No | Not Available | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | Patient implanted with impella left ventricular assist device (lvad). patient was transferred to icu following implant. on arrival in the icu, a leak was discovered near the lower end of the catheter connection location. source of the leak was unknown. device rpms lowered to prevent overheating/total failure of lvad. device was removed from patient once clinically indicated.
38362 | 3904336 | 3904336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3904336&pc=KFM | 01/30/2014 | 06/26/2014 | 06/26/2014 | 06/26/2014 | 06/26/2014 | 07/01/2014 |  | 06/26/2014 |  |  |  |  | 06/26/2014 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA |  |  | KFM |  |  | LP2 |  |  | ABIOMED, INC. |  |  | US |  | 77801 |  | * |  | Device was returned to manufacturer | [6 empty] | Physician |  |  | User Facility report | [8 empty] | 63 YR |  |  | 101 |  |  | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | Once device was started it was found to not perform at the full 2.5 lpm; after troubleshooting and speaking with the rep it was determined to have a clot in the device or that the patient was hypovolemic. the patient was given a fluid bolus. this did not correct the issue so the device was removed from the patient and a new one placed, which worked correctly. the device was inspected and the clot was found in the device. the rep stated that the clot most likely was from a clot in the sheath or an unknown lv thrombus.
38363 | 3890007 | 3890007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3890007&pc=KFM | 02/06/2014 | 06/11/2014 | 06/11/2014 | 06/11/2014 | 06/11/2014 | 06/24/2014 |  | 06/11/2014 |  |  |  |  | 06/11/2014 |  | PUMP, BLOOD |  |  | IMPELLA |  |  | KFM |  |  | * |  |  | ABIOMED, INC |  |  | US |  | * |  | * |  | Device was returned to manufacturer | [6 empty] | Invalid Data |  |  | User Facility report | [8 empty] | 50 YR | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | Patient was in ccu and had an impella (abiomed) device placed for circulatory support. representative for the abiomed device company was present to assist during implantation and management. after device removed from patient the representative expressed concern the device may not have been functioning properly and took the device used for this patient to be tested.the rep independently decided to take the device back to his company.
38364 | 1220648-2014-00002 | 3850522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3850522&pc=KFM | 04/20/2014 | 04/21/2014 |  | 05/21/2014 |  |  |  | 05/21/2014 |  |  | 09/30/2014 |  | 05/21/2014 |  |  |  |  | IMPELLA LP 5.0 PUMP |  |  | KFM |  |  | LP 5.0 |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  |  |  | 2013163199 |  | 4996 |  | Yes | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Other | Foreign | Health Professional | User facility | Company representation |  | 68 YR | [5 empty] | Injury | Yes | Yes | [20 empty] | Required Intervention | [28 empty] | Initial submission | [7 empty] | NA |  | The complaint reported that an obese (b)(6) male patient had an st-segment elevation myocardial infarction (stemi) on (b)(6) 2014, the patient was transferred from a tertiary hosp where the circumflex artery/right coronary artery (cx/rca) was stented. it was determined that the patient's only good flow was coming from the circulation from the intra-aortic balloon pump (iabp) that had been placed on (b)(6) 2014, the patient remained unstable with frequent runs of ventricular tachycardia (v-tach.) consequently, the patient was brought to the operating room on (b)(6) 2014 where an impella lp5.0 was successfully placed. during placement the femoral artery was cut down with a 10 mm graft and sponge plug for hemostasis. there were no vessel loops used, but umbilical tape to hold vessel and maintain hemostasis while sewing the graft was employed. the impella was back-loaded over 0.025 abiomed wire and the impella lp5.0 pump inserted at an approx 45 degree angle. following the successful placement of the pump, issues with pump positioning occurred. the heart was ejecting the pump and physician had to manipulate and reposition the pump multiple times. the pigtail was not pointed towards the apex which led them to torque the catheter to set it in the correct position. while in the operating room the patient had many runs of v-tach requiring both cpr and defibrillation. on (b)(6) 2014 the impella was reported to be working well, with the patient's condition much improved. it was planned to continue impella support for 48 hours. the clinicians were able to wean the patient from inotropes, resulting in the incidents of v-tach and arrythmias ending. on (b)(6) 2014 the clinicians determined that the impella was working very well, with no alarms emitted. the patient was then extubated. an echo was then performed which illustrated that the left ventricle was contracting well, with the patient's condition continuing to improve, and with still no more further episodes of arrhythmias. another echo was performed on (b)(6) 2014. the patient's ef was found to be normal. the impella was then successfully weaned from the patient after 4 days, the case at this point was deemed as a success. it was reported that the image of the device showed it intact as it was removed from the ventricle. as the explanting physician first started to remove the impella no resistance was felt, but as the catheter was being removed resistance began to be felt. the physician increased the arteriotomy approx 5 cm to get a better look at what was causing the resistance. once the arteriotomy was enlarged, plaque was noticed on one side of the artery. stenosis in the artery was also noted proximal to the arteriotomy. the physician continued to remove the device by using his left hand to pull the catheter and his right to gently "milk" the catheter through the resistance. he did not use excessive force and used the "milking" technique to assist device out of the artery. once the pump was removed, the perfusionist informed the surgeon that the cannula was missing. a fluoroscopy was performed which showed that the cannula was located in the thoracic aorta. contrast agent, in conjunction with an angio balloon were employed to remove the cannula. the balloon was inserted through the patient's femoral artery, then inflated above the cannula and successfully pulled down and out of the thoracic aorta with the balloon. the patient required re-intubation due to becoming "somewhat unstable", during the removal procedure. following the removal of the cannula the patient was reported have been extubated and "doing fine." | [5 empty] | The device is currently undergoing evaluation. upon the conclusion of the evaluation, a supplemental medwatch report will be submitted. (b)(4).
38365 | 1220648-2014-00003 | 3819965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3819965&pc=KFM | 04/13/2014 | 04/13/2014 |  | 05/13/2014 |  |  | 04/13/2014 | 05/13/2014 |  |  | 08/04/2014 |  | 05/13/2014 |  |  |  |  | IMPELLA CP |  |  | KFM |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  | 087738 |  | 0048-3018 |  | Yes | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 54 YR |  |  | 65 |  |  | Injury | Yes | Yes | [20 empty] | Required Intervention | [5 empty] | DERMACARRIER II SKIN GRAFT CARRIER, 3:1 (B)(4) | [22 empty] | Initial submission | [7 empty] | NA |  | The complainant reported that a (b)(6) female pt was admitted to the facility's cath lab with an st-segment elevation myocardial infarction (stemi). the catheterization revealed an acute occlusion of the circumflex artery (x) and a 4 plus mitral regurgitation (4+mr). the cx was stented and on (b)(6) 2013 an impella cp was placed with the aid of an oscor sheath. there were no insertion problems reported, and the pt's groin was reported as looking fine after implant. the following day the insertions site was found to be oozing. in the afternoon of (b)(6) 2014 the pt became agitated and she pulled the catheter and the repositioning sheath back by about 3cm. this led to a "positioning wrong" alarm, as well as multiple suction alarms, and steady bleeding, which was caused by the repositioning sheath being backed out. the suction alarms were attributed to the consistent bleeding from the impella's insertion site. the repositioning sheath was pushed back in 3cm; the insertion site was reported as looking better, but there was still bleeding, though there was no continued oozing. the pt's hemoglobin was found to have dropped from a 10.5 reading on (b)(6) 2014 to 7.2 on (b)(6) 2014. the clinicians discontinued administering systematic heparin and added a sandbag for pressure, but the bleeding continued. the pt underwent transfusion of 2 litres of blood. the clinicians decided to wean the pt and explant the device early because of the continued bleeding and because the pt could not be administered heparin. the impella cp was explanted from the pt on (b)(6) 2014 after >44 hrs of pt support. upon device removal there was no initial bleeding, and no need for pressure to be applied. pedal pulses were lost temporarily. a thrombus was observed on the impella pigtail following explantation, but the pulses were regained shortly thereafter, without any action being required. hemostasis was reported to have been achieved. an echo was performed which revealed moderate mr with an ejection fraction of 40% (from a low of 25% upon admittance). the pt support was reported as successful, and the pt reported to be in stable condition. | [5 empty] | The device is currently undergoing eval. upon the conclusion of the eval a supplemental medwatch report will be submitted. (b)(4).
38366 | MW5035751 | 3770902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3770902&pc=KFM | 04/01/2014 | 04/18/2014 |  | 04/21/2014 |  |  |  | 04/21/2014 |  |  | 08/31/2015 |  | 04/21/2014 |  | IMPELLA |  |  | IMPELLA |  |  | KFM |  |  | 005062 |  |  | ABIOMED |  |  | US | [11 empty] | Y | Health Professional |  |  | Voluntary report | [8 empty] | 67 YR |  |  | 94 |  |  | Malfunction | No | Yes | [49 empty] | Initial submission | [9 empty] | Device broke with insertion.
38367 | 1220648-2014-00001 | 3701213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3701213&pc=KFM | 01/20/2014 | 01/20/2014 |  | 02/19/2014 |  |  | 01/20/2014 | 02/19/2014 |  |  | 12/01/2015 |  | 02/19/2014 |  |  |  |  | IMPELLA CP |  |  | KFM |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  | 2015194165 |  | 0048-0003 |  | Yes | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 60 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [7 empty] | NA |  | The complainant reported that the physician was treating a (b)(6) year old male patient in cardiogenic shock, with a previously stent to the left anterior descending artery. the patient presented in the catheter lab with a ventricular tachycardia heart rhythm, and then arrested and decompensated. a temporary pacer was inserted in the patient, and it was decided to place an impella cp. the impella was placed using the 14fr oscor sheath. the sheath placement was reported to have been accomplished without issue, and the impella was placed with ease. the physician indicated that the patient had calmed considerably, and after about 20 minutes, the physician proceeded with the removal of the peel-away sheath. the peel-way sheath was removed without incident and the repositioning sheath was flushed and advanced toward femoral artery. the physician then felt a slight resistance, and in response he pulled the repositioning sheath back and then noticed that the sheath had an accordion like appearance from buckling. the doctor then attempted to smooth out the tip, but it then separated and split at the tip. the impella cp was then removed from the patient. the patient was receiving blood thinners during this time; consequently, the bleeding was unable to be controlled, so the impella was then removed from the patient. the vascular surgeon then successfully repaired the femoral site, and the patient was administered 4 units of blood because of low hemoglobin due to bleeding at the insertion site. the patient was then treated with inotropic drugs, and is reported to be doing well and in stable condition. | [5 empty] | The impella cp and the impella data logs were returned for evaluation. a review of the data logs from the controller revealed that the impella cp pump was in the patient 1 hour, and was run at performance level p8 for 35 minutes prior to being removed from the patient. the returned pump plug, catheter, pump motor, pump, cannula, pump housing, and inflow cage were examined. the inflow cage and pump housing appeared undamaged. no defects were found on the catheter. the repositioning unit was examined and several small stress cracks were found on the tip of the wound closure along with significant tears. no severe cracks were noted on the proximal end (near the hub). visual inspection of the repositioning unit revealed fractures and tears on the tip of wound closure and a small slit on the tip of the repositioning sheath lifting the edge off the catheter. the catheter was cut post electrical inspection and pump run tests. the repositioning unit was removed and the wound closure and repositioning sheath were separated. an examination of the returned repositioning unit demonstrated several longitudinal tears and cracks along the wound closure sheath and a small slit and circumferential deformations in the repositioning sheath. it was noted that the most significant tears on the wound closure aligned with the current mold flash removal. normal processing of this component includes mold flash removal left behind from the injection molding process by using a scalpel to carefully scrape off the flash. this potentially leaves behind visible marks and small skives. the component is molded from a low-density (b)(4). the imperfections generated by the flash removal process, as seen on the wound closure device returned, was able to be simulated in our lab. the appearance of the test part from the stimulation was consistent with the stress cracks and tears noted in this event. the probable root cause of the stress cracks and tears to the repositioning unit was most likely a combination of material brittleness and the flash removal process-induced defects which made the wound closure more susceptible to stress cracks and tears by insertion forces. a correction action has been initiated to address this issue under document number (b)(4).
38368 | MW5034565 | 3639359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3639359&pc=KFM | 02/09/2014 | 02/14/2014 |  | 02/18/2014 |  |  |  | 02/18/2014 |  |  |  |  | 02/18/2014 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | KFM |  |  | 0042-0020 |  |  | ABIOMED |  |  | US | [11 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 40 YR |  |  | 77 |  |  | Death | Yes | Yes | [20 empty] | Death | Life Threatening | Required Intervention |  |  |  | PLACEMENT OF IMPELLA CP | PLACEMENT OF IABP | MANUAL THROMBECTOMY WITH FETCH CATHETER MID, | MECHANICAL THROMBECTOMY WITH ANGIOJET OF LCX OM3 | PCI (BARE METAL STENT) OF MID LAD AND LCX OM3 | RIGHT HEART CATHETER | REMOVAL OF TEMP PACER | PLACEMENT OF TEMP PACER | REMOVAL OF IABP | LAD/ DIAG/ LCX | [13 empty] | Initial submission | [9 empty] | Abiomed impella device failed; pump stopped and patient arrested resulting in death on (b)(6) 2014.
38369 | MW5032584 | 3455685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3455685&pc=PBL | 10/26/2013 | 11/01/2013 |  | 11/04/2013 |  |  |  | 11/04/2013 |  |  | 07/31/2015 |  | 11/04/2013 |  | HEART PUMP |  |  | IMPELLA CP |  |  | PBL | [5 empty] | ABIOMED, INC. |  |  | US | [11 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 86 YR | [5 empty] | Malfunction | No | Yes | [20 empty] | Other | [5 empty] | DOBUTAMINE | DOPAMINE | [21 empty] | Initial submission | [9 empty] | L sided impella failed. the engineer was called, however, there was no indication of machine failure. the machine consol was changed out which also did not function.
38370 | 1220648-2013-00013 | 3466792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3466792&pc=KFM | 09/24/2013 | 10/04/2013 |  | 11/01/2013 |  |  | 10/04/2013 | 11/01/2013 |  |  | 06/01/2015 |  | 11/01/2013 |  |  |  |  | IMPELLA RECOVER LP2.5 PERC CARDIAC SUPPORT |  |  | KFM |  |  | LP2.5 |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | 2014185491 |  | 005042 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | UNK | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [7 empty] | NA |  | The complainant reported that the pt (age and gender unknown) presented with a myocardial infarction with a dissecting left anterior descending artery. on (b)(6) 2013, at presbyterian hospital of plano an impella lp2.5 pump was placed in the pt, along with 3 stents. after the impella was placed, the physician reported that the pt's st segment got a lot better. the physician was then able to place two additional stents in the pt. the doctor was able to stop the spiral dissection and stabilize the pt. the pt's hemodynamics were reported as stable. the pt was then transported to the complainant's hospital, the heart hospital baylor-plano. approximately 22 hours after implant the pt was reported as doing very well, and was being weaned from the impella in preparation to explant the following day. the pt was off all pressors, and exhibited a blood pressure of 127/85. the physician and staff were reported to be pleased with the impella support provided to the pt. after 44 hours of successful impella support the device was explanted from the pt. after device removal an echo was performed, which revealed an abnormality with the pt's mitral valve. the surgeon reported that the papillary muscle looked good and that there was a torn cordae. a robotic mitral repair was performed. the pt is reported to be doing well, and the staff feel that impella support saved the pt life. the surgeon speculated that the initial implant (done at another hospital) was placed into the mitral apparatus, causing the torn cordae. | [5 empty] | The impella lp2.5 pump was not returned for evaluation, as it was discarded subsequent to use. in addition, the aic data logs have not been returned. the manufacturer will continue to investigate all reasonably obtainable sources of information and will provide results and updated conclusions in a supplemental report if additional information and/or the aic data logs have been received. (b)(4).
38371 | 1220648-2013-00014 | 3444946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3444946&pc=KFM | 09/25/2013 | 09/27/2013 |  | 10/24/2013 |  |  | 09/27/2013 | 10/24/2013 |  |  | 04/01/2015 |  | 10/24/2013 |  |  |  |  | IMPELLA CP |  |  | KFM |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  | 2014181841 |  | 0048-0003 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 73 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [7 empty] | NA |  | The complainant reported that (b)(6) old female patient with chest pain was being treated in the hospital's icu. the patient went into ventricular fibrillation, and was then cardioverted once into a normal sinus rhythm. the patient was then taken to the ccl, where it was revealed that the patient had lm, lad, cx stenosis with rca gone. the patient also had an injection fraction of 20%. the patient and family requested a pci be attempted. the physician implanted the impella cp prophylactically, and reported that the insertion was smooth and the impella support was "phenomenal" throughout the case. during the impella support the lad into the lm was stented, the cx was ballooned and left alone as it closed off. the clinicians transitioned to the repositioning sheath after fully removing oscor from patient and peeling it away, as ot was planned to leave the impella in for 1 day. after 5 minutes, a hematoma was noticed. the heparin was reversed, but the hematoma was growing and the patient was bleeding from the mouth. the act's were noted to be at 1.5 to 2 times the recommended levels. the surgeon noted damage to the artery, and surgical wound closure was performed. the source of the bleeding from patient's gum was not determined, but conjectured to be the result of the patient potentially being a "gum bleeder". no gum treatment was performed. no replacement blood products were needed for the patient. once the surgical wound was closed, the bleeding was stopped. the physician reported that at the start of the case it was not fully expected that the patient would survive the procedure, but he was very pleased with the support. he further stated that it was an overall positive experience, and patient survived a very difficult procedure. | [5 empty] | Neither the 0.018 guidewire or impella pump were returned for evaluation. the complainant reported that the guidewire was discarded subsequent to use. the root cause of this event of a hematoma at the insertion site requiring surgical intervention and removal of the impella cp, is unable to be determined without the impella pump or introducer. (b)(4).
38372 | 1220648-2013-00012 | 3402809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3402809&pc=KFM | 09/01/2013 | 09/03/2013 |  | 10/03/2013 |  |  | 09/03/2013 | 10/03/2013 |  |  | 03/01/2015 |  | 10/03/2013 |  |  |  |  | IMPELLA CP |  |  | KFM |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  | 2014179838 |  | 0048-0003 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 68 YR | [5 empty] | Injury | Yes | Yes | [20 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | The complainant reported that a (b)(6) male patient underwent the emergent implant of an impella cp on (b)(6) 2013. the physician reported that they had used the facility's regular sheath to dilate up to 9fr after using a micro-puncture technique to gain access; they then used abiomed dilators. the femoral artery was reported to be smooth without calcium or plaque, and there was no subcutaneous scar. the catheter was "advanced nicely without any difficulty." the impella was placed and the sheath was removed; the repositioning sheath was advanced easily without any resistance. the physician reported that initially the groin looked fine and that approx 30 minutes had passed while the clinicians instituted the cooling venous sheath. the pt was the transported to the icu. after the patient was placed into the icu bed, the physician received a call that the patient had developed a large hematoma and that there was pulsatile bleeding from the patient's groin. the cath lab clinicians responded to the physician's inquiry that there had been no inadvertent movement of the device during transport and that it had not been moved or touched. the clinicians tried to hold pressure initially, but when they saw that there was no decrease in the pulsatile bleeding, they removed entire impella cp and held pressure for about an hour until hemostasis had been achieved. upon removal of the device several tears were noted on the repositioning sheath. an hour later the patient began bleeding again, and pressure was again held for another hour. the patient was administered 3 units of blood due to the hematoma and the blood loss around the sheath. the patient was in stable condition and was free of bleeding thence forth. revascularization of the patient successfully completed. in a subsequent conversation between abiomed and the physician, she was asked what may have contributed to the event, she responded that the impella had been maneuvered several times in an attempt to obtain better flow after they had placed the repositioning sheath, but that was 30 minutes prior to the bleeding, and that "everything had looked great." the physician reported that the pt was fine despite the transfusion and that he had walked out of the hospital and was doing well. | [5 empty] | The device is currently undergoing eval. upon conclusion of the eval a supplemental medwatch report will be submitted. (b)(4).
38373 | 1220648-2013-00011 | 3337658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3337658&pc=KFM | 07/31/2013 | 07/31/2013 |  | 09/03/2013 |  |  | 07/31/2013 | 09/03/2013 |  |  |  |  | 09/03/2013 |  |  |  |  | IMPELLA RECOVER LP2.5 PERC CARDIAC SUPPORT |  |  | KFM |  |  | LP2.5 |  |  | ABIOMED EUROPE, GMBH (GERMANY) |  |  | GM |  | UNK |  | UNK |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Other | Foreign | Company representation | Distributor |  |  | 65 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [7 empty] | NA |  | The abiomed distributor reported that a (b)(6) year old male patient was admitted to their customer's facility with a myocardial infarction. an angiogram was performed, which revealed showing triple vessel disease, including left main with a lesion of 50%, and an ejection fraction 30%. the physician performed a percutaneous coronary intervention (pci). the impella lp2.5 was placed prior to the pci being performed. the impella was placed without complication, and was then removed after almost 24 hours of successfully patient support. it was reported that a few hours later the patient developed severe mitral regurgitation in the tee and had a papillary muscle rupture which lead to mitral valve prolapse. | [5 empty] | The impella lp2.5 pump was not returned for evaluation, as it was discarded subsequent to use. the complainant was unable to give the serial number of the device at issue in this complaint; therefore, the catalog number, log number, manufacture and expiration dates have been unable to be documented in this report. in addition, numerous attempts have been made to obtain additional patient and event information. all attempts have been unsuccessful. an event of this nature, this would be an extremely rare occurrence based on our experience in the united states of greater than 15,000 cases, with an additional 5,000 cases outside of the us. we are aware of one other potential case in which a chordae was found entangled in a pump after explant, but this patient did not experience any untoward issues. there are no moving parts that could possibly come into contact with the mitral apparatus, and insertion of the impella into a subannular position is to be strictly avoided. a mechanical complication is likely to be simply a result of entanglement of the device in the subannular structures. the instruction for use for the impella 2.5 pump contains the following: for optimal positioning of the impella 2.5 catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve. if the impella 2.5 is correctly positioned, echocardiography will likely show the following: catheter in inlet area 3.5 cm below the aortic valve. catheter outlet area well above the aortic valve (frequently not visible on tee or tte images). catheter angled toward the left ventricular apex away from the heart wall and not curled up or blocking the mitral valve. the manufacturer will continue to investigate all reasonably obtainable sources of information and will provide results and updated conclusions in a supplemental report when completed. (b)(4).
38374 | 1220648-2013-00010 | 3337666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3337666&pc=KFM | 07/28/2013 | 07/31/2013 |  | 09/03/2013 |  |  | 07/31/2013 | 09/03/2013 |  |  | 03/01/2015 |  | 09/03/2013 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERC CARDIAC SUPPORT |  |  | KFM |  |  | LP2.5 |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  |  |  | 76367 |  | 0046-3061 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 20 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | Disability | [27 empty] | Initial submission | [7 empty] | NA |  | The complainant reported that on (b)(6) 2013, a female patient was admitted to the facility's emergency room with an altered mental state, in ventricular fibrillation arrest, without a pulse and intubated. the patient was diagnosed with viral myocarditis. the patient was very labile, and her blood pressure was 60/38. the patient was administered multiple high doses of pressers and inotropes in an effort to maintain the blood pressure. an impella lp2.5 was implanted in through the patient's right femoral artery, without the aid of a guide-wire. the 13fr peel-away sheath was left in place. the physician reported that during device placement the clinicians shot the femoral to see flow, administered full drugs and the device was clamped down to maintain pressure. there were no issues the night of implant with the patient's groin or limb. the following day, on (b)(6) 2013, there were multiple "pump position unknown" and "pump position wrong" alarms displayed on the controller. the issues were resolved, by repositioning the catheter. the patient was reported to have shown signs of improvement. on (b)(6) 2013, after 23.03 hours of successful support, the impella was explanted without issue. the following day, on (b)(6) 2013, an issue developed with the patient's right limb that necessitated amputation. | [5 empty] | The impella lp2.5 pump was not returned for evaluation, as it was discarded subsequent to use. there was no indication from the complainant of any malfunction of the device. the physician reported that there were no issues during and after the placement of the device, other than the need to reposition the catheter. the physician reported that the impella lp2.5 support was responsible for keeping this patient alive, and if not for the device this patient would have expired. the device removal label for the provided pouched tyvek-sealed tray that contains the 13fr introducer kit warns the user of the following: attention: the 13f peel-away introducer has to be removed after pump insertion or prior to transportation to transportation to icu/ccu. the manufacturer will continue to investigate all reasonably obtainable sources of information and will provide results and updated conclusions in a supplemental report when completed. (b)(4).
38375 | 1220648-2013-00009 | 3322300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3322300&pc=KFM | 07/23/2013 | 07/24/2013 |  | 08/23/2013 |  |  | 07/24/2013 | 08/23/2013 |  |  | 03/01/2015 |  | 08/23/2013 |  |  |  |  | IMPELLA RECOVER LP 2.5, .018 GUIDEWIRE |  |  | KFM |  |  | .018 IN (0.47MM) |  |  | ABIOMED |  |  | GM |  | 0076367 |  | 0046-3061 | [7 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 53 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [7 empty] | NA |  | The complainant reported that on (b)(6) 2013, the physician placed an impella lp2.5 pump in a (b)(6) year old male patient who had undergone revascularization, and was post lad angioplasty and stent stemi. during device placement of the pump the physician advanced the 0.018" guide-wire into the patient's left ventricle without the aid of fluoroscopic control. during this procedure the physician encountered resistance when pushing the catheter through the sheath. after working through the resistance, a fluoroscopy was performed and revealed the wire through the pericardium. the wire was immediately pulled back and continued to advance the device. after placing the device and the patient was supported two suction alarms occurred. the catheter was adjusted twice, but initially the pressure was in the 120's and then dropped to the 60's followed by the 40's. an echo was performed, which revealed that the left ventricle had been perforated by the 0.018 guide wire. a pericardial window was then performed in the cath lab. the patient was then taken to the operating room where the clinicians found the site and placed a clip to the perforated site of the ventricle. following surgery the patient was in stable condition. | [5 empty] | All impella disposables, including the impella lp2.5 pump and the .018" guide-wire, associated with this event were discarded subsequent to use. abiomed has also not yet received the requested console data log from this event. a definitive root cause of the failure cannot be determined. it appears the guide-wire may have moved with the impella when the physician was having difficulty getting the pump through the sheath, causing it to advance deeply into the ventricle and perforated the ventricle wall. although there is no evidence of a device defect in this case, the product was not returned; consequently, the condition of the impella or the guide-wire cannot be verified. the manufacturer will continue to investigate all reasonably obtainable sources of information and will provide results and updated conclusions in a supplemental report when completed. (b)(4).
38376 | 1220648-2013-00008 | 3261129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3261129&pc=KFM | 07/02/2013 | 07/02/2013 |  | 07/29/2013 |  |  | 07/02/2013 | 07/29/2013 |  |  | 02/01/2015 |  | 07/29/2013 |  |  |  |  | IMPELLA CP |  |  | KFM |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  | 2014173589 |  | 0048-0003 |  | Yes | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 37 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | The complainant reported that the clinician was placing an impella cp in (b)(6) male patient on (b)(6) 2013. the patient's hemoglobin level prior to the procedure was 14.9. the physician reported that during the implant procedure and after removing the peel-away sheath, copious patient bleeding was observed. the patient was reported to have lost approximately 1 liter of blood, and required 2 units of packed red blood cells (prbc) to replace the blood lost. in addition, a vascular surgeon provided vascular intervention to stem the flow of blood via a stent to the right femoral artery. the stent resolved the femoral arterial bleeding. subsequent to the event additional information was received from the clinicians assisting in the procedure. they indicated that a deep incision with the scalpel was made at the access site. the assisting clinicians also reported that the tip of the peel-away sheath was still in the insertion site as it was peeled away, resulting in "significant bleeding." the bleeding was successfully controlled with manual pressure. additional information was received from the clinicians assisting in the procedure indicated that a deep incision with the scalpel at the access site. the assisting clinicians also reported that the tip of the peel-away sheath still in the insertion site as it was peeled away, resulting in "significant bleeding." the bleeding was controlled with manual pressure. the patient was reported to have been successfully supported for 113 minutes with the impella cp, and was reported to be in stable condition following this event and impella cp support. | [5 empty] | Only the impella cp pump was returned for evaluation. the peel-way sheath was not returned. the pump's catheter lumen and wound closure sheath were visually inspected for any kinks or defects. no damage was observed. the wound closure sheath diameter was found to be within specification in all random locations along it's shaft where measurements were taken. the pump serial number was verified. motor resistances were measured and all were found to be within specification. the visual inspection revealed no defects on the impella cp pump. the impella cp was tested and operated as intended. in conclusion, the root cause of the reported patient bleeding after the removal of the peel away sheath could not be determined, as the peel away sheath was not returned for evaluation; however as previously noted in this report, the clinical staff at the complainant facility indicated the tip of the peel away sheath was still in insertion site as it was peeled away. it is believed that user error contributed to this event, but because the peel away repositioning sheath was not returned for evaluation, this cannot be definitively determined. a review of the lot history records for this device did not reveal any anomalies. corrective action: no corrective action is required at this time as the cause of this event couldn't be confirmed.
38377 | MW5030411 | 3149089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3149089&pc=KFM | 05/17/2013 | 05/30/2013 |  | 05/30/2013 |  |  |  | 05/30/2013 |  |  |  |  | 05/30/2013 |  | IABP |  |  | IMPELLA PUMP |  |  | KFM | [5 empty] | BIOMED, INC |  |  | US | [5 empty] | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 58 YR |  |  | 76 |  |  | Injury | No | Yes | [20 empty] | Required Intervention | [28 empty] | Initial submission | [7 empty] | 14 FR 65CM IMPELLA |  | The impella ceased to function and the patient was noted to have an increasing pressor requirement.
38378 | MW5030419 | 3149112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3149112&pc=KFM | 05/17/2013 | 05/30/2013 |  | 05/30/2013 |  |  |  | 05/30/2013 |  |  |  |  | 05/30/2013 |  | IABP |  |  | IMPELLA PUMP 14FR 65CM |  |  | KFM | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | [8 empty] | 58 YR |  |  | 76 |  |  | Injury | No | Yes | [20 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | The impella ceased to function and the pt was noted to have an increasing pressor requirement.
38379 | 1220648-2013-00006 | 3121957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3121957&pc=KFM | 04/15/2013 | 04/17/2013 |  | 05/15/2013 |  |  | 04/17/2013 | 05/15/2013 |  |  | 01/01/2016 |  | 05/15/2013 |  |  |  |  | IMPELLA RECOVER LP 2.5, .018 GUIDEWIRE |  |  | KFM |  |  | .018 IN (0.47MM) |  |  | ABIOMED |  |  | GM |  | 10220028 |  | 0046-0043 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 68 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [7 empty] | NA |  | The complainant reported that the clinicians were performing high-risk percutaneous coronary intervention on a (b)(6) male patient. during the preparation to implant an impella lp2.5 pump the clinician employed a perclose x2 wired insertion, but the 0.018 guidewire looped around the impella going up, and then kinked. the clinicians then found that they were unable to either advance or remove the impella. a vascular surgeon then performed a femoral cut down and successfully removed the impella, guidewire and sheath. the vessel was then repaired and the impella was placed. the right iliac dissected, so the physicians had to go back into the left brachial for intervention. two stents were placed in the patient's circumflex artery, and one stent was placed in the right coronary artery. the patient was successfully supported with the impella lp2.5 pump for 3.75 hours. the impella was then removed from the patient per protocol and without issue. the patient's groin was repaired, and the patient was released from the hospital the following day. | [5 empty] | The complainant reported that there were no further adverse effects to the patient as a result of the reported problem. the model # of the pump: impella recover lp2.5, catalog #: 0046-3061, serial #: (b)(4), batch #: 0075373-16. neither the 0.018 guidewire or impella pump were returned for evaluation. the complainant reported that the guidewire was discarded subsequent to use. th mfr will continue to investigate all reasonably obtainable sources of information and will provide results and updated conclusions in a supplemental report when received. (b)(4).
38380 | 3030139 | 3030139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3030139&pc=DSP | 12/30/2012 | 03/25/2013 | 03/25/2013 | 03/25/2013 | 03/25/2013 | 04/02/2013 |  | 03/25/2013 |  |  |  |  | 03/25/2013 |  | CATHETER, CARDIAC ASSIST PUMP |  |  | IMPELLA 2.5 |  |  | DSP |  |  | * |  |  | ABIOMED |  |  | US |  | 0042-0020 |  | * |  | Yes | [5 empty] | No | Physician |  |  | User Facility report | [8 empty] | 81 YR |  |  | 69 |  |  | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | Prepared impella 2.5 for use during coronary intervention. while attempting to initiate support, the impella console working screen froze. abiomed help line contacted and suggested that "the green button be depressed". screen remained frozen. device removed and procedure completed without use of the device. no adverse pt outcome.what was the original intended procedure?attempted placement of an impella left ventricular assist catheter.device #1is this a laboratory device or laboratory test?no
38381 | 1220648-2013-00005 | 3023790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3023790&pc=KFM | 02/21/2013 | 02/21/2013 |  | 03/21/2013 |  |  | 02/21/2013 | 03/21/2013 |  |  | 12/01/2014 |  | 03/21/2013 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERC CARDIAC SUPPORT |  |  | KFM |  |  | LP2.5 |  |  | ABIOMED |  |  | GM |  |  |  | 0046-3061 |  | Yes | [5 empty] | Y | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 26 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [7 empty] | NA |  | The complainant reported that a (b)(6) female pt with down's syndrome was admitted to the hospital on (b)(6) 2013, with respiratory failure requiring intubation. an echo was performed which indicated that the pt had severe mitral regurgitation. an impella lp2.5 was implanted in the pt on (b)(6) 2013, as a bridge to surgery (mitral valve repair/replacement). two hours post impella placement, the pt's fio2 requirements were down to 60% and the pt was maintaining spo2 of 94%. however, her urinary color had gone from clear yellow to gross bleeding. the device position confirmed. the catheter was pulled back 2 cm, but there was no change in the pt's urine after 30 minutes. after an additional 30 minutes post impella placement, the pt's urine remained bloody, although the device was in good position. the device speed was decreased to p-4, flows to 1.7. the pt remained on support as of (b)(6) 2013, despite continued hemolysis. the device improved the pt's oxygenation significantly, but the pt was now anuric with worsening azotemia and acute renal failure. the pt's bun and cr levels began climbing. the pt was treated with ivf and multiple blood transfusions. daily dialysis was begun. on (b)(6) 2013, the pt was taken to the operating room for mitral valve repair/replacement. following surgery, the device was removed from the pt. the pt was still intubated. | [5 empty] | As of the last report, the pt remained in critical care, intubated, trached and requiring continued dialysis. (b)(4).
38382 | 1220648-2013-00004 | 3023811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=3023811&pc=KFM | 10/12/2012 | 03/05/2013 |  | 03/21/2013 |  |  | 03/05/2013 | 03/21/2013 |  |  | 08/01/2013 |  | 03/21/2013 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERC CARDIAC SUPPORT |  |  | KFM |  |  | LP2.5 |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  | 069151 |  | 0046-3061 |  | Yes | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 48 YR | [5 empty] | Injury | Yes | Yes | [20 empty] | Required Intervention | [28 empty] | Initial submission | [7 empty] | NA |  | The complainant reported that the clinicians were treating (b)(6) male pt who had presented on (b)(6) 2013 in the emergency room with shortness of breath and troponins. the cardiac catheter revealed severe 3x vessel disease. an intra-aortic balloon (iabp) was placed from the right groin and an impella lp 2.5 pump was percutaneously placed. the pt was then prepped for open heart surgery, during which "a few issues as they raised his leg during the prep" occurred. on the night of (b)(6) 2013, the pt had difficulties with hypercalcemia, which led to bundle branch block and hypovolemia and his cardiac index and output bottomed out. the surgeons performed direct manual heart massage. the impella was his only source of cardiac output. the pt improved significantly. his index was now 3.3 and his output was in the mid 6 l/m and was doing well. the pt's thrombolysis also cleared up. on (b)(6) 2013, the pt's heart showed signs of improvement. the iabp was removed, but the impella could not be removed, due to the pt's full dependency on impella to sustain life. a cardiopulmonary bypass also had to be employed during surgery. the pt went to the operating room (or) where his chest was "cleaned out and closed with sires." while in the operating room, the anesthesiologist checked impella placement with transesophageal (tee). placement was good. impella position was also visualized with chest x-ray, and placement had not changed. post-procedure, the pt was still fully dependent on impella and compromise of the left leg circulation was noted. the left leg ischemia progressed, but the physician reported that despite the ischemia, the pt was dependent on the impella lp 2.5 support, and would have died if it had been removed. the decision was made to continue the impella support. the pt subsequently had to have his left leg amputated on (b)(6) 2012. | [5 empty] | The device is currently undergoing eval. upon the conclusion of the eval, a supplemental medwatch report will be submitted. (b)(4).
38383 | 1220648-2013-00002 | 2994343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2994343&pc=KFM | 02/06/2013 | 02/07/2013 |  | 03/04/2013 |  |  | 02/07/2013 | 03/04/2013 |  |  | 09/01/2014 |  | 03/04/2013 |  |  |  |  | IMPELLA CP |  |  | KFM |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  | 0070559 |  | 0048-0003 |  | No | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 54 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | The complainant reported that the physician was treating (b)(6) male patient with severe mistral regurgitation and pulmonary hypertension. the physician placed an impella cp pump via the patient's left femoral artery. the 14 french oscor sheath was placed without difficulty. a diagnostic catheter with the air of an 0.035 wire was then placed across the aortic valve and advancement was stopped. the 0.035 wire was then exchanged with the abiomed. 0.018 wire and the diagnostic catheter was removed, and the 0.018 wire placement was confirmed via fluoroscopy to be in the ventricle. the impella cp catheter was then advanced without difficulty over-the-wire across the aortic valve to the lft intra-ventricular space. under fluoroscopy the catheter appeared to be correctly positioned. the left ventricle was found to be de-compensated with minimal intra-ventricular space and filling. during the removal of the abiomed 0.018 wire, the catheter was withdrawn with the wire, with minimal resistance, but was unknowingly brought down to the level of the left external iliac artery. upon visualization under fluoroscopy the catheter was found to be folded back upon itself, with the wire anchored in the catheter. the physician determined that the artery required cut-down for removal of the impella cp pump and wire. the artery was repaired by placing a graft from what appeared the external iliac to the common femoral. the placement of an impella ld5.0 was discussed, but due to the patient's peripheral artery disease, the placement of an impella ld5.0 was contraindicated. in addition, a short arch prevented impella ld 5.0 placement. as a result of the patient's elevated pulmonary pressures and dilated right ventricle, a decision was made to place an ecmo. the patient was then transported to another facility and placed on bivad support. the patient was reported to be doing well, and there have been no indication of any additional adverse affects to this patient. | [5 empty] | The impella cp pump has not been returned to abiomed for evaluation and analysis as it is against hospital policy to remove the product from the hospital. the manufacturer will continue to investigate all reasonably obtainable sources of information and will provide results and updated conclusions in a supplemental medwatch report. (b)(4).
38384 | 1220648-2013-00001 | 2930324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2930324&pc=KFM |  |  |  | 01/21/2013 |  |  | 09/26/2012 | 01/21/2013 |  |  | 08/13/2014 |  | 01/21/2013 |  |  |  |  | IMPELLA AIC 2.5 PUMP SET, 13FR INTRODUCER |  |  | KFM |  |  | IMPELLA 2.5 |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  | 2013152931 |  | 0046-0060 |  | Yes | [5 empty] | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | UNK | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [7 empty] | NA |  | The complaint reported that a female patient (age unknown) was being treated for stressed inducted cardiomyopathy following a motor vehicle crash. the patient coded in the hospital, and was brought to the cath lab. the patient's blood pressure was marginal, and she was administered levodopa and lopamine. an impella 2.5 pump was implanted on (b)(6) 2012. the device was implanted via the use of abiomed's 13fr x 13 cm introducer. the peel-away 13fr was left in the place with the repositioning sheath "plugged" into the back of it. on (b)(6) 2012, the drips were discontinued. the patient's blood pressure was stable, but an echo was performed which demonstrated severly reduced ejection fraction. the impella 2.5 support was continued. early on (b)(6) 212, after over 30.5 hours of impella support, the patient experienced acute profuse bleeding from the access site. the physician performed manual compression, which did not help. the patient was taken to the operating room for site repair. during this procedure, the sheath appeared to be cracked and causing the bleeding. approximately four units of prbc's, 2ffp, 6pk of platelets and about a liter of albumin were administered to the patient. the device was removed from the patient, the artery was repaired and the patient was stabilized. the patient was subsequently discharged from the hospital to her home. there have been no additional adverse patient effects reported. | [5 empty] | The 13fr introducer was returned for evaluation. visual investigation revealed evidence of excessive force being applied to the device either during insertion and/or during the over 30.5 hours of recovery. bulges were found along the length of the introducer that indicated excessive force during insertion. however, the introducer did not start leaking until approximately 30 hours after the pump had been implanted in the patient. therefore, it is also possible that patient movement, i.e. rolling over onto the introducer, may have caused the split and subsequent leak. it is also possible that the leak may have been due to a combination of these forces. the root cause of this event has been determined to be the result of excessive force applied during insertion of the introducer, which lead to the subsequent crack in the score line and patient bleeding. (b)(4).
38385 | 2976776 | 2976776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2976776&pc=KFM | 11/18/2012 | 12/12/2012 | 12/12/2012 | 12/12/2012 | 12/12/2012 | 02/25/2013 |  | 12/12/2012 |  |  |  |  | 12/12/2012 |  | PUMP, BLOOD |  |  | IMPELLA |  |  | KFM |  |  | * |  |  | ABIOMED, INC |  |  | US |  | * |  | * | [8 empty] | Invalid Data |  |  | User Facility report | [8 empty] | 67 YR | [5 empty] | Death | Yes | Yes | [20 empty] | Death | [28 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | Impella lvad began alarming high motor current. physicians were paged. art heart and rep notified. high motor current continued for several hours before pump power was ordered turned down to p4 to possible extend pump life. the pump failed. no one was in house to pull the pump. per rep: device needs to be withdrawn within 30 minutes of failure. delay in pulling device.======================manufacturer response for impella 2.5 catheter, 2.5 catheter (per site reporter).======================pump was past it's rated useful life at time of failure.what was the original intended procedure?left ventricular assist.
38386 | 1220648-2012-00010 | 2831605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2831605&pc=KFM | 07/10/2012 | 07/24/2012 |  | 11/07/2012 |  |  | 07/24/2012 | 11/07/2012 |  |  | 02/01/2014 |  | 11/07/2012 |  |  |  |  | IMPELLA LP5.0 PUMP |  |  | KFM |  |  | LP 5.0 |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  | 0062173 |  | 004680 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Other | Foreign | Health Professional | User facility | Company representation | Distributor | 36 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [7 empty] | NA |  | The complainant reported that an impella lp5.0 pump was implanted in a (b)(6) female pt in cardiogenic shock. during pt support, pump dislodgement problems occurred. multiple attempts were made to reposition the pump over a two day period during the weekend period of (b)(6) 2012. the re-positioning technique used was outside of recommended practices. it is believed it was done w/o turning down the speed of the device, w/o visual guidance via an echo or an x-ray. also although x-rays and echo's were taken after the repositioning attempts, the staff was unable to interpret the results. the device supported the pt for over 6 days. after weaning the pt off the pump, she was reported to have experienced major aortic valve leakage. an echo was performed which confirmed a left coronary cusp rupture. the physician felt that the pump had created the mechanical damage to the pt's aortic leaflets. in an effort to resolve the aortic leaflet damage a corevalve was successfully implanted in the pt. | [5 empty] | The complainant reported that the pt support from the device was successful. at the time of the pump removal the pt had recovered and was stable. the pt is currently doing fine. the impella lp5.0 pump was not returned for eval. due to this the analysis of this complaint could only be based upon the available clinical info and the analysis of the automated impella controller (aic) data. the aic data revealed the following observations: during the first 3 days of pump support some positioning issues occurred that were successfully resolved by repositioning the pump. the pump sensor showed some slight drift during this time period but stabilized again. the pump ran fine and w/o any issues for the next 2 days before a sudden sensor failure occurred. during this time an assessment of the pump position, based on the pump sensor signal, was no longer possible w/o taking imaging techniques or motor current values into consideration. this sensor failure occurred on a sunday afternoon at around 3 pm. the handling of the device at the occurrence of such sensor failure has been described before this event and communicated via an abiomed technical bulletin. due to the device being discarded and due to suspected device induced aortic valve damage, the analysis of this event has been augmented with physician interviews in a meeting with abiomed, and the distributor at the clinical site. add'l info obtained during that meeting included the following: several attempts to reposition pump by inexperienced personal were done over the weekend. documentation of miss-placement by means of echo and x-ray were reviewed. unfortunately the "weekend" physicians were not able to interpret the echo results. over the weekend frequent changes of pump position were performed w/o visual guidance, and it also potentially was done w/o reducing the speed of the device. these are both against recommended practices. the add'l problems with the failing pump sensor complicated the issue but there are instructions issued for how to manage the pump with a failed sensor. the valve damage was documented and confirmed by the echo performed after the pump was removed from the pt. with the info gathered during the meeting referenced it has been concluded that the inadequate assessment of the pump position and the inappropriate handling/repositioning of the device during the weekend may have contributed valve damage. finally, and w/o having the device back for analysis, the definitive root cause for the valve damage cannot be determined. as a corrective action customer retraining of the hospital staff is currently being performed. "impella technical bulletin (b)(4): recommended catheter explant procedure" ((b)(4)) has been initiated and disseminated to the abiomed clinical team and abiomed customers. among the actions discussed in the bulletin, the user has been advised to: exercise care while manipulating the impella catheter in a pt on cpb, particularly if the aortic valve is in the closed position throughout the cardiac cycle. use fluoroscopic or echocardiographic imaging to view the impella catheter and aortic valve and confirm that the impella catheter is moving freely and is not distorting the cardiac structures. remove the impella catheter in a controlled and gradual manner. in addition, the instructions for use for the impella lp5.0 pump cautions the user with the following warnings: avoid over inserting the impella 5.0 catheter and possibly impinging the catheter tip against the walls of the vasculature, atria or ventricles. do not advance or withdraw the impella 5.0 catheter against resistance w/o using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or ventricular perforation. (b)(4).
38387 | 1220648-2012-00011 | 2831606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2831606&pc=KFM | 07/30/2012 | 10/04/2012 |  | 11/07/2012 |  |  | 10/04/2012 | 11/07/2012 |  |  |  |  | 11/07/2012 |  |  |  |  | IMPELLA LP5.0 PUMP |  |  | KFM |  |  | LP 5.0 |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  | UNK |  | 004680 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Other | Foreign | Health Professional | User facility | Company representation |  | 24 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | The complainant reported that a (b)(6) male was admitted to the hospital. the pt was a heart failure pt and was going to have a heartware device placed for bridge to transplant therapy (btt). the pt initially was felt to be in decompensated heart failure and support with impella 5.0 was deemed to be appropriate, in order to better assess the pt's status. the impella 5.0 was implanted via the femoral artery. the implant was done under fluoroscopic control and thought to be fairly standard. there was no difficulty noted with insertion of the device. the pt had an excellent course while on 5.0 support. there was no issue noted with position. there was no hemolysis issue. the pt was taken to the operating room after about 5 days and was transitioned to a heartware lvad. the procedure was unremarkable. the pt was transitioned, by first doing the aortic anastomosis, and then performing the left ventricle (lv) apical core maneuver. it was reported that the impella was turned off and withdrawn prior to initiating the insertion of the pump in the apex. the entire procedure was performed w/o cardiopulmonary bypass. the removal of the impella was not visualized on fluoroscopy during removal. the heartware device was performing nicely. there was just a "whiff" of ai (aortic insufficiency) noted at the completion of the procedure, and was not though to be anything out of the ordinary. over the next 4 to 5 days the pt gradually started to exhibit signs of low output. at first there was some question that there could be an lvad thrombosis since the lvad was supposedly delivering a relatively high output, but the pt was exhibiting signs of low systemic output. the physicians decided to re-operate on the pt to establish if the lvad was thrombosed. during this work up of the low output syndrome, there was supposedly at least one echo that did not reveal any significant ai, but there was some technical question as to the quality of the study. at surgery another tee (transesophageal echocardiography) was performed. this did show more ai than was shown on prior occasions, and the lvad was found to be functioning normally. it was noted at this time that the ai had progressed to the point that there was systemic hypoperfusion. the decision was made to explore the aortic valve. the pt was placed on cardiopulmonary bypass and it was noted that the non-coronary cusp was dehisced from the sinus and was incompetent. it appeared to the operating physician as if it had deteriorated over the past week or so and was noted to have some fibrosis on the edges, denoting that it was older than just a day or two, but was on the order of a week or less. in an effort to repair the valve, the flail cusp was then sewn to a remaining aortic leaflet to essentially create a bicuspid but competent valve, and the pt was weaned from bypass with the functioning lvad. the pt did well postoperatively and did not suffer any further complications. in further discussion with the physician it was not known whether there could have been a guidewire induced injury or other source of the injury or, if it was related to impella, why there was a latency to the event of 4-5 days after the heartware device was implanted. | [5 empty] | The complainant was unable to report the serial number of the impella lp 5.0 pump, consequently, the device manufacture and expiration dates were unable to be reported. the impella lp5.0 pump was not returned for eval, as it was discarded at the user facility. the data log files have also not been returned for analysis. it was reported that because of the length of time from the date of the event to the reporting of the event to abiomed, the data log files may not be retrievable. the mfr will continue to investigate all reasonably obtainable sources of info and will provide results and updated conclusions in a supplemental report when rec'd. "impella technical bulletin 59: recommended catheter explant procedure" (hcs-tb1079) has been initiated and disseminated to the abiomed clinical team and abiomed customers. among the actions discussed in the bulletin, the user has been advised to: exercise care while manipulating the impella catheter in a pt on cpb, particularly if the aortic valve is in the closed position throughout the cardiac cycle. used fluoroscopic or echocardiographic imaging to view the impella catheter and aortic valve and confirm that the impella catheter is moving freely and is not distorting the cardiac structures. remove the impella catheter in a controlled and gradual manner. in addition, the instructions for use for the impella lp5.0 pump cautions the user with the following warnings: avoid over inserting the impella 5.0 catheter and possibly impinging the catheter tip against the walls of the vasculature, atria or ventricles. do not advance or withdraw the impella 5.0 catheter against resistance w/o using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or ventricular perforation. (b)(4).
38388 | 2804869 | 2804869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2804869&pc=KFM | 09/27/2012 | 10/08/2012 |  | 10/11/2012 |  | 09/27/2012 |  | 10/11/2012 |  |  |  |  | 10/11/2012 |  | HEART PUMP/PURGE CASSETTE |  |  | AUTOMATED IMPELLA CONTROLLER/PURGE CASSETTE |  |  | KFM |  |  | NO NUMBER |  |  | ABIOMED, INC. |  |  | US |  | CASSETTE 1090486 |  |  |  | Yes |  |  | Not Available |  |  |  | Health Professional |  |  | User Facility report |  |  | Unknown | [5 empty] | 88 YR |  |  | 79 |  |  | Death | Yes | No | [20 empty] | Death | [28 empty] | Initial submission | [5 empty] | N | HOSPITAL |  |  | Pt went to cath lab for high-risk pci with hemodynamic support (impella). attending implanted the impella catheter and had full support before beginning to fix the pt's multiple lesions. during the procedure it was noticed the impella purge cassette was leaking and a new purge cassette was loaded into the impella console. the console message display showed that it did not recognize the cassette. the purge pressure decreased and the pt's conditions deteriorated resulting in hypotension, loss of consciousness and the initiation of cpr.
38389 | 1220648-2012-00008 | 2783338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2783338&pc=KFM | 01/07/2011 | 01/07/2011 |  | 10/02/2012 |  |  | 01/07/2011 | 10/02/2012 |  |  |  |  | 10/02/2012 |  |  |  |  | IMPELLA LP 5.0 PUMP |  |  | KFM |  |  | LP 5.0 |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  |  |  | 004601 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Other | Foreign | Literature | Health Professional | User facility | Company representation | 45 YR | [5 empty] | Injury | Yes | Yes | [20 empty] | Required Intervention | [25 empty] | Other |  |  | Initial submission | [9 empty] | Upon independent review of a journal article entitled severe aortic insufficiency secondary to 5l impella device placement published in 06/2012 was found. the article reported an adverse event that was reported to have occurred with an impella lp5.0 pump. a summary of the article is as follows: a 45 year old female pt who had been implanted with an impella lp 5.0 pump developed a severe aortic insufficiency after insertion and subsequent removal of the pump after 2 weeks, as the cardiac function had recovered. in an effort to resolve the pt problem the clinicians implanted a transcatheter corevalve. the pt was reported to have done well post procedure, and was asymptomatic six months following transcatheter valve implantation with a normal functioning prosthetic valve. please reference the attached journal article for details regarding pt history and comorbidities, this article will also supply implant and explant procedure details. | [5 empty] | Published journal article in 06/2012 (rahul chandra, m.d., robert cusimano, m.d., et al: severe aortic insufficiency secondary to 5l impella device placement, journal of cardiac surgery, 2012; 27; 400-402): from event that occurred at toronto general hosp, division of cardiovascular surgery and division of cardiology. investigation results: of note in the use of the 5.0, the outflow window acts as a shroud over the device outflow so the impeller blades cannot come in contact with the outside environment. therefore if there is damage created by impella 5.0, it would have to be caused by the inflow structure or pigtail of the device. the actual impella lp 5.0 pump used was not returned for evaluation. the mobile pump console (mpc) data was obtained after the paper was received and is the basis of this analysis. the case was identified as occurring between january 7 to january 14th, 2011. we received communication from a physician representing the site stating that there was an event three days into the case where a tee was performed and aortic regurgitation was noted. the paper reflected something happening when the impella was withdrawn after 7 days. because of the "3 days" comment, an examination of the console data on the evening of the (b)(6) 2011 was done. it showed changes to the placement signal that may reflect the manipulation referred to during the transesophageal echo. there appeared to be changes in pump position (beginning at 19:13), including a period when the pump was probably mal-positioned, as evidenced by the low motor current and placement signal pulsatility (beginning at 00:20 on (b)(6) 2010). at the same time numerous manipulations in motor speed were made triggering seven error-28 low flow alarms, probably as the device was being repositioned. there were three other significant events characterized by console alarms, abnormal placement signal and motor current waveforms that occurred during the case. the first episode (beginning at 23:30 on (b)(6) 2011) began as a reduction in placement signal and motor current pulsatility accompanied by reported high flows (+5 liters/minute), probably because the pump was completely inside the ventricle. at 00:27am on (b)(6), motor speed was briefly reduced, followed by a significant increase in placement signal and motor current pulsatility, indicating the device was repositioned. pump flow then dropped to 4 liters/minute, and there were a series of pump position wrong (error-2) alarms from 00:38 to 01:00, indicating the pump outflow was positioned at the level of the aortic valve. the alarms resolved after it appeared the device was repositioned. at the same time there was another brief reduction in motor speed triggering a low flow warning. after the pump was repositioned, flows returned to levels that are more normal (4.8-5.0 liters/minute) for an impella 5.0 running at p9. the second event occurred around 05:20 and lasted until 08:00 on (b)(6). it appeared the pump moved or was repositioned, causing a drop in flow to 4 i/m (with the pump speed remained a constant p9). starting at 06:30 there were fifteen more error-2 pump position warnings, indicating the pump outflow was again positioned at the level of the aortic valve. the alarms were resolved, and pump flows returned to previous levels when the pump was repositioned around 08:00. when the pump was repositioned, the user briefly reduced the p-level. the third event occurred on (b)(6) 2011 at 20:50, there was a reduction in placement signal and motor current pulsatility similar to previous events, probably due to pump movement or repositioning. at 21:15, there was a change in the placement signal when the pump was probably repositioned. this was accompanied many rapid changes in motor speed that occurred over the next hour and fifteen mins. the changes in motor speed were accompanied by twenty-nine low flow (error-29) warnings when pump flows dropped to abnormally low levels. two pump position (error-2) warnings also occurred indicating that the pump was being moved while ramping the operating speed as the pump was being repositioned. after the last pump repositioning at 22:35, pump performance stabilized and was normal for the remainder of the case. the role that the pt condition played in the problem could not be assessed, as the complainant refused to release the pt data for analysis. additional info received indicated that aortic insufficiency was first noticed after manipulating the devic on the third day of support. however, the paper published in the journal of cardiac surgery stated, "the right coronary cusp of the aortic valve seems to have been injured at the time of removal of the device as there was no aortic insufficiency with the impella device in place and the severe aortic regurgitation was found only after removal of the device". the device was not removed until day 7. the reason for the discrepancy is not known, but based on the available info it is likely that the injury to the valve occurred during pt support. this could explain all of the manipulation of the device that is evident and described on days 3 and 5 as the physician could have been maneuvering the device while changing speeds to get better visualization of some event they were observing. assuming the aortic insufficiency was first noted on the third day, as the additional info obtained indicated, the changes in the placement signal and user-commanded motor speed seen in the mpc data on the evening of (b)(6) was probably when this occurred (there were no other changes in the console data on (b)(6) indicating pump manipulation). two earlier episodes of pump manipulation on (b)(6) may have caused damage to the valve as well. in those incidents twenty console warnings for pump position were triggered, probably when the pump outflow was moved through the valve. in the later incident on (b)(6), repositioning was accompanied by more pump position warnings along with twenty-nine low flow warnings that were apparently triggered by rapid multiple reductions in motor speed while the device was being repositioned. in conclusion, the root cause of the damage to the right coronary cusp of the aortic valve that resulted in the aortic insufficiency was probably caused by rapid and repeated movement of the impella pump through the aortic valve while the device was operating and was being manipulated by the user at the same time. this was in at least 4 events and there were alarms accompanying the device events noted. the event that was primarily responsible for the damage was probably the manipulation that occurred on (b)(6) 2011, as aortic insufficiency was first noted on a subsequent transesophageal echo. the reason the device was being manipulated is not understood. the pump appeared to be positioned correctly and no pump position or suction alarms occurred that would require repositioning the pump. similar manipulation of the device before and after the primary event may have been contributing factors to the severe aortic insufficiency seen after removal of the impella. the damage to the right coronary cusp of the aortic valve that resulted in the aortic sufficiency was probably caused by rapid and repeated movement of the impella pump through the aortic valve, while the device was being manipulated by the user. the complainant was unable to report the serial number of the impella lp 5.0 pump, consequently, the device manufacture and expiration dates were unable to be reported. (b)(4).
38390 | 1220648-2012-00009 | 2783341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2783341&pc=KFM | 08/17/2012 | 08/24/2012 |  | 10/02/2012 |  |  | 08/24/2012 | 10/02/2012 |  |  | 11/30/2013 |  | 10/02/2012 |  |  |  |  | IMPELLA LP 5.0 PUMP |  |  | KFM |  |  | LP 5.0 |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  | 0060682 |  | 004680 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Other | Foreign | Health Professional | User facility | Company representation |  | 38 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | The complaint reported that an impella lp5.0 pump was implanted in a (b)(6) female pt on (b)(6) 2012 for cardiomyopathy. there was one issue for purge pressure problems that was resolved by changing the console cable. the pt was on support for 8 days as a "tune up" for transitioning to a heartmate ii (hm ii). on the morning of (b)(6) 2012, the pt was taken to the operating room for the hm ii implant procedure. after the pt was placed on bypass, the impella pump was reportedly turned off by disconnecting the console cable before it was removed from the left ventricle. after the surgeon cored the ventricle, there was a significant amount of blood exiting the heart. after emptying the heart and checking the bypass cannula, it was deduced the aortic valve leaflet was torn. an artificial aortic valve was implanted and the hm ii was successfully placed. the procedure was successful and the pt survived and ultimately did well. | [5 empty] | Investigation results: the pump was not returned for evaluation. the analysis of the reported valvular injury is based on the console data, complaint info records and an interview with the cardiothoracic surgeon present during the case. the cardiothoracic surgeon examined the aortic valve after the case and felt the damage did not appear to be acute but to be consistent with trauma that occurred over time. before the device was pulled however, the surgeon reported that the echocardiogram was reviewed and there was no sign of aortic insufficiency (ai). during the interview the possibility that the damage was caused by removing the device while the pt was on bypass with the aortic valve in a fully closed, collapsed state, was discussed. guidewire damage was also discussed. the mpc console data was examined. other than spikes in purge pressure early in the case (apparently resolved by the cable change) and a gradual rise in purge pressure that began on day 6, the pump appeared to operate normally. the placement signal was relatively constant with a min/max-differential pressure range of 0-mmhg/60-mmhg up to the time the pump was shut off. the mpc logged the pump being disconnected while it was running on the morning of (b)(6). just prior to it being disconnected, there were two wrong position (error-2) warnings that may have been triggered by the aortic valve blocking the pump outlet. it is not known whether the pump was repositioned or if bypass was being initiated, causing movement of the pump. it appears the event that caused the damage to the aortic valve occurred after the impella was turned off, based on the relative constant differential pressure signal up to that point, and the lack of ai just prior to removing the pump. it is believed that the most likely source of the damage was when the pump was being pulled through the aortic valve when the pt was on bypass. the pump may have been operating when reviewing the initial interview with the perfusionist who attended the explant procedure. the pt had been placed on bypass, and the aortic valve was likely in a fully closed state when the pump was withdrawn. because the pump was not returned, no assessment could be made concerning the role it may have played in contributing to the damage to the aortic valve, if any. (b)(4).
38391 | 1220648-2012-00007 | 2770113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2770113&pc=KFM | 07/02/2012 | 07/09/2012 |  | 10/01/2012 |  |  | 07/19/2012 | 10/01/2012 |  |  | 01/26/2014 |  | 10/01/2012 |  |  |  |  | IMPELLA RECOVER LP 2.5, .018 GUIDEWIRE |  |  | KFM |  |  | .018 IN (0.47MM) |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  |  |  | 10088287 |  | 0046-0043 |  | Yes |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 56 YR | [5 empty] | Malfunction | No | Yes | [26 empty] | ANGIOPLASTY/STENT OF THE RIGHT CORONARY ARTERY | PRIMARY PERCUTANEOUS TRANSLUMINAL CORONARY | [21 empty] | Initial submission | [7 empty] | NA |  | The complainant reported that a (b)(6) male patient returned to the hospital's clinical vascular laboratory with chest pains on (b)(6) 2012, after an initial st-segment elevation myocardial infarction treated with a primary percutaneous transluminal coronary angioplasty/stent implant of the right coronary artery (date unknown). an impella recover lp 2.5 percutaneous cardiac support device was implanted in the patient via the left femoral artery to provide partial circulatory support. an .018 guidewire was used to implant the impella 2.5 device. there was no resistance felt during insertion/advancement of the guide wire, or advancement of the impella device over the guide wire. during the subsequent removal of the .018 guide wire from the patient, the clinician encountered resistance, and the tip of the .018 guide wire was observed to have unraveled upon explant. there was no indication to the clinician that any portion of the guide wire was left in the patient. | [5 empty] | Investigation results: the guidewire was returned for investigation. the impella pump was also returned. the investigation did not reveal any damage on the pump. the investigation of the guidewire confirmed the reported problem. the coil of the guidewire was found to have completely unraveled and the guidewire core was found to have fractured. this distance was around 2", indicating that the breakage at the flat/round core junction; the tip of the ball was missing from the guidewire. the guidewire was then evaluated by the original equipment manufacturer (oem). the returned guidewire presented indication of a ductile, tensile overload of the core wire shaft proximal of the flattened distal region, approximately 5.78cm distal of the proximal coil to core wire braze joint. the wire also presented extensive stretch damage to the coil wire terminating in a ductile tensile fracture of the coil wire with torsional loading, an indeterminate distance from the proximal coil to core wire braze joint. the wire shaft also presented numerous bends and kinks scattered over the entire length, beginning 0.24cm proximal of the core wire fracture. the distal tip joint and an indeterminate length of core wires were found to be missing and not accounted for. no other damages or inconsistencies were identified. the proximal coil to core braze joint appears otherwise correct and intact by visual examination and by non-destructive testing. in conclusion, a device history record (dhr) review indicated that this device lot number (10088287) was final inspection tested and deemed acceptable for shipment. a sample of the product from each lot is destructively pull tested. the history records for this lot indicate pull test results of over 1.0 pounds (pull strength specification is 0.5 pounds minimum.) the review of the dhr of this device did not present any indication of manufacturing defect or anomaly that could have had an impact on this event. it appears that procedural, and/or clinical factors may have impacted this event. these observations are based upon the evidence presented by the returned guidewire and the information provided by the supporting documentation. the root cause of this event was determined to be the result of interactive excessive force that occurred in this case between the guidewire and introducer, as the force exceeded the guidewire joint pull strength. the root cause of this event has been determined to be the result of the excessive interactive force that was expended in this case between the guidewire and introducer . that force exceeded the guidewire joint pull strength specification. the following corrective action has been initiated to address this issue: based upon the investigation of the failure a capa (b)(4) was opened to review the need to increase the pull strength of the joint. the investigation also revealed one other similar complaint in the last 12 months for serial number (b)(4). a biomed will continue to monitor the complaint database for similar, future occurrences. (b)(4).
38392 | 1220648-2012-00006 | 2770323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2770323&pc=KFM | 05/13/2012 | 05/14/2012 |  | 10/01/2012 |  |  | 05/14/2012 | 10/01/2012 |  |  | 01/30/2014 |  | 10/01/2012 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERC CARDIAC SUPPORT |  |  | KFM |  |  | LP 2.5 |  |  | ABIOMED |  |  | GM |  | NA |  | 002138 |  | Yes |  |  | Yes |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Other | Foreign | Health Professional | User facility | Distributor | [7 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [7 empty] | NA |  | The complainant alleged that the clinician implanted an impella lp2.5 pump in a female pt who was also being treated with triple platelet aggregation inhibitors and anticoagulation drugs. subsequent to the device being implanted, and during the hours of impella support, the pt experienced leaking/bleeding through the valve at the repositioning sheath. the pt's hemoglobin relevant bleeding was estimated to result in a blood loss of approx 300ml. the pt was administered red blood cells, for the treatment of the blood loss. the pt blood loss was reported to have been exacerbated by the anti-coagulation drugs that had been administered to her. after the clinicians unsuccessfully attempted to stop the bleeding, another impella lp2.5 pump was implanted in the pt. the complainant indicated that the second impella lp2.5 pump was successfully implanted in the pt, and was then in stable condition. there were no further adverse effects to the pt as a result of this reported event. | [5 empty] | Investigation results: the device was returned for evaluation. the problem was reproduced. the root cause of the bleeding was a defective valve seal. this results from high frictional interactive forces between the valve seal and the catheter, leading to either malfunction of the haemostatic valve (the valve being pushed into the distal end of the hub) or valve damage (valve crack, valve leaflet damage). the evaluation determined that the root cause of this event was related to the mechanical design of the valve seal. the following corrective action has been initiated to address this issue: based upon the investigation of the failure a capa (b)(4) was opened to procure and design a new valve seal. internal ref: (b)(4).
38393 | 2802549 | 2802549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2802549&pc=KFM | 09/02/2012 | 09/19/2012 | 09/19/2012 | 09/19/2012 | 09/19/2012 |  |  | 09/19/2012 |  |  |  |  | 09/19/2012 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA 2.5 |  |  | KFM |  |  | 2.5 |  |  | ABIOMED |  |  | US |  | * |  | 005042 |  | Yes | [5 empty] | No | Physician |  |  | User Facility report | [8 empty] | 64 YR |  |  | 121 |  |  | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | During the procedure the surgeon used the wire through the diagnostic catheter to exchange for the impella device. the impella device was advanced across the aortic valve and into the left ventricle. location was confirmed by post-impella pump module as well as by fluoroscopy. the impella device was flushed and initiated. initial impella pump flow was noted to be as high as 1.6 to 1.8. however, the impella device was noted to rapidly decline in its augmented flow to 0.8 to 0.9. the impella device was flushed once again. multiple repositioning of the impella device was performed. echocardiogram was performed with limited views to confirm appropriate placement of the impella device. the patient was monitored closely with impella device running. the patient was noted to have stable pulmonary artery pressures as well as oxygen saturations. impella device was noted to have a decline in augmented output again to 0.8 l per minute. decision was then made to discontinue the impella device. to reverse the heparin previously given, protamine was administered. the impella device was then removed without difficulty. hemostasis was obtained with the use of manual pressure. after approximately 40 minutes of manual compression, the patient was noted to have residual hematoma and excellent hemostasis.
38394 | 2619743 | 2619743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2619743&pc=KFM | 01/19/2012 | 05/30/2012 | 05/30/2012 | 06/18/2012 |  | 01/19/2012 |  | 06/18/2012 |  |  |  |  | 06/18/2012 |  | HEART PUMP |  |  | IMPELLA 2.5 |  |  | KFM | [5 empty] | ABIOMED |  |  | US | [5 empty] | Yes |  |  | Not Available |  |  | No | Other |  |  | User Facility report |  |  | Unknown | [5 empty] | 54 YR |  |  | 60 |  |  | Injury | No | Yes | [20 empty] | Required Intervention | [28 empty] | Initial submission |  |  |  |  | 04/13/2012 | N | HOSPITAL | 54645 |  | On (b)(6) 2012, at approx 10:25 am, while at work as rn special assignment within the hosp, (b)(6) female developed severe sob and presented to ed. pt was later transported to the cardiac catheterization lab in cardiogenic shock. the pt was found to have elevated troponin level, an ef of 10% and coronary spasm. she was intubated in the cath lab before any respiratory arrest occurred. due to the severity of the pt's condition, the impella heart pump device by biomed was selected for use. between the hours of 19:00 and 21:15 the machine was alarming "purge pressure". after attempting to follow the troubleshooting recommendations of the machine, the alarm continued. when the alarming continued, the nurse attempted to purge the system again by inserting a blunt tip catheter in the blue bladder of the device. immediately, the nurse noticed fluid leaking from the bladder. at approx 21:35, the machine began failing (i.e., unable to maintain pressure at a proper level). the warning label was affixed to the underside of the diaphragm which contains the blue bladder rather than the warning label covering the blue bladder. while at approx 21:35, the machine began failing, the pt remained hemodynamically stable throughout.
38395 | 2627548 | 2627548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2627548&pc=KFM | 06/04/2012 | 06/11/2012 | 06/11/2012 | 06/15/2012 | 06/15/2012 |  |  | 06/15/2012 |  |  |  |  | 06/15/2012 |  | IMPELLA CONTROLLER |  |  | CONTROLLER, ABIOMED IMPELLA |  |  | KFM |  |  | R9G0042400 |  |  | ABIOMED |  |  | US |  | * |  | * |  | No | [6 empty] | NURSE |  |  | User Facility report | [8 empty] | 34 YR | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | The impella controller screen froze with a black screen and no options to select. there was a constant hum from the controller console. the impella rotation icon showed no rotation. auscultation of the pt demonstrated no impella rotation noises. a new controller was hooked up with complete waveform operation and rotation icons. auscultation of the pt confirmed impella rotation sounds.
38396 | 2569084 | 2569084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2569084&pc=KFM | 03/28/2012 | 04/16/2012 | 04/16/2012 | 04/19/2012 | 04/19/2012 |  |  | 04/19/2012 | 04/19/2012 |  |  |  | 04/19/2012 |  | HEART PUMP | CONTROL INTERFACE FOR IMPELLA |  | SYSTEM, IMPELLA 2.5 | CONTROLLER, AUTOMATED IMPELLA |  | KFM | KFM |  | * | * |  | ABIOMED | ABIOMED |  | US | US | 56387 | * | * | * | Yes | Yes | [5 empty] | Invalid Data | Invalid Data |  | User Facility report | [8 empty] | 58 YR |  |  | 68 |  |  |  | No | No | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * | * | In the patient's room, the impella device was showing low pulsatile pressure and was alarming to alert for evaluating the placement of the catheter. cardiologist was brought to the bedside along with the echo tech. a stat echocardiogram (echo) was done to evaluate the placement of catheter. the catheter was repostioned by cardiologist and the wave forms returned to proper form and all parameters returned to within normal limits. cardiologist tightened the device and left the room. within appoximatley one to two minutes the device alarmed and the machine showed it went from p8 to p0. stat call was placed to the device rep, and the rep walked the team through the same steps they had already performed prior to the call. the team was unable to restart the machine. the cardiologist and echo tech were called back to the room and a decision was made to pull the catheter back into the aorta and shut the machine off.
38397 | 2512959 | 2512959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2512959&pc=KFM | 02/21/2012 | 03/16/2012 | 03/16/2012 | 03/16/2012 | 03/16/2012 |  |  | 03/16/2012 |  |  |  |  | 03/16/2012 |  | INTRACARDIAC LV PUMP |  |  | SET, IMPELLA 5.0 |  |  | KFM |  |  | * |  |  | ABIOMED |  |  | US |  | * |  | 005062 |  | Yes | [6 empty] | Invalid Data |  |  | User Facility report | [8 empty] | 41 YR |  |  | 119 |  |  | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | The impella device was placed in an attempt to stabilize the pt in preparation for mitral valve repair coronary artery bypass grafting (mvr/cabg). post-op day one, the physician noted that the purge system on the device was not functioning properly. pt was taken back to the operating room where the impella was explanted and exchanged for a new device with a different lot number. excellent hemodynamics were noted after placement of the new impella 5.0. post explantation, the first device was noted to be kinked, and was giving an error.
38398 | 2513458 | 2513458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2513458&pc=KFM | 02/20/2012 | 03/16/2012 | 03/16/2012 | 03/16/2012 | 03/16/2012 |  |  | 03/16/2012 |  |  |  |  | 03/16/2012 |  | INTRACARDIAC LV PUMP KIT |  |  | SET, IMPELLA 2.5 |  |  | KFM |  |  | 005042 |  |  | ABIOMED |  |  | US |  | * |  | * |  | Yes | [6 empty] | Physician |  |  | User Facility report | [8 empty] | 80 YR |  |  | 60 |  |  | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | Because of the patient's degenerating hemodynamic condition, the physician left a six french guideliner wire in and got access for the impella device on the left common femoral artery. physician initially tried to advance the impella over the wire and was unable to cross the aortic valve. he removed the wire and theimpella device and advanced a regular j-wire over a jr4 diagnostic catheter. the md exchanged it over the jr4 diagnostic catheter, the 0.018 wire and reloaded the impella device. after flushing it, the md was successfully able to advance it across the aortic valve and impella was started at 2.5 l. immediately after placement of the impella, it was noted that the pt went into ventricular fibrillation. cardioversion was performed and then it was noted that the patient remained pulseless in spite of 2.5l of flow with the impella. so, cpr was started and acls protocol was continued for more than 20 minutes. after 20 minutes, it was decided to withdraw caths, so the impella was pulled from the left ventricle and the pt regained his pulsatility. acls protocol was continued. the pt at that time had a rhythm and a pulse and was able to be oxygenated.
38399 | 1220648-2012-00002 | 2496080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2496080&pc=KFM | 02/08/2012 | 02/08/2012 |  | 03/08/2012 |  |  | 02/08/2012 | 03/08/2012 |  |  | 10/31/2013 |  | 03/08/2012 |  |  |  |  | IMPELLA LP 5.0 |  |  | KFM |  |  | LP 5.0 |  |  | ABIOMED EUROPE GMBH (GERMANY) | [6 empty] | 004601 |  | Yes |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Consumer | Health Professional | User facility | Company representation | [8 empty] | Malfunction | Yes | Yes | [20 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | The following was reported to abiomed through our rep: a pt was implanted with an abiomed impella 5.0 cardiac assist device providing left ventricular circulatory support at ottawa heart institute, ottawa, ontario. the impella device was inserted through the left common femoral artery on saturday, (b)(6) 2012. it was reported that the implantation was uneventful and no excessive force was required during insertion. this was followed by uneventful circulatory support for 4 days. while on the operating table for an impending heart transplant on (b)(6) 2012, the impella device was explanted. upon explant the surgeon discovered that the distal portion ('pigtail') was missing from the device and was discovered separated from the main body of the device and located inside the pt's external iliac artery through fluoroscopic examination. it was located per the report at about 5 cm above femoral artery clamp site. retroperitoneal dissection of the external iliac artery was performed and the remaining component was brought out through the arteriotomy. the pt was unharmed and underwent a successful transplant after 4 days of successful support from the impella 5.0 device. | [5 empty] | The mfr will continue to investigate all reasonably obtainable sources of info and will provide results and updated conclusions in a supplemental mfr medwatch report. (b)(4).
38400 | 2495444 | 2495444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2495444&pc=KFM | 02/26/2012 | 03/07/2012 | 03/07/2012 | 03/07/2012 | 03/07/2012 |  |  | 03/07/2012 |  |  |  |  | 03/07/2012 |  | IMPELLA CATHETER 2.5 |  |  | IMPELLA 2.5 |  |  | KFM |  |  | * |  |  | ABIOMED |  |  | US |  | * |  | * |  | Yes | [6 empty] | Invalid Data |  |  | User Facility report | [8 empty] | 38 YR | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | There was a sudden cessation of the motor pump on the impella. the rep was called and they advised us to pull it back 30cm into the aorta. because if it was left in the ventricle, there would be 1 l/min of backflow into the heart and the patient may not tolerate it. the mfr response has been none as yet.
38401 | 1220648-2012-00003 | 2481048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2481048&pc=KFM | 02/03/2012 | 02/03/2012 |  | 03/02/2012 |  |  | 02/03/2012 | 03/02/2012 |  |  | 07/31/2013 |  | 03/02/2012 |  |  |  |  | IMPELLA RECOVER LP 2.5, .018" GUIDEWIRE |  |  | KFM |  |  | .018" (0.47MM) |  |  | ABIOMED EUROPE GMBH (GERMANY) |  |  | GM |  | 10026365 |  | 0046-0043 |  | Yes |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Consumer | Health Professional | User facility | Company representation |  |  | 72 YR | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [9 empty] | A pt underwent an elective high risk percutaneous coronary intervention (hrpci). an impella 2.5 device was inserted for prophylactic circulatory support over an .018" guidewire prior to the intervention. the distal floppy tip of the guidewire was unraveled when it was pulled out of the impella device and then the tip fell off outside the pt. the intervention was then completed successfully. no part of wire was left in the pt and the pt did not suffer any harm or injury due to the event. | [5 empty] | The mfr will continue to investigate all reasonably obtainable sources of info and will provide results and updated conclusions in a supplemental mfr medwatch report. (b)(4).
38402 | 2494490 | 2494490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2494490&pc=KFM | 02/29/2012 | 02/29/2012 | 02/29/2012 | 02/29/2012 | 02/29/2012 |  |  | 02/29/2012 |  |  |  |  | 02/29/2012 |  | IMPELLA MOBILE CONSOLE AND POWER SUPPLY |  |  | SYSTEM, CIRCULATORY, SUPPORT, IMPELLA 2.5 |  |  | KFM |  |  | * |  |  | ABIOMED |  |  | US |  | * |  | * |  | Yes | [6 empty] | Invalid Data |  |  | User Facility report | [8 empty] | 38 YR | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | There was a sudden cessation of the motor pump on the impella. we called the rep and they advised us to pull it back 30cm into the aorta. if it was left in the ventricle there would be 1l/min of backflow into the heart and the patient may not have tolerated it. we agreed this was probably happening because as soon as the pump stopped working the patient had multiple premature ventricular contractions (pvc's). the dr did this for us. the rep requested it to be d/c'd fully within two hours.
38403 | 1220648-2011-00001 | 2397555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2397555&pc=KFM | 10/21/2011 | 10/21/2011 |  | 11/30/2011 |  |  | 11/01/2011 | 11/30/2011 |  |  | 07/31/2014 |  | 11/30/2011 |  |  |  |  | IMPELLA RECOVER LP 2.5, .018 GUIDEWIRE |  |  | KFM |  |  | .018 IN (0.47 MM) |  |  | ABIOMED |  |  | GM |  | 10028100 |  | 0046-0043 |  | Yes |  |  | Yes |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Consumer | Health Professional | User facility | Company representation |  |  | 73 YR | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [9 empty] | An impella 2.5 device was inserted and advanced into the left ventricle over a .018" guidewire for prophylactic support during a high risk percutaneous coronary intervention (hrpci). the distal tip of the .018" guidewire appeared 'split' when withdrawn from the pt. pt vasculature and left ventricle were evaluated under fluoroscopy and no part of the wire was observed. the impella device was left in the pt and the physician continued with the intervention; lad and circumflex arteries were ballooned and stented. the impella 2.5 device was removed and pt was doing fine following the procedure. the .018" guidewire was returned to abiomed. a visual inspection under magnification revealed that the distal tip was missing. follow up with the physician confirmed that the pt was still doing fine and displayed no side effects. the physician did not believe that any portion of the wire was in the pt because the vasculature and left ventricle were evaluated under fluoroscopy during the case and nothing was observed. | [5 empty] | Evaluation conclusions: conclusions not yet available. evaluation in progress. the failure investigation is currently in progress; therefore, results are not yet available and no conclusion can be drawn at this time. the mfr will continue to investigate all reasonably obtainable sources of info and will provide results and updated conclusion in a supplemental mfr medwatch report.
38404 | 2341332 | 2341332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2341332&pc=KFM | 10/22/2011 | 10/28/2011 | 10/28/2011 | 11/02/2011 | 11/02/2011 |  |  | 11/02/2011 |  |  |  |  | 11/02/2011 |  | IMPELLA CONTROLLER |  |  | IMPELLA CONTROLLER |  |  | KFM |  |  | * |  |  | ABIOMED, INC. |  |  | US |  | * |  | * |  | No | [6 empty] | UNKNOWN |  |  | User Facility report | [8 empty] | 58 YR | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | Impella 5.0 controller failure resulted in a pulseless electrical activity (pea) arrest. the controller failure alarm was on the screen with no movement of the pump and no motor current was shown. pulselessness was verified on the monitor screen and by palpation of pulses. the controller was restarted. the position of catheter was verified within the patient. the patient was intubated and cpr was initiated. contacted abiomed rep. all functions of the controller were normalized after restart.
38405 | 2341348 | 2341348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2341348&pc=KFM | 10/20/2011 | 10/28/2011 | 10/28/2011 | 11/02/2011 | 11/02/2011 |  |  | 11/02/2011 |  |  |  |  | 11/02/2011 |  | ABIOMED IMPELLA PURGE CASSETTE |  |  | IMPELLA PURGE CASSETTE |  |  | KFM |  |  | 0043-0001 |  |  | ABIOMED, INC. |  |  | US |  | * |  | * |  | No | [6 empty] | UNKNOWN |  |  | User Facility report | [8 empty] | 61 YR | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | The cassette would not snap into place and prime. a new cassette was obtained and it worked.
38406 | 2305274 | 2305274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2305274&pc=KFM | 09/01/2011 | 09/22/2011 | 09/22/2011 | 09/30/2011 | 09/30/2011 |  |  | 09/30/2011 |  |  |  |  | 09/30/2011 |  | CATHETER TO INTRAVASCULAR BLOOD PUMP |  |  | CATHETER, IMPELLA 2.5 |  |  | KFM |  |  | 52277 - ONLY NUMBER LOCATED DUE TO HEAVY BIOBURDEN |  |  | ABIOMED |  |  | US |  | * |  | * |  | Yes | [6 empty] | Invalid Data |  |  | User Facility report | [8 empty] | 50 YR |  |  | 82 |  |  | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | The impella device was inserted and the pt returned to the nursing unit. during the nurse's assessment, it was noted that there was a leak at the distal end of the catheter and there was a crack at the catheter hub. the device was removed and a second device was inserted.
38407 | 2271995 | 2271995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2271995&pc=KFM | 09/01/2011 | 09/06/2011 | 09/06/2011 | 09/06/2011 | 09/06/2011 |  |  | 09/06/2011 |  |  |  |  | 09/06/2011 |  | IMPELLA ASSIST DEVICE |  |  | IMPELLA PUMP KIT 2.5 |  |  | KFM |  |  | * |  |  | ABIOMED |  |  | US |  | * |  | 004672 |  | Yes | [6 empty] | Invalid Data |  |  | User Facility report | [8 empty] | 82 YR |  |  | 71 |  |  | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | The pt had an impella placed five weeks ago for the treatment of cardiogenic shock. three days later, the device gave a message that the motor had stopped in the afternoon. the pt was emergently taken to the cath lab to replace the impella device, which was replaced without further incident. however, the pt did develop ventricular tachycardia and required increasing pressor suppport. the patient did expire the next day.
38408 | 2214047 | 2214047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2214047&pc=KFM | 05/10/2011 | 06/06/2011 | 06/06/2011 | 08/09/2011 | 08/09/2011 |  |  | 08/09/2011 |  |  |  |  | 08/09/2011 |  | IMPELLA |  |  | IMPELLA |  |  | KFM |  |  | 2.5 |  |  | ABIOMED |  |  | US |  | 2012112043 |  | 004672 |  | Device was returned to manufacturer | [6 empty] | Physician |  |  | User Facility report | [8 empty] | 77 YR | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | I was informed that impella was defective and leaking at the connection and removed from the pt. the pt was brought to the cath lab for an impella insertion. when the original impella catheter was inserted, the connection was leaking and they could not get the pressure to rise. the rep from abiomed determined the catheter was defective. the catheter was removed and a new one was inserted without incident. the rep took the catheter with him to be evaluated by the company.
38409 | 2201262 | 2201262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2201262&pc=KFM | 06/24/2011 | 07/08/2011 | 07/08/2011 | 07/22/2011 | 07/22/2011 |  |  | 07/22/2011 |  |  |  |  | 07/22/2011 |  | MEDICAL DEVICE (IABP) |  |  | SYSTEM, IMPELLA 2.5 |  |  | KFM |  |  | * |  |  | ABIOMED, INC. |  |  | US |  | 004672 |  | * |  | Yes | [6 empty] | Invalid Data |  |  | User Facility report | [8 empty] | 76 YR |  |  | 105 |  |  | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | The impella guide wire was advanced into the left ventricle and the catheter system was advanced over the wire into the ventricle and positioned. the impella left ventricular assistive device was ramped up to p4, but flows were registering in the 0.2 range. after trouble shooting which included repositioning of the catheter the decision was made then to completely remove the impella device. after the device was removed, it was tested outside the body, where it was submerged in water and again ramped to p4, but the catheter registered 0.2 flow.
38410 | 2194391 | 2194391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2194391&pc=KFM | 07/13/2011 | 07/20/2011 | 07/20/2011 | 07/20/2011 | 07/20/2011 |  |  | 07/20/2011 |  |  |  |  | 07/20/2011 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS |  |  | IMPELLA 2.5 |  |  | KFM |  |  | * |  |  | ABIOMED |  |  | US |  | 215168515130406 |  | 004672 |  | Yes | [6 empty] | Invalid Data |  |  | User Facility report | [8 empty] | 70 YR |  |  | 76 |  |  | Malfunction | No | Yes | [26 empty] | NONE KNOWN | [22 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | Hemodynamic support was provided with the impella catheter at a performance level of "6" throughout the catheterization procedure. at the end of the procedure, there was noted to be leakage from the bacteriostatic filter on the flush line of the impella catheter. the low pressure alarm sounded. a leak was found between the pressure reservoir and the air/bacterial trap. the physician believes there was a fracture and breakage at the hub of the filter. the catheter had to be removed and a balloon pump was placed in the opposite groin. the patient remained stable.
38411 | 2115378 | 2115378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2115378&pc=KFM | 03/08/2011 | 04/12/2011 | 04/12/2011 | 05/27/2011 | 05/27/2011 |  |  | 05/27/2011 |  |  |  |  | 05/27/2011 |  | HEART PUMP |  |  | SYSTEM, IMPELLA 2.5 |  |  | KFM |  |  | 49845A |  |  | ABIOMED INC. |  |  | US |  | UNK |  | UNK |  | Yes | [6 empty] | Invalid Data |  |  | User Facility report | [8 empty] | 85 YR |  |  | 78 |  |  | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | The impella device was advanced through the right femoral sheath and into the left ventricle without difficulty. the impella device was initially set at p9 and then decreased to p8 and maintained at p8 throughout the procedure. while the impella device was in use, it alarmed. the alarm indicated that the motor would stop. the device continued to be used and a second alarm indicated that the motor had stopped. the impella device was removed and a second one was placed. the impella device was tried to start again after removal and it would not.
38412 | 2068449 | 2068449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=2068449&pc=KFM | 03/16/2011 | 04/05/2011 | 04/05/2011 | 04/05/2011 | 04/05/2011 |  |  | 04/05/2011 |  |  |  |  | 04/05/2011 |  | VENTRICULAR ASSIST DEVICE |  |  | IMPELLA 5.0 |  |  | KFM |  |  | * |  |  | ABIOMED |  |  | US |  | * |  | * |  | Device was returned to manufacturer | [6 empty] | Invalid Data |  |  | User Facility report | [8 empty] | 67 YR | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | An inspection revealed two severe kinks in the steering catheter. a review of the console data revealed that the catheter was most likely kinked from the start of the case, as high purge pressures were observed shortly after pump insertion. this review also confirmed that the console shut down the pump due to high motor current levels. the kinks in the steering catheter were the primary root cause for the pump stopping. these kinks restricted the purge flow in the inner purge lumen, causing the observed high purge pressures and allowing blood to enter the motor, resulting in an increase in motor current. the console eventually shut down the pump as a precaution once the motor current reached unsafe levels. to address this issue, abiomed has re-designed the impella 5.0/ld pumps and the next generation 5.0/ld incorporates a nitinol re-inforced outer catheter and improved inner purge lumen. the mfr response was a thorough evaluation and a report to our facility.
38413 | 1931454 | 1931454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1931454&pc=KFM | 12/04/2010 | 12/09/2010 | 12/09/2010 | 12/09/2010 | 12/09/2010 |  |  | 12/09/2010 |  |  |  |  | 12/09/2010 |  | LEFT VENTRICULAR ASSIST DEVICE - TEMPORARY |  |  | ABIOMED IMPELLA 5.0 |  |  | KFM |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | * |  | 004 683 |  | No | [6 empty] | UNKNOWN |  |  | User Facility report | [8 empty] | * | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | While changing the purge fluid tubing for the pt's impella vad, the clear connector that was connected to the yellow connector broke. this left the male end of the connector in the blue disk. this disk was removed in an attempt to attatch the dextrose fluid to the bacterial disk, but it also fell apart in my hand. this left the male end of the connector in the most proximal disk. i attached the fluid tubing into the most proximal disk as firmly as i could and called the doctor. the doctor was in the room and ordered the epinephrine increased and bivalirudin increased. the pt's vad flows have not dropped, the motor power remains unchanged, and the pt is hemodynamically stable. prior to switching the fluid tubing, i reviewed the process with the abiomed representative and read the procedure manual.
38414 | 1220648-2010-00009 | 1907993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1907993&pc=KFM | 10/08/2010 | 10/18/2010 |  | 11/17/2010 |  |  | 10/18/2010 | 11/17/2010 |  |  |  |  | 11/17/2010 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERC CARDIAC SUPPORT |  |  | KFM | [5 empty] | ABIOMED |  |  | GM | [5 empty] | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 63 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | The complainant has reported that a pt underwent a high risk percutaneous coronary intervention procedure involving an impella recover 2.5 percutaneous catheter device. the impella pump was implanted emergently in response to cardiac arrest. shortly thereafter, hematuria was reported along with elevated plasma free hemoglobin which raised suspicion for hemolysis caused by the impella device. the device and its position were re-evaluated using echo; however, the plasma free hemoglobin levels remained unaffected despite repositioning of the device. the pci procedure with impella support was described as successful and the device was explanted following approx 27 hrs of use. the pt's condition required hemodialysis treatment. no other complications were reported to have occurred. | [5 empty] | Conclusions: (device discarded, unable to f/u on device involved), (no conclusion can be drawn). the device was discarded by the user and therefore cannot be investigated. nonetheless a failure investigation of the available data is currently in progress; therefore, results are not yet available and no conclusion can be drawn at this time. the mfr will continue to investigate all reasonably obtainable sources of info and will provide results and updated conclusions in a supplemental mfr medwatch report. (b)(4).
38415 | 1220648-2010-00007 | 1786171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1786171&pc=KFM | 06/30/2010 | 06/30/2010 |  | 07/30/2010 |  |  | 06/30/2010 | 07/30/2010 |  |  | 04/01/2013 |  | 07/30/2010 |  |  |  |  | IMPELLA RECOVER LP 2.5, .018 GUIDEWIRE |  |  | KFM |  |  | .018 IN (0.47MM) |  |  | ABIOMED |  |  | GM |  | P1421061 |  | 0046-0043 |  | Yes |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Consumer | Health Professional | User facility | Company representation |  |  | 59 YR | [5 empty] | Malfunction | No | Yes | [26 empty] | IAB (MFR UNK) PRIOR TO IMPELLA | [22 empty] | Initial submission | [9 empty] | It was reported that an acute myocardial/cardiogenic shock pt, failing post coronary artery bypass, began to further decompensate despite being treated with a conventional intra-aortic balloon pump. the physician elected to exchange the iab for an impella 2.5 device. after successful insertion and placement of the impella pump, the physician reported resistance while attempting to remove the .018 guidewire through the left femoral access sheath. a fluoroscopic assessment revealed that the distal tip of the wire had become freely detached from the parent wire: it had become apposed to the distal end of the introducer sheath within the pt¿s left femoral artery. one attempt to retrieve the detached guidewire tip through the sheath was unsuccessful due to the location target; however, a contralateral approach was successful via the existing right femoral access site, used during earlier percutaneous coronary intervention, to retrieve the detached tip using a snare device and angioplasty balloon catheter in tandem. the impella therapy ensued thereafter. the pt sustained no injury or adverse effects as a result of this reported device malfunction. | [5 empty] | "o2 device evaluation anticipated." the device involved in the event was received on (b)(6). a failure investigation is currently in process. evaluation: conclusions: the failure investigation is currently in progress; therefore, results are not yet available and no conclusion can be drawn at this time. the mfr will continue to investigate all reasonably obtainable sources of info and will provide results and updated conclusions in a supplemental mfr medwatch report. internal ref: (b)(4).
38416 | 1220648-2010-00006 | 1773144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1773144&pc=KFM | 06/09/2010 | 06/09/2010 |  | 07/09/2010 |  |  | 06/09/2010 | 07/09/2010 |  |  | 07/24/2010 |  | 07/09/2010 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERCUTANEOUS CS |  |  | KFM |  |  | LP 2.5 |  |  | ABIOMED |  |  | GM |  | 0025251 |  | 004671 |  | Yes |  |  | Yes |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 82 YR | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [9 empty] | It was reported that an impella 2.5 device was attempted to be inserted via right femoral access (rfa) for additional support during an emergent high risk percutaneous coronary intervention (pci). after successfully inserting the device's introducer sheath, the physician was unable to pass the impella device from the iliac artery into the descending aorta as he encountered resistance due to vessel tortuosity. he attempted to advance the device four or five times without success. upon removal of the impella and introducer sheath together, the pigtail assembly detached from the impella catheter but remained over the guidewire in the iliac artery. the physician elected to allow the detached pigtail to safely reside undisturbed in the iliac artery while cardiac intervention via left femoral access lfa ensued. pci was completed on all target vessels with excellent results. following the completion of pci, the impella pigtail was successfully retrieved via rfa using a snare device. no other intervention or treatment was required and/or no injury or adverse effects resulted from this device incident. | [5 empty] | Follow-up with the physician and clinical personnel supporting the case, confirm that the impella procedure was aborted due to tortuous vascular anatomy. the diagnostic imaging indicated tortuous vasculature prior to the procedure and the physician conveyed his "understanding of the situation considering the intensity of the case and the risk he was taking trying to pass through his tortuous vessel". a thorough failure investigation was conducted which included a visual inspection of the returned product, mechanical testing of reserve samples, and review of all related manufacturing and complaint records. the device involved in the event was returned in two components; the pigtail assembly and the parent catheter up to and including all inlet cage struts. the struts of the pump's inflow cage were bent and twisted. the struts were disjoined from the pigtail teardrop clearly at the point of each weldment. no other abnormalities were identified. the manufacturing traveler did not show any manufacturing problems associated with the pump. studies on similar failure modes, including deformation, buckling, or breaking of the struts and the weld strength of the struts to the teardrop ball, have been carried out extensively on representative device samples and have demonstrated that the failure susceptibility is related to the strength of the inflow cage struts, among other contributing factors. the most obvious is excessive loads or strains imposed on the distal end of the device during handling and implantation. high axial and bending forces can occur when the device is loaded on the guidewire and inserted into the introducer sheath. axial and bending strains can occur as the device is advanced through the sheath and arteries during placement, especially in patients with tortuous vasculature. a fourteen month complaint history review reveals a total of (b)(4) cases involving detachment of the pigtail tip due to deformation of struts during clinical use. this represents an extremely low failure rate at (b)(4) of all clinical use. in order to prevent future inadvertent damage as a result of excessive force and unusual loads applied to the device, corrective action has been implemented and includes a redesign of the device to improve its buckling resistance (b)(4). neither of the two complaint events resulted in injury to the patient. the overall failure rate is determined to be within acceptable limits: no further corrective actions are warranted at this time. this event will be maintained and monitored within our quality system as part of our design control and continuous improvement processes. (b)(4).
38417 | 1752441 | 1752441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1752441&pc=KFM | 06/17/2010 | 06/24/2010 | 06/24/2010 | 06/24/2010 | 06/24/2010 |  |  | 06/24/2010 | 06/24/2010 |  |  |  | 06/24/2010 |  | CATHETER,PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYP | CONSOLE, PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYP |  | IMPELLA 2.5 | CONSOLE FOR IMPELLA 2.5 |  | KFM | KFM |  | IMPELLA 2.5 | * |  | ABIOMED, INC | ABIOMED |  | US | US | * | * | * | * | Device was returned to manufacturer | Device was returned to manufacturer | [5 empty] | Physician | Invalid Data |  | User Facility report | [8 empty] | 57 YR | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y |  | * | * | The impella device was placed while patient was in the cath lab. patient was returned to an icu bed, still in cath lab. the impella console began to alarm and showed "monitor motor current, restart pump." it was the same problem that we reported 10 days ago on another impella device. after the alarm showed, the staff continued to troubleshoot to restart the pump but were unsuccessful in doing so. the impella did not look like it was positioned correctly by viewing on echo, but the console was ok with the position. the decision was then made to pull back the impella out of the ventricle and treat the patient medically. the impella was later removed. the console was pulled out of service as well.
38418 | 1749767 | 1749767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1749767&pc=KFM | 05/31/2010 | 06/14/2010 | 06/14/2010 | 06/14/2010 | 06/14/2010 |  |  | 06/14/2010 |  |  |  |  | 06/14/2010 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA 2.5 |  |  | KFM |  |  | IMPELLA 2.5 |  |  | ABIOMED |  |  | US |  | * |  | * |  | Yes | [6 empty] | Physician |  |  | User Facility report | [8 empty] | 55 YR | [6 empty] | No | Not Available | [49 empty] | Initial submission | [5 empty] | Y |  | * |  | A critically ill patient was admitted directly to the heart cath lab because of severe cardiogenic shock with an ejection fraction of 10-15%. an impella 2.5 was placed, but would not function. the motor current became so high that the pump kept shutting off. unable to keep the pump running. each time staff tried to restart it, it would immediately shut down. after being unsuccessful in trouble shooting, the abiomed representative was called. several other things were tried. the blue cable was unplugged and then replugged and an attempt was made to restart the machine. the blue cable was also exchanged for a different cable. all attempts were unsuccessful and the abiomed rep said that it was a defective pump. the pump was turned off and the doctor removed the impella. there is a machine and a braun vista pump that connect externally to the catheter to run the pump, but that equipment was functioning properly. there are a couple of ideas for pump malfunction. one is that the pump was defective and the other was that the patient had a clot in his left ventricle. the device was defective. the manufacturer initially stated that the device was defective.
38419 | 1220648-2010-00005 | 1710590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1710590&pc=KFM | 04/20/2010 | 04/20/2010 |  | 06/03/2010 |  |  | 05/04/2010 | 06/03/2010 |  |  | 01/11/2012 |  | 06/03/2010 |  |  |  |  | IMPELLA RECOVER LP 2.5 CARDIAC SUPPORT SYSTEM |  |  | KFM |  |  | LP2.5 |  |  | ABIOMED, INC. |  |  | GM |  | 39297 |  |  |  | Yes |  |  | Yes |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | [20 empty] | Other | Required Intervention | [27 empty] | Initial submission | [5 empty] | N |  |  |  | It was reported that a patient, admitted with acute posterior infarction, was supported using an impella 2.5 device for approx one day subsequent to post cardiac catheter intervention. while attempting to pull back the device at the end of the procedure, the user incorrectly pulled the white sheath which separated the repositioning unit and resulted in bleeding. the bleeding was controlled when the repositioning unit was reassembled correctly by the user. the pt's hemodynamics remained acceptable, however, a hemotoma had formed at the groin site and the pt was administered a blood transfusion. two days after removal of the impella device, the patient experienced right leg ischemia which required thrombectomy of the popliteal artery. despite this intervention, the limb remained ischemic and signs of necrosis had developed at the right foot. on the same day, the patient had also reportedly developed signs of systemic inflammatory response syndrome and septic shock which were confirmed to have been attributed to preexisting pneumonia and not related to the impella device. the pt ultimately expired five days post impella explant due to multi organ failure caused by septic shock, due to the unrelated pneumonia. the physician has reported that no device malfunction is suspected to be related to the patient's death. | [5 empty] | The device involved in the event was returned to the mfr and a complete failure analysis was conducted. a review of the lot process traveler showed no mfg problems with this device. a batch analysis of the components of the repositioning unit was performed: all parts (including lots of tapered wound closures and repositioning units) successfully passed incoming inspections. f/u with the physician revealed that a fellow assisting in the procedure first damaged the repositioning sheath after removing the 13 french peel-away sheath when he disconnected the "yellow from the white part of the repositioning sheath". this inadvertent disconnection required "some manipulation" in order to reattach the repositioning sheath correctly. additionally, the physician reported that an attempt was made to pull the impella catheter back through the yellow sheath. the repositioning unit of the returned pump was observed to have severe damage, including multiple cracks in the tapered wound closure, caused by the user's attempt to retract the pump through the repositioning sheath. separation of components at the yellow hub on the repositioning sheath can only be achieved if excess force is applied to the collar while the hub of the sheath is restrained at the same time. this has been simulated using several representative repositioning units and is demonstrated only during abnormal handling of the device. nevertheless, if this type of separation does occur, the junction may be manually reconnected in order to maintain successful hemostasis. the root cause is concluded to be a separation of the repositioning unit due to improper handling by the user. the clinic staff at the user facility was retrained with the proper handling of the impella device. a twelve month complaint history review was conducted and revealed that a total of 2 occurrences exist for this failure mode (inadvertent yellow hub disconnection by the user); however, this is the first occurrence that involves adverse impact to the patient. the rate of occurrence and associated risks are currently within acceptable tolerances; therefore, the corrective actions in place are considered acceptable and no further action is warranted at this time.
38420 | 1707243 | 1707243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1707243&pc=KFM | 05/19/2010 | 05/21/2010 | 05/21/2010 | 05/21/2010 | 05/21/2010 |  |  | 05/21/2010 |  |  |  |  | 05/21/2010 |  | PUMP, BLOOD, CARDIOPULMONARY BYPASS, NON-ROLLER TYPE |  |  | IMPELLA 2.5 |  |  | KFM |  |  | * |  |  | ABIOMED, INC |  |  | US |  | * |  | * |  | Device was returned to manufacturer | [6 empty] | Invalid Data |  |  | User Facility report | [8 empty] | 78 YR | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | The placement signal failed. the catheter was removed and replaced with a new catheter. no harm to the patient. we have asked the manufacturer to share the results of the analysis with our facility.
38421 | 1220648-2010-00004 | 1678258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1678258&pc=KFM | 12/11/2009 | 12/13/2009 |  | 05/07/2010 |  |  | 12/13/2009 | 05/07/2010 |  |  | 08/21/2011 |  | 05/07/2010 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERCUTANEOUS CS |  |  | KFM | [5 empty] | ABIOMED |  |  | GM |  | 0035654 |  |  |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Consumer | Health Professional | User facility | Company representation |  |  | 63 YR |  |  | 100 |  |  | Injury | Yes | No | [20 empty] | Other | Required Intervention | [27 empty] | Initial submission | [9 empty] | It was reported that a patient underwent high risk percutaneous coronary intervention, using an impella 2.5 pump as a partial cardiac assist device. the impella, including the introducer sheath and catheter, were placed without issue and the intervention lasted approximately two hours during which three target coronary vessels were treated under the planned approach. following successful pci, the physician elected to continue impella support by moving the patient to the icu for continued therapy. in order to prepare the patient for transfer to the icu, the peel-away sheath was removed and the repositioning sheath was advanced into the arteriotomy but without success. the physician was unable to occlude the arteriotomy using the impella's repositioning sheath and the patient developed a large hematoma at the lfa site which was confirmed by a ct scan. a femostop was attempted to gain hemostasis, but it was unsuccessful due to poor placement. manual pressure was applied to the groin site, the impella was discontinued, and the patient was taken to the operating room where the lfa was repaired after locating a "clean looking hole" in the artery. the patient received six units of packed red blood cells in response to the blood loss related to the reported hematoma. the repair to the lfa was reportedly noninvasive and achieved with a few sutures. | [5 empty] | Evaluation: analysis of labeling performed, (the device was discarded by the user). results - bleeding at access site. additional manufacturer narrative: the device involved in the reported event was discarded by the user; therefore, a complete investigation of the suspect device cannot be conducted by the manufacturer and a conclusive root cause cannot be determined. however, a comprehensive evaluation of the device's labeling, device history record, and complaint history review was conducted in concert with follow-up interviews with the user for additional information related to the event. the event was reported for a hematoma that developed following the 13 fr peel-away sheath removal and repositioning sheath advancement. follow-up interviews with the site, reveal that a fellow participating in the case had broken away the 13 fr sheath prior to removing it from the patient's vessel which is the most likely root cause of the damage and resultant bleeding at the arteriotomy. separate from this peel-away sheath issue, the attending physician reported that he believes a portion of the re-positioning sheath did not reach into the femoral artery and that it did not have an "area" large enough to occlude vessel: that the re-positioning sheath overall was simply not long enough by design. nevertheless, the physician did report that "the [impella] device did what it had to do to support the patient during its use." the ifu(0046-9005 rev. b), mentions on page 5.16 "remove the 13 fr peel-away introducer completely from the artery over the catheter shaft" and states "grasp the two "wings" and bend back until the valve assembly comes apart. continue to peel the two wings until the introducer is completely separated from the catheter shaft". the clinical staff was trained on the proper removal technique of the introducer during the case and in follow on training sessions. lastly, bleeding at the groin site is a known, inherent risk of the procedure. a review of the complaint database shows this to be the only occurrence of this event type within the previous 12 consecutive months. additionally, a review of the pump device history record showed that is was manufactured and tested in accordance with all specifications. all known and reasonably obtainable information has been evaluated and reported herein. no further action is warranted at this time.
38422 | 1220648-2010-00003 | 1656853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1656853&pc=KFM | 03/11/2010 | 03/11/2010 |  | 04/09/2010 |  |  | 03/11/2010 | 04/09/2010 |  |  | 09/04/2011 |  | 04/09/2010 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERCUTANEOUS CS |  |  | KFM |  |  | LP 2.5 |  |  | ABIOMED |  |  | GM |  | 36232 |  |  |  | Yes |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Consumer | Health Professional | User facility | Company representation |  |  | 81 YR | [5 empty] | Injury | Yes | No | [20 empty] | Other | Required Intervention |  |  |  |  | IABP (UNK MFR) | BALLOON CATHETER (UNK MFR) | CARDIAC STENT (UNK MFR) | [20 empty] | Initial submission | [5 empty] | N |  |  |  | Reportedly, a patient, underwent cardiac catheterization involving an impella 2.5 partial assist device. an iabp, supporting the patient prior to the pci intervention, was removed and the impella system was placed in the left ventricle. impella support was initiated by increasing the flow rate incrementally. the rate was increased until a suction event was detected by the device and provided an alarm to the user. the flow was decreased in response to the alarm, but the device was removed as the patient developed severe intractable hypotension and bradycardia. the patient stabilized and the physician continued with successful coronary intervention including stent deployment and balloon angioplasty to the lmca. at the end of the procedure, normal coronary perfusion was noted; however, the patient remained severely hypotensive. an echo was performed which revealed pericardial effusion and tamponade physiology. a pericardial catheter was placed to decompress the effusion which resulted in instantaneous recovery of the patient's hemodynamic stability. the patient was then transferred to the operating room where a small perforation in the ventricle was identified and repaired. the patient's condition was reported to be stable at the end of the event. | [5 empty] | Evaluation, conclusions: the investigation is currently in progress. the device has been received by the manufacturer and a failure investigation is in process, but is not yet completed; therefore, no conclusions can be drawn at this time. the manufacturer will continue to investigate all reasonably obtainable sources of information and will provide results and conclusions in a supplemental manufacturer medwatch report. (b) (4).
38423 | 1220648-2010-00002 | 1611146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1611146&pc=KFM | 01/17/2010 | 01/18/2010 |  | 02/17/2010 |  |  | 01/18/2010 | 02/17/2010 |  |  | 10/13/2011 |  | 02/17/2010 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERCUTANEOUS CS |  |  | KFM | [5 empty] | ABIOMED, INC. |  |  | GM |  | 36650 |  |  |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Consumer | Health Professional | User facility | Company representation |  |  | 34 YR |  |  | 120 |  |  | Death | Yes | No | [20 empty] | Death | Other | Required Intervention |  |  |  | IABP (DATASCOPE) | [22 empty] | Initial submission | [9 empty] | It was reported that an acute myocardial infarct and cardiogenic shock patient treated with an iabp (not manufactured by abiomed) developed right leg ischemia requiring extensive intervention to repair. the intervention lasted approximately 5-8 hours during which no other cardiac support was administered. following the attempted repair, the physician elected to continue treatment using an impella recover lp 2.5 cardiac assist device via left groin access. the impella implant procedure was described to have demonstrated a slight "ooze" from the access site which had resolved once the repositioning unit was fully advanced into the arteriotomy and secured at the site. the patient's hemodynamics improved while the impella was functioning the impella support continued for over an hour without incident. a vascular surgeon then arrived for specialized consult on the right leg injury resulting from the earlier iabp incident. upon examination, the vascular surgeon ordered emergent surgical repair of the right leg: the patient was immediately transferred to the or. upon being moved onto the or table, the patient rapidly developed a large hematoma localized to the impella insertion site as no pressure was being applied to the groin during transfer. the surgeon removed the device, and the patient ultimately expired 20-30 minutes later. | [5 empty] | Conclusions: (iabp involved in incident is not an abiomed device). the user facility has discarded the device involved in the event; therefore, a failure investigation cannot be completed on the device involved in the event and a conclusion cannot be drawn. the manufacturer will continue to investigate all other reasonably obtainable sources of information and will provide results and conclusions in a supplemental manufacturer medwatch report. internal ref: (b) (4).
38424 | 1220648-2010-00001 | 1604521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1604521&pc=KFM | 01/12/2010 | 01/12/2010 |  | 02/11/2010 |  |  | 01/12/2010 | 02/11/2010 |  |  | 08/05/2011 |  | 02/11/2010 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERCUTANEOUS CS |  |  | KFM |  |  | LP 2.5 |  |  | ABIOMED, INC. |  |  | GM |  | 35222 |  |  |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Consumer | Health Professional | User facility | Company representation |  |  | 84 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [5 empty] | CARDIAC ECTOPY | "TO CALM HER DOWN" PRIOR TO CASE. | CARDENE FOR SPASMS | [20 empty] | Initial submission | [9 empty] | It was reported that a patient underwent a non-emergent, high risk pci procedure involving additional cardiac support using an impella 2.5 partial assist device. the patient was administered nicardipine hydrochloride (cardene) prior to the intervention in order to manage hypertension, spasms, and cardiac ectopy throughout the procedure which induced low blood pressure observations throughout the case. the physician reported that while inserting the impella device, the guidewire was advanced too deeply into the ventricle because the technician, handling the guidewire in tandem, had lost control of it. the procedure continued and successful pci treatment with impella support was accomplished. all devices were removed without suspect; however, the blood pressure remained low. in response, an echocardiogram and a ventriculogram were ordered which revealed a small dissection in the left ventricle in the lateral aspect of the apex. this was believed to have been caused by a perforation of the 0.018" guidewire used during placement of the impella device. the perforation was described to be "very small" and "would have likely healed without intervention"; however, the acting surgeon elected to perform a full exploratory via sternotomy in order to rule out any other possible injuries and/or sources of bleeding. blood was drained from the pericardial sac, the perforation was closed with a single patch on the pericardium, and no other injuries or adverse effects were identified or experienced. the patient has made a full recovery and was expected to be discharged home. | [5 empty] | The device was discarded by the user. the device involved in the reported event was discarded by the user. therefore, it is not available for analysis. nevertheless, a thorough investigation was conducted with all reasonably known information and information that could be obtained through clinical follow-up from the user. the size of the perforation was too small to have been caused by a 0.35" guidewire, diagnostic catheter or the impella device. it was, however, consistent with a 0.18" guidewire. the 0.018 guidewire used in the case was a boston scientific platinum plus. a review of records shows the suspect lot had passed all inspection criteria. no other problem was with this lot of wires is known to have occurred. the manufacturing records for the pump were also examined and no manufacturing problems were identified. follow-up information form the physician confirms that the guidewire involved in the case was inspected by the physician and the clinical representative present during the case; no defects were observed upon their visual inspection after withdrawal from the patient. the physician believes the technician, operating the guidewire in tandem, had lost control of the wire while the physician was advancing the impella through the sheath. as a result, the guidewire was also pushed along with the impella catheter through the sheath rather than allowing the device to track over it. because of this, the wire was advanced too far into the ventricle which caused the guidewire to prolapse. this was later confirmed on fluoroscopic images which were reviewed in retrospect. consequently, the over-inserted, prolapsed guidewire is what the physician believes to be related to the actual cause of the perforation. he does not suspect a device, component, or material related problem. overall, the physician believes the problem was isolated to operator error rather than a device related issue. this was an apparent case of operator error when the technician restraining the guidewire lost control of the wire as the physician advanced the impella through the sheath. a comprehensive review of the abiomed's complaint handling database was conducted to identify similar events; no previous incidents of vascular or ventricular injury due to the guidewire were found. no corrective actions are recommended at this time. abiomed will continue to monitor the complaint database for similar, future occurrences. (b) (4).
38425 | 1603554 | 1603554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1603554&pc=KFM | 01/29/2010 | 02/05/2010 | 02/05/2010 | 02/05/2010 | 02/05/2010 |  |  | 02/05/2010 |  |  |  |  | 02/05/2010 |  | CATHETER, INTRAVASCULAR, DIAGNOSTIC |  |  | IMPELLA |  |  | KFM |  |  | * |  |  | ABIOMED, INC. |  |  | US |  | * |  | * |  | Yes | [6 empty] | Invalid Data |  |  | User Facility report | [8 empty] | 51 YR | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | An impella 2.5 left ventricular device catheter pump was inserted through a hemashield graft, sewed to the right axillary artery. the placement sensor signal failed to work. the use of echo verification of catheter position was required. approximately five (5) days later, this catheter was removed and replaced with bivad pump due to the patient's hemodynamic status.facility has asked the manufacturer to share with us information relating to their inspection of the device.
38426 | 1220648-2009-00010 | 1471061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1471061&pc=KFM | 08/08/2009 | 08/08/2009 |  | 09/10/2009 |  |  | 08/08/2009 | 09/10/2009 |  |  | 04/08/2010 |  | 09/10/2009 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERCUTANEOUS CS |  |  | KFM |  |  | LP 2.5 |  |  | ABIOMED |  |  | GM |  | 0021119 |  | 002137 |  | Yes |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Foreign | Consumer | Health Professional | User facility | Company representation |  | 37 YR |  |  | 75 |  |  | Injury | Yes | No | [20 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | The complainant reported that after three days of successful support by an impella 2.5 device, purge pressure complications were identified by impella controller alarms which the user did not resolve. the manufacturer recommended removal and 3 to 4 hours later, the physician decided to discontinue impella support and removed the pump from the patient. while withdrawing the catheter, the physician reported that the pump head became "caught" at the bifurcation of the femoral and iliac arteries. when the device was retracted from the groin, the distal tip of the catheter remained at the site of impediment while the patient catheter exited the body. a surgical cutdown was then performed in order to preserve the patient's vasculature while successfully removing the tip of the device. no other intervention or treatment was required. no other injury or adverse effects were involved during or as a result of this incident. the physician reported that he was otherwise very pleased with the performance of the device and impressed by its overall therapeutic effectiveness in the patient. he felt the device saved the patient's life by reversing the heart failure from acute to chronic. | [5 empty] | Device evaluation anticipated but not yet begun, device returned from foreign site for evaluation. the device was returned to the manufacturer from a foreign site and a failure investigation is currently in progress; therefore, a root cause for the reported event is not yet available. results of failure investigation and any other pertinent information obtained will be provided in a supplemental medwatch report.
38427 | 1220648-2009-00009 | 1425338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1425338&pc=KFM | 04/16/2009 | 04/16/2009 |  | 07/29/2009 |  |  | 04/16/2009 | 07/29/2009 |  |  | 12/12/2010 |  | 07/29/2009 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERCUTANEOUS CS |  |  | KFM |  |  | LP 2.5 |  |  | ABIOMED |  |  | GM |  | 0029364 |  | 004671 |  | Yes |  |  | Yes |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | Company representation |  |  |  |  | 55 YR | [5 empty] | Injury | Yes | No | [20 empty] | Required Intervention | [5 empty] | NEO | PRIMACOR ALONG WITH A DATASCOPE IAB IN 1:1 ASSIST | LEVOPHED | RATIO | [19 empty] | Initial submission | [9 empty] | It was reported that a female patient underwent a 6 hr coronary artery bypass graft procedure with subsequent cardiac tamponade. the patient was returned to the operating room for a clean out during which she was noted to be struggling hemodynamically. in response, the physician elected to try an impella 2.5 pump to achieve further support of the patient. while attempting to introduce the impella via the groin site, the physician reported resistance due to vessel calcification and tortuosity. repeated attempts to insert the device were unsuccessful and finally resulted in deformation of the device so a second impella 2.5 pump was prepared in order to resume implant attempts. after the insertion site was traversed with the second device the physician was then unable to overcome further advancement and placement difficulties due to the patient's condition. despite the difficulty, the pump was forcibly maneuvered in the patient's aorta without success until ultimately the pigtail detached from the parent device. the detached pigtail was located on fluoroscopy and after discussing several remedial actions, the physician successfully retrieved the component altogether through an abdominal incision. no other device related complications were reported. according to the physician, neither of the two impella devices involved in the procedure was suspect for causing or contributing to the death of the patient. | [5 empty] | Follow-up with individuals present during the event reveal that the surgeon used aggressive methods during insertion attempts. the first impella device attempted to be implanted and the guidewire being used were both bent upon insertion due to the handling of the device in the presence of severe calcification and tortuous vasculature. clinical representatives attending the procedure advised the surgeon on proper techniques and reiterated caution against advancing the device with force prior to use of the second device. the actual device involved in the event was returned in two pieces; the handle and catheter up to the pump motor was in one piece and the pump, cannula, inflow cage and pigtail comprised the other. a thorough failure investigation was conducted which included visual inspection of the returned product, mechanical testing of reserve samples, and review of all related manufacturing and complaint records. a visual examination of the returned catheter was performed. the tip assembly was separated from the parent device at the junction between the pump assembly and the catheter. the actual separation was at the distal side of the reinforcing band through the epo-tek joint. the face of the fractured surface was planar with the end of the catheter for approximately 60% of the cross section (starting on the pressure lumen side) and then extends at a 45 degree angle to the other side. the purge lumen tube and motor wires were pulled out of the joint. the shape of the motor side of the break reflects the fracture on the catheter side. the catheter had multiple kinks starting where the epo-tek joint inside the catheter ends and extending back to the 11 cm mark. kinks at the 1 and 7 cm marks were severe with less pronounced kinks at 5 and 11 cm. the kink at 1 cm consists of multiple deformations on opposite sides of the catheter. lateral scrapes or gouges in the catheter just behind the reinforcing band were also noted. the motor housing, pump outflow, and impeller appeared normal. the cage was relatively undamaged. the pigtail appears to have been bent slightly out of line but otherwise undamaged. in order to re-create the failure, representative devices were subjected to tensile challenge which successfully reproduced failures that resembled those of the clinical device. all test devices required forces exceeding the specified tolerance in order to induce a bond failure and all demonstrated kinking of the catheter as a result of the stress applied. the kinking observed in the suspect catheter was much more pronounced than all test catheters including those that required the highest exertion forces (forces exceeding 170% of the tolerance limit) to induce failure. the kinks on the clinical failure indicate that significant pulling force was applied to the catheter at an angle relative to the pump body itself. the investigation concludes that the failure was the direct result of aggressive manipulation of the impella catheter by the user as evidenced by the severe kinking of the catheter proximal to the break. a review of the lot process traveler showed no manufacturing problems with this device. a review of the complaint history revealed no other similar events involving separation of the pigtail at the distal side of the reinforcing band through the epo-tek joint. this is the first report of this failure mode. in order to prevent future inadvertent damage as a result of excessive force and unusual loads applied to the device, we have initiated a project to evaluate the design of the device in order to maximize resilience of the tip bond, inflow cage, and catheter body. this event will be maintained and monitored within our quality system as part of our design control and continuous improvement processes. no further corrective actions are warranted at this time.
38428 | 1220648-2009-00008 | 1408663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1408663&pc=KFM | 01/27/2009 | 01/30/2009 |  | 06/26/2009 |  |  | 01/30/2009 | 06/26/2009 |  |  | 07/18/2010 |  | 06/26/2009 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERC CARDIAC SUPPORT |  |  | KFM | [5 empty] | ABIOMED |  |  | GM |  | 0025066 |  |  |  | Yes |  |  | Yes |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | [20 empty] | Hospitalization | [28 empty] | Initial submission | [9 empty] | The complainant has reported that following 13 days of hospitalization for chronic diabetes and heart failure, a patient underwent pci treatment, during which partial cardiac assist using an impella 2.5, device was employed. the planned duration of the use of impella was 24 hours following the completion of the pci procedure. after approximately 8 hours following the initiation of impella support, the patient's plasma free hemoglobin was noted to be remarkably elevated. the physician elected to continue the use of the impella for the full 24 hours of support as originally planned prior to explanting the device. reportedly, the patient subsequently developed acute renal failure for which he was treated with dialysis. the patient's plasma free hemoglobin concentrations had returned to baseline within 48 hours of explanting the impella device. it was later discovered that the device was mispositioned by the user throughout the duration of impella support. no other complications were reported to have occurred in relation to the use of the impella device. | [5 empty] | (evaluation codes): results: device incorrectly positioned, hemolysis is not unanticipated, no device failure. conclusions: device involved in event within specifications-no blood damage observed, device incorrectly positioned, incorrect position contributed to increased hemolysis, relationship of hemolysis to acute renal failure unknown- see narrative. a thorough failure investigation of the event and suspect device was conducted by the manufacturer. the evaluation involved visual, electrical, and functional examination of the returned pump; review of device manufacturing history records, console log data from the procedure, and patient's clinical data; and a follow-up interview with the physician. a review of the device history records and complaint data was performed. the affected pump #41053 was manufactured within lot #0025066. of all units manufactured in this lot, 3 were not released for clinical use for unrelated reasons. no anomalies were revealed during the manufacture of this device. no other similar complaints have been received on any other device from the same lot. a visual inspection of the device was conducted using a borescope. no conspicuous details could be found. functional testing of the pump was conducted in vitro multiple times, using a continuous flow blood loop designed to mimic in vivo flow characteristics. several impella pumps from reserve inventory were tested in parallel as control devices. additionally, the modified index of hemolysis (mih) was measured serially throughout the duration of this testing. the results of the in vitro testing were unremarkable, as the device demonstrated mih values of less than 50% of the acceptable limit. the pump was run without any technical issue and worked as intended throughout testing. no deviation from any performance and/or hemolysis specification for the device was identified. the data recorded by the console throughout the clinical procedure in question was obtained and examined. the logged data shows the aortic pressure reading on the console showed high systolic pressures. within the first hour of impella support, the system had detected a decrease in motor current pulsatility and had alerted the user with multiple "pump position wrong" alarms accordingly. the clinical event details, including patient history and hematology parameters during the procedure, were reviewed. the patient had been hospitalized for 13 days prior to the procedure involved in the complaint and had undergone several diagnostic procedures including contrast enhanced computerized tomography. the dates and details of procedures within the timeframe leading up to the impella procedure were not provided, although, the physician stated that the patient presented a "very ill status". in fact, the patient required several blood transfusions prior to and during the use of the impella device. at the time of the use of the impella, the clinical lab values demonstrated a dramatic increase in the patient's plasma free hemoglobin (pfhb) of as much as 560 times the initial value as well as a reduction of hematocrit (hct) by approximately 25%. follow-up interviews with the physician and attending staff were conducted to discuss details of the clinical event. the pump position alarms identified and recorded by console data were specifically addressed. the physician indicated that he believed the device to have been in proper position and offered fluoroscopic images of the implanted device for review. in contrast, all angiographic images revealed that the pump was too deep within the ventricle, which rendered the outlet area to be at the level of the aortic valve, under the valve leaflets. the physician had admittedly misinterpreted the geometry and location of the device under fluoroscopy at the time of placement, as he misidentified the 12 french section of the device to be the area of the outflow windows. effects of a pump malpositioned in this way were discussed with the physician: the relative position of the device outflow could have resulted in high shear and therefore, could have contributed to an increased rate of hemolysis. the influence of the blood transfusions prior to and during the use of the impella is unknown. the relationship between the hemolysis and the acute renal failure in this particular patient remains speculative because of the patient's prior conditions and the multiple uses of contrast used prior to and during the impella procedure. furthermore, the physician also reported that upon reviewing the case retrospectively, it was concluded that the patient's hemopathy was likely treatable prior to progressing to acute renal failure had in-house nephrology been consulted. the results of the investigation conclude that a technical failure or a design issue, e.g. sharp edges or sucked particles, was not present and therefore, no malfunction of the device has occurred. clinical event details and follow-up information obtained directly from the physician lead to the conclusion that a malpositioned implant contributed to the change in patient condition. the relationship of the hemolysis to the reported acute renal failure cannot be substantiated; therefore, a conclusive root cause of this cannot be determined. all impella pumps are inspected by an endoscope as final production steps before they are packed to ensure they are free of sharp edges or fibers, which could present a risk for increased hemolysis. as corrective and preventive measures, further internal testing will be conducted to investigate the effect of various positions of the pump at the aortic valve on blood shear rates and hemolysis. additionally, the customer will be re-trained on correct positioning of the pump while using fluoroscopic guidance. no further corrective action is warranted at this time.
38429 | 1220648-2009-00006 | 1424687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1424687&pc=KFM | 10/01/2008 | 10/01/2008 |  | 04/03/2009 |  |  | 10/01/2008 | 04/03/2009 |  |  | 08/01/2010 |  | 04/03/2009 |  |  |  |  | IMPELLA RECOVER LP2.5 CARDIAC SUPPORT SYSTEM |  |  | KFM |  |  | 2.5 |  |  | ABIOMED EUROPE |  |  | GM |  | 0025589 |  |  |  | Yes |  |  | Yes |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Consumer | Health Professional | User facility | Company representation |  |  | 89 YR | [5 empty] | Malfunction | No | Yes | [26 empty] | CORONARY STENTS, UNKNOWN MANUFACTURER | [22 empty] | Initial submission | [9 empty] | It was reported that while a patient underwent percutaneous coronary intervention involving the placement of multiple stents (manufacture unknown), an impella 2.5 device was used for additional support throughout the intervention. the impella device was successfully inserted and therapy was initiated unremarkably. following several hours of therapy, it was decided to transfer the patient to the intensive care unit to monitor the scheduled removal of the impella device. in order to prepare the patient for transfer, the peel-away introducer sheath of the impella required removal. when the physician attempted to retract the introducer sheath by grasping the hub, the hub assembly "snapped apart" from the body of the sheath. it was then difficult to remove the resultant sheath since bleeding from the site of detachment prevented adequate grip on the peel-away wings. as a result, approximately 100cc was estimated to have bled from the site. the physician was able to remove the sheath component by peeling it apart as intended, but the simultaneous retraction from the groin site. the repositioning unit of the impella catheter was adjusted to achieve hemostasis at the insertion site, by design. no treatment was required for the bleeding event and no other impella related difficulties were experienced throughout the procedure. the patient sustained no adverse effects or injury as a result of the reported event. | [5 empty] | Method: microscopic evaluation, ftir chemical analysis of valve bonds. results: silicone contamination, silicone cross linked to bond material. the actual device involved in the event was returned to the manufacturer and a thorough failure investigation was conducted. a visual inspection confirmed that the hemostatic valve was detached from the parent peel-away sheath. no adhesive remained on the wings of the sheath. the peel-away sheath was submitted for third-party analysis, which included microscopic and chemical evaluations. the device was examined under a stereomicroscope at up to a magnification of 100x and fragments of a thin film were noted on the inner diameter surface of wing components. these thin layer fragments were collected from the two halves of the wings and were analyzed by fourier transform infrared spectroscopy (ftir). the results concluded that silicone contamination of the bonding surface was present. the silicone was existed as a moiety of the whole contaminant (thin film) and was cross-linked to epoxy resins that are used prior to the bonding process. the final lubrication step of the sheath manufacturing process involves silicone; however, it is conducted in a separate room than the bonding operation. no evidence of silicone contamination was detected in any manufacturing areas involved in the bonding operation. the sheath vendor was notified, but was ruled out as the source of the contamination. ultimately, the actual means of contamination could not be determined. the device history record was also reviewed for the reported introducer kit lot, which revealed no material issues during production. a total of eight (8) samples were pulled from introducer lot 0025589 and subjected to a 90 degrees pull test. all samples passed the minimum 15 n tensile specification. data was comparable to other lots that were subjected to the same testing. bleeding is an inherent and labeled risk involved with the use of impella device and specifically the introducer sheath. these risks, as well as proper handling of the introducer sheath, are described in the product's instructions for use. in addition, additional communication has been distributed in order to reinforce proper handling of the introducer sheath, which includes precautions such as "holding the back of the sheath, locking mechanism, hemostatic valve, or sheath body could be damaging to the product and/or may cause bleeding". the root cause of the reported failure has been determined to be silicone contamination of the bond between the bayonet of the hemostatic valve and the inside diameter of the wings of the introducer assembly. this is concluded to be an isolated occurrence as all other reserve samples from the same manufacturing lot successfully passed additional tensile evaluation. corrective and preventative actions include improvements to the bonding process to ensure stringent manufacturing controls; however, no further corrective action is warranted at this time.
38430 | 1220648-2009-00003 | 1329391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1329391&pc=KFM | 11/04/2008 | 11/04/2008 |  | 02/26/2009 |  |  | 11/04/2008 | 02/26/2009 |  |  |  |  | 02/26/2009 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERCUTANEOUS CS |  |  | KFM | [5 empty] | ABIOMED |  |  | GM |  | 0024693 |  | 004671 |  | Yes |  |  | Yes |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | 77 YR |  |  | 112 |  |  | Injury | Yes | Yes | [20 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | It was reported that an impella 2.5 device was attempted to be used as a peripheral transition during the completion of an open coronary artery bypass surgical procedure. the physician experienced insertion difficulties which resulted in damage to the device and consequently required subsequent intervention through the existing surgical site. the physician was unable to advance the guide catheter through the aortic valve. transesophageal echocardiogram (tee) confirmed the valve to be closed; therefore, the physician excluded the use of the guide catheter and advanced the guide wire and the impella through the valve and into the target site. transition from bypass was begun successfully. the impella procedure was continued for approximately 12 minutes at which point the device was ejected into the aorta by the native ventricle. the physician attempted to re-advance the device into the ventricle without the use of the guide wire by pushing and twisting the distal pigtail of the impella approximately 50 to 80 times against the aortic valve leaflets. direct manual manipulation to control the device through the existing open surgical site was also employed. ultimately, the distal end of the device was driven across the aortic valve which resulted in the detachment of the pigtail into the ventricle. fluoroscopy and tee were used to verify the location and nature of the detached subcomponent. the parent catheter of the device was found to have returned to the aorta while the pigtail and teardrop were rotated within the ventricle approximately 180 degrees. because direct access was possible through the existing sternotomy from the bypass procedure, a puncture site was created in the patient's aorta and secured with purse-string sutures for successful retrieval of the complete pigtail assembly. the physician elected to close the sternum prior to removing the parent catheter of the impella device which was subsequently removed through the femoral access site. a cut-down was performed in the groin without dissection of the artery as a precaution to avoid injury by any potential deformations of the resultant distal end of the device. the sheath and the pump were withdrawn together. no ill effects or permanent injury were reported to have occurred as a result of this event. no fragments or device components remained in the body at the end of the procedure. the procedure was completed without further impact to the patient. | [5 empty] | Evaluation: actual device was evaluated, device from reserve sample evaluated, mechanical tests, dynamic - fatigue test, visual examination. results: misapplication/misuse of device, caused by operational context, component/ subassembly failure; prong, incorrect technique/ procedure. conclusions: device failure related to user handling, user error caused event, operational context caused event, unusual event. manufacturer narrative: a visual examination of the catheter was done at the cardiac surgery pathology lab. the pigtail and teardrop were separated from the inflow cage at the strut-teardrop attachments. the struts/prongs had been "pinched" in. buckling of several struts and severe deformation of the distal end of the cannula was observed. in addition, delamination of the cannula at the attachment to the pump motor was noted. an inspection of the inlet cage and detached pigtail was conducted at the manufacturer site. all four struts/prongs were detached at the weld to the teardrop. the fracture was irregular and planar. one strut was twisted at the end. there was significant deformation of the cannula proximal to the suction cage. a dimensional analysis of the struts and residual weldment on the teardrop was performed using a video microscope and micrometer. the thickness at the weldment for all four struts was all above the minimum specification for over polishing. performance and/or mechanical testing of the device involved in the incident was impossible due to the condition in which the device was received; therefore, mechanical testing was conducted on representative impella 2.5 pump from a reserve lot to attempt to re-create the failure mode. pumps with undamaged and pre-damaged struts were tested. simulation of the incident was conducted using a mock aorta with a simulated aortic valve. the device was subject to manual manipulation (direct and peripherally) and multiple "attempts" to cross the aortic valve. peripheral manipulation of the device resulted in no damage to the suction cage. repeated attempts to cross the aortic valve were conducted with no strut damage to either the undamaged or pre-damaged test cage. ultimately, the failure mode was recreated only after excessive manipulation was employed to cross the valve similar to the techniques in the reported event and not during the usual and instructed procedural techniques. pump was manufactured as part of lot 0024693. this lot consisted of total assemblies, some were released for distribution. the failed pumps were ruled out for the following reasons; one damaged motor cable, one failed an early leakage current test, one stopped, and one showed high inner leakage current. nothing conspicuous was noted. no other pump from this lot have been involved in a product complaint. nevertheless, a complete review of all complaints received for the impella 2.5 product line was conducted for the reported failure mode (inflow cage breakage) and revealed a total incidents; one of these showed the complete inflow cage detachment from the cannula due to bonding failure after the user deformed the cage manually without patient involvement, two complaints involved breakage at the welding area of demo pumps which had been handled frequently and/or by inexperienced clinicians; and lastly, one complaint involved breakage after a physician noticed damage post-procedure so, inspected the device while manipulating the distal tip until the inflow cage detached. the root cause was determined to be user handling and not product related. the cause of this failure was direct excessive manipulation of the catheter through the ascending aorta to attempt to cross the aortic valve. this included peripheral manipulation and resulted in loads being applied to the suction cage that were higher than could be applied under normal operation. these high loads, combined with multiple attempts to cross the aortic valve, led to fatigue failure of the struts at the welds to the teardrop and separation of the pigtail assembly. in addition, the use of transesophageal echocardiograph y (tee) during the case (instead of fluoroscopy) was contributed to the failure. technical bulletins have been released to cover the issues identified and have been distributed to field personnel. no unanticipated on unacceptable risk has been identified and further corrective actions are not required at this time.
38431 | 1220648-2009-00002 | 1333680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1333680&pc=KFM | 07/01/2008 | 07/01/2008 |  | 02/26/2009 |  |  | 07/01/2008 | 02/26/2009 |  |  |  |  | 02/26/2009 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERC CARDIAC SUPPORT |  |  | KFM | [5 empty] | ABIOMED |  |  | GM |  | 0022474 |  | 004671 |  | Yes |  |  | Yes |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [9 empty] | It has been reported that while attempting to implant a patient with an impella 2.5, the physician was "unable to access the ventricle with the device". according to the physician, the difficulty was directly attributable to the "patient's aorta [being] very tortuous" near the insertion site. the physician therefore removed the device from the patient. upon visual inspection, the "inlet port prongs" were observed to be "frayed open and detaching from the pigtail". the physician continued to examine in the inlet port outside of the patient until ultimately the pigtail detached from the parent device. the physician elected to use an intra-aortic balloon pump device and the patient was successfully supported without adverse effects. no impact was sustained by the patient and ultimately the physician was satisfied with the performance of the device. | [5 empty] | Evaluation results: (extensive tortuosity precluded placement). conclusion: (tortuous precluding placement and inducing resistance). the pump was withdrawn from the patient due to the inability to advance the pump across the aortic valve. the user reported that the "frayed" port prongs were not discovered until after the device was removed from the patient. the placement difficulty was directly attributable to the patient's tortuous aorta. it is unknown how much and what type of force was exerted on the device during the procedure and in the presence of the resistance during the placement attempt. it is known, however, that the device was inspected by the user prior to insertion and found to be without defect. detachment of the pigtail did not occur until after the device was removed from the patient and was being handled by the user. the device involved in the incident was returned to the manufacturer for evaluation and was subjected to visual, microscopic, optical, and dimensional analyses. a preliminary inspection on the returned device confirmed that the pigtail assembly was detached from the parent catheter. the detachment was observed to have occurred at the welds between the struts of the suction cage and metal ball on the pigtail. evaluations of the detachments on a microscopic level revealed that the cross sectional strut areas adjacent to the breaks and the geometries of the wire ends varied among struts. additionally, optical and dimensional measurements detected some variation in the height of the residual weldments on the pigtail ball that was consistent with the variation in the thickness of the broken struts. in order to re-create the failure, four representative devices were subjected to fracture simulation. all four devices were flexed until failure (strut breakage) occurred and were subsequently analyzed with the same methods as the suspect device. the number of flexes observed to induce strut fracture during simulation was greater than the number that is expected to occur during clinical use of the device (zero). the reserve samples demonstrated similar detachment locations. the cross sectional areas at the breakage points were thicker than the device involved in the incident and showed less variation among struts. additionally, the heights at the weldments were found to be greater and with less variation. the failure of the device is ultimately concluded to be a combination of; the tortuous vasculature, a thin section of one of the struts attaching the pigtail assembly to the cannula, and the user's manipulation of the device. the thin strut is suspected to have been exposed to insertion forces beyond its tolerance and ultimately fractured while passing through the tortuous vasculature. one broken strut can compromise the stability of the suction basket which can subsequently result in breakage of the remaining struts if additional manipulation is employed. the remaining struts were reported to have broken after the device was explanted from the patient and during direct manipulation by the user. a device history record review was conducted for the affected pump lot. no anomalies related to the reported failure mode were identified. no other devices from the lot have been involved in a complaint event. a similar complaint review of all incidents reported prior to and including the date of the event was conducted and revealed that of the total number of complaints involving inlet cage damage/deformity during use, 75% were reported by the institution involved in this event. the high rate of damaged inflow cages at this particular user facility indicates the device is not being used in an appropriate manner and that additional training is needed at the site. because the procedural factors during this case are unknown, a conclusive root cause for the strut fracture and subsequent pigtail separation cannot be determined; nevertheless, furth corrective action will be initiated by the manufacturer to ensure that the risk of detachment and/or strut fracture is minimized. current manufacturing procedures cover fabrication of the pigtail assembly and final inspection of the non-sterile pump. the final inspection procedure calls for a visual inspection but no dimensional check. corrective actions will ensure quality of strut fabrication through new dimensional specifications and corresponding final inspection processes. a new dimensional specification for cross sections of the finished struts at the weld to the pigtail ball will be created along with a revised final inspection procedure to verify the welded struts meet the dimensional specification. also, additional training in the use of the impella 2.5 will also be conducted at the user facility. the training module will emphasize that excessive force during the use of the device should be avoided. no further corrective action is warranted at this time.
38432 | 1328986 | 1328986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1328986&pc=KFM | 02/15/2009 | 02/23/2009 | 02/23/2009 | 02/23/2009 | 02/23/2009 |  |  | 02/23/2009 |  |  |  |  | 02/23/2009 |  | CATHETER, CARDIAC ASSIST PUMP |  |  | IMPELLA 2.5 VAD |  |  | KFM |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 41891 |  | 004671 |  | Yes | [5 empty] | No | Invalid Data |  |  | User Facility report | [8 empty] | * | [5 empty] | Malfunction | No | Yes | [49 empty] | Initial submission | [5 empty] | Y | HOSPITAL | * |  | An impella 2.5 vad was placed into the patient. the device was left in as a bridge to surgery in 48 hours after glycemic control was maintained. device motor failed around 34 hours later. narrow pulse pressure noted the day after implant. continued to narrow throughout the day. impella switched out. problem incurred and was switched back to original. that evening, pressures increased and continual purges until the console stated that the motor wasn't functioning. md stat called to pull device. iabp at bedside as a precaution. although the device ifu states that the catheter is meant only for 6 hours of use, the company had stated that it was possible to leave it in for 5 days. device was explanted at time of failure and not replaced since patient had stabilized.====================== manufacturer response for catheter, cardiac assist pump, impella 2.5======================they are picking up the device in a few days for evaluation.
38433 | 1220648-2008-00010 | 1275036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1275036&pc=KFM | 11/14/2008 | 11/26/2008 |  | 12/10/2008 |  |  | 11/26/2008 | 12/10/2008 |  |  | 12/01/2009 |  | 12/10/2008 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERCUTANEOUS CS |  |  | KFM | [5 empty] | ABIOMED |  |  | GM | [5 empty] | Yes |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Study | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | Yes | [20 empty] | Death | Required Intervention |  |  |  |  | GROIN FEMORAL ACCESS AND SUCCESSFUL AT RIGHT GROIN | FEMORAL ACCESS PRIOR TO IMPELLA USE ON | PERCLOSE SUTURING DEVICE: FAILED ATTEMPTED AT LEFT | [20 empty] | Initial submission | [9 empty] | Abiomed, inc. was officially notified on 26 nov, 2008 by hospital that a clinical trial unanticipated adverse device effect report was submitted by the user facility to western institutional review board. the report describes unsuccessful attempts at implanting the impella 2.5 device during the clinical trial, and that the procedure resulted in a hypovolemic shock and patient death. with the same notification, the manufacturer was also made aware that the user facility has submitted a medical device report to the fda's medwatch program. the user facility report number was not provided to the manufacturer. the procedure involved two devices. this medwatch report #1220648-2008-00010) is a supplemental report to the user facility report and is for the first device used in the procedure. | [5 empty] | The device has been received and an investigation is currently in process. conclusions: no conclusion can be drawn (device investigation in progress). additional manufacturer narrative: further information was obtained through follow-up directly with the user facility upon notification that an unanticipated adverse device effect had been reported. the information obtained, in addition to the account of the case provided by an abiomed clinical representative present at the case, reveals the following details and clinical sequence: a written informed consent was obtained from the patient prior to enrollment in the study. after femoral access was obtained, the treating physician performed a bilateral femoral angiogram prior to device insertion to assess the iliac artery anatomy and disease. the left side was chosen for impella insertion. a double pre-close suturing was attempted on the left side of the groin using a perclose device; however, the sutures could not be deployed due to severe calcification. the physician proceeded with pre-closing a femoral access site on the right side of groin successfully. appropriate sheaths were upsized in order to initiate the impella insertion through the right groin. approximately one hour into the procedure, an impella 2.5 device was attempted to be inserted using a guidewire with the monorail technique, but the physician was unable to advance the device beyond the distal abdominal aorta. a significant amount of resistance was encountered at the junction of the aorto-iliac bifurcation. blood loss was noted to have occurred at the origin of the femoral access. the impella catheter was removed and the guidewire was left in situ. an arteriography was subsequently performed which showed presence of ectasia in the right proximal common iliac artery. the physician elected to re-insert the same impella device through the existing sheath but with a second guidewire thereby attempting to sue the existing guidewire as a "buddy" wire for added support and navigation. resistance was still encountered and the catheter could not be advanced. the catheter was withdrawn and the inlet cage was observed to be bent. the physician attributed the damage to the number of attempts to advance the catheter against resistance without success. at that time, after 20 minutes of unsuccessful attempts to implant the impella device, the patient became hypotensive and required atropine, dopamine, levophed, and fluids. as a direct result of this treatment, the patient's heart rate and blood pressure improved promptly, but then the heart rate continued to increase until supraventricular tachycardia with aberration was suspected. the tachycardia spontaneously resolved and then reemerged but was ultimately resolved with one direct current synchronized shock of 120 joules. resumption of normal sinus rhythm was obtained. following this, the physician elected to move forward with attempting to implant a second impella device. repeat attempts at advancing the second device were unsuccessful through the pre-existing sheath as well as after exchanging the sheath to a fresh 13-french impella sheath. the second impella was finally withdrawn and was observed to have experienced similar damage to the inlet cage as the first impella device. insertion attempts had persisted for an additional 20 minutes, during which dopamine was continuously administered. blood loss throughout the procedure eventuated hypovolemia and the patient's blood pressure declined to a critical state. chest compressions were initiated and while in progress, the physician advanced an intra-aortic balloon pump through the left femoral access site to stabilize patient's hemodynamics. resuscitative efforts continued for an additional 48 minutes and included endotracheal intubation, close chest massage, prbc transfusions, and pressors. no spontaneous respiration, blood pressure, or rhythm could be achieved and the patient ultimately expired. an autopsy was requested per the clinical trial protocol, but was not performed by the institution. by manufacturer request, an external review of all details and imaging obtained throughout the procedure was conducted by an independent interventional cardiologist. the review has concluded that "the patient [involved] has severe pvd [peripheral vascular disease] which appears to be severe stenosis at the origin of the right common iliac" and that the "patient should not have been included [in the trial]." significant calcification and peripheral vascular disease are contraindications for the use of the impella device. bleeding was noted on adverse event reports completed by the user facility. bleeding is addressed in the device's instructions for use as a potential adverse event and is a labeled precaution, and an inherent risk of this procedure. the impella devices have been returned to the manufacturer and a full failure investigation is currently in progress. a preliminary visual evaluation of the returned devices confirms the report of bent inflow cages. the catheters show a series of kinks which are demonstrative of excessive insertion forces. additional information and conclusions will be provided in a supplemental manufacturer medwatch report following the completion of the investigation.
38434 | 1220648-2008-00011 | 1275038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1275038&pc=KFM | 11/14/2008 | 11/26/2008 |  | 12/10/2008 |  |  | 11/26/2008 | 12/10/2008 |  |  | 11/01/2009 |  | 12/10/2008 |  |  |  |  | IMPELLA RECOVER LP 2.5 PERCUTANEOUS CS |  |  | KFM | [5 empty] | ABIOMED |  |  | GM | [5 empty] | Yes |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Study | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | Yes | [20 empty] | Death | Required Intervention |  |  |  |  | FEMORAL ACCESS PRIOR TO IMPELLA USE ON | GROIN FEMORAL ACCESS AND SUCCESSFUL AT RIGHT GROIN | PERCLOSE SUTURING DEVICE: FAILED ATTEMPT AT LEFT | [20 empty] | Initial submission | [9 empty] | Abiomed, inc. was officially notified on 26 nov, 2008 by hospital that a clinical trial unanticipated adverse device effect report was submitted by the user facility to western institutional review board. the report describes unsuccessful attempts at implanting the impella 2.5 device during the clinical trial and that the procedure resulted in a hypovolemic shock and patient death. with the same notification, the manufacturer was also made aware that the user facility has submitted a medical device report to the fda's medwatch program. the user facility report number was not provided to the manufacturer. the procedure involved two devices. this medwatch report #1220648-2008-00011) is a supplemental report to the user facility report and is for the second device used in the procedure. | [5 empty] | The device has been received and an investigation is currently in process. conclusions: no conclusion can be drawn (device investigation in progress). additional manufacturer narrative: further information was obtained through follow-up directly with the user facility upon notification that an unanticipated adverse device effect had been reported. the information obtained, in addition to the account of the case provided by an abiomed clinical representative present at the case, reveals the following details and clinical sequence: a written informed consent was obtained from the patient prior to enrollment in the study. after femoral access was obtained, the treating physician performed a bilateral femoral angiogram prior to device insertion to assess the iliac artery anatomy and disease. the left side was chosen for impella insertion. a double pre-close suturing was attempted on the left side of the groin using a perclose device; however, the sutures could not be deployed due to severe calcification. the physician proceeded with pre-closing a femoral access site on the right side of groin successfully. appropriate sheaths were upsized in order to initiate the impella insertion through the right groin. approximately one hour into the procedure, an impella 2.5 device was attempted to be inserted using a guidewire with the monorail technique, but the physician was unable to advance the device beyond the distal abdominal aorta. a significant amount of resistance was encountered at the junction of the aorto-iliac bifurcation. blood loss was noted to have occurred at the origin of the femoral access. the impella catheter was removed and the guidewire was left in situ. an arteriography was subsequently performed which showed presence of ectasia in the right proximal common iliac artery. the physician elected to re-insert the same impella device through the existing sheath but with a second guidewire thereby attempting to sue the existing guidewire as a "buddy" wire for added support and navigation. resistance was still encountered and the catheter could not be advanced. the catheter was withdrawn and the inlet cage was observed to be bent. the physician attributed the damage to the number of attempts to advance the catheter against resistance without success. at that time, after 20 minutes of unsuccessful attempts to implant the impella device, the patient became hypotensive and required atropine, dopamine, levophed, and fluids. as a direct result of this treatment, the patient's heart rate and blood pressure improved promptly, but then the heart rate continued to increase until supraventricular tachycardia with aberration was suspected. the tachycardia spontaneously resolved and then reemerged but was ultimately resolved with one direct current synchronized shock of 120 joules. resumption of normal sinus rhythm was obtained. following this, the physician elected to move forward with attempting to implant a second impella device. repeat attempts at advancing the second device were unsuccessful through the pre-existing sheath as well as after exchanging the sheath to a fresh 13-french impella sheath. the second impella was finally withdrawn and was observed to have experienced similar damage to the inlet cage as the first impella device. insertion attempts had persisted for an additional 20 minutes, during which dopamine was continuously administered. blood loss throughout the procedure eventuated hypovolemia and the patient's blood pressure declined to a critical state. chest compressions were initiated and while in progress, the physician advanced an intra-aortic balloon pump through the left femoral access site to stabilize patient's hemodynamics. resuscitative efforts continued for an additional 48 minutes and included endotracheal intubation, close chest massage, prbc transfusions, and pressors. no spontaneous respiration, blood pressure, or rhythm could be achieved and the patient ultimately expired. an autopsy was requested per the clinical trial protocol, but was not performed by the institution. by manufacturer request, an external review of all details and imaging obtained throughout the procedure was conducted by an independent interventional cardiologist. the review has concluded that "the patient [involved] has severe pvd [peripheral vascular disease] which appears to be severe stenosis at the origin of the right common iliac" and that the "patient should not have been included [in the trial]." significant calcification and peripheral vascular disease are contraindications for the use of the impella device. bleeding was noted on adverse event reports completed by the user facility. bleeding is addressed in the device's instructions for use as a potential adverse event and is a labeled precaution and an inherent risk of this procedure. the impella devices have been returned to the manufacturer and a full failure investigation is currently in progress. a preliminary visual evaluation of the returned devices confirms the report of bent inflow cages. the catheters show a series of kinks which are demonstrative of excessive insertion forces. additional information and conclusions will be provided in a supplemental manufacturer medwatch report following the completion of the investigation.
38435 | 1259550 | 1259550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1259550&pc=KFM | 11/14/2008 | 11/20/2008 | 11/20/2008 | 11/24/2008 | 11/20/2008 | 11/20/2008 |  | 11/24/2008 | 11/24/2008 |  | 12/01/2009 | 11/01/2009 | 11/24/2008 |  | IMPELLA CATHETER | IMPELLA CATHETER |  | IMPELLA LP 2.5 | IMPELLA LP 2.5 |  | KFM | KFM |  |  |  |  | ABIOMED | ABIOMED |  | US | US | CATHETER PUMP #39450 | CATHETER PUMP #39351 | 004610 | 004610 |  |  |  | Not Available | Not Available |  | No | Health Professional | Health Professional |  | User Facility report |  |  | Unknown | [5 empty] | 82 YR |  |  | 68 |  |  | Death | Yes | Yes | [20 empty] | Death | [28 empty] | Initial submission |  |  |  |  | 11/14/2008 | Y | HOSPITAL |  |  | High risk percutaneous coronary intervention attempted. randomized to impella device. attempted unsuccessfully twice to insert device. pt became hemodynamically compromised. adverse events all serious, included bleeding from sheath during insertion hypovolemia, hypotension, supraventricular tachycardia, cardioversion, bradycardia, respiratory failure. ventricular fibrillation and asystole. resuscitation attempts with defibrillation, fluids, vasopressors, mechanical ventilation and cpr were unsuccessful. the cages of each appear to have been bent during insertion attempts.
38436 | 1220648-2008-00009 | 1111016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1111016&pc=KFM | 07/11/2008 | 07/13/2008 |  | 08/12/2008 |  |  | 07/13/2008 | 08/12/2008 |  |  |  |  | 08/12/2008 |  |  |  |  | IMPELLA RECOVER LP 2.5 PECUTANEOUS CARDIAC S |  |  | KFM | [5 empty] | ABIOMED |  |  | GM | [5 empty] | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Health Professional | User facility | Company representation |  |  |  | UNK | [5 empty] | Injury | Yes | Yes | [20 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | The complainant has reported that a patient underwent a procedure involving an impella 2.5 device. during the initial insertion technique of the device, the 13fr peel-away introducer, an accessory of the device's introducer kit, was observed to be leaking from the hemostatic valve around the impella catheter. the bleeding was discovered while the physician was backloading the 7fr sheath. the sheath continued to leak around the impella catheter. an estimated blood loss of 600cc was observed from the hemostatic valve, so a blood transfusion was performed intra-operatively. it was decided to give 2 units of packaged red blood cells before removal of pump. the patient was also given 40mg iv lasix intra-procedure, and total uop for the procedure was 2 litres. the patient seemed much improved after the procedure (sbp staying around 115, pap decreased to 59/26). the impella was successfully weaned over approximately 30 min, without hemodynamic compromise, and the patient was subsequently discharged. | [5 empty] | The device was not returned to the manufacturer; therefore, a failure analysis cannot be completed to determine the root cause of the reported device defect or the relationship of the device to the reported event (blood loss requiring transfusion). a supplemental medwatch report will be submitted if additional information can be obtained.
38437 | 1220648-2008-00007 | 1082142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=1082142&pc=KFM | 06/20/2008 | 06/24/2008 |  | 07/25/2008 |  |  | 06/24/2008 | 07/25/2008 |  |  |  |  | 07/25/2008 |  |  |  |  | IMPELLA RECOVER LP 2.5 CARDIAC SUPPORT SYSTEM |  |  | KFM |  |  | LP2.5 |  |  | ABIOMED, INC. |  |  | GM | [5 empty] | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report |  |  | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | Yes | Yes | [20 empty] | Other | [28 empty] | Initial submission | [9 empty] | The complainant has reported to abiomed, inc that a patient (age and gender unknown) had undergone a procedure involving the use of an impella 2.5 device. during the procedure, there was a problem during the implantation" with the sheath. it was reported that while the sheath was "inside the punctuation", the "head part of the sheath loosened from the shaft unexpectedly and without any mechanical influence". the problem was not noticed immediately by the physician. the complaint reported that the patient "lost quite a lot of blood". the bleeding was stopped using a "backup 13f sheath". the incident had no further impact on the patient's health. | [5 empty] | Device discarded, unable to follow-up. a supplemental medwatch report will be submitted if additional relevant information can be reasonably obtained. (b) (4).